## Final_cold-B
ELECTRICAL TESTS OF DIPOLES
COLD MASS
(WITH END CAPS)
MAIN CONTRACTOR : |  |  | ANSALDO |  | CONTRACT NUMBER :
SUB CONTRACTOR : |  |  |  |  | CONTRACT NUMBER :
COMPONENT NAME : |  |  | COLDMASS
CERN PART NUMBER : |  |  | LHCMBP2A1 |  | Serial Number : |  | 000000
CERN INSPECTION & TEST PLAN NUMBER : |  |  | ITP Step 25a to 25p |  | TEST NUMBER : |  | 25a to 25p
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number
NAME OF TEST OPERATOR: |  |  | M. Cornelis
NAME OF RESPONSIBLE: |  |  | M. Cornelis
DATE OF TEST : |  |  | 2002-02-08 |  | TIME : |  | 9.3
R[dc] of the Dipoles |  |  | ITP 25a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 23.9 | °C
Humidite: | 34 | %
D1 & D2 | [20°C] | * | 3022.37 | mW | * | 3031.5 | mW
Magnet | [20°C] |  | * | 6053.87 | mW
Poles Nr.
R[dc] of the Quench Heaters |  |  | ITP 25b/d
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
| Apt.1 & Apt. 2 |  | [W] | [W] |  | [W] | [W]
|  | * | 19.29 | 19.12 |  | 17.97 | 18.79
|  | * | 19.08 | 18.59 |  | 19.19 | 18.56
R[dc] Spot Heaters |  |  | Upper | Lower
|  |  | [W] | [W]
|  | YT181 | 4.28
|  | YT182 | 4.35 | 4.39 |  | YT184
Cryogenic heater | [5 empty] | not available
Cold mass temperature |  | Nr. :CX_LS_X | TT441 | 51.016 | W | R[iso] >40MW
|  |  | TT461 | 52.806 | W | R[iso] >40MW
R[dc] of the Spools |  |  | ITP 25c
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
| Lyra Side |  | [mW] |  |  | [mW]
| Circuit MCS (Apt.1 + Apt.2) |  | 313.7 | Bus 18/EE831 |  | 2085.6 | Bus 18/EE931
MCS Nr.
ELECTRICAL TESTS OF DIPOLES
COLD MASS
(WITH END CAPS)
Voltage Taps |  |  | ITP 25e
|  |  |  | [mV] |  | [mV]
I = | 1 | A | EE015 |  |  |  | EE211
|  |  | EE013 | 6053.8 |  |  | EE219
|  |  | EE111 | 0.03 |  | 4552.8 | EE212
|  |  | EE112 | 1503.4 |  | 3022.4 | EE213
|  |  | EE119 |  |  |  | EE012
|  |  | EE113 | 1503.4 |  |  | EE014
Capacitance |  |  | ITP 25h
[ Megger 100V ; > 1'] |  |  |  |  | MAGNET
Magnet | --> Ground |  | * | 420 | nF
INDUCTANCE L/Q[10Hz] |  |  | ITP 25g
[Value from gain phase analyser] |  |  |  |  | MAGNET
Magnet |  |  | (L) | 100.25 | mH
|  |  | (Q) | 0.991
INDUCTANCE L/Q[100Hz] |  |  | ITP 25g
[Value from gain phase analyser] |  |  |  |  | MAGNET
Magnet | * |  | (L) | 96.24 | mH
|  |  | (Q) | 4.16
INDUCTANCE L/Q[1.0KHz] |  |  | ITP 25g
[Value from gain phase analyser] |  |  |  |  | MAGNET
Magnet | * |  | (L) | 61.45 | mH
|  |  | (Q) | 1.57
ELECTRICAL TESTS OF DIPOLES
COLD MASS
(WITH END CAPS)
INSULATION RESISTANCE |  |  | ITP 25o
Magnet + QHs + Bus MB ---> Ground |  |  |  | 5.2 | GW |  | >1000MW
[V=500V, 30'']
| QHs(All) Magnet |  | *
|  |  |  | 4.7 | GW |  | >1000MW
[V=500V, 30'']
Circuit spools [Bus 18/19] ---> Ground |  |  |  | 25.2 | GW |  | >1000MW
Bus [11/12] --> Ground |  |  |  | 121 | GW |  | >1000MW
[V=500V, 30'']
DIELECTRIC TESTS |  |  | ITP 25i/n
Magnet + QHs + Bus MB ---> Ground |  |  |  | 0.09 | µA |  | <25µA
[V=500V, 2']
| QHs(All) Magnet |  |  | 0.1 | µA |  | <20µA
[V=500V, 2']
Circuit spools [Bus 18/19] ---> Ground |  |  |  | 0.02 | µA
Bus [11/12] --> Ground |  |  |  | 0.004 | µA
[V=500V, 30'']
DISCHARGE TEST |  |  | ITP 25m
V = | 25V/turn |  |  |  | MAGNET
[C= 2µF]
| Magnet [4KV] |  |  | 2.78 | mS
| Quench Heaters |  | YT111 & 112 | YT121 & 122 |  | YT211 & 212 | YT221 & 222
| [V=1.600KV] |  |  |  | µS |  |  | µS
| [C= 2µF] | Current | 49 | 49 | A | 49 | 49 | A
Electrical Continuty of Coldmass |  |  | ITP 25p
I = | 20 | A |  |  | COLD MASS
Temp.* : | 23.9 | °C
| Coldmass |  | * |  | mW
LHC/MMS TRACEABILITY INFORMATION

## Internal Vtaps
MBP2 COLD MASS
Temp. |  | 19 | °C |  |  |  |  | INTERNAL Vtaps |  |  | CERN: | 2002-02-08
Hy. |  | 34 | % |  |  |  |  | A1 with end covers
I : | 1 |  | A
[7 empty] | [mV] |  |  |  |  | [mV]
|  |  |  | I | EE | 1219 |  |  | I | EE | 2209 | 3391.6
|  |  |  | H | EE | 1218 |  |  | H | EE | 2208 | 4001.1
|  |  |  | SD | EE | 188 |  |  | SD | EE | 288
|  |  |  | SC | EE | 187 | 2029.8 |  | SC | EE | 287 | 4021.4
|  |  |  | B | EE | 1212 |  |  | B | EE | 2202 | 4041.8
|  |  |  | A | EE | 1211 |  |  | A | EE | 2201 | 4042
|  |  |  | SB | EE | 186 | 1993.2 |  | SB | EE | 286 | 4058.3
|  |  |  | SA | EE | 185 |  |  | SA | EE | 285
UPPER
|  |  |  | C | EE | 1213 |  |  | C | EE | 2203 | 4075.3
|  |  |  | D | EE | 1214 |  |  | D | EE | 2204
|  |  |  | E | EE | 1215 |  |  | E | EE | 2205
|  |  |  | F | EE | 1216 |  |  | F | EE | 2206
|  |  |  | G | EE | 1217 |  |  | G | EE | 2207
|  |  |  | G | EE | 1207 |  |  | G | EE | 2217
|  |  |  | F | EE | 1206 |  |  | F | EE | 2216
|  |  |  | E | EE | 1205 |  |  | E | EE | 2215
|  |  |  | D | EE | 1204 |  |  | D | EE | 2214
|  |  |  | C | EE | 1203 |  |  | C | EE | 2213 | 5025.5
LOWER
|  |  |  | SA | EE | 181 |  |  | SA | EE | 281
|  |  |  | SB | EE | 182 |  |  | SB | EE | 282
|  |  |  | A | EE | 1201 |  |  | A | EE | 2211 | 5058.6
|  |  |  | B | EE | 1202 |  |  | B | EE | 2212 | 5058.7
|  |  |  | SC | EE | 183 | 977.4 |  | SC | EE | 283 | 5078.3
|  |  |  | SD | EE | 184 | 977.3 |  | SD | EE | 284 | 5078.4
|  |  |  | H | EE | 1208 |  |  | H | EE | 2218 | 5098.4
|  |  |  | I | EE | 1209 |  |  | I | EE | 2219 | 5693.5
I. 1.00A
OUTER
INNER
INNER
OUTER
OUTER
INNER
INNER
OUTER
D1
D2
1.00
C0M. mesure

## External Vtaps
[7 empty] | EXTERNAL VOLTAGE TAPS CONTROL |  |  |  |  | Date: | 2002-02-08
MBP2 A1_Cold Mass with end covers
| Temperature: | 19 | °C
| Humidity: | 34 | %
| Current: | 1 | A |  |  |  |  | MB
[6 empty] | V[mV] |  |  |  |  | V[mV]
[7 empty] | EE513 |  |  | EE511
|  |  | Bus MQ | [5 empty] | diode
[10 empty] | EE012 | 0
[10 empty] | EE111 | 0.03
V[mV] |  | QD |  | V[mV] | QF
EE711 |  |  |  | EE611
0 |  |  |  | 0
0 |  |  |  | 0
EE712 |  |  |  | EE612
[10 empty] | EE113 | 1503.4
[10 empty] | EE112 | 1503.4
[10 empty] | EE114 | 3022.4
[10 empty] | EE211 | 3022.4
EE791 |  |  |  | EE691
EE792 |  |  |  | EE692 | [5 empty] | EE213
[10 empty] | EE212 | 4552.8
[10 empty] | EE214 | 6053.9
[10 empty] | EE013 | 6053.9
diode
EE512
[7 empty] | EE592 |  |  | EE591
Upper Pole
inner
outer
inner
inner
inner
outer
outer
outer
Upper Pole
Lower Pole
Lower Pole
External Bus
External Bus
External Bus
DIPOLE D1 Left
DIPOLE D2 Right
Internal Bus
Internal Bus
External Bus
External Bus
Ap1
int
ext
int
ext
QF
QD
int
ext
MB
Ap2
Connection side
Non Accessible