## 2066
|  |  | ELECTRICAL TESTS OF DIPOLES | [8 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
|  |  | (WITH BEAM TUBES) | [6 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ANSALDO SUPERCONDUTTORI s.p.a. |  | CONTRACT NUMBER : F462 |  |  |  | YT111 | 22.8836 | Magnet | 6.0654 | 6065.416 | 6164.0407 | 10 | 0.9 | 21.8 |  | 0.881 | 21.8
SUB CONTRACTOR : | [8 empty] | YT112 | 22.954 | Aperture1 | 3.0311 | 3031.117 | 3080.4035 | 100 | 4.68 | 21.4 |  | 4.67 | 21.5
[9 empty] | YT121 | 21.3204 | Aperture2 | 3.0343 | 3034.306 | 3083.6443 | 1000 | 1.61 | 14.9 |  | 1.61 | 14.9
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 21.6993 | Apt1 Pup | 1.5146 | 1514.641 | 1539.2693
CERN PART NUMBER : |  |  | HCMBB_A001 |  | Serial Number : |  | 2066 |  | YT211 | 22.4011 | Apt1 Plw | 1.5165 | 1516.472 | 1541.1301 | 10 | 0.89 | 21.6 |  | 0.891 | 21.8
[9 empty] | YT212 | 22.653 | Apt2 Pup | 1.5154 | 1515.414 | 1540.0549 | 100 | 4.72 | 21.5 |  | 4.63 | 21.5
CERN INSPECTION & TEST PLAN NUMBER : 9a To 9i |  |  |  |  | TEST NUMBER : |  | 9a to 9i |  | YT221 | 20.9903 | Apt2 Plw | 1.5189 | 1518.888 | 1543.5854 | 1000 | 1.63 | 14.7 |  | 1.62 | 14.6
NON CONFORMITY: |  |  | No |  | NON CONFORMITY: |  | Number |  | YT222 | 21.5865
[15 empty] | 10 | 0.888 | 43.3 |  | 0.898 | 43.6
NAME OF TEST OPERATOR: |  |  | M.Levrero | [11 empty] | 100 | 4.73 | 42.8 |  | 4.65 | 42.9
NAME OF RESPONSIBLE: |  |  | B.Caserza | [11 empty] | 1000 | 1.61 | 29.5 |  | 1.61 | 29.5
DATE OF TEST : |  |  | 2003-12-22 |  | TIME : |  | 7:30
[15 empty] | 10 | 0.903 | 87
[15 empty] | 100 | 4.67 | 85.7
[15 empty] | 1000 | 1.61 | 59
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 16 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 30 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1514.641 | 1516.472 |  | 1515.414 | 1518.888
Poles [20°C] |  | * | 1539.2693 | 1541.1301 |  | 1540.0549 | 1543.5854
D1 & D2 |  | * | 3080.4035 | mW |  | 3083.6443 | mW
Magnet |  |  | * | 6164.0407 | mW
|  | Poles Nr. | 02P00262 | 02P00263 |  | 02P00264 | 02P00265
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  | * | 22.8836 | 21.3204 |  | 22.4011 | 20.9903
|  | * | 22.954 | 21.6993 |  | 22.653 | 21.5865
Capacitance |  |  | ITPstep 9d/e
[ Megger 100V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs | --> Dipoles |  | 200 | nF |  | 200 | nF
D1 & D2 | --> Masse |  | 200 | nF |  | 200 | nF
Magnet | --> Masse |  | * | 400 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBE)
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.8 | 21.6 |  | 21.8 | 21.8
|  | (Q) | 0.9 | 0.89 |  | 0.881 | 0.891
D1 & D2 |  | (L) | 43.3 | mH |  | 43.6 | mH
|  | (Q) | 0.888 |  |  | 0.898
Magnet |  |  | (L) | 87 | mH
|  |  | (Q) | 0.903
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES | * | (L) | 21.4 | 21.5 |  | 21.5 | 21.5
|  | (Q) | 4.68 | 4.72 |  | S16 | 4.63
D1 & D2 | * | (L) | 42.8 | mH |  | 42.9 | mH
|  | (Q) | 4.73 |  |  | 4.65
Magnet | * |  | (L) | 85.7 | mH
|  |  | (Q) | 4.67
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES | * | (L) | 14.9 | 14.7 |  | 14.9 | 14.6
|  | (Q) | 1.61 | 1.63 |  | 1.61 | 1.62
D1 & D2 | * | (L) | 29.5 | mH |  | 29.5 | mH
|  | (Q) | 1.61 |  |  | 1.61
Magnet | * |  | (L) | 59 | mH
|  |  | (Q) | 1.61
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBE)
INSULATION RESISTANCE
Dipole D1 + QHs D1 ----> Masse * |  |  |  | 35.2 | GW |  | >1000MW
[V=1.0.KV, 60'']
Dipole D2 + QHs D2 ----> Masse * |  |  |  | 36 | GW |  | >1000MW
[V=1.0.KV, 60'']
| QHs D1(8x) D1 |  | * | 18.5 | GW |  | >1000MW
[V=1.0.KV, 60'']
| QHs D2(8x) D2 |  | * | 18.7 | GW |  | >1000MW
[V=1.0.KV, 60'']
DIPOLE 1
| POLE UP -----> POLE LW |  | * | 61.5 | GW |  | >1000MW
[V=1.0.KV, 60'']
DIPOLE 2
| POLE UP -----> POLE LW |  | * | 61.5 | GW |  | >1000MW
DIELECTRIC TESTS |  |  | ITP step 9f/h
Dipole D1 + QHs D1 ----> Masse * |  |  |  | 0.1 | µA |  | < 20 µA
[V=5.0KV 5' ]
Dipole D2 + QHs D2 ----> Masse * |  |  |  | 0.1 | µA |  | < 20 µA
[V=5.0KV 5' ]
| QHs D1(8x) D1 |  | * | 0.1 | µA |  | < 20 µA
[V=3.0.KV 5' ]
| QHs D2(8x) D2 |  | * | 0.1 | µA |  | < 20 µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  | * | 0.1 | µA |  | < 20 µA
[V=3.0.KV 5' ]
DIPOLE 2
| POLE UP -----> POLE LW |  | * | 0.1 | µA |  | < 20 µA
[V=3.0.KV 5' ]
DISCHARGE TEST |  |  | ITP step 9i
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] | * | 1.028 | 0.974 |  | 0.972 | 0.968
| Dipoles D1 & D2 [4KV] * |  | 1.71 | mS | * | 1.7 | mS
| Magnet [4KV] |  | * | 2.54 | mS
| Quench Heaters |  | YT111 & 112 | YT121 & 122 |  | YT211 & 212 | YT221 & 222
| [V=850V] |  | 1211/1214 | 1151/1146 | µS | 1184/1181 | 1159/1129 | µS
| [C= 7mF] | Current | 35,8/36,5 | 38,8/38,1 | A | 37,0/36,5 | 39,2/38,0 | A
Insul. Resistance All QHs --> Masse |  |  |  | 32.8 | GW |  | >1000MW
[V=3.0.KV, 60'']
Insul. Resistance All QHs ---> Magnet |  |  |  | 10.3 | GW |  | >1000MW
[V=3.0.KV, 60'']
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION

## Check sheet
Quenchs | Poles | ……………..….Induttanza |  | Resistenza Isolamento | Alta Tensione
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK | OK | OK
OK | OK | OK | OK
OK |  | OK | OK |  | Discharge test
|  | OK | OK |  | OK
|  | OK |  |  | OK
|  | OK |  |  | OK
|  | OK |  |  | OK
Nº poli | Data | Nº magnete | Temperatura | Umidità | OK
OK | OK | OK | OK | OK | OK
OK
OK |  | Discharge test su quenchs
OK |  | OK | OK | OK | OK
|  | OK | OK | OK | OK