| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B. Caserza – F. Terzi | 5. Date: 13/05/2003 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: HCMBBRA001-02000009 |
| 3. Contract No: F302/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary | 8. Dwg No: LHCMBB_A0002/F (683RM08451) |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection |  Final inspection |
|  In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During He leak test on CM 2009, a leak has been detected between cold mass pressurized at 26 bar and insulating vacuum. The calculated leak rate is 5.92E-9 mbarl/s (nom. value 1.0E-9 mbarl/s max.). As the leak is very small, we believe it comes from an "Helicoflex" seal. See next page for details.

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) - The results of the leak test have been communicated to AT-VAC at CERN (E.Mahner and J.Hansen) for checking and they confirmed that such non-conformity cannot be accepted. AT-VAC proposed to check the leak-tightness of the test equipment. - Ansaldo should then redo the test after having replaced the "Helicoflex" seals. - CERN informed Ansaldo that an alternative design of the end flanges using welded caps in lieu of bolted joints is being evaluated. More information on such solution will be communicated in the next days. |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary)

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: F.Savary | Date: 13/05/03 |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: F.Terzi Date: 13/05/03 | Name: Date: |

10. Description of non conformity (continuation page)

Please find here below the template about pressure/leak test performed on CM 2009.

F. Terzi