| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B. Caserza – F. Terzi | 5. Date: 18/07/2003 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: HCMBARA001-02000020 |
| 3. Contract No: F302/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary | 8. Dwg No: LHCMBB_A0002/F (683RM08451) |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection |  Final inspection |
|  In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During He leak test on CM 2020, a leak has been detected between cold mass pressurized to 26 bar and insulating vacuum. The test has been performed again under CERN witness (F.Savary and J.Hansen), on 18/07/2003, with similar result (see next page for details): calculated leak rate = 3.70E-9 mbarl/s (nom. value 1.0E-9 mbarl/s max.). Based on previous experience (ref. to NC CERN 106), ASG suggests to weld a s.s. ring in order to close the flange of the capillary tube cold head and to repeat the test.

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) - The proposal of Ansaldo to close the top box of the IFS by welding the cover flange with a stainless steel ring in lieu of the 8 clamps can be accepted. The pressure and helium leak test shall be redone. - For the execution of the final electrical tests, Ansaldo will remove the cover flange by grinding the weld. - After the execution of the electrical tests, Ansaldo will close the top box of the IFS with a “Viton” seal by spot-welding the cover flange with the same stainless steel ring. - In order to protect the instrumentation wiring during the subsequent work packages at CERN (the top box has to be cut off), Ansaldo is requested to put inside the top box a stainless steel sleeve of sufficient height (more than 45 mm). |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary)

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: F.Savary | Date: July, 25th 2003 |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: F. Terzi Date: 25/07/2003 | Name: Date: |

10. Description of non conformity (continuation page)

18/07/2003

Please find here below the template about pressure/leak test performed on CM 2020.

This test has been performed under CERN witness, after opening the flange of the capillary tube cold head and substitution of the Helicoflex ring.

F. Terzi