| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B.Caserza – F. Terzi | 5. Date: 15/01/2004 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: HCMBALA001-02000046 |
| 3. Contract No: F462/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary | 8. Dwg No: LHCMBB_A0002/F (683RM08451) |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection |  Final inspection |
|  In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During electrical test performed before welding of end covers on CM 2046, the temperature sensor TT821 (serial nr. CX_LS_X14243) showed an electrical resistance R=133 Ohm (at T=17°C) – (reference value R=46.19 Ohm at 22°C).

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary)

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: | Date: |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: Date: | Name: Date: |