| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B.Caserza, F. Terzi | 5. Date: 18/10/2004 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: MCS-MA-E0324 |
| 3. Contract No: F462/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary, M. Bajko | 8. Dwg No: 683RM08450/B |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection |  Final inspection |
|  In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During electrical test at 1.5 kV, performed before welding of end covers on CM 2118, a low insulation resistance to ground has been measured on corrector magnet MCS-MA-E0324. Measured values: leakage current I=10μA (nom. value 2μA) Corrector magnet has been disinstalled from CM and tested again: the same value of leakage current (10μA) was found.

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) Corrector magnet id nr. MCS-MA- E0324 has been dismantled from CM and will be sent to CERN for inspection. A new corrector magnet has been installed on CM 2118 Apt.2: id. nr. MCS-MA-E0709 |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary) CERN will make investigation on the problem together with our colleagues AT MEL.

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: Marta Bajko | Date: 22.10.2004 |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: F. Terzi Date: 25/10/2004 | Name: Date: |