| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B.Caserza | 5. Date: 15/01/2003 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: pole P092 |
| 3. Contract No: F302/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary | 8. Dwg No: LHCMB__A0001/G (620RM08234/L) |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection | Final inspection |
| In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During discharge test performed on c.c. 2023 after collaring, an inter-turn discharge occurred on pole nr. P092 (aperture 2 – upper pole).

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) - The pole cannot be used “as is”. The assembly has to be de-collared for investigation and repair. |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary)

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: F.Savary | Date: 4th June 2003 |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated (see next page) |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: F. Terzi Date: 04/06/2003 | Name: Date: |

CLOSURE OF THE NON CONFORMITY

04/06/2003

Here is a short description of  the corrective actions performed on CC 2023:

- de-collaring

- visual inspection of ground insulation and QHs about dipole D1 (aperture 1).  As they were good, in agreement with CERN personnel during a visit in ASG on 18/01/2003, we didn't changed any components.

- on dipoleD2 (aperture 2), all the ground insulation and the QHs have been changed

- re-collaring

- electrical test: OK

F. Terzi