| – |
| --- |
| Non conformity Report |  |
| IDENTIFICATION |  |
| 1. Originator’s Name: B. Caserza – F. Terzi | 5. Date: 09/04/2003 |
| 2. Contractor/Supplier: Ansaldo Superconduttori S.p.A. | 6. Part description: HCMBB_A001-02000021 |
| 3. Contract No: F302/LHC/LHC | 7. Qty: 1 |
| 4. Project Engineer: F. Savary | 8. Dwg No: LHCMBB_A0002/F (683RM08451) |

| 9. Found during what activity: |  |
| --- | --- |
|  Incoming inspection |  Final inspection |
|  In-process inspection |  Other: |

10. Description of non conformity (use continuation page if necessary)

During the welding of the end plate on CS of cold mass nr. 2021, the connecting wire of temperature sensor TT821 was cut by the edge of the end plate and the shell. Now the length of the cable, out of the end plate, is reduced to 1m about, instead of 4m nominal value. After the electrical test, that the temperature sensor (CX_LS_X13727) is working.

11. Action taken to prevent misuse (use continuation page if necessary)

| IMPORTANCE |  |  |
| --- | --- | --- |
| 12. |  Non critical |  Critical |

| DISPOSITION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 13. |  Use-as-is |  Repair |  Reject |  Rework |  Return to supplier |
| Description of proposed action (use continuation page if necessary) The length of cabling remaining available (1 m) can be extended to the nominal value, 4 m, by welding/soldering a new piece of cabling of about 3 m (as long as necessary to reconstruct it according to nominal length for a correct assembly in the instrumentation feed-through system). There are 4 wires to be lengthened. A heat-shrinkable sleeve (type RT 102, qualified for low temperature application) shall protect each connection. The 4 connections shall be shifted one with respect to the other so that they do not overlap. Then, the 4 connections shall be protected in a common heat-shrinkable sleeve. The section of repairing shall be totally contained in the volume of the end cover (this repairing-zone cannot go into the IFS). Finally, the standard electrical tests shall be repeated and results reported in the traveller. t |  |  |  |  |  |

CORRECTIVE/PREVENTIVE ACTION

14. Description of proposed action (use continuation page if necessary)

| APPROVAL OF NON CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 15 | Project Engineer: F.Savary | Date: 9th of April 2003 |

| APPROVAL OF CRITICAL NON CONFORMITIES |  |  |
| --- | --- | --- |
| 16 | Project Management: | Date: |

| CLOSURE OF THE NON CONFORMITY |  |  |
| --- | --- | --- |
| Planned actions have been completed and corrective/preventive actions have been initiated |  |  |
|  | For non critical non conformities Quality Manager or Project Engineer | For critical non conformities Project Engineer |
| 17 | Name: F. Terzi Date: 05/05/2003 | Name: Date: |