## 2002
ITP Nr.
| LHC - Cold masses: HELIUM MASS SPECTROMETER LEAK TEST REPORT | [14 empty] | 23
24
Cold Mass Nr. |  | 2023
Step Nr. |  | 1 |  |  |  | 2 |  |  |  | 3 |  |  |  | 4
Volume / Volume to be tested |  | CM -> Vacuum |  |  |  | CM -> cold bore tubes |  |  |  | CM -> Heat Exch. |  |  |  | Heat Exch -> Vacuum
Fuga calibrata / Calibrated leak parameter
Calibrated leak N°. |  | 4011007195 |  |  |  | 4011007225 |  |  |  | 4011007195 |  |  |  | 4011007195
Data calibr. / calibration date |  | 2002-10-08 |  |  |  | 2002-10-08 |  |  |  | 2002-10-08 |  |  |  | 2002-10-08
Temp. calibrazione fuga / Calibration Temp. | T | 23 | °C |  | T | 23 | °C |  | T | 23 | °C |  | T | 23 | °C
Valore nom. fuga calibrata / Calibrated leak nom. value | qFR | 3e-08 | mbar l s-1 |  | qFR | 3.3e-08 | mbar l s-1 |  | qFR | 3e-08 | mbar l s-1 |  | qFR | 3e-08 | mbar l s-1
Calibrazione del sistema / System calibration
Conc. He nelle linee di test (100%) / Volumetric fraction of tracer gas in the injection envelope | C | 1 |  |  | C | 1 |  |  | C | 1 |  |  | C | 1
T ambiente / Test temp. | T | 31 | °C |  | T | 31 | °C |  | T | 30.1 | °C |  | T | 31 | °C
Fuga calibrata con correz. T ed età / Size of calib. leak after corr. for ageing and T) | qFR | 3.84e-08 | mbar l s-1 |  | qFR | 4.224e-08 | mbar l s-1 |  | qFR | 3.7455e-08 | mbar l s-1 |  | qFR | 3.84e-08 | mbar l s-1
Segnale residuo prima delle misure di SFR / Residual signal prior SFR meas. | RFR | 4.79e-09 | mbar l s-1 |  | RFR | 4.09e-10 | mbar l s-1 |  | RFR | 3.34e-10 | mbar l s-1 |  | RFR | 4.79e-09 | mbar l s-1
Segnale del LD / Signal given by the calibrated leak | SFR | 4.48e-08 | mbar l s-1 |  | SFR | 4.25e-08 | mbar l s-1 |  | SFR | 4.89e-08 | mbar l s-1 |  | SFR | 4.48e-08 | mbar l s-1
Min. dev. segnale (=2x amp. segn. residuo) / Smallest read. signal dev. (= 2 x ampl. of RFR noise) | Sm | 2e-11 | mbar l s-1 |  | Sm | 2e-12 | mbar l s-1 |  | Sm | 2e-12 | mbar l s-1 |  | Sm | 2e-11 | mbar l s-1
Tempo di attesa stabiliz. segnale / Time to achieve stabilised leak signal | 3t | 180 | sec |  | 3t | 180 | sec |  | 3t | 180 | sec |  | 3t | 180 | sec
SENSIBILITA' DEL TEST / Sensitivity of the leak test | qGm | 1.9195e-11 | mbar l s-1 |  | qGm | 2.0071e-12 | mbar l s-1 |  | qGm | 1.5424e-12 | mbar l s-1 |  | qGm | 1.9195e-11 | mbar l s-1
Condizioni del test / Leak test conditions
Pressione del sistema / System pressure | P | 4.9e-05 | mbar |  | P |  | mbar |  | P |  | mbar |  | P | 4.9e-05 | mbar
Segnale residuo del cercafughe ad inizio test / Residual signal prior to SF measurement | RF | 4.81e-09 | mbar l s-1 |  | RF | 4.65e-10 | mbar l s-1 |  | RF | 3.59e-10 | mbar l s-1 |  | RF | 4.94e-09 | mbar l s-1
Segnale del LD a fine test / Signal given by the leak after 30 min. (>3t) | SF | 5.23e-09 | mbar l s-1 |  | SF | 4.39e-10 | mbar l s-1 |  | SF | 3.38e-10 | mbar l s-1 |  | SF | 5.11e-09 | mbar l s-1
CALCOLO DELLA FUGA / Leak evaluation | qG | 4.031e-10 | mbar l s-1 |  | qG | <1.0E-10 | mbar l s-1 |  | qG | <1.0E-05 | mbar l s-1 |  | qG | 1.6316e-10 | mbar l s-1
VALORE DI RIFERIMENTO / REF. VALUE (MAX) |  | 1e-09 | mbar l s-1 | at 26 bar |  | 1e-10 | mbar l s-1 | at 26 bar |  | 1e-05 | mbar l s-1 | at 26 bar |  | 1e-09 | mbar l s-1 | at 5 bar
CONFORMANCE |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES
Doc. di riferimento / Ref. documents
CERN contract number: F302/LHC/LHC
CERN technical spec.: LHC MMS-98-198 rev.2
Leak test procedure (Ref. N°, Revision): 780RM09442 rev.0
Strumentazione / Test equipment | on vessel |  |  |  | on c.b.t. lines |  |  |  | on heat exchanger line |  |  |  | on vessel
Helium Mass Spectrometer type: | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260
Pressure gauge type: | full range compact PFEIFFER PKR 251 | [11 empty] | full range compact PFEIFFER PKR 251
Pumping group: | turbo pump LEYBOLD PT 360 lt/s | [11 empty] | turbo pump LEYBOLD PT 360 lt/s
| rotary vane pump PFEIFFER DUO 65 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 65 m3/h
Prepared by: Name / Date | Note / Remarks
Caserza - 08/07/03
Approved by: Name / Date
Terzi - 08/07/03
Checked by: Name / Date
Gagliardi - 08/07/03
Checked at CERN by / Signature / Date