## Collegamenti
Magnete N. | 2076 |  | test ok
Temperatura °C | 20
Calibrazione fuga camera
Valore inizio | 5.52e-09
Valore fine | 2.65e-08
Minimo segnale | 2e-11
Tempo di risposta | 900
Calibrazione fuga linea "X"
Valore inizio | 8.81e-10
Valore fine | 2.92e-07
Minimo segnale | 2e-12
Tempo di risposta | 1800
Calibrazione fuga linea V1 e V2
Valore inizio | 5.28e-10
Valore fine | 3.01e-08
Minimo segnale | 2e-12
Tempo di risposta | 900
Test del magnete 26 bar
Pressione nella camera | 5.2e-05
Valore cercafughe inizio test | 5.57e-09
Valore cercafughe fine test 30 min | 5.27e-09
Test V1 V2
Pressione nella camera
Valore cercafughe inizio test | 5.13e-10
Valore cercafughe fine test 30 min | 5.22e-10
Test della linea "X"
Pressione nella camera
Valore cercafughe inizio test | 8.43e-10
Valore cercafughe fine test 30 min | 7.9e-10
Test della linea "X" 5 bar
5.2e-05
Valore cercafughe inizio test | 5.07e-09
Valore cercafughe fine test 30 min | 4.97e-09
,,,,,,,,,

## 2076
ITP Nr.
| LHC - Cold masses: HELIUM MASS SPECTROMETER LEAK TEST REPORT | [14 empty] | 23
24
Cold Mass Nr. |  | 2076
Step Nr. |  | 1 |  |  |  | 2 |  |  |  | 3 |  |  |  | 4
Volume / Volume to be tested |  | CM -> Vacuum |  |  |  | CM -> cold bore tubes |  |  |  | CM -> Heat Exch. |  |  |  | Heat Exch -> Vacuum
Fuga calibrata / Calibrated leak parameter
Calibrated leak N°. |  | 4011007195 |  |  |  | 4011007225 |  |  |  | FC03000505 |  |  |  | 4011007195
Data calibr. / calibration date |  | 2002-10-08 |  |  |  | 2002-10-08 |  |  |  | 2003-05-20 |  |  |  | 2002-10-08
Temp. calibrazione fuga / Calibration Temp. | T | 23 | °C |  | T | 23 | °C |  | T | 20 | °C |  | T | 23 | °C
Valore nom. fuga calibrata / Calibrated leak nom. value |  | 3e-08 | mbar l s-1 |  |  | 3.3e-08 | mbar l s-1 |  |  | 1.4e-07 | mbar l s-1 |  |  | 3e-08 | mbar l s-1
Calibrazione del sistema / System calibration
Conc. He nelle linee di test (100%) / Volumetric fraction of tracer gas in the injection envelope | C | 1 |  |  | C | 1 |  |  | C | 1 |  |  | C | 1
T ambiente / Test temp. | T | 20 | °C |  | T | 20 | °C |  | T | 20 | °C |  | T | 20 | °C
Fuga calibrata con correz. T ed età / Size of calib. leak after corr. for ageing and T) | qFR | 2.685e-08 | mbar l s-1 |  | qFR | 2.9535e-08 | mbar l s-1 |  | qFR | 1.4e-07 | mbar l s-1 |  | qFR | 2.685e-08 | mbar l s-1
Segnale residuo prima delle misure di SFR / Residual signal prior SFR meas. | RFR | 5.52e-09 | mbar l s-1 |  | RFR | 5.28e-10 | mbar l s-1 |  | RFR | 8.81e-10 | mbar l s-1 |  | RFR | 5.52e-09 | mbar l s-1
Segnale del LD / Signal given by the calibrated leak | SFR | 2.65e-08 | mbar l s-1 |  | SFR | 3.01e-08 | mbar l s-1 |  | SFR | 2.92e-07 | mbar l s-1 |  | SFR | 2.65e-08 | mbar l s-1
Min. dev. segnale (=2x amp. segn. residuo) / Smallest read. signal dev. (= 2 x ampl. of RFR noise) | Sm | 2e-11 | mbar l s-1 |  | Sm | 2e-12 | mbar l s-1 |  | Sm | 2e-12 | mbar l s-1 |  | Sm | 2e-11 | mbar l s-1
Tempo di attesa stabiliz. segnale / Time to achieve stabilised leak signal | 3t | 900 | sec |  | 3t | 900 | sec |  | 3t | 1800 | sec |  | 3t | 900 | sec
SENSIBILITA' DEL TEST / Sensitivity of the leak test | qGm | 2.5596e-11 | mbar l s-1 |  | qGm | 1.9975e-12 | mbar l s-1 |  | qGm | 9.6181e-13 | mbar l s-1 |  | qGm | 2.5596e-11 | mbar l s-1
Condizioni del test / Leak test conditions
Pressione del sistema / System pressure | P | 5.2e-05 | mbar |  | P |  | mbar |  | P |  | mbar |  | P | 5.2e-05 | mbar
Segnale residuo del cercafughe ad inizio test / Residual signal prior to SF measurement | RF | 5.57e-09 | mbar l s-1 |  | RF | 5.13e-10 | mbar l s-1 |  | RF | 8.43e-10 | mbar l s-1 |  | RF | 5.07e-09 | mbar l s-1
Segnale del LD a fine test / Signal given by the leak after 30 min. (>3t) | SF | 5.27e-09 | mbar l s-1 |  | SF | 5.22e-10 | mbar l s-1 |  | SF | 7.9e-10 | mbar l s-1 |  | SF | 4.97e-09 | mbar l s-1
CALCOLO DELLA FUGA / Leak evaluation | qG | <1.0E-09 | mbar l s-1 |  | qG | 8.9887e-12 | mbar l s-1 |  | qG | <1.0E-09 | mbar l s-1 |  | qG | <1.0E-09 | mbar l s-1
VALORE DI RIFERIMENTO / REF. VALUE (MAX) |  | 1e-09 | mbar l s-1 | at 26 bar |  | 1e-10 | mbar l s-1 | at 26 bar |  | 1e-05 | mbar l s-1 | at 26 bar |  | 1e-09 | mbar l s-1 | at 5 bar
CONFORMANCE |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES
Doc. di riferimento / Ref. documents
CERN contract number: F302/LHC/LHC
CERN technical spec.: LHC MMS-98-198 rev.2
Leak test procedure (Ref. N°, Revision): 780RM09442 rev.0
Strumentazione / Test equipment | on vessel |  |  |  | on c.b.t. lines |  |  |  | on heat exchanger line |  |  |  | on vessel
Helium Mass Spectrometer type: | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260
Pressure gauge type: | full range compact PFEIFFER PKR 251 | [11 empty] | full range compact PFEIFFER PKR 251
Pumping group: | turbo pump LEYBOLD PT 360 lt/s | [11 empty] | turbo pump LEYBOLD PT 360 lt/s
| rotary vane pump PFEIFFER DUO 65 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 65 m3/h
Prepared by: Name / Date | Note / Remarks
PIU S. - Caserza B. 30/04/2004
Approved by: Name / Date
F. Terzi 30/04/2004
P. Gagliardi 30/04/2004
Checked at CERN by / Signature / Date