## Alarm sheet
File name | HCMB__A001-02000004_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 2000004
Date of test Ap 1 | 2001-12-20 |  |  |  | Date of test Ap 2 | 2001-12-20
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | yellow alarm | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | yellow alarm | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
20 | HCMB__A001 | 2000004 |  | 1 |  |  | 3
COMMENTS:
| HOLDING POINT | Ok | Data at CERN |  | 2001-12-20 | Analysis by |  | E. Todesco | Answer to MMS-LD |  |  |  | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
| Pre-stress (estimate, Mpa) | [6 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
| Inner average |  |  | Outer average |  |  |  | Seen from conection side
| Used | Nominal | Differ. | Used | Nominal | Differ.
| 75 | 75 | 0 | 75 | 75 | 0
| Peak pre-stress during collaring |  |  |  |  | 110
| Collared coil template - version 11.02.02 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans Assembly Resume et Original Data
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet

## Original data
File name |  | HCMB__A001-02000004_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 2000004 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Ansaldo | [9 empty] | File name for aperture 1 |  |  | ansaldo4cct1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | ansaldo4cct2.txt |  |  |  | a even | 1
Test Operator |  | Cid/Reynes |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2001-12-20 |  |  | Time of test |  |  | 8:00/12:00 |  |  |  | Date of test |  | 2001-12-20 |  |  | Time of test |  |  | 13:00/15:00
Mole name |  | Dipole 5 |  |  | Calibration date |  |  | 1999-07-20 |  |  |  | Mole name |  | Dipole 5 |  |  | Calibration date |  |  | 1999-07-20
Coil length |  | 0.7499 |  |  | Coil temperature |  |  | 19.9 |  |  |  | Coil length |  | 0.7499 |  |  | Coil temperature |  |  | 20
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 7.2414e-05 |  |  | Dy average |  |  | 0.00070729 |  |  |  | Dx average |  | 0.00013302 |  |  | Dy average |  |  | 0.0006997
Magnetic length |  | 14.4408 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0065 |  |  |  | Magnetic length |  | 14.4452 |  |  | Tranfs. Function (mT/KA) |  |  | 595.8693
Phase |  | -11.1922 |  |  | Integral field (T.m) |  |  | 0.073159 |  |  |  | Phase |  | -11.1167 |  |  | Integral field (T.m) |  |  | 0.073164
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003076 | 0.005069 | 0.005066 | 0.005065 | 0.005065 | 0.005065 | 0.005065 | 0.005065 | 0.005065 | 0.005066 | 0.005066 | 0.005066 | 0.005066 | 0.005066 | 0.005067 | 0.005066 | 0.005066 | 0.005066 | 0.005069 | 0.003292 | 0.073159 | C1 (T) | 0.003144 | 0.005069 | 0.005065 | 0.005064 | 0.005065 | 0.005064 | 0.005064 | 0.005063 | 0.005064 | 0.005065 | 0.005064 | 0.005064 | 0.005064 | 0.005065 | 0.005065 | 0.005065 | 0.005065 | 0.005065 | 0.005068 | 0.003252 | 0.073164
Angle (mrad) | 2.1232 | 1.3235 | 0.52553 | 0.41118 | 0.009324 | -0.009344 | 0.1915 | 0.049078 | 0.065431 | -0.084046 | 0.18153 | 0.19472 | 0.25371 | -0.43444 | -0.6988 | -0.62901 | -0.40738 | -0.36716 | -0.65022 | -1.7844 | -11.1922 | Angle (mrad) | 1.5847 | 0.1413 | 0.56245 | 0.45143 | 0.30726 | 0.49992 | 0.41241 | 0.08202 | 0.23428 | -0.086669 | -0.13579 | 0.10402 | -0.12007 | -0.25864 | -0.52835 | -0.52312 | -0.25603 | -0.54123 | -0.48565 | -1.1427 | -11.1167
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.6312 | -0.4897 | -0.22437 | 0.57412 | 0.65489 | 0.35449 | 0.0038643 | 0.23333 | -0.2897 | -0.1202 | 1.0351 | 0.69329 | 0.23484 | -0.2283 | 0.21416 | 0.14872 | -0.10125 | -0.003138 | -0.31469 | -0.17714 | 0.034379 | b2 | -1.7944 | 0.33468 | -0.20909 | 0.97011 | 0.9437 | 0.64888 | 0.46816 | 0.48281 | 0.36343 | 0.88469 | 1.531 | 0.31178 | 1.0461 | 0.43972 | 1.2388 | 1.8145 | 1.1079 | 1.094 | -0.24739 | -0.94798 | 0.59706
b3 | 26.2526 | 1.7597 | 1.8167 | 1.7139 | 1.2791 | 0.99871 | 1.2171 | 1.6988 | 0.8368 | 1.2151 | 1.5926 | 1.6026 | 1.6351 | 1.6436 | 1.9075 | 1.5698 | 1.714 | 1.3284 | 1.891 | -5.1391 | 2.0782 | b3 | 30.1868 | 1.4132 | 0.6209 | 0.61752 | 1.1935 | 0.6229 | 0.43703 | 0.60204 | 0.33783 | 0.75779 | 0.39446 | 0.40427 | 1.3112 | 1.3722 | 1.7065 | 1.0861 | 0.88564 | 1.0895 | 1.3903 | -7.4286 | 1.5683
b4 | -0.13344 | 0.015718 | 0.0097606 | 0.037067 | -0.023635 | -0.0096274 | -0.049132 | 0.071265 | -0.29281 | -0.06562 | 0.069218 | 0.012976 | -0.25314 | -0.29294 | -0.10908 | -0.11602 | -0.16327 | -0.05141 | -0.14143 | -0.123 | -0.078567 | b4 | -0.065603 | 0.21286 | 0.27219 | 0.31958 | 0.2815 | 0.099397 | -0.051529 | 0.19973 | 0.021522 | 0.1015 | 0.010081 | 0.071581 | -0.18704 | -0.083769 | 0.13246 | 0.27077 | 0.26059 | 0.19393 | 0.19145 | -0.15154 | 0.11312
b5 | -6.2355 | 0.71884 | 0.70506 | 0.75538 | 0.61862 | 0.91215 | 0.89521 | 1.1192 | 1.0239 | 0.74136 | 0.90768 | 0.86593 | 0.98985 | 0.93206 | 0.96591 | 0.7425 | 0.85489 | 0.89006 | 0.88336 | -3.5355 | 0.49013 | b5 | -4.7108 | 0.77443 | 0.93887 | 1.0025 | 0.70933 | 0.78707 | 0.9181 | 0.98942 | 0.82475 | 0.84379 | 0.97941 | 0.8467 | 0.83969 | 0.69709 | 0.78634 | 0.68975 | 0.89379 | 0.89284 | 0.88945 | -3.8945 | 0.51284
b6 | -0.1411 | 0.019796 | 0.02659 | 0.073412 | 0.056993 | 0.041887 | 0.030508 | 0.049173 | -0.011536 | 0.024093 | -0.020637 | -0.031035 | -0.056113 | -0.019567 | -0.019764 | -0.0078346 | -0.0014554 | -0.0017382 | 0.0062411 | -0.062755 | 0.0016894 | b6 | -0.13264 | 0.023124 | 0.050231 | 0.070686 | 0.045561 | 0.029857 | 0.0092111 | 0.035966 | -0.03243 | 0.0065196 | -0.046305 | -0.050689 | -0.087913 | -0.12134 | -0.028818 | 0.02204 | 0.0020423 | -0.0075183 | 0.034263 | -0.096446 | -0.00985
b7 | 1.7656 | 0.30479 | 0.35402 | 0.34359 | 0.31984 | 0.3109 | 0.32082 | 0.40137 | 0.31742 | 0.33086 | 0.32892 | 0.32618 | 0.32978 | 0.33721 | 0.35382 | 0.35761 | 0.31533 | 0.34086 | 0.38575 | -0.21865 | 0.36397 | b7 | 1.8464 | 0.33545 | 0.34885 | 0.27371 | 0.32557 | 0.34832 | 0.34649 | 0.3187 | 0.29963 | 0.28764 | 0.29086 | 0.2996 | 0.36549 | 0.39261 | 0.33876 | 0.34585 | 0.35324 | 0.34833 | 0.33667 | -0.31573 | 0.35816
b8 | -0.084719 | -0.00027725 | 0.005366 | 0.047614 | 0.016459 | -0.03439 | -0.0083954 | 0.010453 | 0.021397 | 0.031356 | 0.0051008 | 0.012031 | 0.033721 | 0.0038826 | 0.04327 | 0.0043391 | -0.012989 | 0.022256 | 0.028124 | -0.014336 | 0.0087549 | b8 | -0.079059 | -0.0098834 | -0.0042641 | 0.032603 | 0.022707 | -0.011525 | -0.044922 | -0.00057596 | -0.0009953 | 0.028387 | -0.020072 | -0.017468 | -0.0031972 | -0.0065073 | 0.048265 | 0.0057394 | 0.0065432 | 0.019303 | 0.03432 | 0.012998 | 0.0019612
b9 | 0.18967 | 0.30996 | 0.28897 | 0.30423 | 0.28048 | 0.29888 | 0.28288 | 0.29781 | 0.28375 | 0.28691 | 0.29391 | 0.28994 | 0.29804 | 0.30603 | 0.30883 | 0.27767 | 0.30203 | 0.30657 | 0.29182 | 0.19278 | 0.28816 | b9 | 0.17556 | 0.27841 | 0.26224 | 0.2793 | 0.26564 | 0.25702 | 0.26213 | 0.27492 | 0.26711 | 0.27161 | 0.27883 | 0.27328 | 0.27209 | 0.26354 | 0.27584 | 0.26599 | 0.27506 | 0.2773 | 0.27815 | 0.15316 | 0.26402
b10 | -0.0064539 | -0.020662 | 0.024672 | 0.07273 | -0.01161 | -0.10121 | -0.062339 | 0.020811 | -0.011455 | 0.054659 | -0.045328 | -0.026584 | -0.033002 | 0.00117 | 0.11526 | 0.015551 | -0.021123 | 0.047426 | 0.072554 | -0.10521 | 0.0010022 | b10 | 0.021761 | 0.0057125 | 0.038124 | 0.094707 | 0.063439 | 0.0026461 | -0.10212 | 0.014405 | -0.043137 | 0.05063 | -0.037742 | 0.0013848 | -0.0095036 | -0.010249 | 0.077542 | -0.080903 | -0.061278 | -0.009278 | 0.039829 | -0.028151 | 0.0015415
b11 | 0.59464 | 0.73019 | 0.72753 | 0.73031 | 0.72569 | 0.73694 | 0.7323 | 0.73702 | 0.73776 | 0.74201 | 0.74006 | 0.73874 | 0.7395 | 0.73895 | 0.7425 | 0.73388 | 0.7414 | 0.7385 | 0.72722 | 0.59735 | 0.72647 | b11 | 0.61055 | 0.73281 | 0.72696 | 0.73843 | 0.736 | 0.72607 | 0.72328 | 0.73022 | 0.73728 | 0.7407 | 0.73706 | 0.73321 | 0.7317 | 0.7288 | 0.73666 | 0.73232 | 0.73504 | 0.73747 | 0.72747 | 0.59145 | 0.7242
b12 | -0.0058762 | -0.0064052 | 0.0022226 | 0.012375 | -0.0027425 | -0.017496 | -0.0082462 | 0.0047741 | 0.0020478 | 0.010724 | -0.0049339 | -0.0012597 | 0.0011121 | 0.00082699 | 0.019531 | 0.0039487 | -0.0043147 | 0.012307 | 0.01405 | -0.0040272 | 0.0016796 | b12 | -0.0026107 | -0.00021084 | 0.0018624 | 0.013834 | 0.010509 | 0.00041406 | -0.015075 | 0.0022003 | -0.0047385 | 0.0083528 | -0.0027394 | 0.0019941 | 0.0069401 | 0.0024166 | 0.018896 | -0.0053539 | -0.0028962 | 0.0037872 | 0.0080053 | -0.00062255 | 0.0023973
b13 | 0.099209 | 0.086575 | 0.087987 | 0.088166 | 0.086428 | 0.083903 | 0.083613 | 0.085186 | 0.08521 | 0.087521 | 0.088557 | 0.087782 | 0.085879 | 0.086477 | 0.08698 | 0.091475 | 0.089032 | 0.087932 | 0.085573 | 0.052435 | 0.086129 | b13 | 0.095465 | 0.084402 | 0.08139 | 0.080517 | 0.082233 | 0.080941 | 0.08069 | 0.080404 | 0.081861 | 0.082056 | 0.082435 | 0.080569 | 0.085293 | 0.085543 | 0.085613 | 0.0879 | 0.085655 | 0.084824 | 0.083541 | 0.041761 | 0.082124
b14 | 0.0073857 | -0.0061117 | -0.0016869 | 0.0019612 | -0.004828 | -0.012413 | -0.009707 | -0.0025157 | -0.0031315 | 0.00039328 | -0.0084146 | -0.0059127 | -0.0038189 | -0.00037501 | 0.0077766 | -0.0037722 | -0.0070849 | -0.00021499 | 0.0027005 | -0.0057589 | -0.0029298 | b14 | 0.011967 | -0.0060694 | -0.0025584 | 0.002231 | -0.000366 | -0.0043387 | -0.014947 | -0.0036841 | -0.0092658 | -0.002223 | -0.010508 | -0.0062353 | -0.0050076 | -0.0057378 | -0.00014436 | -0.012683 | -0.011226 | -0.0071644 | -0.0034745 | -0.0097938 | -0.0053101
b15 | 0.0074209 | 0.043394 | 0.040002 | 0.042428 | 0.045227 | 0.041276 | 0.045847 | 0.045014 | 0.04562 | 0.048574 | 0.046163 | 0.046906 | 0.044357 | 0.04286 | 0.047467 | 0.045501 | 0.046199 | 0.044615 | 0.043077 | 0.02126 | 0.04273 | b15 | 0.0069528 | 0.043135 | 0.044498 | 0.046399 | 0.046899 | 0.045507 | 0.042192 | 0.044279 | 0.048514 | 0.048789 | 0.044278 | 0.044929 | 0.040375 | 0.040016 | 0.045147 | 0.04201 | 0.044497 | 0.042789 | 0.041616 | 0.019989 | 0.042206
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 21.2316 | 13.2346 | 5.2553 | 4.1118 | 0.093237 | -0.093438 | 1.915 | 0.49078 | 0.65431 | -0.84046 | 1.8153 | 1.9472 | 2.537 | -4.3444 | -6.988 | -6.2901 | -4.0738 | -3.6716 | -6.5022 | -17.8439 | 0 | a1 | 15.8466 | 1.413 | 5.6245 | 4.5143 | 3.0726 | 4.9992 | 4.1242 | 0.8202 | 2.3428 | -0.86669 | -1.3579 | 1.0402 | -1.2007 | -2.5864 | -5.2835 | -5.2312 | -2.5603 | -5.4123 | -4.8565 | -11.4266 | 1.1855e-13
a2 | 17.069 | 0.089642 | 0.26508 | 0.47823 | 0.66362 | 0.037446 | 0.37481 | 0.49206 | -0.16025 | 0.97355 | 0.67814 | 0.90395 | 0.049529 | -0.65831 | -0.64556 | 0.14203 | 0.70067 | 1.3043 | 1.0387 | -15.8693 | 0.35192 | a2 | 9.7344 | 1.0833 | 1.2957 | 0.80761 | 1.062 | 0.85345 | 0.81289 | 0.53979 | 0.68648 | 0.37193 | 1.143 | 1.543 | 1.4681 | 1.2408 | 1.2283 | 1.2884 | 1.0428 | 1.0217 | 0.78302 | 0.80693 | 1.2891
a3 | -5.1224 | -0.63977 | -0.5763 | -0.5818 | -0.57892 | -0.7757 | -0.77195 | -0.69907 | -0.50819 | -0.81865 | -0.58198 | -0.37708 | -0.49701 | -0.17922 | -0.38008 | -0.42407 | -0.13329 | -0.35179 | -0.11756 | -0.091555 | -0.63155 | a3 | -4.2765 | -1.1248 | -1.2927 | -1.1861 | -1.0584 | -0.94888 | -1.2382 | -0.9709 | -0.51904 | -0.92767 | -0.89767 | -1.0236 | -0.53924 | -0.77017 | -0.77075 | -0.64038 | -1.0327 | -1.1299 | -1.38 | -0.95375 | -1.0756
a4 | 4.9526 | 0.16357 | 0.19674 | 0.18606 | 0.34115 | 0.36937 | 0.65323 | 0.54881 | 0.75427 | 0.69293 | 0.41016 | 0.46459 | 0.41685 | 0.33387 | 0.39849 | 0.38564 | 0.74174 | 0.66231 | 0.62506 | -0.20639 | 0.58251 | a4 | 3.2822 | 0.17008 | 0.34491 | 0.21206 | -0.083098 | -0.067548 | 0.12015 | -0.10823 | 0.3187 | 0.24961 | 0.39713 | -0.064492 | 0.020543 | 0.28109 | 0.14325 | -0.086349 | 0.22942 | 0.29218 | 0.056651 | 0.57739 | 0.25095
a5 | 1.7696 | -0.17379 | -0.1319 | -0.1379 | -0.16927 | -0.2257 | -0.1087 | -0.15619 | 0.050918 | -0.10769 | -0.022298 | -0.018952 | -0.047569 | 0.098347 | -0.0726 | -0.022711 | 0.0044333 | -0.12641 | -0.033021 | 0.0079042 | -0.016702 | a5 | 1.7171 | -0.19748 | -0.25143 | -0.30072 | -0.12322 | -0.25045 | -0.27643 | -0.20181 | -0.046462 | -0.15844 | -0.25153 | -0.21006 | -0.017776 | 0.057276 | -0.015156 | -0.1584 | -0.30071 | -0.33396 | -0.33778 | -0.34122 | -0.13125
a6 | 0.39057 | -0.12167 | -0.11864 | -0.078782 | 0.0049019 | -0.029445 | -0.0013605 | 0.091324 | -0.13224 | 0.087316 | 0.067016 | 0.082169 | -0.0047203 | -0.027943 | -0.10144 | -0.073791 | -0.030497 | 0.074729 | -0.0073522 | -0.21756 | -0.011672 | a6 | 0.31886 | -0.040468 | -0.029468 | -0.18413 | -0.01242 | 0.010227 | -0.0066191 | -0.02761 | 0.055687 | -0.022594 | 0.031723 | 0.036425 | 0.024399 | 0.040742 | -0.029289 | -0.00012408 | 0.0024431 | -0.033994 | -0.11015 | -0.19359 | -0.011509
a7 | 1.4345 | -0.027002 | -0.070219 | -0.064367 | -0.03534 | 0.007538 | -0.001376 | -0.0095593 | 0.044407 | -0.033884 | 0.0034484 | 0.05007 | 0.083916 | 0.094768 | 0.091113 | 0.04518 | 0.02942 | -0.039256 | -0.048749 | 0.12518 | 0.055679 | a7 | 1.3953 | -0.034498 | -0.075438 | -0.031297 | -0.022926 | -0.0058649 | -0.026072 | -0.05584 | 0.04952 | -0.019606 | -0.072777 | -0.046798 | 0.032042 | 0.045684 | 0.058157 | 0.014071 | -0.050895 | -0.075326 | -0.038161 | 0.067674 | 0.028718
a8 | -0.010101 | 0.030238 | 0.034534 | 0.027653 | 0.0018184 | 0.014988 | 0.020935 | 0.057934 | 0.034757 | 0.053599 | 0.060023 | 0.03001 | 0.054758 | 0.066249 | 0.026697 | 0.034952 | 0.075391 | 0.088298 | 0.068801 | 0.029372 | 0.041264 | a8 | 0.055015 | 0.00070606 | 0.036914 | 0.046495 | -0.029724 | 0.016472 | -0.0022676 | -0.024788 | 0.0018609 | 0.016561 | 0.024161 | 0.015741 | -0.0029813 | 0.022788 | 0.034451 | 0.014035 | 0.036871 | 0.062616 | 0.036103 | 0.01029 | 0.018003
a9 | -0.30309 | 0.015823 | 0.01893 | 0.022111 | 0.010963 | -0.0089021 | -0.012113 | 0.0153 | 0.0099535 | -0.0035533 | 0.0025667 | 0.0058878 | 0.012937 | 0.019602 | 0.0030033 | 0.0025815 | 0.016077 | 0.020767 | 0.028135 | -0.0063507 | -0.00042552 | a9 | -0.26285 | 0.027581 | 0.049126 | 0.019101 | 0.03 | 0.025614 | 0.027175 | 0.048988 | 0.053301 | 0.030689 | 0.043715 | 0.038085 | 0.041725 | 0.030291 | -0.001445 | -0.025587 | -0.011516 | -0.020673 | -0.020004 | -0.028946 | 0.010597
a10 | -0.10034 | 0.016719 | 0.059322 | 0.024168 | 0.003148 | 0.01594 | -0.015806 | 0.021127 | 0.0070022 | -0.019181 | -0.0011249 | -0.01886 | 0.027181 | 0.028401 | -0.04013 | -0.017738 | -0.0092511 | 0.036377 | 0.016249 | -0.026947 | 0.0028619 | a10 | -0.09282 | 0.0076454 | 0.012489 | 0.0071501 | -0.018043 | -0.0082285 | -0.011952 | 0.0054282 | -0.027812 | -0.022849 | 0.0071304 | 0.0072965 | 0.043886 | 0.03729 | -0.00087818 | 0.014502 | 0.009689 | 0.036424 | 0.036788 | -0.046061 | 0.0025333
a11 | 0.13761 | -0.028753 | -0.039201 | -0.041791 | -0.045247 | -0.038426 | -0.030729 | -0.029273 | -0.022292 | -0.032942 | -0.029425 | -0.024724 | -0.020883 | -0.024692 | -0.030633 | -0.033438 | -0.030047 | -0.034506 | -0.039962 | -0.025823 | -0.026517 | a11 | 0.14861 | -0.034859 | -0.037719 | -0.03842 | -0.037835 | -0.033414 | -0.036461 | -0.037663 | -0.029283 | -0.0398 | -0.048071 | -0.036694 | -0.029596 | -0.023109 | -0.027317 | -0.039143 | -0.043775 | -0.045108 | -0.043196 | -0.011579 | -0.029982
a12 | -0.019442 | -0.0054737 | 0.0022312 | -0.0013311 | 8.3461e-05 | -0.0040369 | -0.0054807 | 0.0013573 | -0.0068422 | -0.0035078 | 0.0011026 | 0.00024268 | 0.0024929 | -0.0017582 | -0.012445 | -0.007331 | -0.0049436 | 0.0038167 | 6.3329e-05 | -0.013335 | -0.0032317 | a12 | -0.026154 | -0.0021347 | -0.0046988 | -0.012353 | -0.0076413 | -0.0053606 | -0.0035496 | -0.0043386 | -0.0045263 | -0.0045666 | -0.0027053 | -0.0005676 | 0.0042362 | 0.0024607 | -0.0049185 | 0.0011007 | 0.00018324 | 6.6703e-05 | -0.0041503 | -0.013295 | -0.0040615
a13 | -0.037347 | -0.0024095 | -0.0061936 | -0.0077519 | -0.0079182 | -0.0062733 | -0.0038737 | -0.0041061 | -0.0031078 | -0.0075468 | -5.0066e-05 | -0.00067563 | -0.0040507 | -0.0021375 | -0.0076773 | -0.0069051 | -0.0051756 | -0.0070641 | -0.0068752 | -0.0080345 | -0.0061146 | a13 | -0.034299 | -0.0033568 | -0.0016282 | -0.006384 | -0.0054117 | -0.004579 | -0.0026109 | -0.0026062 | 0.00012209 | -0.0040913 | -0.0013352 | -0.00070756 | -0.0036015 | -0.0026017 | -0.0089111 | -0.0059008 | -0.006284 | -0.0073261 | -0.010832 | -0.0061011 | -0.0053665
a14 | 0.004714 | -0.016298 | -0.013433 | -0.015932 | -0.019336 | -0.014711 | -0.01812 | -0.015821 | -0.016253 | -0.019603 | -0.017773 | -0.018789 | -0.014293 | -0.013375 | -0.019641 | -0.01902 | -0.013683 | -0.012508 | -0.015205 | -0.0019136 | -0.015172 | a14 | 0.0045441 | -0.018238 | -0.018255 | -0.018535 | -0.023778 | -0.021443 | -0.019159 | -0.020897 | -0.023398 | -0.02134 | -0.019022 | -0.021311 | -0.01653 | -0.016435 | -0.01884 | -0.018797 | -0.017669 | -0.014519 | -0.016569 | -0.0032182 | -0.017857
a15 | 0.00022283 | -0.0062637 | -0.0032374 | 0.0011017 | -0.0067795 | -0.014243 | -0.011935 | -0.0045836 | -0.0056735 | -0.0020237 | -0.0098363 | -0.00779 | -0.0064958 | -0.0049603 | 0.0047303 | -0.0053239 | -0.0066955 | -0.0012158 | 0.0003316 | -0.013672 | -0.0051759 | a15 | 0.00021518 | -0.0088799 | -0.0062917 | -0.001183 | -0.0031282 | -0.0075776 | -0.018543 | -0.0068167 | -0.011599 | -0.0058283 | -0.012147 | -0.0081735 | -0.0077292 | -0.0078827 | -0.0017616 | -0.014203 | -0.013647 | -0.0093425 | -0.0069321 | -0.012594 | -0.0082899
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 8.0527e-05 | 4.956e-05 | -5.0036e-05 | -0.00016554 | 2.7551e-05 | 0.00023476 | 0.00014293 | -4.5995e-05 | 2.686e-05 | -0.00012693 | 0.00010386 | 5.9655e-05 | 7.7568e-05 | -5.0786e-07 | -0.00026722 | -3.7818e-05 | 4.7497e-05 | -0.00010503 | -0.00016702 | 0.00029554 | 7.2414e-05 | Dx (m) | 2.1862e-06 | -1.238e-05 | -8.7402e-05 | -0.00021659 | -0.00014828 | -7.0671e-06 | 0.000238 | -3.2797e-05 | 9.6764e-05 | -0.00011868 | 8.775e-05 | -2.3582e-06 | 2.6162e-05 | 2.6637e-05 | -0.00017878 | 0.00018907 | 0.00014255 | 2.6424e-05 | -8.7663e-05 | 7.8293e-05 | 0.00013302
Dy (m) | 0.00026824 | -3.6972e-05 | -0.00014131 | -6.573e-05 | -5.6568e-06 | -2.4531e-05 | 4.2691e-05 | -5.0558e-05 | -1.5323e-05 | 3.8309e-05 | 6.7133e-06 | 4.5398e-05 | -6.0294e-05 | -6.5355e-05 | 8.0855e-05 | 3.9367e-05 | 2.3137e-05 | -8.8645e-05 | -4.7162e-05 | 8.9463e-05 | 0.00070729 | Dy (m) | 0.00025792 | -1.8325e-05 | -3.3741e-05 | -2.7729e-05 | 3.4053e-05 | 1.8941e-05 | 4.0089e-05 | -1.4329e-05 | 6.7971e-05 | 4.6065e-05 | -1.0723e-05 | -1.7036e-05 | -0.0001009 | -8.6137e-05 | -4.6028e-06 | -2.3558e-05 | -1.3919e-05 | -8.2348e-05 | -9.1174e-05 | 0.00013393 | 0.0006997
Magnetic length (m) |  | 14.4408 | [19 empty] | Magnetic length (m) |  | 14.4452
Transfer funct. (mT/KA) |  | 596.0065 | [19 empty] | Transfer funct. (mT/KA) |  | 595.8693
Coil positioning (mm) |  | 0.01556 | [19 empty] | Coil positioning (mm) |  | 0.016748
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMB__A001-2000004 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-2000004 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 361.8824 | 596.3529 | 596 | 595.8824 | 595.8824 | 595.8824 | 595.8824 | 595.8824 | 595.8824 | 596 | 596 | 596 | 596 | 596 | 596.1176 | 596 | 596 | 596 | 596.3529 | 387.2941 | 0.073159 | C1/i (mT/KA) | 369.8824 | 596.3529 | 595.8824 | 595.7647 | 595.8824 | 595.7647 | 595.7647 | 595.6471 | 595.7647 | 595.8824 | 595.7647 | 595.7647 | 595.7647 | 595.8824 | 595.8824 | 595.8824 | 595.8824 | 595.8824 | 596.2353 | 382.5882 | 0.073164
Angle (mrad) | 2.1232 | 1.3235 | 0.52553 | 0.41118 | 0.009324 | -0.009344 | 0.1915 | 0.049078 | 0.065431 | -0.084046 | 0.18153 | 0.19472 | 0.25371 | -0.43444 | -0.6988 | -0.62901 | -0.40738 | -0.36716 | -0.65022 | -1.7844 | -11.1922 | Angle (mrad) | 1.5847 | 0.1413 | 0.56245 | 0.45143 | 0.30726 | 0.49992 | 0.41241 | 0.08202 | 0.23428 | -0.086669 | -0.13579 | 0.10402 | -0.12007 | -0.25864 | -0.52835 | -0.52312 | -0.25603 | -0.54123 | -0.48565 | -1.1427 | -11.1167
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.6312 | -0.4897 | -0.22437 | 0.57412 | 0.65489 | 0.35449 | 0.0038643 | 0.23333 | -0.2897 | -0.1202 | 1.0351 | 0.69329 | 0.23484 | -0.2283 | 0.21416 | 0.14872 | -0.10125 | -0.003138 | -0.31469 | -0.17714 | 0.034379 | b2 | -1.7944 | 0.33468 | -0.20909 | 0.97011 | 0.9437 | 0.64888 | 0.46816 | 0.48281 | 0.36343 | 0.88469 | 1.531 | 0.31178 | 1.0461 | 0.43972 | 1.2388 | 1.8145 | 1.1079 | 1.094 | -0.24739 | -0.94798 | 0.59706
b3 | 26.2526 | 1.7597 | 1.8167 | 1.7139 | 1.2791 | 0.99871 | 1.2171 | 1.6988 | 0.8368 | 1.2151 | 1.5926 | 1.6026 | 1.6351 | 1.6436 | 1.9075 | 1.5698 | 1.714 | 1.3284 | 1.891 | -5.1391 | 2.0782 | b3 | 30.1868 | 1.4132 | 0.6209 | 0.61752 | 1.1935 | 0.6229 | 0.43703 | 0.60204 | 0.33783 | 0.75779 | 0.39446 | 0.40427 | 1.3112 | 1.3722 | 1.7065 | 1.0861 | 0.88564 | 1.0895 | 1.3903 | -7.4286 | 1.5683
b4 | -0.13344 | 0.015718 | 0.0097606 | 0.037067 | -0.023635 | -0.0096274 | -0.049132 | 0.071265 | -0.29281 | -0.06562 | 0.069218 | 0.012976 | -0.25314 | -0.29294 | -0.10908 | -0.11602 | -0.16327 | -0.05141 | -0.14143 | -0.123 | -0.078567 | b4 | -0.065603 | 0.21286 | 0.27219 | 0.31958 | 0.2815 | 0.099397 | -0.051529 | 0.19973 | 0.021522 | 0.1015 | 0.010081 | 0.071581 | -0.18704 | -0.083769 | 0.13246 | 0.27077 | 0.26059 | 0.19393 | 0.19145 | -0.15154 | 0.11312
b5 | -6.2355 | 0.71884 | 0.70506 | 0.75538 | 0.61862 | 0.91215 | 0.89521 | 1.1192 | 1.0239 | 0.74136 | 0.90768 | 0.86593 | 0.98985 | 0.93206 | 0.96591 | 0.7425 | 0.85489 | 0.89006 | 0.88336 | -3.5355 | 0.49013 | b5 | -4.7108 | 0.77443 | 0.93887 | 1.0025 | 0.70933 | 0.78707 | 0.9181 | 0.98942 | 0.82475 | 0.84379 | 0.97941 | 0.8467 | 0.83969 | 0.69709 | 0.78634 | 0.68975 | 0.89379 | 0.89284 | 0.88945 | -3.8945 | 0.51284
b6 | -0.1411 | 0.019796 | 0.02659 | 0.073412 | 0.056993 | 0.041887 | 0.030508 | 0.049173 | -0.011536 | 0.024093 | -0.020637 | -0.031035 | -0.056113 | -0.019567 | -0.019764 | -0.0078346 | -0.0014554 | -0.0017382 | 0.0062411 | -0.062755 | 0.0016894 | b6 | -0.13264 | 0.023124 | 0.050231 | 0.070686 | 0.045561 | 0.029857 | 0.0092111 | 0.035966 | -0.03243 | 0.0065196 | -0.046305 | -0.050689 | -0.087913 | -0.12134 | -0.028818 | 0.02204 | 0.0020423 | -0.0075183 | 0.034263 | -0.096446 | -0.00985
b7 | 1.7656 | 0.30479 | 0.35402 | 0.34359 | 0.31984 | 0.3109 | 0.32082 | 0.40137 | 0.31742 | 0.33086 | 0.32892 | 0.32618 | 0.32978 | 0.33721 | 0.35382 | 0.35761 | 0.31533 | 0.34086 | 0.38575 | -0.21865 | 0.36397 | b7 | 1.8464 | 0.33545 | 0.34885 | 0.27371 | 0.32557 | 0.34832 | 0.34649 | 0.3187 | 0.29963 | 0.28764 | 0.29086 | 0.2996 | 0.36549 | 0.39261 | 0.33876 | 0.34585 | 0.35324 | 0.34833 | 0.33667 | -0.31573 | 0.35816
b8 | -0.084719 | -0.00027725 | 0.005366 | 0.047614 | 0.016459 | -0.03439 | -0.0083954 | 0.010453 | 0.021397 | 0.031356 | 0.0051008 | 0.012031 | 0.033721 | 0.0038826 | 0.04327 | 0.0043391 | -0.012989 | 0.022256 | 0.028124 | -0.014336 | 0.0087549 | b8 | -0.079059 | -0.0098834 | -0.0042641 | 0.032603 | 0.022707 | -0.011525 | -0.044922 | -0.00057596 | -0.0009953 | 0.028387 | -0.020072 | -0.017468 | -0.0031972 | -0.0065073 | 0.048265 | 0.0057394 | 0.0065432 | 0.019303 | 0.03432 | 0.012998 | 0.0019612
b9 | 0.18967 | 0.30996 | 0.28897 | 0.30423 | 0.28048 | 0.29888 | 0.28288 | 0.29781 | 0.28375 | 0.28691 | 0.29391 | 0.28994 | 0.29804 | 0.30603 | 0.30883 | 0.27767 | 0.30203 | 0.30657 | 0.29182 | 0.19278 | 0.28816 | b9 | 0.17556 | 0.27841 | 0.26224 | 0.2793 | 0.26564 | 0.25702 | 0.26213 | 0.27492 | 0.26711 | 0.27161 | 0.27883 | 0.27328 | 0.27209 | 0.26354 | 0.27584 | 0.26599 | 0.27506 | 0.2773 | 0.27815 | 0.15316 | 0.26402
b10 | -0.0064539 | -0.020662 | 0.024672 | 0.07273 | -0.01161 | -0.10121 | -0.062339 | 0.020811 | -0.011455 | 0.054659 | -0.045328 | -0.026584 | -0.033002 | 0.00117 | 0.11526 | 0.015551 | -0.021123 | 0.047426 | 0.072554 | -0.10521 | 0.0010022 | b10 | 0.021761 | 0.0057125 | 0.038124 | 0.094707 | 0.063439 | 0.0026461 | -0.10212 | 0.014405 | -0.043137 | 0.05063 | -0.037742 | 0.0013848 | -0.0095036 | -0.010249 | 0.077542 | -0.080903 | -0.061278 | -0.009278 | 0.039829 | -0.028151 | 0.0015415
b11 | 0.59464 | 0.73019 | 0.72753 | 0.73031 | 0.72569 | 0.73694 | 0.7323 | 0.73702 | 0.73776 | 0.74201 | 0.74006 | 0.73874 | 0.7395 | 0.73895 | 0.7425 | 0.73388 | 0.7414 | 0.7385 | 0.72722 | 0.59735 | 0.72647 | b11 | 0.61055 | 0.73281 | 0.72696 | 0.73843 | 0.736 | 0.72607 | 0.72328 | 0.73022 | 0.73728 | 0.7407 | 0.73706 | 0.73321 | 0.7317 | 0.7288 | 0.73666 | 0.73232 | 0.73504 | 0.73747 | 0.72747 | 0.59145 | 0.7242
b12 | -0.0058762 | -0.0064052 | 0.0022226 | 0.012375 | -0.0027425 | -0.017496 | -0.0082462 | 0.0047741 | 0.0020478 | 0.010724 | -0.0049339 | -0.0012597 | 0.0011121 | 0.00082699 | 0.019531 | 0.0039487 | -0.0043147 | 0.012307 | 0.01405 | -0.0040272 | 0.0016796 | b12 | -0.0026107 | -0.00021084 | 0.0018624 | 0.013834 | 0.010509 | 0.00041406 | -0.015075 | 0.0022003 | -0.0047385 | 0.0083528 | -0.0027394 | 0.0019941 | 0.0069401 | 0.0024166 | 0.018896 | -0.0053539 | -0.0028962 | 0.0037872 | 0.0080053 | -0.00062255 | 0.0023973
b13 | 0.099209 | 0.086575 | 0.087987 | 0.088166 | 0.086428 | 0.083903 | 0.083613 | 0.085186 | 0.08521 | 0.087521 | 0.088557 | 0.087782 | 0.085879 | 0.086477 | 0.08698 | 0.091475 | 0.089032 | 0.087932 | 0.085573 | 0.052435 | 0.086129 | b13 | 0.095465 | 0.084402 | 0.08139 | 0.080517 | 0.082233 | 0.080941 | 0.08069 | 0.080404 | 0.081861 | 0.082056 | 0.082435 | 0.080569 | 0.085293 | 0.085543 | 0.085613 | 0.0879 | 0.085655 | 0.084824 | 0.083541 | 0.041761 | 0.082124
b14 | 0.0073857 | -0.0061117 | -0.0016869 | 0.0019612 | -0.004828 | -0.012413 | -0.009707 | -0.0025157 | -0.0031315 | 0.00039328 | -0.0084146 | -0.0059127 | -0.0038189 | -0.00037501 | 0.0077766 | -0.0037722 | -0.0070849 | -0.00021499 | 0.0027005 | -0.0057589 | -0.0029298 | b14 | 0.011967 | -0.0060694 | -0.0025584 | 0.002231 | -0.000366 | -0.0043387 | -0.014947 | -0.0036841 | -0.0092658 | -0.002223 | -0.010508 | -0.0062353 | -0.0050076 | -0.0057378 | -0.00014436 | -0.012683 | -0.011226 | -0.0071644 | -0.0034745 | -0.0097938 | -0.0053101
b15 | 0.0074209 | 0.043394 | 0.040002 | 0.042428 | 0.045227 | 0.041276 | 0.045847 | 0.045014 | 0.04562 | 0.048574 | 0.046163 | 0.046906 | 0.044357 | 0.04286 | 0.047467 | 0.045501 | 0.046199 | 0.044615 | 0.043077 | 0.02126 | 0.04273 | b15 | 0.0069528 | 0.043135 | 0.044498 | 0.046399 | 0.046899 | 0.045507 | 0.042192 | 0.044279 | 0.048514 | 0.048789 | 0.044278 | 0.044929 | 0.040375 | 0.040016 | 0.045147 | 0.04201 | 0.044497 | 0.042789 | 0.041616 | 0.019989 | 0.042206
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 21.2316 | 13.2346 | 5.2553 | 4.1118 | 0.093237 | -0.093438 | 1.915 | 0.49078 | 0.65431 | -0.84046 | 1.8153 | 1.9472 | 2.537 | -4.3444 | -6.988 | -6.2901 | -4.0738 | -3.6716 | -6.5022 | -17.8439 | 0 | a1 | 15.8466 | 1.413 | 5.6245 | 4.5143 | 3.0726 | 4.9992 | 4.1242 | 0.8202 | 2.3428 | -0.86669 | -1.3579 | 1.0402 | -1.2007 | -2.5864 | -5.2835 | -5.2312 | -2.5603 | -5.4123 | -4.8565 | -11.4266 | 1.1855e-13
a2 | 17.069 | 0.089642 | 0.26508 | 0.47823 | 0.66362 | 0.037446 | 0.37481 | 0.49206 | -0.16025 | 0.97355 | 0.67814 | 0.90395 | 0.049529 | -0.65831 | -0.64556 | 0.14203 | 0.70067 | 1.3043 | 1.0387 | -15.8693 | 0.35192 | a2 | 9.7344 | 1.0833 | 1.2957 | 0.80761 | 1.062 | 0.85345 | 0.81289 | 0.53979 | 0.68648 | 0.37193 | 1.143 | 1.543 | 1.4681 | 1.2408 | 1.2283 | 1.2884 | 1.0428 | 1.0217 | 0.78302 | 0.80693 | 1.2891
a3 | -5.1224 | -0.63977 | -0.5763 | -0.5818 | -0.57892 | -0.7757 | -0.77195 | -0.69907 | -0.50819 | -0.81865 | -0.58198 | -0.37708 | -0.49701 | -0.17922 | -0.38008 | -0.42407 | -0.13329 | -0.35179 | -0.11756 | -0.091555 | -0.63155 | a3 | -4.2765 | -1.1248 | -1.2927 | -1.1861 | -1.0584 | -0.94888 | -1.2382 | -0.9709 | -0.51904 | -0.92767 | -0.89767 | -1.0236 | -0.53924 | -0.77017 | -0.77075 | -0.64038 | -1.0327 | -1.1299 | -1.38 | -0.95375 | -1.0756
a4 | 4.9526 | 0.16357 | 0.19674 | 0.18606 | 0.34115 | 0.36937 | 0.65323 | 0.54881 | 0.75427 | 0.69293 | 0.41016 | 0.46459 | 0.41685 | 0.33387 | 0.39849 | 0.38564 | 0.74174 | 0.66231 | 0.62506 | -0.20639 | 0.58251 | a4 | 3.2822 | 0.17008 | 0.34491 | 0.21206 | -0.083098 | -0.067548 | 0.12015 | -0.10823 | 0.3187 | 0.24961 | 0.39713 | -0.064492 | 0.020543 | 0.28109 | 0.14325 | -0.086349 | 0.22942 | 0.29218 | 0.056651 | 0.57739 | 0.25095
a5 | 1.7696 | -0.17379 | -0.1319 | -0.1379 | -0.16927 | -0.2257 | -0.1087 | -0.15619 | 0.050918 | -0.10769 | -0.022298 | -0.018952 | -0.047569 | 0.098347 | -0.0726 | -0.022711 | 0.0044333 | -0.12641 | -0.033021 | 0.0079042 | -0.016702 | a5 | 1.7171 | -0.19748 | -0.25143 | -0.30072 | -0.12322 | -0.25045 | -0.27643 | -0.20181 | -0.046462 | -0.15844 | -0.25153 | -0.21006 | -0.017776 | 0.057276 | -0.015156 | -0.1584 | -0.30071 | -0.33396 | -0.33778 | -0.34122 | -0.13125
a6 | 0.39057 | -0.12167 | -0.11864 | -0.078782 | 0.0049019 | -0.029445 | -0.0013605 | 0.091324 | -0.13224 | 0.087316 | 0.067016 | 0.082169 | -0.0047203 | -0.027943 | -0.10144 | -0.073791 | -0.030497 | 0.074729 | -0.0073522 | -0.21756 | -0.011672 | a6 | 0.31886 | -0.040468 | -0.029468 | -0.18413 | -0.01242 | 0.010227 | -0.0066191 | -0.02761 | 0.055687 | -0.022594 | 0.031723 | 0.036425 | 0.024399 | 0.040742 | -0.029289 | -0.00012408 | 0.0024431 | -0.033994 | -0.11015 | -0.19359 | -0.011509
a7 | 1.4345 | -0.027002 | -0.070219 | -0.064367 | -0.03534 | 0.007538 | -0.001376 | -0.0095593 | 0.044407 | -0.033884 | 0.0034484 | 0.05007 | 0.083916 | 0.094768 | 0.091113 | 0.04518 | 0.02942 | -0.039256 | -0.048749 | 0.12518 | 0.055679 | a7 | 1.3953 | -0.034498 | -0.075438 | -0.031297 | -0.022926 | -0.0058649 | -0.026072 | -0.05584 | 0.04952 | -0.019606 | -0.072777 | -0.046798 | 0.032042 | 0.045684 | 0.058157 | 0.014071 | -0.050895 | -0.075326 | -0.038161 | 0.067674 | 0.028718
a8 | -0.010101 | 0.030238 | 0.034534 | 0.027653 | 0.0018184 | 0.014988 | 0.020935 | 0.057934 | 0.034757 | 0.053599 | 0.060023 | 0.03001 | 0.054758 | 0.066249 | 0.026697 | 0.034952 | 0.075391 | 0.088298 | 0.068801 | 0.029372 | 0.041264 | a8 | 0.055015 | 0.00070606 | 0.036914 | 0.046495 | -0.029724 | 0.016472 | -0.0022676 | -0.024788 | 0.0018609 | 0.016561 | 0.024161 | 0.015741 | -0.0029813 | 0.022788 | 0.034451 | 0.014035 | 0.036871 | 0.062616 | 0.036103 | 0.01029 | 0.018003
a9 | -0.30309 | 0.015823 | 0.01893 | 0.022111 | 0.010963 | -0.0089021 | -0.012113 | 0.0153 | 0.0099535 | -0.0035533 | 0.0025667 | 0.0058878 | 0.012937 | 0.019602 | 0.0030033 | 0.0025815 | 0.016077 | 0.020767 | 0.028135 | -0.0063507 | -0.00042552 | a9 | -0.26285 | 0.027581 | 0.049126 | 0.019101 | 0.03 | 0.025614 | 0.027175 | 0.048988 | 0.053301 | 0.030689 | 0.043715 | 0.038085 | 0.041725 | 0.030291 | -0.001445 | -0.025587 | -0.011516 | -0.020673 | -0.020004 | -0.028946 | 0.010597
a10 | -0.10034 | 0.016719 | 0.059322 | 0.024168 | 0.003148 | 0.01594 | -0.015806 | 0.021127 | 0.0070022 | -0.019181 | -0.0011249 | -0.01886 | 0.027181 | 0.028401 | -0.04013 | -0.017738 | -0.0092511 | 0.036377 | 0.016249 | -0.026947 | 0.0028619 | a10 | -0.09282 | 0.0076454 | 0.012489 | 0.0071501 | -0.018043 | -0.0082285 | -0.011952 | 0.0054282 | -0.027812 | -0.022849 | 0.0071304 | 0.0072965 | 0.043886 | 0.03729 | -0.00087818 | 0.014502 | 0.009689 | 0.036424 | 0.036788 | -0.046061 | 0.0025333
a11 | 0.13761 | -0.028753 | -0.039201 | -0.041791 | -0.045247 | -0.038426 | -0.030729 | -0.029273 | -0.022292 | -0.032942 | -0.029425 | -0.024724 | -0.020883 | -0.024692 | -0.030633 | -0.033438 | -0.030047 | -0.034506 | -0.039962 | -0.025823 | -0.026517 | a11 | 0.14861 | -0.034859 | -0.037719 | -0.03842 | -0.037835 | -0.033414 | -0.036461 | -0.037663 | -0.029283 | -0.0398 | -0.048071 | -0.036694 | -0.029596 | -0.023109 | -0.027317 | -0.039143 | -0.043775 | -0.045108 | -0.043196 | -0.011579 | -0.029982
a12 | -0.019442 | -0.0054737 | 0.0022312 | -0.0013311 | 8.3461e-05 | -0.0040369 | -0.0054807 | 0.0013573 | -0.0068422 | -0.0035078 | 0.0011026 | 0.00024268 | 0.0024929 | -0.0017582 | -0.012445 | -0.007331 | -0.0049436 | 0.0038167 | 6.3329e-05 | -0.013335 | -0.0032317 | a12 | -0.026154 | -0.0021347 | -0.0046988 | -0.012353 | -0.0076413 | -0.0053606 | -0.0035496 | -0.0043386 | -0.0045263 | -0.0045666 | -0.0027053 | -0.0005676 | 0.0042362 | 0.0024607 | -0.0049185 | 0.0011007 | 0.00018324 | 6.6703e-05 | -0.0041503 | -0.013295 | -0.0040615
a13 | -0.037347 | -0.0024095 | -0.0061936 | -0.0077519 | -0.0079182 | -0.0062733 | -0.0038737 | -0.0041061 | -0.0031078 | -0.0075468 | -5.0066e-05 | -0.00067563 | -0.0040507 | -0.0021375 | -0.0076773 | -0.0069051 | -0.0051756 | -0.0070641 | -0.0068752 | -0.0080345 | -0.0061146 | a13 | -0.034299 | -0.0033568 | -0.0016282 | -0.006384 | -0.0054117 | -0.004579 | -0.0026109 | -0.0026062 | 0.00012209 | -0.0040913 | -0.0013352 | -0.00070756 | -0.0036015 | -0.0026017 | -0.0089111 | -0.0059008 | -0.006284 | -0.0073261 | -0.010832 | -0.0061011 | -0.0053665
a14 | 0.004714 | -0.016298 | -0.013433 | -0.015932 | -0.019336 | -0.014711 | -0.01812 | -0.015821 | -0.016253 | -0.019603 | -0.017773 | -0.018789 | -0.014293 | -0.013375 | -0.019641 | -0.01902 | -0.013683 | -0.012508 | -0.015205 | -0.0019136 | -0.015172 | a14 | 0.0045441 | -0.018238 | -0.018255 | -0.018535 | -0.023778 | -0.021443 | -0.019159 | -0.020897 | -0.023398 | -0.02134 | -0.019022 | -0.021311 | -0.01653 | -0.016435 | -0.01884 | -0.018797 | -0.017669 | -0.014519 | -0.016569 | -0.0032182 | -0.017857
a15 | 0.00022283 | -0.0062637 | -0.0032374 | 0.0011017 | -0.0067795 | -0.014243 | -0.011935 | -0.0045836 | -0.0056735 | -0.0020237 | -0.0098363 | -0.00779 | -0.0064958 | -0.0049603 | 0.0047303 | -0.0053239 | -0.0066955 | -0.0012158 | 0.0003316 | -0.013672 | -0.0051759 | a15 | 0.00021518 | -0.0088799 | -0.0062917 | -0.001183 | -0.0031282 | -0.0075776 | -0.018543 | -0.0068167 | -0.011599 | -0.0058283 | -0.012147 | -0.0081735 | -0.0077292 | -0.0078827 | -0.0017616 | -0.014203 | -0.013647 | -0.0093425 | -0.0069321 | -0.012594 | -0.0082899
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.080527 | -0.04956 | 0.050036 | 0.16554 | -0.027551 | -0.23476 | -0.14293 | 0.045995 | -0.02686 | 0.12693 | -0.10386 | -0.059655 | -0.077568 | 0.00050786 | 0.26722 | 0.037818 | -0.047497 | 0.10503 | 0.16702 | -0.29554 |  | Dx (mm) | -0.0021862 | 0.01238 | 0.087402 | 0.21659 | 0.14828 | 0.0070671 | -0.238 | 0.032797 | -0.096764 | 0.11868 | -0.08775 | 0.0023582 | -0.026162 | -0.026637 | 0.17878 | -0.18907 | -0.14255 | -0.026424 | 0.087663 | -0.078293
Dy (mm) | -0.26824 | 0.036972 | 0.14131 | 0.06573 | 0.0056568 | 0.024531 | -0.042691 | 0.050558 | 0.015323 | -0.038309 | -0.0067133 | -0.045398 | 0.060294 | 0.065355 | -0.080855 | -0.039367 | -0.023137 | 0.088645 | 0.047162 | -0.089463 |  | Dy (mm) | -0.25792 | 0.018325 | 0.033741 | 0.027729 | -0.034053 | -0.018941 | -0.040089 | 0.014329 | -0.067971 | -0.046065 | 0.010723 | 0.017036 | 0.1009 | 0.086137 | 0.0046028 | 0.023558 | 0.013919 | 0.082348 | 0.091174 | -0.13393
Mag. len. (m) | 14.4408 | [20 empty] | Mag. len. (m) | 14.4452
TF (mT/KA) | 596.0065 | [20 empty] | TF (mT/KA) | 595.8693
Coil pos. (mm) | 0.01556 | [20 empty] | Coil pos. (mm) | 0.016748
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.034379 | 0.72942 | 0.127 | 0.41461 | 0.35192 | 5.3684 | 0.36993 | 0.54296 | 0.59706 | 0.84113 | 0.73041 | 0.552 | 1.2891 | 1.9751 | 1.0133 | 0.30968 |  | 0
3 | 2.0782 | 5.8093 | 1.5245 | 0.30788 | -0.63155 | 1.0619 | -0.50028 | 0.21126 | 1.5683 | 6.9134 | 0.90165 | 0.42633 | -1.0756 | 0.77576 | -0.96975 | 0.24901 |  | 0
4 | -0.078567 | 0.11074 | -0.077582 | 0.11899 | 0.58251 | 1.0384 | 0.46002 | 0.19421 | 0.11312 | 0.15408 | 0.12703 | 0.13339 | 0.25095 | 0.72314 | 0.13114 | 0.16778 |  | 0
5 | 0.49013 | 1.8265 | 0.86197 | 0.12834 | -0.016702 | 0.42032 | -0.077658 | 0.085793 | 0.51284 | 1.5942 | 0.84939 | 0.096403 | -0.13125 | 0.44322 | -0.18609 | 0.11761 |  | 0
6 | 0.0016894 | 0.048513 | 0.0065141 | 0.036016 | -0.011672 | 0.12619 | -0.019735 | 0.077635 | -0.00985 | 0.059087 | -0.0026784 | 0.047914 | -0.011509 | 0.10196 | -0.01867 | 0.057657 |  | 0
7 | 0.36397 | 0.34954 | 0.33736 | 0.025594 | 0.055679 | 0.32285 | 0.0061018 | 0.053709 | 0.35816 | 0.37648 | 0.33039 | 0.031112 | 0.028718 | 0.31865 | -0.020054 | 0.043228 |  | 0
8 | 0.0087549 | 0.029535 | 0.011378 | 0.012563 | 0.041264 | 0.025033 | 0.040518 | 0.022228 | 0.0019612 | 0.028819 | 0.0038205 | 0.015775 | 0.018003 | 0.024217 | 0.014535 | 0.021935 |  | 0
9 | 0.28816 | 0.033344 | 0.29235 | 0.0099082 | -0.00042552 | 0.070692 | 0.0099501 | 0.009961 | 0.26402 | 0.033673 | 0.26826 | 0.008417 | 0.010597 | 0.068494 | 0.02136 | 0.025703 |  | 0
10 | 0.0010022 | 0.056461 | 3.8745e-09 | 1.5699e-08 | 0.0028619 | 0.034069 | -4.1397e-10 | 2.8149e-09 | 0.0015415 | 0.051494 | -1.4882e-09 | 1.4972e-08 | 0.0025333 | 0.031846 | -3.6311e-10 | 4.756e-09 |  | 0
11 | 0.72647 | 0.043272 | 0.73497 | 0.0054097 | -0.026517 | 0.038438 | -0.032032 | 0.0070059 | 0.7242 | 0.040967 | 0.73234 | 0.004926 | -0.029982 | 0.041977 | -0.036655 | 0.0064269 |  | 0
12 | 0.0016796 | 0.0089587 | 0.0013506 | 0.0027199 | -0.0032317 | 0.0059835 | -0.0032675 | 0.0031611 | 0.0023973 | 0.0075183 | 0.0024476 | 0.0031865 | -0.0040615 | 0.0066593 | -0.0039441 | 0.0031139 |  | 0
13 | 0.086129 | 0.0085135 | 0.086365 | 0.0023728 | -0.0061146 | 0.0076178 | -0.004802 | 0.0020931 | 0.082124 | 0.010024 | 0.082528 | 0.0020356 | -0.0053665 | 0.0072337 | -0.0042031 | 0.0027708 |  | 0
14 | -0.0029298 | 0.0052226 | -0.0036611 | 0.0012321 | -0.015172 | 0.005999 | -0.017298 | 0.0014201 | -0.0053101 | 0.0059801 | -0.0060545 | 0.00085788 | -0.017857 | 0.0066526 | -0.020113 | 0.0013261 |  | 0
15 | 0.04273 | 0.0098309 | 0.044696 | 0.002195 | -0.0051759 | 0.0049389 | -0.0050496 | 0.0046138 | 0.042206 | 0.0099794 | 0.044215 | 0.0024956 | -0.0082899 | 0.0047107 | -0.0084259 | 0.0044098 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.74556 | -0.88321
3 | [12 empty] | 3 | 9 | -1.3664 | -1.1214
4 | [12 empty] | 4 | 16 | -1.8838 | -1.8998
5 | [12 empty] | 5 | 25 | -2.2544 | -2.2398
6 | [12 empty] | 6 | 36 | -2.9398 | -2.9458
7 | [12 empty] | 7 | 49 | -3.2948 | -3.3057
8 | [12 empty] | 8 | 64 | -4.0917 | -3.9845
9 | [12 empty] | 9 | 81 | -4.6117 | -4.2193
10 | [12 empty] | sigma (mm) |  | 0.01556 | 0.016748
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.61258 | 0.15315 | 0.036755 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.35886 | 0.089715 | 0.021532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.20767 | 0.051917 | 0.01246 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.11871 | 0.029678 | 0.0071228 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.067037 | 0.016759 | 0.0040222 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.037395 | 0.0093488 | 0.0022437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.020607 | 0.0051516 | 0.0012364 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.011217 | 0.0028042 | 0.00067302 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0060316 | 0.0015079 | 0.0003619 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0032039 | 0.00080097 | 0.00019223 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0016811 | 0.00042029 | 0.00010087 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0008714 | 0.00021785 | 5.2284e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -2.2208 | -0.48222 | -0.24465 | 0.53625 | 0.65865 | 0.37983 | 0.028207 | 0.21998 | -0.28797 | -0.13466 | 1.055 | 0.70655 | 0.24623 | -0.22978 | 0.15781 | 0.1437 | -0.09131 | -0.023221 | -0.3525 | -0.35486 | 0.034379
3 | 26.0163 | 1.7609 | 1.8216 | 1.7149 | 1.2793 | 0.99991 | 1.211 | 1.7032 | 0.83745 | 1.2119 | 1.5934 | 1.5991 | 1.636 | 1.6475 | 1.9069 | 1.5679 | 1.7096 | 1.3397 | 1.9003 | -5.1423 | 2.0782
4 | -0.36328 | 0.022588 | -0.0029258 | 0.0055122 | -0.01985 | 0.039454 | -0.017934 | 0.057295 | -0.28616 | -0.086791 | 0.091433 | 0.025333 | -0.23575 | -0.29154 | -0.16843 | -0.12242 | -0.15374 | -0.076043 | -0.17652 | -0.36902 | -0.078567
5 | -6.2696 | 0.7178 | 0.69974 | 0.75028 | 0.61909 | 0.91483 | 0.89651 | 1.1199 | 1.0232 | 0.73948 | 0.90691 | 0.86429 | 0.98848 | 0.93153 | 0.96987 | 0.74344 | 0.85507 | 0.89206 | 0.88295 | -3.5352 | 0.49013
6 | -0.22673 | 0.024775 | 0.016836 | 0.051845 | 0.060033 | 0.067712 | 0.046712 | 0.042487 | -0.0082864 | 0.0097288 | -0.0085883 | -0.02497 | -0.045299 | -0.017442 | -0.055735 | -0.013236 | 0.0035904 | -0.015602 | -0.017309 | -0.089515 | 0.0016894
7 | 1.7639 | 0.30524 | 0.35591 | 0.34109 | 0.32003 | 0.30773 | 0.31996 | 0.40237 | 0.31788 | 0.32838 | 0.32898 | 0.32591 | 0.33221 | 0.33899 | 0.34817 | 0.35698 | 0.31436 | 0.34312 | 0.38515 | -0.22148 | 0.36397
8 | -0.039272 | 0.007227 | -0.00017941 | 0.024599 | 0.020125 | -0.0014748 | 0.010874 | 0.0043714 | 0.025055 | 0.014283 | 0.019457 | 0.020045 | 0.044967 | 0.0044123 | 0.0043193 | -0.00065022 | -0.0064133 | 0.0079696 | 0.0058128 | 0.012742 | 0.0087549
9 | 0.20364 | 0.30974 | 0.29276 | 0.2987 | 0.28032 | 0.28651 | 0.27852 | 0.29787 | 0.28365 | 0.28363 | 0.29142 | 0.28955 | 0.29755 | 0.30701 | 0.29425 | 0.27772 | 0.30161 | 0.30564 | 0.28581 | 0.17759 | 0.28816
10 | 5.6561e-09 | -2.0976e-09 | -3.1966e-10 | 1.0438e-08 | 6.9274e-09 | 4.6954e-08 | 1.1989e-08 | -1.3126e-09 | -4.8927e-09 | -1.9769e-08 | -7.3887e-09 | 5.3233e-11 | -2.2411e-09 | -1.5318e-10 | 3.3904e-08 | -3.2017e-09 | 8.6874e-10 | 1.2508e-08 | -1.2524e-08 | 5.1276e-08 | 0.0010022
11 | 0.5977 | 0.72986 | 0.72766 | 0.72893 | 0.72564 | 0.73422 | 0.73169 | 0.73692 | 0.73772 | 0.74121 | 0.73973 | 0.73869 | 0.73966 | 0.73887 | 0.73977 | 0.73397 | 0.74134 | 0.73789 | 0.72571 | 0.59735 | 0.72647
12 | 0.068346 | -0.034394 | -0.015028 | 0.017134 | -0.010932 | -0.037011 | 0.0030613 | 0.018618 | 0.036297 | 0.030863 | 0.015584 | 0.024585 | 0.056419 | 0.0069738 | 0.035622 | 0.016986 | -0.012451 | 0.053457 | 0.037324 | 0.07419 | 0.016796 | ten times!
13 | 0.98697 | 0.85883 | 0.866 | 0.87116 | 0.86243 | 0.81399 | 0.83143 | 0.84663 | 0.84955 | 0.88057 | 0.8798 | 0.88165 | 0.84993 | 0.85808 | 0.86605 | 0.92157 | 0.89017 | 0.87101 | 0.8468 | 0.51264 | 0.86129 | ten times!
14 | 0.078286 | -0.045314 | -0.037119 | -0.037631 | -0.038334 | -0.047213 | -0.03891 | -0.044115 | -0.02194 | -0.046203 | -0.04412 | -0.03317 | -0.013079 | -0.006599 | -0.029843 | -0.050167 | -0.051502 | -0.041628 | -0.032115 | 0.0042279 | -0.029298 | ten times!
15 | 0.074209 | 0.43394 | 0.40002 | 0.42428 | 0.45227 | 0.41276 | 0.45847 | 0.45014 | 0.4562 | 0.48574 | 0.46163 | 0.46906 | 0.44357 | 0.4286 | 0.47467 | 0.45501 | 0.46199 | 0.44615 | 0.43077 | 0.2126 | 0.4273 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 17.8489 | 0.078258 | 0.23827 | 0.47631 | 0.66089 | 0.01314 | 0.36795 | 0.48574 | -0.16337 | 0.99125 | 0.67228 | 0.90986 | 0.033396 | -0.67093 | -0.61547 | 0.15119 | 0.70459 | 1.2948 | 1.0305 | -15.9266 | 0.35192
3 | -5.0584 | -0.63844 | -0.57828 | -0.58766 | -0.57724 | -0.76036 | -0.75584 | -0.70416 | -0.50382 | -0.83461 | -0.57438 | -0.37209 | -0.48861 | -0.17588 | -0.40043 | -0.42745 | -0.12774 | -0.36326 | -0.1348 | -0.10426 | -0.63155
4 | 4.5926 | 0.15529 | 0.17484 | 0.17974 | 0.33923 | 0.35164 | 0.65857 | 0.53719 | 0.7509 | 0.70283 | 0.41105 | 0.47358 | 0.40194 | 0.31952 | 0.42143 | 0.39272 | 0.74645 | 0.64687 | 0.61656 | -0.28026 | 0.58251
5 | 1.7677 | -0.17578 | -0.13126 | -0.13548 | -0.16933 | -0.22804 | -0.10838 | -0.15816 | 0.049925 | -0.11068 | -0.020292 | -0.017925 | -0.046681 | 0.098727 | -0.065097 | -0.021981 | 0.0039973 | -0.12867 | -0.032747 | -0.012658 | -0.016702
6 | 0.5985 | -0.12612 | -0.13506 | -0.082992 | 0.0039197 | -0.031512 | 0.0034041 | 0.084317 | -0.13353 | 0.093308 | 0.067922 | 0.088449 | -0.0094406 | -0.035738 | -0.099931 | -0.069425 | -0.027429 | 0.06552 | -0.010899 | -0.21141 | -0.011672
7 | 1.4248 | -0.026381 | -0.071243 | -0.06754 | -0.035358 | 0.0093342 | -0.00029149 | -0.010874 | 0.044657 | -0.03619 | 0.0060294 | 0.051032 | 0.084828 | 0.09465 | 0.089616 | 0.044706 | 0.030771 | -0.043887 | -0.054027 | 0.12822 | 0.055679
8 | 0.0023549 | 0.025215 | 0.014872 | 0.01652 | 0.0012139 | 0.010554 | 0.025803 | 0.050517 | 0.032837 | 0.058984 | 0.061077 | 0.036369 | 0.046774 | 0.056832 | 0.03807 | 0.04005 | 0.079039 | 0.074483 | 0.060113 | 0.036605 | 0.041264
9 | -0.30829 | 0.016666 | 0.015513 | 0.017462 | 0.011043 | -0.0056066 | -0.014718 | 0.014229 | 0.010146 | -0.0011559 | 0.0023437 | 0.0046532 | 0.015107 | 0.019553 | 0.013614 | 0.0032607 | 0.015585 | 0.016518 | 0.024887 | -0.01555 | -0.00042552
10 | -2.5437e-09 | 2.482e-09 | -2.7258e-09 | -6.2084e-09 | 2.187e-10 | -8.452e-10 | 5.2861e-10 | 4.8437e-09 | -7.5106e-10 | 4.4071e-10 | 2.2819e-10 | 1.4237e-09 | -3.2836e-09 | -7.3299e-10 | 1.4684e-09 | -6.399e-09 | -2.5269e-10 | 1.6423e-10 | 1.949e-09 | 4.1456e-09 | 0.0028619
11 | 0.13558 | -0.028776 | -0.039476 | -0.042175 | -0.045235 | -0.038762 | -0.031464 | -0.02947 | -0.022431 | -0.032388 | -0.029372 | -0.024752 | -0.020801 | -0.024726 | -0.02746 | -0.033158 | -0.030263 | -0.035471 | -0.040398 | -0.028607 | -0.026517 | ten times!
12 | -0.027799 | -0.078175 | -0.063267 | -0.04516 | -0.0041564 | -0.065293 | -0.033519 | -0.015495 | -0.078228 | -0.0046493 | 0.015186 | 0.030273 | -0.013839 | -0.057469 | -0.060325 | -0.046047 | -0.036631 | -0.011617 | -0.019749 | -0.117 | -0.032317 | ten times!
13 | -0.35542 | -0.028544 | -0.054973 | -0.058337 | -0.083047 | -0.086813 | -0.061711 | -0.034524 | -0.034049 | -0.056325 | -0.015048 | -0.01738 | -0.047224 | -0.021136 | -0.031829 | -0.064686 | -0.057979 | -0.06045 | -0.050307 | -0.088609 | -0.061146 | ten times!
14 | 0.063681 | -0.17875 | -0.17955 | -0.18379 | -0.19701 | -0.18298 | -0.17913 | -0.17522 | -0.16954 | -0.17859 | -0.18359 | -0.17418 | -0.16911 | -0.15679 | -0.17521 | -0.17379 | -0.13064 | -0.15659 | -0.16923 | -0.036749 | -0.15172 | ten times!
15 | 0.0022283 | -0.062637 | -0.032374 | 0.011017 | -0.067795 | -0.14243 | -0.11935 | -0.045836 | -0.056735 | -0.020237 | -0.098363 | -0.077899 | -0.064958 | -0.049603 | 0.047303 | -0.053239 | -0.066955 | -0.012158 | 0.003316 | -0.13672 | -0.051759 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.6569 | 0.3302 | -0.2206 | 0.9505 | 0.92712 | 0.65048 | 0.48623 | 0.47885 | 0.37143 | 0.87914 | 1.5339 | 0.30962 | 1.0438 | 0.43621 | 1.2025 | 1.8369 | 1.1211 | 1.0865 | -0.27653 | -1.0014 | 0.59706
3 | 30.0374 | 1.4132 | 0.61876 | 0.60634 | 1.1867 | 0.623 | 0.43402 | 0.60061 | 0.33438 | 0.75364 | 0.39537 | 0.40405 | 1.3107 | 1.376 | 1.7024 | 1.0948 | 0.89276 | 1.0947 | 1.3883 | -7.4443 | 1.5683
4 | -0.17223 | 0.20975 | 0.25089 | 0.26653 | 0.25774 | 0.099205 | 0.002492 | 0.19142 | 0.041043 | 0.07966 | 0.029668 | 0.070269 | -0.18229 | -0.078239 | 0.099367 | 0.30057 | 0.28959 | 0.19301 | 0.16586 | -0.21253 | 0.11312
5 | -4.7351 | 0.77413 | 0.93728 | 0.99648 | 0.70747 | 0.78695 | 0.91882 | 0.98896 | 0.82272 | 0.84387 | 0.97832 | 0.84692 | 0.83974 | 0.69717 | 0.78782 | 0.69098 | 0.89389 | 0.89196 | 0.88562 | -3.8891 | 0.51284
6 | -0.25823 | 0.021435 | 0.038571 | 0.049456 | 0.028798 | 0.029027 | 0.038685 | 0.031994 | -0.023385 | -0.0052097 | -0.037572 | -0.05122 | -0.083397 | -0.11626 | -0.050099 | 0.045236 | 0.019565 | -0.006459 | 0.022619 | -0.10837 | -0.00985
7 | 1.8405 | 0.3355 | 0.34952 | 0.27134 | 0.3246 | 0.34822 | 0.34213 | 0.31856 | 0.29954 | 0.28594 | 0.29024 | 0.29973 | 0.36533 | 0.39335 | 0.33528 | 0.34643 | 0.35384 | 0.35066 | 0.33678 | -0.31588 | 0.35816
8 | -0.046975 | -0.011268 | -0.01427 | 0.0043852 | 0.0036899 | -0.012608 | -0.016076 | -0.0044888 | 0.009463 | 0.012553 | -0.0083373 | -0.017466 | 0.0021339 | -0.0019759 | 0.025055 | 0.029122 | 0.02492 | 0.02195 | 0.021987 | 0.020466 | 0.0019612
9 | 0.18826 | 0.27845 | 0.2607 | 0.26855 | 0.26099 | 0.25709 | 0.24951 | 0.27471 | 0.2659 | 0.26899 | 0.27712 | 0.27334 | 0.27431 | 0.2651 | 0.2685 | 0.25807 | 0.27051 | 0.27876 | 0.27808 | 0.15526 | 0.26402
10 | -1.2212e-09 | 1.3625e-09 | -1.1385e-09 | -1.5331e-08 | -7.6367e-09 | 2.6131e-10 | 5.1261e-08 | -4.1062e-09 | 3.4023e-09 | -1.6568e-08 | 2.1749e-09 | -3.3779e-11 | -1.3748e-09 | -5.1743e-10 | -1.5287e-08 | -1.9411e-08 | -2.1483e-09 | -1.0438e-09 | -6.5247e-10 | 7.4934e-09 | 0.0015415
11 | 0.61491 | 0.73279 | 0.72675 | 0.73627 | 0.73516 | 0.72613 | 0.72105 | 0.73014 | 0.73718 | 0.74019 | 0.73688 | 0.7332 | 0.7321 | 0.72898 | 0.73446 | 0.73168 | 0.73478 | 0.73754 | 0.72677 | 0.59257 | 0.7242
12 | 0.03781 | -0.0099186 | -0.031978 | 0.013988 | 0.020317 | 0.00071505 | -0.014455 | 0.0031254 | 0.0084705 | 0.016118 | 0.023566 | 0.018515 | 0.082587 | 0.038669 | 0.080635 | 0.06279 | 0.056611 | 0.049435 | 0.021387 | 0.022622 | 0.023973 | ten times!
13 | 0.94589 | 0.84204 | 0.8109 | 0.79755 | 0.82894 | 0.81275 | 0.78557 | 0.80268 | 0.82392 | 0.8301 | 0.81574 | 0.80302 | 0.83917 | 0.84343 | 0.85567 | 0.85728 | 0.84243 | 0.83765 | 0.82618 | 0.41504 | 0.82124 | ten times!
14 | 0.11934 | -0.066432 | -0.059361 | -0.06072 | -0.060054 | -0.044853 | -0.06065 | -0.049605 | -0.047505 | -0.067703 | -0.074158 | -0.064372 | -0.0478 | -0.054191 | -0.067979 | -0.064177 | -0.061587 | -0.068668 | -0.069994 | -0.071159 | -0.053101 | ten times!
15 | 0.069528 | 0.43135 | 0.44498 | 0.46399 | 0.46899 | 0.45507 | 0.42192 | 0.44279 | 0.48514 | 0.48789 | 0.44278 | 0.44929 | 0.40375 | 0.40016 | 0.45147 | 0.4201 | 0.44497 | 0.42789 | 0.41616 | 0.19989 | 0.42206 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.6492 | 1.0819 | 1.3065 | 0.83582 | 1.0852 | 0.85563 | 0.78028 | 0.54252 | 0.68328 | 0.38899 | 1.1332 | 1.5425 | 1.4509 | 1.2245 | 1.2436 | 1.2712 | 1.024 | 1.0076 | 0.78234 | 0.6811 | 1.2891
3 | -4.2782 | -1.1259 | -1.2996 | -1.1958 | -1.0546 | -0.94847 | -1.2336 | -0.97078 | -0.51334 | -0.93207 | -0.89154 | -1.0238 | -0.53581 | -0.76758 | -0.77537 | -0.64438 | -1.0276 | -1.1314 | -1.3839 | -0.94935 | -1.0756
4 | 2.9972 | 0.16732 | 0.34263 | 0.22084 | -0.073116 | -0.063624 | 0.11333 | -0.11 | 0.33083 | 0.26318 | 0.38947 | -0.067769 | 0.00049784 | 0.26732 | 0.14304 | -0.097219 | 0.21641 | 0.27281 | 0.044537 | 0.44838 | 0.25095
5 | 1.7072 | -0.19746 | -0.25117 | -0.28957 | -0.12222 | -0.2503 | -0.27678 | -0.20169 | -0.045526 | -0.15757 | -0.25056 | -0.20983 | -0.014979 | 0.060669 | -0.013577 | -0.15856 | -0.30061 | -0.33405 | -0.33586 | -0.34948 | -0.13125
6 | 0.48802 | -0.042487 | -0.031295 | -0.18442 | -0.0073075 | 0.01257 | -0.0039067 | -0.028576 | 0.064566 | -0.017097 | 0.028368 | 0.034663 | 0.011679 | 0.029236 | -0.033509 | -0.0020607 | -0.0018529 | -0.04482 | -0.11981 | -0.20664 | -0.011509
7 | 1.387 | -0.034427 | -0.076707 | -0.035816 | -0.020793 | -0.0060028 | -0.027035 | -0.055501 | 0.049567 | -0.019876 | -0.071816 | -0.046691 | 0.032143 | 0.046164 | 0.055529 | 0.015108 | -0.048769 | -0.075299 | -0.040753 | 0.068723 | 0.028718
8 | 0.076053 | -0.0018555 | 0.030729 | 0.040904 | -0.027561 | 0.018678 | 0.0057212 | -0.027398 | 0.012832 | 0.020736 | 0.024559 | 0.013508 | -0.015388 | 0.012485 | 0.033975 | 0.0088092 | 0.034297 | 0.051613 | 0.024994 | 0.018876 | 0.018003
9 | -0.25999 | 0.027476 | 0.047867 | 0.016891 | 0.03256 | 0.025672 | 0.023502 | 0.048784 | 0.050324 | 0.033359 | 0.04426 | 0.038064 | 0.042841 | 0.031284 | -0.0015508 | -0.023127 | -0.010333 | -0.019759 | -0.023634 | -0.032851 | 0.010597
10 | 4.4528e-09 | -1.5818e-09 | 3.2757e-09 | 2.9772e-09 | -1.7024e-09 | 9.774e-10 | 1.6184e-09 | 9.9311e-10 | -1.3775e-09 | -3.3216e-09 | -2.2429e-09 | 9.5528e-10 | -1.7602e-08 | 3.2794e-09 | 6.4138e-10 | 4.431e-10 | 3.6994e-10 | 3.4152e-09 | 2.3457e-09 | 1.4125e-08 | 0.0025333
11 | 0.14813 | -0.034839 | -0.037494 | -0.036937 | -0.03687 | -0.033384 | -0.037399 | -0.037591 | -0.029775 | -0.0392 | -0.048206 | -0.036715 | -0.029978 | -0.023202 | -0.026804 | -0.038927 | -0.043732 | -0.045309 | -0.043433 | -0.012307 | -0.029982
12 | -0.088267 | -0.031971 | -0.065369 | -0.12953 | -0.050982 | -0.042556 | -0.017049 | -0.050915 | -0.0059029 | -0.015556 | -0.03412 | -0.015353 | -0.019055 | -0.027894 | -0.040721 | -0.011486 | -0.012907 | -0.050006 | -0.088566 | -0.096845 | -0.040615 | ten times!
13 | -0.31931 | -0.030991 | -0.0034205 | -0.033614 | -0.027249 | -0.045259 | -0.06556 | -0.020417 | -0.020909 | -0.022329 | -0.025255 | -0.005879 | -0.035459 | -0.025585 | -0.06335 | -0.083901 | -0.080906 | -0.071683 | -0.094793 | -0.072968 | -0.053665 | ten times!
14 | 0.060212 | -0.18798 | -0.19038 | -0.19384 | -0.22081 | -0.20689 | -0.214 | -0.21236 | -0.21607 | -0.1892 | -0.20291 | -0.21926 | -0.20052 | -0.19447 | -0.18751 | -0.21824 | -0.19781 | -0.17624 | -0.19193 | -0.018255 | -0.17857 | ten times!
15 | 0.0021518 | -0.088799 | -0.062917 | -0.01183 | -0.031282 | -0.075776 | -0.18543 | -0.068167 | -0.11599 | -0.058283 | -0.12147 | -0.081735 | -0.077292 | -0.078827 | -0.017616 | -0.14203 | -0.13647 | -0.093425 | -0.069321 | -0.12594 | -0.082899 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00058121 | -1.0966e-05 | -0.00020836 | -0.00020836 | -0.00020836 | -0.00020836 | -0.00020836 | -0.00020836 | -1.0966e-05 | -1.0966e-05 | -1.0966e-05 | -1.0966e-05 | -1.0966e-05 | 0.00018643 | -1.0966e-05 | -1.0966e-05 | -1.0966e-05 | 0.00058121 |  | 596.0065
|  | 0.00081169 | 2.1938e-05 | -0.0001755 | 2.1938e-05 | -0.0001755 | -0.0001755 | -0.00037294 | -0.0001755 | 2.1938e-05 | -0.0001755 | -0.0001755 | -0.0001755 | 2.1938e-05 | 2.1938e-05 | 2.1938e-05 | 2.1938e-05 | 2.1938e-05 | 0.00061425 |  | 595.8693
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3714.5748
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.9632 | -3927.7739 | 6.5391 | 0.6169 | -1.3572 | -1.3572 | -1.3572 | -1.3572 | -1.3572 | -1.3572 | 0.6169 | 0.6169 | 0.6169 | 0.6169 | 0.6169 | 2.591 | 0.6169 | 0.6169 | 0.6169 | 6.5391 | -3501.3757
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.0041642 | 2.1273 | 1.3276 | 0.52969 | 0.41534 | 0.013488 | -0.0051798 | 0.19567 | 0.053242 | 0.069595 | -0.079882 | 0.18569 | 0.19888 | 0.25787 | -0.43028 | -0.69463 | -0.62485 | -0.40322 | -0.363 | -0.64606 | -1.7802
db2 | 0.13197 | -2.6312 | -0.62167 | -0.35634 | 0.44215 | 0.52292 | 0.22252 | -0.12811 | 0.10136 | -0.42167 | -0.25217 | 0.90314 | 0.56132 | 0.10287 | -0.36027 | 0.08219 | 0.016752 | -0.23322 | -0.13511 | -0.44666 | -0.17714
db3 | 1.5234 | 26.2526 | 0.23635 | 0.29339 | 0.1905 | -0.24427 | -0.52466 | -0.30622 | 0.17549 | -0.68657 | -0.30829 | 0.069261 | 0.079274 | 0.11169 | 0.12026 | 0.38409 | 0.04648 | 0.19063 | -0.19501 | 0.36759 | -5.1391
db4 | -0.075118 | -0.13344 | 0.090835 | 0.084878 | 0.11218 | 0.051483 | 0.06549 | 0.025986 | 0.14638 | -0.2177 | 0.0094975 | 0.14434 | 0.088094 | -0.17802 | -0.21782 | -0.033961 | -0.040902 | -0.088152 | 0.023707 | -0.066317 | -0.123
db5 | 0.86233 | -6.2355 | -0.14349 | -0.15727 | -0.10695 | -0.24371 | 0.049823 | 0.032883 | 0.25686 | 0.16156 | -0.12097 | 0.045348 | 0.0036018 | 0.12752 | 0.069734 | 0.10358 | -0.11983 | -0.0074433 | 0.027734 | 0.021028 | -3.5355
db6 | 0.0088341 | -0.1411 | 0.010962 | 0.017756 | 0.064578 | 0.048159 | 0.033053 | 0.021674 | 0.040339 | -0.02037 | 0.015259 | -0.029471 | -0.039869 | -0.064947 | -0.028401 | -0.028598 | -0.016669 | -0.01029 | -0.010572 | -0.002593 | -0.062755
db7 | 0.33773 | 1.7656 | -0.032938 | 0.016289 | 0.005865 | -0.017887 | -0.026824 | -0.016903 | 0.06364 | -0.020301 | -0.006863 | -0.0088024 | -0.011549 | -0.0079506 | -0.00051712 | 0.016096 | 0.019887 | -0.022397 | 0.0031348 | 0.04802 | -0.21865
db8 | 0.01274 | -0.084719 | -0.013017 | -0.007374 | 0.034874 | 0.0037192 | -0.04713 | -0.021135 | -0.0022865 | 0.0086571 | 0.018616 | -0.0076392 | -0.00070852 | 0.020981 | -0.0088574 | 0.03053 | -0.0084009 | -0.025729 | 0.009516 | 0.015384 | -0.014336
db9 | 0.29493 | 0.18967 | 0.015031 | -0.0059559 | 0.0093037 | -0.014445 | 0.0039508 | -0.012052 | 0.0028858 | -0.011176 | -0.0080149 | -0.0010161 | -0.004993 | 0.0031088 | 0.011104 | 0.013906 | -0.017263 | 0.0070971 | 0.011641 | -0.0031119 | 0.19278
db10 | 0.0050844 | -0.0064539 | -0.025746 | 0.019588 | 0.067645 | -0.016695 | -0.1063 | -0.067423 | 0.015727 | -0.01654 | 0.049575 | -0.050412 | -0.031668 | -0.038086 | -0.0039144 | 0.11017 | 0.010467 | -0.026208 | 0.042342 | 0.06747 | -0.10521
db11 | 0.73558 | 0.59464 | -0.0053913 | -0.0080561 | -0.0052697 | -0.0098985 | 0.0013548 | -0.003284 | 0.0014361 | 0.0021726 | 0.0064248 | 0.0044787 | 0.0031575 | 0.0039196 | 0.003363 | 0.0069169 | -0.0017005 | 0.0058168 | 0.0029194 | -0.0083607 | 0.59735
db12 | 0.00214 | -0.0058762 | -0.0085452 | 8.258e-05 | 0.010235 | -0.0048825 | -0.019636 | -0.010386 | 0.0026341 | -9.2259e-05 | 0.0085838 | -0.0070739 | -0.0033997 | -0.001028 | -0.0013131 | 0.017391 | 0.0018086 | -0.0064547 | 0.010167 | 0.01191 | -0.0040272
db13 | 0.086904 | 0.099209 | -0.00032904 | 0.0010826 | 0.0012613 | -0.00047611 | -0.0030009 | -0.0032915 | -0.001718 | -0.0016941 | 0.00061648 | 0.0016524 | 0.00087823 | -0.0010252 | -0.00042748 | 7.5477e-05 | 0.0045712 | 0.0021282 | 0.0010275 | -0.0013311 | 0.052435
db14 | -0.0031753 | 0.0073857 | -0.0029364 | 0.0014885 | 0.0051365 | -0.0016527 | -0.0092381 | -0.0065316 | 0.00065967 | 4.3785e-05 | 0.0035686 | -0.0052393 | -0.0027373 | -0.00064359 | 0.0028003 | 0.010952 | -0.00059687 | -0.0039095 | 0.0029603 | 0.0058759 | -0.0057589
db15 | 0.044696 | 0.0074209 | -0.0013018 | -0.004694 | -0.0022682 | 0.00053133 | -0.0034201 | 0.0011514 | 0.00031768 | 0.00092382 | 0.003878 | 0.001467 | 0.0022103 | -0.00033919 | -0.0018361 | 0.0027709 | 0.00080553 | 0.0015031 | -8.0685e-05 | -0.0016189 | 0.02126
da2 | 0.37376 | 17.069 | -0.28412 | -0.10867 | 0.10447 | 0.28986 | -0.33631 | 0.0010555 | 0.1183 | -0.53401 | 0.59979 | 0.30438 | 0.53019 | -0.32423 | -1.0321 | -1.0193 | -0.23173 | 0.32691 | 0.93057 | 0.66491 | -15.8693
da3 | -0.49958 | -5.1224 | -0.14019 | -0.076719 | -0.082217 | -0.079341 | -0.27612 | -0.27237 | -0.1995 | -0.0086072 | -0.31907 | -0.082397 | 0.1225 | 0.0025659 | 0.32035 | 0.1195 | 0.075511 | 0.36629 | 0.14779 | 0.38202 | -0.091555
da4 | 0.4636 | 4.9526 | -0.30003 | -0.26687 | -0.27755 | -0.12245 | -0.094232 | 0.18963 | 0.085207 | 0.29067 | 0.22933 | -0.053442 | 0.00098975 | -0.046755 | -0.12974 | -0.06511 | -0.077962 | 0.27814 | 0.19871 | 0.16146 | -0.20639
da5 | -0.077833 | 1.7696 | -0.095955 | -0.054064 | -0.060065 | -0.091439 | -0.14787 | -0.030869 | -0.078357 | 0.12875 | -0.029858 | 0.055535 | 0.058881 | 0.030265 | 0.17618 | 0.0052336 | 0.055123 | 0.082267 | -0.048573 | 0.044812 | 0.0079042
da6 | -0.017801 | 0.39057 | -0.10387 | -0.10084 | -0.060981 | 0.022703 | -0.011644 | 0.01644 | 0.10912 | -0.11444 | 0.10512 | 0.084817 | 0.09997 | 0.013081 | -0.010142 | -0.083634 | -0.05599 | -0.012696 | 0.09253 | 0.010449 | -0.21756
da7 | 0.0066728 | 1.4345 | -0.033675 | -0.076892 | -0.07104 | -0.042013 | 0.00086524 | -0.0080487 | -0.016232 | 0.037735 | -0.040556 | -0.0032243 | 0.043397 | 0.077244 | 0.088096 | 0.08444 | 0.038507 | 0.022747 | -0.045929 | -0.055422 | 0.12518
da8 | 0.043424 | -0.010101 | -0.013186 | -0.0088899 | -0.015771 | -0.041606 | -0.028436 | -0.022489 | 0.01451 | -0.0086673 | 0.010175 | 0.016599 | -0.013414 | 0.011334 | 0.022825 | -0.016727 | -0.0084718 | 0.031967 | 0.044873 | 0.025376 | 0.029372
da9 | 0.010004 | -0.30309 | 0.0058193 | 0.008926 | 0.012107 | 0.00095872 | -0.018906 | -0.022117 | 0.0052967 | -5.0293e-05 | -0.013557 | -0.0074371 | -0.004116 | 0.0029332 | 0.0095978 | -0.0070004 | -0.0074223 | 0.006073 | 0.010763 | 0.018131 | -0.0063507
da10 | 0.007419 | -0.10034 | 0.0092996 | 0.051903 | 0.016749 | -0.0042709 | 0.0085215 | -0.023225 | 0.013708 | -0.00041678 | -0.0266 | -0.0085439 | -0.026279 | 0.019762 | 0.020982 | -0.047549 | -0.025157 | -0.01667 | 0.028958 | 0.0088298 | -0.026947
da11 | -0.032054 | 0.13761 | 0.0033003 | -0.0071471 | -0.0097375 | -0.013193 | -0.0063727 | 0.0013242 | 0.0027804 | 0.0097617 | -0.00088825 | 0.0026288 | 0.0073292 | 0.011171 | 0.0073619 | 0.0014202 | -0.0013849 | 0.0020067 | -0.0024521 | -0.0079085 | -0.025823
da12 | -0.00232 | -0.019442 | -0.0031538 | 0.0045512 | 0.00098889 | 0.0024035 | -0.0017169 | -0.0031607 | 0.0036773 | -0.0045222 | -0.0011878 | 0.0034226 | 0.0025627 | 0.0048129 | 0.00056176 | -0.010125 | -0.005011 | -0.0026236 | 0.0061367 | 0.0023833 | -0.013335
da13 | -0.0049884 | -0.037347 | 0.0025789 | -0.0012051 | -0.0027634 | -0.0029297 | -0.0012848 | 0.0011147 | 0.00088231 | 0.0018807 | -0.0025583 | 0.0049384 | 0.0043128 | 0.00093778 | 0.0028509 | -0.0026889 | -0.0019166 | -0.00018713 | -0.0020757 | -0.0018868 | -0.0080345
da14 | -0.016322 | 0.004714 | 2.3786e-05 | 0.0028887 | 0.00039015 | -0.0030144 | 0.0016113 | -0.0017983 | 0.00050065 | 6.9086e-05 | -0.0032811 | -0.0014515 | -0.0024671 | 0.0020285 | 0.0029472 | -0.003319 | -0.0026985 | 0.0026391 | 0.0038143 | 0.0011171 | -0.0019136
da15 | -0.0050496 | 0.00022283 | -0.0012141 | 0.0018122 | 0.0061513 | -0.0017299 | -0.0091932 | -0.0068853 | 0.00046604 | -0.00062392 | 0.0030259 | -0.0047867 | -0.0027403 | -0.0014461 | 8.9319e-05 | 0.0097799 | -0.00027428 | -0.0016459 | 0.0038338 | 0.0053812 | -0.013672
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3685.3797
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.8162 | -3792.0055 | 9.0089 | 1.1107 | -0.86387 | 1.1107 | -0.86387 | -0.86387 | -2.8384 | -0.86387 | 1.1107 | -0.86387 | -0.86387 | -0.86387 | 1.1107 | 1.1107 | 1.1107 | 1.1107 | 1.1107 | 7.0343 | -3578.7538
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.0078025 | 1.5925 | 0.1491 | 0.57026 | 0.45923 | 0.31507 | 0.50772 | 0.42022 | 0.089822 | 0.24208 | -0.078867 | -0.12798 | 0.11182 | -0.11226 | -0.25083 | -0.52055 | -0.51532 | -0.24823 | -0.53343 | -0.47785 | -1.1349
db2 | 0.73466 | -1.7944 | -0.39998 | -0.94375 | 0.23544 | 0.20904 | -0.085778 | -0.26651 | -0.25185 | -0.37123 | 0.15003 | 0.79635 | -0.42288 | 0.31148 | -0.29495 | 0.50418 | 1.0798 | 0.37323 | 0.35938 | -0.98205 | -0.94798
db3 | 0.90238 | 30.1868 | 0.51077 | -0.28148 | -0.28487 | 0.29114 | -0.27948 | -0.46535 | -0.30034 | -0.56455 | -0.14459 | -0.50792 | -0.49811 | 0.40886 | 0.46978 | 0.80412 | 0.18374 | -0.01674 | 0.18712 | 0.48792 | -7.4286
db4 | 0.12871 | -0.065603 | 0.084147 | 0.14348 | 0.19087 | 0.15279 | -0.029314 | -0.18024 | 0.071023 | -0.10719 | -0.027207 | -0.11863 | -0.057131 | -0.31575 | -0.21248 | 0.003749 | 0.14205 | 0.13188 | 0.065218 | 0.062739 | -0.15154
db5 | 0.85018 | -4.7108 | -0.075753 | 0.088684 | 0.1523 | -0.14085 | -0.063113 | 0.067914 | 0.13924 | -0.025432 | -0.0063956 | 0.12923 | -0.0034856 | -0.010496 | -0.1531 | -0.063844 | -0.16043 | 0.043609 | 0.042653 | 0.039269 | -3.8945
db6 | -0.0025282 | -0.13264 | 0.025652 | 0.052759 | 0.073214 | 0.048089 | 0.032385 | 0.011739 | 0.038494 | -0.029902 | 0.0090479 | -0.043776 | -0.048161 | -0.085384 | -0.11881 | -0.02629 | 0.024569 | 0.0045705 | -0.00499 | 0.036792 | -0.096446
db7 | 0.33088 | 1.8464 | 0.0045723 | 0.017974 | -0.057162 | -0.0053057 | 0.017442 | 0.015616 | -0.012176 | -0.031241 | -0.043239 | -0.040016 | -0.031275 | 0.03461 | 0.061733 | 0.0078884 | 0.014972 | 0.022367 | 0.017451 | 0.0057909 | -0.31573
db8 | 0.0043587 | -0.079059 | -0.014242 | -0.0086229 | 0.028245 | 0.018348 | -0.015884 | -0.04928 | -0.0049347 | -0.005354 | 0.024029 | -0.024431 | -0.021826 | -0.007556 | -0.010866 | 0.043906 | 0.0013806 | 0.0021845 | 0.014944 | 0.029961 | 0.012998
db9 | 0.27103 | 0.17556 | 0.0073882 | -0.0087897 | 0.0082736 | -0.0053856 | -0.014007 | -0.0088989 | 0.0038959 | -0.0039148 | 0.00058882 | 0.0078033 | 0.0022535 | 0.0010668 | -0.0074846 | 0.004813 | -0.0050379 | 0.0040344 | 0.0062774 | 0.0071234 | 0.15316
db10 | 0.0019005 | 0.021761 | 0.003812 | 0.036223 | 0.092806 | 0.061539 | 0.00074552 | -0.10402 | 0.012505 | -0.045037 | 0.048729 | -0.039643 | -0.00051573 | -0.011404 | -0.012149 | 0.075642 | -0.082804 | -0.063179 | -0.011179 | 0.037929 | -0.028151
db11 | 0.73286 | 0.61055 | -4.7189e-05 | -0.0059044 | 0.0055756 | 0.0031361 | -0.0067927 | -0.0095759 | -0.0026361 | 0.0044223 | 0.0078374 | 0.0041992 | 0.00034601 | -0.0011578 | -0.0040614 | 0.0037966 | -0.00054039 | 0.002183 | 0.0046137 | -0.0053942 | 0.59145
db12 | 0.0026777 | -0.0026107 | -0.0028885 | -0.00081528 | 0.011156 | 0.0078313 | -0.0022636 | -0.017753 | -0.00047732 | -0.0074162 | 0.0056752 | -0.0054171 | -0.00068356 | 0.0042625 | -0.00026106 | 0.016219 | -0.0080316 | -0.0055739 | 0.0011096 | 0.0053276 | -0.00062255
db13 | 0.083104 | 0.095465 | 0.0012986 | -0.0017134 | -0.0025865 | -0.00087047 | -0.0021629 | -0.0024141 | -0.0026994 | -0.0012422 | -0.0010477 | -0.00066911 | -0.0025352 | 0.0021892 | 0.002439 | 0.0025095 | 0.0047962 | 0.0025515 | 0.0017201 | 0.00043683 | 0.041761
db14 | -0.0057446 | 0.011967 | -0.00032482 | 0.0031862 | 0.0079756 | 0.0053786 | 0.0014059 | -0.0092022 | 0.0020605 | -0.0035212 | 0.0035216 | -0.0047636 | -0.00049068 | 0.00073701 | 6.7954e-06 | 0.0056002 | -0.0069387 | -0.0054814 | -0.0014198 | 0.0022701 | -0.0097938
db15 | 0.044215 | 0.0069528 | -0.0010804 | 0.00028313 | 0.0021838 | 0.0026842 | 0.0012921 | -0.0020233 | 6.3876e-05 | 0.0042994 | 0.0045739 | 6.2826e-05 | 0.00071393 | -0.0038397 | -0.0041993 | 0.00093228 | -0.0022047 | 0.00028197 | -0.0014255 | -0.0025985 | 0.019989
da2 | 1.0151 | 9.7344 | 0.068198 | 0.28055 | -0.20751 | 0.046832 | -0.16167 | -0.20223 | -0.47533 | -0.32864 | -0.6432 | 0.12786 | 0.52791 | 0.453 | 0.22566 | 0.21315 | 0.2733 | 0.027631 | 0.006587 | -0.2321 | 0.80693
da3 | -0.96951 | -4.2765 | -0.15532 | -0.32316 | -0.21663 | -0.088933 | 0.020632 | -0.26872 | -0.0013886 | 0.45047 | 0.041847 | 0.071846 | -0.0541 | 0.43028 | 0.19934 | 0.19877 | 0.32914 | -0.063234 | -0.16038 | -0.41045 | -0.95375
da4 | 0.13478 | 3.2822 | 0.035303 | 0.21013 | 0.077276 | -0.21788 | -0.20233 | -0.014635 | -0.24301 | 0.18392 | 0.11483 | 0.26235 | -0.19927 | -0.11424 | 0.14631 | 0.008468 | -0.22113 | 0.094643 | 0.1574 | -0.078131 | 0.57739
da5 | -0.18747 | 1.7171 | -0.010008 | -0.063954 | -0.11325 | 0.064258 | -0.062973 | -0.088955 | -0.014334 | 0.14101 | 0.02903 | -0.064055 | -0.022585 | 0.1697 | 0.24475 | 0.17232 | 0.029074 | -0.11323 | -0.14649 | -0.15031 | -0.34122
da6 | -0.016401 | 0.31886 | -0.024067 | -0.013067 | -0.16773 | 0.0039811 | 0.026628 | 0.0097823 | -0.011209 | 0.072088 | -0.0061931 | 0.048124 | 0.052826 | 0.040801 | 0.057143 | -0.012887 | 0.016277 | 0.018845 | -0.017592 | -0.093753 | -0.19359
da7 | -0.019779 | 1.3953 | -0.014719 | -0.055659 | -0.011518 | -0.003147 | 0.013914 | -0.0062924 | -0.03606 | 0.0693 | 0.00017368 | -0.052998 | -0.027019 | 0.051821 | 0.065463 | 0.077936 | 0.03385 | -0.031116 | -0.055547 | -0.018382 | 0.067674
da8 | 0.017001 | 0.055015 | -0.016295 | 0.019913 | 0.029495 | -0.046725 | -0.00052882 | -0.019268 | -0.041789 | -0.01514 | -0.00043934 | 0.0071603 | -0.0012601 | -0.019982 | 0.0057871 | 0.01745 | -0.0029662 | 0.01987 | 0.045615 | 0.019102 | 0.01029
da9 | 0.021454 | -0.26285 | 0.0061276 | 0.027672 | -0.0023528 | 0.0085464 | 0.0041605 | 0.0057218 | 0.027534 | 0.031847 | 0.0092349 | 0.022261 | 0.016632 | 0.020272 | 0.0088372 | -0.022899 | -0.047041 | -0.03297 | -0.042126 | -0.041458 | -0.028946
da10 | 0.0075531 | -0.09282 | 9.2246e-05 | 0.0049361 | -0.00040303 | -0.025596 | -0.015782 | -0.019505 | -0.0021249 | -0.035365 | -0.030402 | -0.00042275 | -0.00025664 | 0.036333 | 0.029736 | -0.0084313 | 0.0069486 | 0.0021359 | 0.028871 | 0.029235 | -0.046061
da11 | -0.036748 | 0.14861 | 0.0018891 | -0.00097081 | -0.0016719 | -0.0010867 | 0.0033343 | 0.00028708 | -0.00091499 | 0.0074646 | -0.0030518 | -0.011323 | 5.3756e-05 | 0.0071518 | 0.013639 | 0.0094313 | -0.0023955 | -0.007027 | -0.0083605 | -0.0064484 | -0.011579
da12 | -0.0029702 | -0.026154 | 0.0008355 | -0.0017286 | -0.0093825 | -0.0046712 | -0.0023904 | -0.00057941 | -0.0013684 | -0.0015561 | -0.0015964 | 0.00026487 | 0.0024026 | 0.0072063 | 0.0054309 | -0.0019483 | 0.0040708 | 0.0031534 | 0.0030369 | -0.0011802 | -0.013295
da13 | -0.0043359 | -0.034299 | 0.00097908 | 0.0027077 | -0.0020481 | -0.0010758 | -0.00024306 | 0.001725 | 0.0017297 | 0.004458 | 0.00024463 | 0.0030007 | 0.0036283 | 0.00073435 | 0.0017342 | -0.0045752 | -0.0015649 | -0.0019481 | -0.0029902 | -0.0064963 | -0.0061011
da14 | -0.019152 | 0.0045441 | 0.00091406 | 0.00089723 | 0.00061676 | -0.0046264 | -0.0022911 | -6.7744e-06 | -0.0017452 | -0.0042465 | -0.0021881 | 0.00013003 | -0.0021592 | 0.0026218 | 0.0027169 | 0.00031241 | 0.00035459 | 0.0014831 | 0.0046332 | 0.0025834 | -0.0032182
da15 | -0.0084259 | 0.00021518 | -0.000454 | 0.0021342 | 0.007243 | 0.0052977 | 0.00084831 | -0.010117 | 0.0016092 | -0.0031735 | 0.0025976 | -0.0037211 | 0.00025245 | 0.00069673 | 0.00054326 | 0.0066643 | -0.0057769 | -0.0052215 | -0.0009166 | 0.0014938 | -0.012594
Field Colinearity | -0.075526
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 40.6 | 4.6 | 3.5 | 7 | -3.47 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.34 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -2.45 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.73 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.61 | 0.3 | 3.5 | 7 | 0.62 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.33 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.266 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 24.5 | 56.7 | 8.4 | 72.8 | -12.22 | 5.28 | -20.97 | 14.03
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.41 | 3.09 | -2.16 | 4.84 | -0.64 | 0.64 | -1.28 | 1.28 | -7 | 2.1 | -11.55 | 6.65 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.415 | 1.045 | 0.1 | 1.36 | -0.216 | 0.216 | -0.432 | 0.432 | 1.56 | 3.66 | 0.51 | 4.71 | -0.08 | 1.32 | -0.78 | 2.02
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.225 | 0.435 | 0.12 | 0.54 | -0.08 | 0.08 | -0.16 | 0.16 | 0.091 | 0.441 | -0.084 | 0.616 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 24 | 4.6 | 3.5 | 7 | -6.652 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.43 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -5.57 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.54 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 1.62 | 0.3 | 3.5 | 7 | -0.0495 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.28 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.147 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 7.9 | 40.1 | -8.2 | 56.2 | -15.402 | 2.098 | -24.152 | 10.848
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.32 | 3.18 | -2.07 | 4.93 | -0.64 | 0.64 | -1.28 | 1.28 | -10.12 | -1.02 | -14.67 | 3.53 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.225 | 0.855 | -0.09 | 1.17 | -0.216 | 0.216 | -0.432 | 0.432 | 0.57 | 2.67 | -0.48 | 3.72 | -0.7495 | 0.6505 | -1.4495 | 1.3505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.175 | 0.385 | 0.07 | 0.49 | -0.08 | 0.08 | -0.16 | 0.16 | -0.028 | 0.322 | -0.203 | 0.497 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 34.5 | 4.6 | 3.5 | 7 | 2.43 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.74 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.44 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.7 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.54 | 0.3 | 3.5 | 7 | 0.389 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.26 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.282 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 18.4 | 50.6 | 2.3 | 66.7 | -6.32 | 11.18 | -15.07 | 19.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.01 | 2.49 | -2.76 | 4.24 | -0.64 | 0.64 | -1.28 | 1.28 | -8.99 | 0.11 | -13.54 | 4.66 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.385 | 1.015 | 0.07 | 1.33 | -0.216 | 0.216 | -0.432 | 0.432 | 1.49 | 3.59 | 0.44 | 4.64 | -0.311 | 1.089 | -1.011 | 1.789
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.155 | 0.365 | 0.05 | 0.47 | -0.08 | 0.08 | -0.16 | 0.16 | 0.107 | 0.457 | -0.068 | 0.632 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.13197 | 1.5234 | -0.075118 | 0.86233 | 0.33773 | -0.041642 | 0.37376 | -0.49958 | 0.4636 | 15.56 | 0.59601 | 0.73466 | 0.90238 | 0.12871 | 0.85018 | 0.33088 | -0.078024 | 1.0151 | -0.96951 | 0.13478 | 16.7479 | 0.59587 | 0.034379 | 2.0782 | -0.078567 | 0.49013 | 0.36397 | 0 | 0.35192 | -0.63155 | 0.58251 | 14440.839 | 0.59706 | 1.5683 | 0.11312 | 0.51284 | 0.35816 | 1.1855e-13 | 1.2891 | -1.0756 | 0.25095 | 14445.163
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8