## 2362
ITP Nr.
| LHC - Cold masses: HELIUM MASS SPECTROMETER LEAK TEST REPORT | [14 empty] | 23
24
Cold Mass Nr. |  | 2362
Step Nr. |  | 1 |  |  |  | 2 |  |  |  | 3 |  |  |  | 4
Volume / Volume to be tested |  | CM -> Vacuum |  |  |  | CM -> cold bore tubes |  |  |  | CM -> Heat Exch. |  |  |  | Heat Exch -> Vacuum
Fuga calibrata / Calibrated leak parameter
Calibrated leak N°. |  | 4011007195 |  |  |  | 4011007225 |  |  |  | FC03000505 |  |  |  | 4011007195
Data calibr. / calibration date |  | 2002-10-08 |  |  |  | 2002-10-08 |  |  |  | 2003-05-20 |  |  |  | 2002-10-08
Temp. calibrazione fuga / Calibration Temp. | T | 23 | °C |  | T | 23 | °C |  | T | 20 | °C |  | T | 23 | °C
Valore nom. fuga calibrata / Calibrated leak nom. value |  | 3e-08 | mbar l s-1 |  |  | 3.3e-08 | mbar l s-1 |  |  | 1.4e-07 | mbar l s-1 |  |  | 3e-08 | mbar l s-1
REMARKS: check calibration date (updating leak rate by ageing + leak re-calibration)
Calibrazione del sistema / System calibration
Conc. He nelle linee di test (100%) / Volumetric fraction of tracer gas in the injection envelope | C | 1 |  |  | C | 1 |  |  | C | 1 |  |  | C | 1
T ambiente / Test temp. | T | 27.5 | °C |  | T | 27.5 | °C |  | T | 27.5 | °C |  | T | 27.5 | °C
Fuga calibrata con correz. T ed età / Size of calib. leak after corr. for ageing and T) | qFR | 3.4531e-08 | mbar l s-1 |  | qFR | 3.7968e-08 | mbar l s-1 |  | qFR | 1.6121e-07 | mbar l s-1 |  | qFR | 3.4531e-08 | mbar l s-1
Segnale residuo prima delle misure di SFR / Residual signal prior SFR meas. | RFR | 5.98e-09 | mbar l s-1 |  | RFR | 3.92e-10 | mbar l s-1 |  | RFR | 8.26e-10 | mbar l s-1 |  | RFR | 5.98e-09 | mbar l s-1
Segnale del LD / Signal given by the calibrated leak | SFR | 4.78e-08 | mbar l s-1 |  | SFR | 4.76e-08 | mbar l s-1 |  | SFR | 1.95e-07 | mbar l s-1 |  | SFR | 4.78e-08 | mbar l s-1
Min. dev. segnale (=2x amp. segn. residuo) / Smallest read. signal dev. (= 2 x ampl. of RFR noise) | Sm | 6e-11 | mbar l s-1 |  | Sm | 1.8e-11 | mbar l s-1 |  | Sm | 1.2e-11 | mbar l s-1 |  | Sm | 6e-11 | mbar l s-1
Tempo di attesa stabiliz. segnale / Time to achieve stabilised leak signal | 3t | 2300 | sec |  | 3t | 2100 | sec |  | 3t | 2700 | sec |  | 3t | 2300 | sec
SENSIBILITA' DEL TEST / Sensitivity of the leak test | qGm | 4.9542e-11 | mbar l s-1 |  | qGm | 1.4477e-11 | mbar l s-1 |  | qGm | 9.9628e-12 | mbar l s-1 |  | qGm | 4.9542e-11 | mbar l s-1
Condizioni del test / Leak test conditions
Pressione del sistema / System pressure | P | 5.1e-05 | mbar |  | P |  | mbar |  | P |  | mbar |  | P | 5.1e-05 | mbar
Segnale residuo del cercafughe ad inizio test / Residual signal prior to SF measurement | RF | 5.97e-09 | mbar l s-1 |  | RF | 4.26e-10 | mbar l s-1 |  | RF | 8.91e-10 | mbar l s-1 |  | RF | 5.98e-09 | mbar l s-1
Segnale del LD a fine test / Signal given by the leak after 30 min. (>3t) | SF | 5.95e-09 | mbar l s-1 |  | SF | 5.03e-10 | mbar l s-1 |  | SF | 9.12e-10 | mbar l s-1 |  | SF | 5.91e-09 | mbar l s-1
CALCOLO DELLA FUGA / Leak evaluation | qG | <1.0E-09 | mbar l s-1 |  | qG | 6.1929e-11 | mbar l s-1 |  | qG | 1.7435e-11 | mbar l s-1 |  | qG | <1.0E-09 | mbar l s-1
VALORE DI RIFERIMENTO / REF. VALUE (MAX) |  | 1e-09 | mbar l s-1 | at 26 bar |  | 1e-10 | mbar l s-1 | at 26 bar |  | 1e-05 | mbar l s-1 | at 26 bar |  | 1e-09 | mbar l s-1 | at 5 bar
CONFORMANCE |  | YES |  |  |  | YES |  |  |  | YES |  |  |  | YES
Doc. di riferimento / Ref. documents
CERN contract number: F302/LHC/LHC
CERN technical spec.: LHC MMS-98-198 rev.2
Leak test procedure (Ref. N°, Revision): 780RM09442 rev.0
Strumentazione / Test equipment | on vessel |  |  |  | on c.b.t. lines |  |  |  | on heat exchanger line |  |  |  | on vessel
Helium Mass Spectrometer type: | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260 |  |  |  | PFEIFFER HLT 260
Pressure gauge type: | full range compact PFEIFFER PKR 251 | [11 empty] | full range compact PFEIFFER PKR 251
Pumping group: | turbo pump LEYBOLD PT 360 lt/s | [11 empty] | turbo pump LEYBOLD PT 360 lt/s
| rotary vane pump PFEIFFER DUO 65 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 20 m3/h |  |  |  | rotary vane pump PFEIFFER DUO 65 m3/h
Prepared by: Name / Date | Note / Remarks
Gnassi M. - Piu S. - Caserza B. 15/07/2006
Approved by: Name / Date
F. Terzi 15/07/2006
P. Gagliardi 15/07/2006

## Collegamenti e Dati
Magnete N. | 2362
Temperatura °C | 27.5
Calibrazione fuga camera
Valore inizio | 5.98e-09
Valore fine | 4.78e-08
Minimo segnale | 6e-11
Tempo di risposta | 2300
Calibrazione fuga linea "X"
Valore inizio | 8.26e-10
Valore fine | 1.95e-07
Minimo segnale | 1.2e-11
Tempo di risposta | 2700
Calibrazione fuga linea V1 e V2
Valore inizio | 3.92e-10
Valore fine | 4.76e-08
Minimo segnale | 1.8e-11
Tempo di risposta | 2100
Test del magnete 26 bar
Pressione nella camera | 5.1e-05
Valore cercafughe inizio test | 5.97e-09
Valore cercafughe fine test 30 min | 5.95e-09
Test V1 V2
Pressione nella camera
Valore cercafughe inizio test | 4.26e-10
Valore cercafughe fine test 30 min | 5.03e-10
Test della linea "X"
Pressione nella camera
Valore cercafughe inizio test | 8.91e-10
Valore cercafughe fine test 30 min | 9.12e-10
Test della linea "X" 5 bar
5.1e-05
Valore cercafughe inizio test | 5.98e-09
Valore cercafughe fine test 30 min | 5.91e-09

## Dati Fughe Calibrate
Calibrated leak data correction |  |  | Calibrated leak id. nr.
|  |  | 4011007195 | 4011007225 | FC03000505
qnom | nominal value of calibrated leak / val. nom. fuga cal. |  | 3e-08 | 3.3e-08 | 1.4e-07
c | correction factor for ageing / fatt. corr. per età | % | 0.0014 | 0.0015 | 0.02
y | current year / anno corrente |  | 2006 | 2006 | 2006
y0 | last calibration year / ultima calibrazione |  | 2002 | 2002 | 2003
k | correction factor for test temperature / fatt. corr. per temperatura | 1/Kelvin | 0.035 | 0.035 | 0.03
T | test temperature / temp. di test |  | 27.5 | 27.5 | 27.5
T0 | last calibration temperature / temp. ultima calibrazione | °C | 23 | 23 | 20
Formula
qFR = qnom * (1-c*(y-y0)) * (1+k*(T-T0)) |  |  | 3.4531e-08 | 3.7968e-08 | 1.6121e-07
=caselle da compilare