Jacques Billan

##### Instrumentation for the measurement of the cold bore geometry

and magnetic axis

### History

#### Demonstration of the laser tracker possibilities		june 1998

Study of magnetic axis measurement with an array of fixed coils in AC field

Development of a mole measuring cold bore tube axis and magnetic axis

Development of two techniques : 	self centering 	purely mechanical

optic		LED + CCD camera

optic mole selected		1 prototype in construction			test in January

10 (16) series, order in preparation

by-product : purely geometric mole	self-centering

1 twin set prototype in test 			since september 2000

8 twin sets series ordered			march 2001

The equipment

## The global mole

Reading with a CCD camera of the reflection of four LED’s to measure the position of the mole w.r.t. the cold bore tube

in the same cross-section as the reflector and the centre of the coils

Motorized levelling

## Processing of the CCD camera image

Image 768 x 538 pixels  (350 kpixels)

Camera tilted by 45° to optimise the display inside the rectangular image

# Spots		reflection of the LEDs on the cold bore tube wall		measure the relative distance between

# the mole and the tube at four points

direct vision of the LEDs						“finger print” of the mole

other reflections ( suppressed by soft)

Image processing

Determination of two circles		inner - intrisic to the mole	related to the corner cube centre

outer – centered on the tube

Calculation of				excentration of the mole at four points		distance corner cube centre – tube axis

diameter of the tube in two directions

The simplified geometric mole

Designed for a rapid measurement of the cold bore tube axis at different stages of the fabrication

Verticalization assured by the pendulum effect of the tunsten counter-weight (no motorization)

Total weight in equilibrium around the two roller balls

The corner cube reflector centered between the four centering pads

In the same cross-section (no paralax error)

No weight on the centering pads

Questions and answers

- At which stage of the magnets construction do you plan to perform this measurement?

This system needs a weak modulated AC current for working. We superimpose a fixed DC current to this AC current. This approach allows to measure the magnet at any stage of the construction. In particular, we can measure the magnet when it is finished with its protection diodes. We already measured a quadrupole in this configuration and we plan to measure a dipole next year.

- Is the residual field dangerous for your measurements?

No, the residual field is not important when an AC current is used. Furthermore, the iron yoke is sufficiently far for not influencing the measurement.

# Text boxes

###### The coil assembly

## Four coils

Four LEDs

One corner cube

centred in the mole

In the same mole cross-section

The coils

100mm long

600 turns

0.6m2