## Geometrical data
LHC Part Id
HCLQXA_001_FL000007
Operation
|  | Laser Tracker Nr.: |  | 939 |  | Used AC Mole Calibration Values
|  | AC Mole Number |  |  |  | Optic X | Optic Z | Magn. X | Magn. Z
GEO Mole Number
|  | Operator |  | SIMCO |  | The AC Mole data Magnetic and Mechanic are added using a bestfit on the fiducials
|  | Date of treatment |  | 2005-11-09
|  | Date of measurement |  | 2005-11-04
|  | RMS Bundle [mm] |  | 0.039
Coldmass Temp. [°C]
|  | Best fit done on |  | GeoMole shape
- Coldmass Reference Points (expressed in RST Coordinate system)
|  | CM references |  |  |  | INFO | Differences to CM references FERMILAB
| NIP side |  | CM reference points |  |  | IP side |  | CM reference points
|  |  | X | Y | Z |  |  | DX | DY | DZ
|  | Upper | 144.832 | -1056.49 | 39.784 |  | Upper | 0.0012 | 0.29698 | 0.76706
|  | Lower | 144.968 | -1056.491 | -35.213 |  | Lower | 0.00258 | 0.34932 | 0.93882
| IP side |  | CM reference points |  |  | NIP side |  | CM reference points
|  |  | X | Y | Z |  |  | DX | DY | DZ
|  | Upper | -142.065 | 6679.137 | 24.157 |  | Upper | -0.12726 | -0.30896 | 1.0811
|  | Lower | -141.926 | 6677.604 | -50.867 |  | Lower | -0.1178 | -0.3205 | 1.0836
- Extremity flanges (Values expressed in RST Coordinate system)
COLD BORE TUBE EXTREMITIES
| NIP side | CBT FLANGE
Center (Outer diameter)
|  | DX | DY | DZ | Tol(Y)
|  | -0.418 | -6.39 | -0.175 | ± 1.0
|  | DR |  | Tol(R)
|  | 0.45315 |  | ± 1.4
| IP side | CBT FLANGE
Center (Outer diameter)
|  | DX | DY | DZ | Tol(Y)
|  | 0.201 | -9.329 | 0.641 | ± 1.0
|  | DR (mm) |  | Tol(R)
|  | 0.67178 |  | ± 1.4
ANGLES OF COLD BORE TUBE EXTREMITIES W.R.T. GA
| NIP side | Lead End side
CBT FLANGE
|  | Angle/Cmeas (mrad) |  | Tol(gV1c)
|  | 2.3 |  | ±4.5
| IP side | Non Lead End side
CBT FLANGE
|  | Angle/Cmeas (mrad) |  | Tol(gV1c)
|  | 2.4 |  | ±4.5

## RST parameters
LHC part ID
HCLQXA_001_FL000007
- Fiducials RST parameters
| Id | R | S | T | RMS R | RMS S | RMS T | Total RMS | INFO | DR | DS | DT
| NIP_A | 424.44 | 822.635 | 537.992 | 0.009 | 0.006 | 0.006 | 0.012 | NIP_A | -0.19 | -0.145 | 0.952
| NIP_B | 324.3 | 1342.365 | 572.835 | 0.008 | 0.004 | 0.042 | 0.043 | NIP_B | -0.19 | -0.145 | 0.935
| NIP_C | -325.17 | 1339.924 | 574.838 | 0.021 | 0.008 | 0.001 | 0.022 | NIP_C | -0.18 | -0.146 | 0.918
| NIP_D | -424.205 | 819.212 | 540.496 | 0.011 | 0.015 | 0.001 | 0.019 | NIP_D | -0.135 | -0.138 | 0.956
| IP_A | 425.7 | 5027.477 | 537.318 | 0.021 | 0.006 | 0.007 | 0.023 | IP_A | -0.28 | -0.323 | 0.998
| IP_B | 325.734 | 5537.2 | 572.555 | 0.008 | 0.013 | 0.001 | 0.016 | IP_B | -0.306 | -0.37 | 0.895
| IP_C | -322.667 | 5537.58 | 574.527 | 0.004 | 0.004 | 0.007 | 0.009 | IP_C | -0.287 | -0.38 | 0.987
| IP_D | -423.214 | 5028.573 | 540.002 | 0.033 | 0.009 | 0.011 | 0.036 | IP_D | -0.234 | -0.337 | 0.942

## Cartography
LHC part ID
HCLQXA_001_FL000007
- Lead End side
| Point | X (mm) |  |  |  | Y (mm) |  |  |  | Z (mm) |  |  |  | Delta Radial (mm)
|  | Measured | Theo | Delta | Tolerance | Measured | Theo | Delta | Tolerance | Measured | Theo | Delta | Tolerance | Delta | Tolerance
| EE | -389.027 | -389.6 | 0.573 | ±8 | -1280.838 | -1273.65 | -7.188 | 9/-6 | 68.781 | 75 | -6.219 | ±8 | 6.2453 | ±11.3
| LD1 | -240.668 | -240 | -0.668 | ±8 | -1284.639 | -1273.65 | -10.989 | 9/-6 | 279.475 | 277.5 | 1.975 | ±8 | 2.0849 | ±11.3
| L | -1.848 | 0 | -1.848 | ±5 | -1280.634 | -1273.65 | -6.984 | ±10 | 380.472 | 376.9 | 3.572 | ±5 | 4.0217 | ±7.1
| XBt | 239.014 | 240 | -0.986 | ±5 | -1278.108 | -1273.65 | -4.458 | 15/-12 | 278.005 | 277.5 | 0.505 | ±5 | 1.1078 | ±7.1
| FF | 386.103 | 389.6 | -3.497 | ±8 | -1278.799 | -1273.65 | -5.149 | 9/-6 | 68.29 | 75 | -6.71 | ±8 | 7.5666 | ±11.3
| M1 (1) |  |  |  | ±5 |  |  |  | ±20 |  |  |  | ±5 |  | ±7.1
| M2 (1) |  |  |  | ±5 |  |  |  | ±20 |  |  |  | ±5 |  | ±7.1
| M4 (1) |  |  |  | ±5 |  |  |  | ±20 |  |  |  | ±5 |  | ±7.1
| V | -0.418 | 0 | -0.418 | ±1 | -1057.54 | -1051.15 | -6.39 | ±1 | -0.175 | 0 | -0.175 | ±1 | 0.45315 | ±1.4
| W' | -4.565 | 0 | -4.565 | ±5 | -931.1877 | -923.65 | -7.5377 | ±10 | 79.279 | 75 | 4.279 | ±5 | 6.2569 | ±7.1
- Non Lead End side
| Point | X (mm) |  |  |  | Y (mm) |  |  |  | Z (mm) |  |  |  | Delta Radial (mm)
|  | Measured | Theo | Delta | Tolerance | Measured | Theo | Delta | Tolerance | Measured | Theo | Delta | Tolerance | Delta | Tolerance
EE
LD1
L
XBt
FF
M1 (2)
M2 (2)
M4 (2)
| V | 0.201 | 0 | 0.201 | ±1 | 6694.521 | 6703.85 | -9.329 | ±1 | 0.641 | 0 | 0.641 | ±1 | 0.67178 | ±1.4
| W' | -0.301 | 0 | -0.301 | ±5 | 6567.2567 | 6576.35 | -9.0933 | ±10 | 75.903 | 75 | 0.903 | ±5 | 0.95185 | ±7.1

## Mechanical shape data
LHC Part Id
HCLQXA_001_FL000007
V1
Point Id | dx | y | dz
OL2 | 0.30988 | -881.0791 | 0.15779
OL3 | 0.6055 | -703.2811 | 0.39529
OC23 | 0.51952 | -663.2471 | 0.5214
OL4 | 0.78735 | -331.3071 | 0.58233
OC22 | 0.75313 | -322.4041 | 0.68835
OC21 | 0.57154 | 26.5859 | 0.58932
OL5 | 0.60028 | 37.3059 | 0.47835
OC20 | 0.43304 | 372.2329 | 0.52129
OL6 | 0.48065 | 387.7929 | 0.41833
OC19 | 0.14712 | 734.6639 | 0.4283
OL7 | 0.1591 | 735.2909 | 0.2723
OC18 | -0.14638 | 1079.8589 | 0.25726
OL8 | -0.13658 | 1088.1569 | 0.08229
OC17 | -0.35581 | 1422.4209 | -0.02078
OL9 | -0.32317 | 1437.0179 | -0.20074
OC16 | -0.41368 | 1783.0499 | -0.28978
OL10 | -0.33276 | 1786.2559 | -0.47077
OC15 | -0.44815 | 2126.9899 | -0.41382
OL11 | -0.34848 | 2140.4419 | -0.56978
OC14 | -0.45698 | 2485.7619 | -0.44482
OL12 | -0.33321 | 2495.1379 | -0.62679
OC13 | -0.25528 | 2823.2029 | -0.57588
OL13 | -0.1707 | 2840.3579 | -0.75783
OC12 | -0.05403 | 3178.5059 | -0.66688
OL14 | 0.05659 | 3194.0569 | -0.81684
OC11 | 0.04035 | 3528.7369 | -0.68591
OL15 | 0.17105 | 3541.1969 | -0.82987
OC10 | -0.00452 | 3889.5219 | -0.7119
OL16 | 0.11243 | 3891.3029 | -0.8419
OC9 | -0.08416 | 4240.5679 | -0.62892
OL17 | 0.00481 | 4241.8069 | -0.76192
OC8 | -0.23363 | 4584.4789 | -0.60796
OL18 | -0.16882 | 4592.2459 | -0.70694
OC7 | -0.2904 | 4940.9549 | -0.40997
OL19 | -0.21356 | 4947.3509 | -0.53295
OC6 | -0.20889 | 5285.8469 | -0.16901
OL20 | -0.16219 | 5298.0309 | -0.32197
OC5 | -0.02167 | 5642.8459 | 0.13699
OL21 | 0.05713 | 5650.7399 | -0.00699
OC4 | 0.24791 | 5984.7089 | 0.49894
OL22 | 0.3077 | 5993.4819 | 0.33497
OL23 | 0.44497 | 6266.8239 | 0.64473
OC3 | 0.36714 | 6341.3369 | 0.82994
OC2 | 0.19054 | 6528.0709 | 0.82046
average | 0.050106 |  | -0.10028
max | 0.78735 |  | 0.82994
min | -0.45698 |  | -0.8419
rms | 0.34506 |  | 0.5411

## Magnetic shape data
LHC Part Id
HCLQXA_001_FL000007
V1
Point Id | dx | y | dz | ANGLE | Polarity | Mounting Angle
OC21 | -0.17446 | 26.5859 | 0.83832
OL5 | -0.39472 | 37.3059 | 0.66235
OC20 | -0.17596 | 372.2329 | 0.81029
OL6 | -0.38335 | 387.7929 | 0.66833
OC19 | -0.17688 | 734.6639 | 0.8313
OL7 | -0.3949 | 735.2909 | 0.6603
OC18 | -0.24838 | 1079.8589 | 0.83926
OL8 | -0.48558 | 1088.1569 | 0.65029
OC17 | -0.13581 | 1422.4209 | 0.71922
OL9 | -0.36617 | 1437.0179 | 0.54126
OC16 | -0.14768 | 1783.0499 | 0.58722
OL10 | -0.35976 | 1786.2559 | 0.42723
OC15 | -0.12315 | 2126.9899 | 0.46018
OL11 | -0.31448 | 2140.4419 | 0.32222
OC14 | -0.12498 | 2485.7619 | 0.36018
OL12 | -0.29821 | 2495.1379 | 0.19721
OC13 | -0.11528 | 2823.2029 | 0.22212
OL13 | -0.3197 | 2840.3579 | 0.03517
OC12 | -0.11203 | 3178.5059 | 0.10512
OL14 | -0.28341 | 3194.0569 | -0.05084
OC11 | -0.08065 | 3528.7369 | 0.05509
OL15 | -0.24595 | 3541.1969 | -0.11487
OC10 | -0.02452 | 3889.5219 | 0.0281
OL16 | -0.20657 | 3891.3029 | -0.1439
OC9 | -0.01216 | 4240.5679 | 0.16008
OL17 | -0.22219 | 4241.8069 | -0.02492
OC8 | -0.00963 | 4584.4789 | 0.28504
OL18 | -0.23682 | 4592.2459 | 0.14906
OC7 | -0.0464 | 4940.9549 | 0.47703
OL19 | -0.24256 | 4947.3509 | 0.34205
OC6 | 0.01011 | 5285.8469 | 0.67199
OL20 | -0.21319 | 5298.0309 | 0.53403
OC5 | -0.06367 | 5642.8459 | 0.84499
OL21 | -0.23287 | 5650.7399 | 0.70901
OC4 | -0.09909 | 5984.7089 | 1.0029
OL22 | -0.2763 | 5993.4819 | 0.84197
OC3 | -0.12386 | 6341.3369 | 1.2289

## Shape data Geo Mole
LHC Part Id
HCLQXA_001_FL000007
V1
Point Id | dx | y | dz
Centre_Conn_V | -0.418 | -1057.54 | -0.175
NIP2 | -0.109 | -1022.154 | -0.042
NIP3 | 0.844 | -355.691 | 0.749
IP12 | 0.655 | -166.61 | 0.674
NIP4 | 0.559 | 420.706 | 0.627
IP11 | 0.268 | 536.66 | 0.604
NIP5 | -0.081 | 1092.398 | 0.254
IP10 | -0.422 | 1184.942 | 0.252
IP9 | -0.55 | 1822.434 | -0.165
NIP6 | -0.369 | 1840.677 | -0.303
NIP7 | -0.318 | 2449.135 | -0.556
IP8 | -0.479 | 2486.319 | -0.357
NIP8 | -0.062 | 2930.552 | -0.643
IP7 | -0.123 | 3145.223 | -0.569
NIP9 | 0.131 | 3447.285 | -0.754
IP6 | -0.003 | 3834.954 | -0.563
NIP10 | 0.173 | 3862.747 | -0.696
NIP11 | -0.109 | 4459.658 | -0.624
IP5 | -0.266 | 4494.254 | -0.483
NIP12 | -0.123 | 5047.082 | -0.302
IP4 | -0.21 | 5105.762 | -0.07
IP3 | -0.025 | 5739.048 | 0.392
NIP13 | 0.165 | 5751.031 | 0.182
IP2 | 0.312 | 6474.102 | 1.048
NIP14 | 0.37 | 6499.18 | 0.868
Centre_Lyre_V | 0.201 | 6694.521 | 0.641
average | 0.00042308 |  | -0.00042308
max | 0.844 |  | 1.048
min | -0.55 |  | -0.754
rms | 0.35646 |  | 0.5573

## cold_parameters
Low-Beta Quadropole Type: |  |  | Q1
Positon in LHC: | [5 empty] | of installation in the LHC Tunnel
Number of thermal cycles
done befor Fiducialisation: |  |  | 0 | (In case of more than 3 thermal cycles please give number 3)
Main magnet type: |  |  | MQXA
Warm in Fidu co-ordinates system | [6 empty] | Cold in director beam B1 co-ordinates system
Point ID | R_Fidu | S_Fidu | T_Fidu | R_Depla | S_Depla | T_Depla | Fiducials | R | S | T
NIP_A | 424.44 | 822.635 | 537.992 | 0 |  | -0.45 |  |  |  | 0.53844
NIP_B | 324.3 | 1342.365 | 572.835 | 0 |  | -0.45 |  |  |  | 0.57329
NIP_C | -325.17 | 1339.924 | 574.838 | 0 |  | -0.45 |  |  |  | 0.57529
NIP_D | -424.205 | 819.212 | 540.496 | 0 |  | -0.45 |  |  |  | 0.54095
IP_A | 425.7 | 5027.477 | 537.318 | 0 |  | -0.45 |  |  |  | 0.53777
IP_B | 325.734 | 5537.2 | 572.555 | 0 |  | -0.45 |  |  |  | 0.573
IP_C | -322.667 | 5537.58 | 574.527 | 0 |  | -0.45 |  |  |  | 0.57498 | Geo-Ref-Angle:
IP_D | -423.214 | 5028.573 | 540.002 | 0 |  | -0.45 |  |  |  | 0.54045
Valid for all Low-Beta magnets (Q1 to Q3):
1.Case: | the magnet is installed in the same direction than the fiducialisation Coordinate system.
- All Low-Beta magnets installed on the left side of the IP
2.Case: | the magnet is installed in the opposite direction than the fiducialisation Coordinate system.
- All Low-Beta magnets installed on the right side of the IP
Deplacement values:
| R_depla: | 0
| S_depla: | 0.34 | for Q1 in 1.Case
|  | 11 | for Q2 in 1.Case
|  | 12.36 | for Q3 in 1.Case
|  | 6370.34 | for Q1 in 2.Case
|  | 5511 | for Q2 in 2.Case
|  | 6382.36 | for Q3 in 2.Case
| T_depla: | -0.45 | in case of 0 thermal cycles befor Fiducialisation
|  | -0.15 | in case of 1 thermal cycles befor Fiducialisation
|  | -0.1 | in case of 2 thermal cycles befor Fiducialisation
|  | 0 | in case of 3 or more thermal cycles befor Fiducialisation

## parameters_insertion
| Magnet ID | Class | Pos. | Point | Code | R | S | T | BR | BL | BV | BT | Angle
LHC | HCLQXA_001_FL000007 | MQXA | 0 | E | 1
LHC | HCLQXA_001_FL000007 | MQXA | 0 | A | 1
LHC | HCLQXA_001_FL000007 | MQXA | 0 | D | 1
LHC | HCLQXA_001_FL000007 | MQXA | 0 | T | 1
LHC | HCLQXA_001_FL000007 | MQXA | 0 | B | 1
LHC | HCLQXA_001_FL000007 | MQXA | 0 | S | 1

## Installation_deviations
FID ID | Point ID | calc. coordinates at cold |  |  | theor. coordinates at cold |  |  | Diff. theoretical to real
|  | R | S | T | R(theo) | S(theo) | T(theo) | delta_R | delta_S | delta_T | dR_bump | dS_bump | dT_bump
| NIP_A |  |  | 0.53844
| NIP_B |  |  | 0.57329
| NIP_C |  |  | 0.57529
| NIP_D |  |  | 0.54095
| IP_A |  |  | 0.53777
| IP_B |  |  | 0.573
| IP_C |  |  | 0.57498
| IP_D |  |  | 0.54045
Values of theoretical coordiates for the different LowBeta Typs in all possible Positions:
| Q1 | [5 empty] | Q3
| Main magnet type MQXA in 1L = 1. Case |  |  | Main magnet type MQXA in 1R = 2. Case |  |  | Main magnet type MQXA in 3L = 1. Case |  |  | Main magnet type MQXA in 3R = 2. Case
A | 0.425 | 0.8299 | 0.54 | 0.325 | 0.8246 | 0.575 | 0.425 | 0.8484 | 0.54 | 0.325 | 0.8161 | 0.575
E | 0.325 | 1.3454 | 0.575 | 0.425 | 1.3301 | 0.54 | 0.325 | 1.3539 | 0.575 | 0.425 | 1.3216 | 0.54
B | 0.425 | 5.0399 | 0.54 | 0.325 | 5.0246 | 0.575 | 0.425 | 5.0484 | 0.54 | 0.325 | 5.0161 | 0.575
S | 0.325 | 5.5454 | 0.575 | 0.425 | 5.5401 | 0.54 | 0.325 | 5.5539 | 0.575 | 0.425 | 5.5216 | 0.54
T | -0.325 | 5.5454 | 0.575 | 0.425 | 5.5401 | 0.54 | -0.325 | 5.5539 | 0.575 | 0.425 | 5.5216 | 0.54
D | -0.425 | 5.0399 | 0.54 | 0.325 | 5.0246 | 0.575 | -0.425 | 5.0484 | 0.54 | 0.325 | 5.0161 | 0.575
Q2
| Main magnet type MQXB in B2L = 1. Case |  |  | Main magnet type MQXB in B2R = 2. Case
A | 0.425 | 1.2148 | 0.54 | 0.325 | -5.3003 | 0.575 | Q1 | are always in Position 1L or 1R and of the type MQXA!
E | 0.325 | 1.7203 | 0.575 | 0.425 | -4.7948 | 0.54 | Q2 | are always in Position B2L or B2R and of the type MQXB!
C | 0.425 | 5.7548 | 0.54 | 0.325 | -0.7603 | 0.575 | Q3 | are always in Position 3L or 3R and of the type MQXA!
M | 0.325 | 6.2603 | 0.575 | 0.425 | -0.2548 | 0.54
B | 0.425 | 10.2948 | 0.54 | 0.325 | 3.7797 | 0.575
S | 0.325 | 10.8003 | 0.575 | 0.425 | 4.2852 | 0.54
T | -0.325 | 10.8003 | 0.575 | -0.425 | 4.2852 | 0.54
D | -0.425 | 10.2948 | 0.54 | -0.325 | 3.7797 | 0.575
Note on L2: ce calcul est fait avec l ' hypothese que les bump sont les memes en Entrée qu'en sortie
Note on L3: le bump en R est entré avec la convention inverse de signe du R
Note on M2: ce calcul est fait avec l ' hypothese que les bump sont les memes en Entrée qu'en sortie
Note on N2: ce calcul est fait avec l ' hypothese que les bump sont les memes en Entrée qu'en sortie