## N-0213I
ELECTRICAL TESTS OF INNER LAYERS / innere Lagen
LHC/MMS TRACEABILITY INFORMATION
MAIN CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | F303
SUB CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | ----
[5 empty] | SPECIFICATION: |  | LHC-MMS/98-198 Rev. 2.0
COMPONENT NAME : |  |  | INNER LAYER |  | BNN SERIAL NUMBER : |  | N-0213I
CERN PART NUMBER : |  |  | HCMB__A011 |  | SERIAL NUMBER : |  | ----
BNN DOC ID NUMBER : |  |  | CF0013_B__N-0213I |  | BNN ORDER NUMBER : |  | 91300
CERN INSPECTION & TEST PLAN NUMBER : |  |  | ITP Step 3a, 3b, 3c, 3d, 6 |  | TEST NUMBER : |  | 3 and 6
NON CONFORMITY : |  |  | <No> |  | NON-CONFORMITY NUMBER :
NAME OF TEST OPERATOR : |  |  | Mahler |  | WORKING PLACE : |  | ON WOODEN TABLE
NAME OF RESPONSIBLE : |  |  | Zehner |  | BBP-SDA-CODE : |  | -----
DATE OF TEST : |  | (Starting) | 2003-01-31 |  | TIME :
BNN INT. REG. NO : |  | 'file:///tmp/in/input.xls'#$N-0213I |  |  | DATE OF REPORT :
Cable "I" Number : |  |  | HCMB__A046-01B10097A
REQUIREMENTS OF TEST OBJECT/SPECIFICATION: |  |  |  | CERN: BNN: | LHC-MMS/98-198 Rev. 2.0 Annex B10 TABLE A Arbeitsanweisung CA0012_0
| R[dc] of the INNER LAYER |  | Date of Test: | 2003-01-31 | Equipment: | Kepco Power Supply [BM 00125], Testo 965 [BM 00116], HP34401A [BM 00122]
I = | 1 | A |  |  |  |  | ITP step 3a
Temp.* : | 23.1 | °C
[5 empty] | 496 | mV
| INNER LAYER |  |  |  | 496 | mW | SIGN:
| INNER LAYER [R20] |  |  |  | 490.2251 | mW
| INDUCTANCE |  | Date of Test: | 2003-01-31 | Equipment: | Kepco Power Supply [BM 00125], Solartron Gain Phase Analyser [BM 00124]
| [Value from gain phase analyser] | [5 empty] | ITP step 3b
[5 empty] | L [mH] | Q
| INNER LAYER |  |  | 10Hz | 2.2 | 0.288
|  |  |  | 100Hz | 1.86 | 2.3
SIGN: |  |  |  | 1000Hz | 1.82 | 13.8
| INSULATION RESISTANCE |  | Date of Test: | 2003-01-31 | Equipment: | Megger BM21 [BM 00138]
| [Copper wedges] |  | [V = 500V; 30"] |  |  |  | ITP step 3c
| Left side |  |  |  | 60500 | MW | SIGN:
| Right side |  |  |  | 49600 | MW
| DISCHARGE TEST |  | Date of Test: | 2003-01-31 | Equipment: | Seitz Impulse Tester [BM 00119]
| [Without Strain] |  | [V = 120V/turn; 1.8KV; 10puls] |  |  |  | ITP step 3d
|  |  |  | T | 388 | µS [Reference layer]
| INNER LAYER |  |  | T | 388 | µS SIGN:
| [With Strain] |  | Date of Test: | 2003-01-31
| [30 and 80N/mm² straight part |  | [V = 120V/turn; 1.8KV; 10puls] |  |  |  | ITP step 6
and the stress fades down in the ends] |  |  |  |  | 30N/mm² |  | 80N/mm²
SIGN: |  |  |  | C.side T | 388 | µS | 388
| INNER LAYER |  |  | N.C. side T | 388 | µS | 388
|  | test by |  | supervisor |  | approval
name |  | Mahler |  | Zehner |  | Zehner
company |  | BNM |  | BNN |  | BNN
date |  | 2003-01-31 |  | 2003-02-14 |  | 2003-02-14
signature