## N-3230I
ELECTRICAL TESTS OF INNER LAYERS / innere Lagen
LHC/MMS TRACEABILITY INFORMATION
MAIN CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | F489
SUB CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | ----
[5 empty] | SPECIFICATION: |  | LHC-MMS/98-198 Rev. 2.0
COMPONENT NAME : |  |  | INNER LAYER |  | BNN SERIAL NUMBER : |  | N-3230I
CERN PART NUMBER : |  |  | HCMB__A011 |  | SERIAL NUMBER : |  | ----
BNN DOC ID NUMBER : |  |  | CF0013_C__N-3230I |  | BNN ORDER NUMBER : |  | N.911004
CERN INSPECTION & TEST PLAN NUMBER : |  |  | ITP Step 3a, 3b, 3c, 3d, 6 |  | TEST NUMBER : |  | 3 and 6
NON CONFORMITY : |  |  | <No> |  | NON-CONFORMITY NUMBER :
NAME OF TEST OPERATOR : |  |  | see ITP steps |  | WORKING PLACE : |  | ON WOODEN TABLE
NAME OF RESPONSIBLE : |  |  | Zehner |  | BBP-SDA-CODE : |  | -----
DATE OF TEST : |  | (Starting) | 11.09.03 |  | TIME : |  | -----
BNN INT. REG. NO : |  | 'file:///tmp/in/input.xls'#$N-3230I |  |  | DATE OF REPORT : |  | 11.09.03
Cable "I" Number : |  |  | HCMB__A046-01B10313B
REQUIREMENTS OF TEST OBJECT/SPECIFICATION: |  |  |  | CERN: BNN: | LHC-MMS/98-198 Rev. 2.0 Annex B10 TABLE A Arbeitsanweisung CA0012_0
| R[dc] of the INNER LAYER |  | Date of Test: | 11.09.03 | Equipment: | Kepco Power Supply [BM 125], Testo 965 [BM 116], HP34401A [BM 122]
I = | 1 | A |  |  |  |  | ITP step 3a
Temp.* : | 22.7 | °C
[5 empty] | 495.23 | mV
| INNER LAYER |  |  |  | 495.23 | mW | NAME:
| INNER LAYER [R20] |  |  |  | 490.2005 | mW | Schmitt
| INDUCTANCE |  | Date of Test: | 11.09.03 | Equipment: | Kepco Power Supply [BM 125], Solartron Gain Phase Analyser [BM 124]
| [Value from gain phase analyser] | [5 empty] | ITP step 3b
[5 empty] | L [mH] | Q
| INNER LAYER |  |  | 10Hz | 2.15 | 0.27796
|  |  |  | 100Hz | 1.85 | 2.27
NAME: | Schmitt |  |  | 1000Hz | 1.81 | 13.51
| INSULATION RESISTANCE |  | Date of Test: | 11.09.03 | Equipment: | Megger BM21 [BM 138]
| [Copper wedges] |  | [V = 500V; 30"] |  |  |  | ITP step 3c
| Left side |  |  |  | 36400 | MW | NAME:
| Right side |  |  |  | 41400 | MW | Schmitt
| DISCHARGE TEST |  | Date of Test: | 11.09.03 | Equipment: | Seitz Impulse Tester [BM 119]
| [Without Strain] |  | [V = 120V/turn; 1.8KV; 10puls] |  |  |  | ITP step 3d
|  |  |  | T | 388 | µS [Reference layer]
| INNER LAYER |  |  | T | 388 | µS NAME: | Schmitt
| [With Strain] |  | Date of Test: | 11.09.03
| [30 and 80N/mm² straight part |  | [V = 120V/turn; 1.8KV; 10puls] |  |  |  | ITP step 6
and the stress fades down in the ends] |  |  |  |  | 30N/mm² |  | 80N/mm²
NAME: | INNER |  |  | C.side T | 388 | µS | 388
Schmitt | LAYER |  |  | N.C. side T | 388 | µS | 388
|  | test by |  | supervisor |  | approval
name |  | Schmitt |  | Zehner |  | Zehner
company |  | BNM |  | BNN |  | BNN
date |  | 11.09.03 |  | 2003-09-26 |  | 2003-09-30
signature
Note on H7: Bei einer Wiederholungs-messung ist der Revisions-buchstabe (z.B. A) direkt an die BNN SERIAL NUMBER anzuhängen, z.B. N-0245IA)!