## N-3058E
ELECTRICAL TESTS OF OUTER LAYERS / äußere Lagen
LHC/MMS TRACEABILITY INFORMATION
MAIN CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | F303
SUB CONTRACTOR : |  |  |  |  | CONTRACT NUMBER : |  | ----
[5 empty] | SPECIFICATION: |  | LHC-MMS/98-198 Rev. 2.0
COMPONENT NAME |  |  | OUTER LAYER |  | BNN SERIAL NUMBER : |  | N-3058E
CERN PART NUMBER : |  |  | HCMB__A012 |  | Serial Number : |  | ----
BNN DOC ID NUMBER : |  |  | CF0014_C__N-3058E |  | BNN ORDER NUMBER : |  | 91300
CERN INSPECTION & TEST PLAN NUMBER : |  |  | ITP Step 3a, 3b, 3c, 3d, 6 |  | TEST NUMBER : |  | 3 and 6
NON CONFORMITY : |  |  | <No> |  | NON-CONFORMITY NUMBER :
NAME OF TEST OPERATOR : |  |  | see ITP steps |  | WORKING PLACE : |  | ON WOODEN TABLE
NAME OF RESPONSIBLE : |  |  | Zehner |  | BBP-SDA-CODE : |  | -----
DATE OF TEST : |  | (Starting) | 2003-05-27 |  | TIME : |  | -----
BNN INT. REG. NO : |  | 'file:///tmp/in/input.xls'#$N-3058E |  |  | DATE OF REPORT : |  | 2003-05-27
Cable "O" Number : |  |  | HCMB__A047-02K10601B
REQUIREMENTS OF TEST OBJECT/SPECIFICATION : |  |  |  | CERN: BNN: | LHC-MMS/98-198 Rev. 2.0 Annex B10 TABLE A Arbeitsanweisung CA0012_0
| R[dc] of the OUTER LAYER |  | Date of Test: | 27.05.003 | Equipment: | Kepco Power Supply [BM 00297], Testo 965 [BM 00116], HP34401A [BM 00294]
I = | 1 | A |  |  |  |  | ITP step 3a
Temp.* : | 23.7 | °C
[5 empty] | 1048 | mV
| OUTER LAYER |  |  |  | 1048 | mW | SIGN:
| OUTER LAYER [R20] |  |  |  | 1033.4694 | mW | Ortega
| INDUCTANCE |  | Date of Test: | 2003-05-27 | Equipment: | Kepco Power Supply [BM 00297], Solartron Gain Phase Analyser [BM 00298]
| [Value from gain phase analyser] | [5 empty] | ITP step 3b
[5 empty] | L [mH] | Q
| OUTER LAYER |  |  | 10Hz | 6.85 | 0.40389
|  |  |  | 100Hz | 6.63 | 3.83
SIGN: | Ortega |  |  | 1000Hz | 6.62 | 30.16
| INSULATION RESISTANCE |  | Date of Test: | 2003-05-27 | Equipment: | Megger BM21 [BM 00138]
| [Copper wedges] |  | [V = 500V; 30"] |  |  |  | ITP step 3c
| Left side |  |  |  | 71000 | MW | SIGN:
| Right side |  |  |  | 69000 | MW | Ortega
| DISCHARGE TEST |  | Date of Test: | 2003-05-27 | Equipment: | Seitz Impulse Tester [BM 00296]
| [Without Strain] |  | [V = 120V/turn; 3.0KV; 10puls] |  |  |  | ITP step 3d
|  |  |  | T | 704 | µS [Reference layer]
| OUTER LAYER |  |  | T | 704 | µS SIGN: | Ortega
| [With Strain] |  | Date of Test:
| [Whith Strain 30 and 80N/mm² |  | [V = 120V/turn; 3.0KV; 10puls] |  |  |  | ITP step 6
straight part and the stress fades down in the ends] |  |  |  |  | 30N/mm² |  | 80N/mm²
SIGN: | OUTER |  |  | C.side T | 704 | µS | 704
Ortega | LAYER |  |  | N.C. side T | 704 | µS | 704
|  | test by |  | supervisor |  | approval
name |  | Ortega |  | Zehner |  | Zehner
company |  | BNM |  | BNN |  | BNN
date |  | 2003-05-27 |  | 2003-06-03 |  | 2003-06-04
signature
Note on H7: Bei einer Wiederholungs-messung ist der Revisions-buchstabe (z.B. A) direkt an die BNN SERIAL NUMBER anzuhängen, z.B. N-0245EA)!