## Coils-Coll.
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 21.3397 | Magnet | 5.9586 | 5958.589 | 5982.5191 | 10 | 0.889 | 21.8 |  | 0.884 | 21.8
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 22.257 | Aperture1 | 2.9748 | 2974.774 | 2986.7209 | 100 | 3.08 | 20.5 |  | 3.07 | 20.5
[9 empty] | YT121 | 21.5295 | Aperture2 | 2.9838 | 2983.796 | 2995.7791 | 1000 | 0.657 | 6.19 |  | 0.737 | 6.68
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 22.3562 | Apt1 Pup | 1.4879 | 1487.914 | 1493.8896
CERN PART NUMBER : |  |  | HCMBAL_A00101 |  | Serial Number : |  | 000082bis |  | YT211 | 21.4933 | Apt1 Plw | 1.4868 | 1486.837 | 1492.8082 | 10 | 0.889 | 21.8 |  | 0.888 | 21.8
[9 empty] | YT212 | 22.2822 | Apt2 Pup | 1.4958 | 1495.767 | 1501.7741 | 100 | 3.06 | 20.5 |  | 3.07 | 20.4
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 21.4732 | Apt2 Plw | 1.488 | 1487.991 | 1493.9669 | 1000 | 0.637 | 6 |  | 0.785 | 6.8
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 22.3
[15 empty] | 10 | 0.893 | 43.7 |  | 0.89 | 43.7
NAME OF TEST OPERATOR: |  |  | JARDOT/GUTOWSKI | [11 empty] | 100 | 3.07 | 41.1 |  | 3.07 | 41
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.651 | 12.3 |  | 0.759 | 13.5
DATE OF TEST : |  |  | 2004-01-08 |  | TIME :
[15 empty] | 10 | 0.892 | 87.4
[15 empty] | 100 | 3.09 | 82.1
[15 empty] | 1000 | 0.697 | 25.5
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 19 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 33 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1487.914 | 1486.837 |  | 1495.767 | 1487.991
Poles [20°C] |  |  | 1493.8896 | 1492.8082 |  | 1501.7741 | 1493.9669
D1 & D2 |  |  | 2986.7209 | mW |  | 2995.7791 | mW
Magnet |  |  |  | 5982.5191 | mW
|  | Poles Nr. | JP 326 | JP 328 |  | JP 327 | JP 329
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 21.3397 | 21.5295 |  | 21.4933 | 21.4732
|  |  | 22.257 | 22.3562 |  | 22.2822 | 22.3
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 200 | nF |  | 200 | nF
D1 & D2 -----> Ground |  |  | 200 | nF |  | 210 | nF
Magnet -----> Ground |  |  |  | 330 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000082bis |  |  |  |  | Date: | 2004-01-08
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.8 | 21.8 |  | 21.8 | 21.8
|  | (Q) | 0.889 | 0.889 |  | 0.884 | 0.888
D1 & D2 |  | (L) | 43.7 | mH |  | 43.7 | mH
|  | (Q) | 0.893 |  |  | 0.89
Magnet |  |  | (L) | 87.4 | mH
|  |  | (Q) | 0.892
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 20.5 | 20.5 |  | 20.5 | 20.4
|  | (Q) | 3.08 | 3.06 |  | 3.07 | 3.07
D1 & D2 |  | (L) | 41.1 | mH |  | 41 | mH
|  | (Q) | 3.07 |  |  | 3.07
Magnet |  |  | (L) | 82.1 | mH
|  |  | (Q) | 3.09
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 6.19 | 6 |  | 6.68 | 6.8
|  | (Q) | 0.657 | 0.637 |  | 0.737 | 0.785
D1 & D2 |  | (L) | 12.3 | mH |  | 13.5 | mH
|  | (Q) | 0.651 |  |  | 0.759
Magnet |  |  | (L) | 25.5 | mH
|  |  | (Q) | 0.697
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000082bis |  |  |  |  | Date: | 2004-01-08
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 12.5 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 12.5 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 3.36 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 3.18 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 11.1 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 10.6 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.2 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.2 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.2 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.2 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.95 | 0.948 |  | 0.95 | 0.95
| Dipoles D1 & D2 [4KV] |  | 1.75 | mS |  | 1.75 | mS
| Magnet [4KV] |  |  | 2.62 | mS
| Quench Heaters |  | YT111 & 112 | YT121 & 122 |  | YT211 & 212 | YT221 & 222
| [V=1600KV] |  | 365 | 363 | µS | 365 | 360 | µS
| [C= 2µF] | Current | 37 | 36.5 | A | 36.5 | 37 | A
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION