## Coils-Coll.
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 21.9722 | Magnet | 6.0723 | 6072.331 | 6000.3271 | 10 | 0.873 | 21.9 |  | 0.873 | 21.9
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 22.0069 | Aperture1 | 3.0382 | 3038.161 | 3002.1354 | 100 | 3.08 | 20.6 |  | 3.08 | 20.6
[9 empty] | YT121 | 21.9506 | Aperture2 | 3.0341 | 3034.135 | 2998.1571 | 1000 | 0.673 | 6.29 |  | 0.752 | 6.83
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 22.0446 | Apt1 Pup | 1.5209 | 1520.947 | 1502.9121
CERN PART NUMBER : |  |  | HCMBAL_A00101 |  | Serial Number : |  | 000160 |  | YT211 | 21.8787 | Apt1 Plw | 1.5172 | 1517.181 | 1499.1907 | 10 | 0.872 | 21.8 |  | 0.874 | 21.8
[9 empty] | YT212 | 21.9069 | Apt2 Pup | 1.5197 | 1519.682 | 1501.6621 | 100 | 3.06 | 20.6 |  | 3.08 | 20.5
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 21.8678 | Apt2 Plw | 1.5144 | 1514.427 | 1496.4694 | 1000 | 0.657 | 6.17 |  | 0.787 | 6.85
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 21.9416
[15 empty] | 10 | 0.876 | 43.8 |  | 0.877 | 43.8
NAME OF TEST OPERATOR: |  |  | JARDOT | [11 empty] | 100 | 3.08 | 41.2 |  | 3.09 | 41.1
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.658 | 12.4 |  | 0.76 | 13.6
DATE OF TEST : |  |  | 13/9/2004 |  | TIME :
[15 empty] | 10 | 0.876 | 87.5
[15 empty] | 100 | 3.09 | 82.4
[15 empty] | 1000 | 0.704 | 26
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 23 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 47 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1520.947 | 1517.181 |  | 1519.682 | 1514.427
Poles [20°C] |  |  | 1502.9121 | 1499.1907 |  | 1501.6621 | 1496.4694
D1 & D2 |  |  | 3002.1354 | mW |  | 2998.1571 | mW
Magnet |  |  |  | 6000.3271 | mW
|  | Poles Nr. | JP 638 | JP 639 |  | JP 640 | JP 641
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 21.9722 | 21.9506 |  | 21.8787 | 21.8678
|  |  | 22.0069 | 22.0446 |  | 21.9069 | 21.9416
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 210 | nF |  | 210 | nF
D1 & D2 -----> Ground |  |  | 200 | nF |  | 200 | nF
Magnet -----> Ground |  |  |  | 330 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000160 |  |  |  |  | Date: | 13/9/2004
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.9 | 21.8 |  | 21.9 | 21.8
|  | (Q) | 0.873 | 0.872 |  | 0.873 | 0.874
D1 & D2 |  | (L) | 43.8 | mH |  | 43.8 | mH
|  | (Q) | 0.876 |  |  | 0.877
Magnet |  |  | (L) | 87.5 | mH
|  |  | (Q) | 0.876
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 20.6 | 20.6 |  | 20.6 | 20.5
|  | (Q) | 3.08 | 3.06 |  | 3.08 | 3.08
D1 & D2 |  | (L) | 41.2 | mH |  | 41.1 | mH
|  | (Q) | 3.08 |  |  | 3.09
Magnet |  |  | (L) | 82.4 | mH
|  |  | (Q) | 3.09
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 6.29 | 6.17 |  | 6.83 | 6.85
|  | (Q) | 0.673 | 0.657 |  | 0.752 | 0.787
D1 & D2 |  | (L) | 12.4 | mH |  | 13.6 | mH
|  | (Q) | 0.658 |  |  | 0.76
Magnet |  |  | (L) | 26 | mH
|  |  | (Q) | 0.704
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000160 |  |  |  |  | Date: | 13/9/2004
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 13.5 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 13.3 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 3.38 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 3.24 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 9.45 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 9.85 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.952 | 0.95 |  | 0.95 | 0.95
Verification absence court-circuit entre spires sur chaque pole (comparaison entre chocs 1KV et 4KV)
|  |  | OK | OK |  | OK | OK
| Dipoles D1 & D2 [4KV] |  | 1.75 | mS |  | 1.75 | mS
| Magnet [4KV] |  |  | 2.64 | mS
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000160 |  |  |  |  | Date: | 13/9/2004
DIQH TEST
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 21.9722 | 21.9506 |  |  | 21.8787 | 21.8678
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.0069 | 22.0446 |  |  | 21.9069 | 21.9416
Insulation Resistance
QHs -----> Masse |  | 3.74 |  | GW |  | 3.88 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 2.98 |  | GW |  | 2.8 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 9.55 | 9.35 | GW |  | 8.7 | 9.05 | GW
[V=3.0.KV, 30'']
Discharge 850 V 2.5 KJ
|  | I (A) | T (ms) |  |  | I (A) | T (ms)
| YT 111 | 38 | 150 |  | YT 211 | 37.8 | 149
| YT 112 | 37.6 | 150 |  | YT 212 | 38.1 | 150
| YT 121 | 37.7 | 149 |  | YT 221 | 37.5 | 149
| YT 122 | 37.2 | 151 |  | YT 222 | 38.2 | 149
Insulation Resistance
QHs -----> Masse |  | 5.6 |  | GW |  | 5.5 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.26 |  | GW |  | 3.1 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 9.45 | 9.25 | GW |  | 8.8 | 9.05 | GW
[V=3.0.KV, 30'']
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 21.98 | 21.95 |  |  | 21.91 | 21.89
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.03 | 22.07 |  |  | 21.94 | 21.98
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION