## Coils-Coll.
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 22.1058 | Magnet | 5.9894 | 5989.444 | 6001.4469 | 10 | 0.893 | 21.9 |  | 0.893 | 21.9
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 22.0012 | Aperture1 | 2.9917 | 2991.732 | 2997.7275 | 100 | 3.08 | 20.6 |  | 3.08 | 20.6
[9 empty] | YT121 | 20.7409 | Aperture2 | 2.9977 | 2997.672 | 3003.6794 | 1000 | 0.664 | 6.21 |  | 0.746 | 6.74
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 20.9111 | Apt1 Pup | 1.497 | 1496.958 | 1499.9579
CERN PART NUMBER : |  |  | HCMBAL_A00101 |  | Serial Number : |  | 175 |  | YT211 | 21.5401 | Apt1 Plw | 1.4948 | 1494.798 | 1497.7936 | 10 | 0.891 | 21.8 |  | 0.892 | 21.9
[9 empty] | YT212 | 20.9588 | Apt2 Pup | 1.5009 | 1500.937 | 1503.9449 | 100 | 3.07 | 20.5 |  | 3.08 | 20.5
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 21.097 | Apt2 Plw | 1.4967 | 1496.73 | 1499.7295 | 1000 | 0.655 | 6.09 |  | 0.779 | 6.76
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 20.7717
[15 empty] | 10 | 0.896 | 43.8 |  | 0.896 | 43.8
NAME OF TEST OPERATOR: |  |  | GUTOWSKI / WENDEL | [11 empty] | 100 | 3.09 | 41.2 |  | 3.09 | 41.1
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.653 | 12.2 |  | 0.762 | 13.5
DATE OF TEST : |  |  | 18/10/2004 |  | TIME :
BAIE 2 | [14 empty] | 10 | 0.897 | 87.6
[15 empty] | 100 | 3.09 | 82.2
[15 empty] | 1000 | 0.703 | 25.7
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 19.5 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 47.5 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1496.958 | 1494.798 |  | 1500.937 | 1496.73
Poles [20°C] |  |  | 1499.9579 | 1497.7936 |  | 1503.9449 | 1499.7295
D1 & D2 |  |  | 2997.7275 | mW |  | 3003.6794 | mW
Magnet |  |  |  | 6001.4469 | mW
|  | Poles Nr. | JP 698 | JP 699 |  | JP 700 | JP 701
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 22.1058 | 20.7409 |  | 21.5401 | 21.097
|  |  | 22.0012 | 20.9111 |  | 20.9588 | 20.7717
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 210 | nF |  | 210 | nF
D1 & D2 -----> Ground |  |  | 200 | nF |  | 200 | nF
Magnet -----> Ground |  |  |  | 340 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 175 |  |  |  |  | Date: | 18/10/2004
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.9 | 21.8 |  | 21.9 | 21.9
|  | (Q) | 0.893 | 0.891 |  | 0.893 | 0.892
D1 & D2 |  | (L) | 43.8 | mH |  | 43.8 | mH
|  | (Q) | 0.896 |  |  | 0.896
Magnet |  |  | (L) | 87.6 | mH
|  |  | (Q) | 0.897
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 20.6 | 20.5 |  | 20.6 | 20.5
|  | (Q) | 3.08 | 3.07 |  | 3.09 | 3.08
D1 & D2 |  | (L) | 41.2 | mH |  | 41.1 | mH
|  | (Q) | 3.09 |  |  | 3.09
Magnet |  |  | (L) | 82.2 | mH
|  |  | (Q) | 3.09
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 6.21 | 6.09 |  | 6.74 | 6.76
|  | (Q) | 0.664 | 0.655 |  | 0.746 | 0.779
D1 & D2 |  | (L) | 12.2 | mH |  | 13.5 | mH
|  | (Q) | 0.653 |  |  | 0.762
Magnet |  |  | (L) | 25.7 | mH
|  |  | (Q) | 0.703
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 175 |  |  |  |  | Date: | 18/10/2004
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 16.6 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 57.5 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 4.92 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 4.14 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 9.65 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 9.6 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.958 | 0.95 |  | 0.95 | 0.956
Verification absence court-circuit entre spires sur chaque pole (comparaison entre chocs 1KV et 4KV)
|  |  | OK | OK |  | OK | OK
| Dipoles D1 & D2 [4KV] |  | 1.76 | mS |  | 1.76 | mS
| Magnet [4KV] |  |  | 2.65 | mS
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 175 |  |  |  |  | Date: | 18/10/2004
DIQH TEST
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 22.1058 | 20.7409 |  |  | 21.5401 | 21.097
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.0012 | 20.9111 |  |  | 20.9588 | 20.7717
Insulation Resistance
QHs -----> Masse |  | 4 |  | GW |  | 3.96 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.32 |  | GW |  | 3.68 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 10.4 | 10.3 | GW |  | 9.45 | 9 | GW
[V=3.0.KV, 30'']
Discharge 850 V 2.5 KJ
|  | I (A) | T (ms) |  |  | I (A) | T (ms)
| YT 111 | 37.1 | 150 |  | YT 211 | 38.2 | 146
| YT 112 | 37.6 | 149 |  | YT 212 | 39.5 | 142
| YT 121 | 39.6 | 141 |  | YT 221 | 39.2 | 144
| YT 122 | 39.5 | 142 |  | YT 222 | 39.9 | 141
Insulation Resistance
QHs -----> Masse |  | 5.85 |  | GW |  | 5.65 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.82 |  | GW |  | 3.32 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 10.7 | 10.7 | GW |  | 9.6 | 9.8 | GW
[V=3.0.KV, 30'']
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 22.15 | 20.82 |  |  | 21.59 | 21.15
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.1 | 20.98 |  |  | 21.17 | 20.83
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION