## Coils-Coll.
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 21.0328 | Magnet | 5.9847 | 5984.741 | 5890.4931 | 10 | 0.887 | 21.8 |  | 0.887 | 21.8
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 20.5907 | Aperture1 | 2.9912 | 2991.171 | 2944.0659 | 100 | 3.05 | 20.6 |  | 3.06 | 20.5
[9 empty] | YT121 | 21.1975 | Aperture2 | 2.9935 | 2993.549 | 2946.4065 | 1000 | 0.655 | 6.11 |  | 0.737 | 6.57
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 20.7696 | Apt1 Pup | 1.496 | 1495.984 | 1472.4252
CERN PART NUMBER : |  |  | HCMBBL_A00101 |  | Serial Number : |  | 000099 |  | YT211 | 22.0384 | Apt1 Plw | 1.4952 | 1495.176 | 1471.6299 | 10 | 0.887 | 21.8 |  | 0.887 | 21.8
[9 empty] | YT212 | 22.4829 | Apt2 Pup | 1.4975 | 1497.512 | 1473.9291 | 100 | 3.04 | 20.5 |  | 3.04 | 20.5
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 21.9669 | Apt2 Plw | 1.496 | 1495.954 | 1472.3957 | 1000 | 0.641 | 5.96 |  | 0.775 | 6.65
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 22.4114
[15 empty] | 10 | 0.891 | 43.7 |  | 0.891 | 43.7
NAME OF TEST OPERATOR: |  |  | GUTOWSKIWENDEL/HENRY | [11 empty] | 100 | 3.06 | 41.1 |  | 3.06 | 40.9
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.644 | 12 |  | 0.754 | 13.2
DATE OF TEST : |  |  | 2004-06-07 |  | TIME :
[15 empty] | 10 | 0.892 | 87.5
[15 empty] | 100 | 3.05 | 82.1
[15 empty] | 1000 | 0.697 | 25.4
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 24 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 50 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1495.984 | 1495.176 |  | 1497.512 | 1495.954
Poles [20°C] |  |  | 1472.4252 | 1471.6299 |  | 1473.9291 | 1472.3957
D1 & D2 |  |  | 2944.0659 | mW |  | 2946.4065 | mW
Magnet |  |  |  | 5890.4931 | mW
|  | Poles Nr. | JP394 | JP395 |  | JP 396 | JP 397
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 21.0328 | 21.1975 |  | 22.0384 | 21.9669
|  |  | 20.5907 | 20.7696 |  | 22.4829 | 22.4114
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 200 | nF |  | 200 | nF
D1 & D2 -----> Ground |  |  | 200 | nF |  | 200 | nF
Magnet -----> Ground |  |  |  | 330 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000099 |  |  |  |  | Date: | 2004-06-07
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.8 | 21.8 |  | 21.8 | 21.8
|  | (Q) | 0.887 | 0.887 |  | 0.887 | 0.887
D1 & D2 |  | (L) | 43.7 | mH |  | 43.7 | mH
|  | (Q) | 0.891 |  |  | 0.891
Magnet |  |  | (L) | 87.5 | mH
|  |  | (Q) | 0.892
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 20.6 | 20.5 |  | 20.5 | 20.5
|  | (Q) | 3.05 | 3.04 |  | 3.06 | 3.04
D1 & D2 |  | (L) | 41.1 | mH |  | 40.9 | mH
|  | (Q) | 3.06 |  |  | 3.06
Magnet |  |  | (L) | 82.1 | mH
|  |  | (Q) | 3.05
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 6.11 | 5.96 |  | 6.57 | 6.65
|  | (Q) | 0.655 | 0.641 |  | 0.737 | 0.775
D1 & D2 |  | (L) | 12 | mH |  | 13.2 | mH
|  | (Q) | 0.644 |  |  | 0.754
Magnet |  |  | (L) | 25.4 | mH
|  |  | (Q) | 0.697
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000099 |  |  |  |  | Date: | 2004-06-07
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 7.55 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 12.3 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 4.26 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 5.45 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 9.15 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 6.45 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.1 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.956 | 0.952 |  | 0.952 | 0.958
Verification absence court-circuit entre spires sur chaque pole (comparaison entre chocs 1KV et 4KV)
|  |  | OK | OK |  | OK | OK
| Dipoles D1 & D2 [4KV] |  | 1.77 | mS |  | 1.77 | mS
| Magnet [4KV] |  |  | 2.65 | mS
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000099 |  |  |  |  | Date: | 2004-06-07
DIQH TEST
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 21.0328 | 21.1975 |  |  | 22.0384 | 21.9669
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 20.5907 | 20.7696 |  |  | 22.4829 | 22.4114
Insulation Resistance
QHs -----> Masse |  | 3.24 |  | GW |  | 4.02 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.12 |  | GW |  | 3.36 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 10.3 | 10.4 | GW |  | 10 | 10.2 | GW
[V=3.0.KV, 30'']
Discharge 850 V 2.5 KJ
|  | I (A) | T (ms) |  |  | I (A) | T (ms)
| YT 111 | 39.6 | 144 |  | YT 211 | 37.8 | 149
| YT 112 | 38.8 | 146 |  | YT 212 | 38.9 | 146
| YT 121 | 39.4 | 144 |  | YT 221 | 37.7 | 150
| YT 122 | 39.1 | 146 |  | YT 222 | 37.6 | 149
Insulation Resistance
QHs -----> Masse |  | 3.44 |  | GW |  | 5.4 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.52 |  | GW |  | 3.7 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 8.15 | 9.5 | GW |  | 10.1 | 9.8 | GW
[V=3.0.KV, 30'']
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 21.12 | 21.12 |  |  | 21.8 | 22.05
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 21.42 | 21.29 |  |  | 21.45 | 21.85
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION