## Coll.119
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 22.4192 | Magnet | 6.0572 | 6057.239 | 6021.1123 | 10 | 0.875 | 21.9 |  | 0.875 | 21.9
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 22.6806 | Aperture1 | 3.0296 | 3029.634 | 3011.5646 | 100 | 3.09 | 20.6 |  | 3.09 | 20.6
[9 empty] | YT121 | 22.4846 | Aperture2 | 3.0276 | 3027.608 | 3009.5507 | 1000 | 0.759 | 7.05 |  | 0.797 | 7.17
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 22.8029 | Apt1 Pup | 1.5151 | 1515.106 | 1506.0696
CERN PART NUMBER : |  |  | HCMBBL_A00101 |  | Serial Number : |  | 000119 |  | YT211 | 22.4104 | Apt1 Plw | 1.5145 | 1514.524 | 1505.4911 | 10 | 0.874 | 21.9 |  | 0.875 | 21.9
[9 empty] | YT212 | 22.7297 | Apt2 Pup | 1.5151 | 1515.079 | 1506.0427 | 100 | 3.09 | 20.5 |  | 3.08 | 20.6
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 22.3813 | Apt2 Plw | 1.5125 | 1512.518 | 1503.497 | 1000 | 0.764 | 7.02 |  | 0.816 | 7.14
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 22.7054
[15 empty] | 10 | 0.878 | 43.8 |  | 0.879 | 43.9
NAME OF TEST OPERATOR: |  |  | JARDOT/GUTOWSKI | [11 empty] | 100 | 3.1 | 41.2 |  | 3.11 | 41.2
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.756 | 14 |  | 0.804 | 14.3
DATE OF TEST : |  |  | 2004-02-04 |  | TIME :
[15 empty] | 10 | 0.879 | 87.7
[15 empty] | 100 | 3.1 | 82.4
[15 empty] | 1000 | 0.783 | 28.4
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 21.5 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 30 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1515.106 | 1514.524 |  | 1515.079 | 1512.518
Poles [20°C] |  |  | 1506.0696 | 1505.4911 |  | 1506.0427 | 1503.497
D1 & D2 |  |  | 3011.5646 | mW |  | 3009.5507 | mW
Magnet |  |  |  | 6021.1123 | mW
|  | Poles Nr. | JP 474 | JP 476 |  | JP 475 | JP 477
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 22.4192 | 22.4846 |  | 22.4104 | 22.3813
|  |  | 22.6806 | 22.8029 |  | 22.7297 | 22.7054
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 220 | nF |  | 220 | nF
D1 & D2 -----> Ground |  |  | 210 | nF |  | 210 | nF
Magnet -----> Ground |  |  |  | 350 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000119 |  |  |  |  | Date: | 2004-02-04
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.9 | 21.9 |  | 21.9 | 21.9
|  | (Q) | 0.875 | 0.874 |  | 0.875 | 0.875
D1 & D2 |  | (L) | 43.8 | mH |  | 43.9 | mH
|  | (Q) | 0.878 |  |  | 0.879
Magnet |  |  | (L) | 87.7 | mH
|  |  | (Q) | 0.879
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 20.6 | 20.5 |  | 20.6 | 20.6
|  | (Q) | 3.09 | 3.09 |  | 3.1 | 3.08
D1 & D2 |  | (L) | 41.2 | mH |  | 41.2 | mH
|  | (Q) | 3.1 |  |  | 3.11
Magnet |  |  | (L) | 82.4 | mH
|  |  | (Q) | 3.1
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 7.05 | 7.02 |  | 7.17 | 7.14
|  | (Q) | 0.759 | 0.764 |  | 0.797 | 0.816
D1 & D2 |  | (L) | 14 | mH |  | 14.3 | mH
|  | (Q) | 0.756 |  |  | 0.804
Magnet |  |  | (L) | 28.4 | mH
|  |  | (Q) | 0.783
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000119 |  |  |  |  | Date: | 2004-02-04
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 14 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 16.8 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 4 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 4.14 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 12.5 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 12.6 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.2 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.1 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.5 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.5 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.966 | 0.964 |  | 0.964 | 0.98
| Dipoles D1 & D2 [4KV] |  | 1.79 | mS |  | 1.79 | mS
| Magnet [4KV] |  |  | 2.7 | mS
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000119 |  |  |  |  | Date: | 2004-02-04
DIQH TEST
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 22.4192 | 22.4846 |  |  | 22.4104 | 22.3813
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.6806 | 22.8029 |  |  | 22.7297 | 22.7054
Insulation Resistance
QHs -----> Masse |  | 4.06 |  | GW |  | 4.6 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 3.98 |  | GW |  | 3.82 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 12.2 | 12.1 | GW |  | 12.1 | 12.2 | GW
[V=3.0.KV, 30'']
Discharge 850 V 2.5 KJ
|  | I (A) | T (ms) |  |  | I (A) | T (ms)
| YT 111 | 37.2 | 152 |  | YT 211 | 37.2 | 152
| YT 112 | 36.9 | 153 |  | YT 212 | 36.6 | 154
| YT 121 | 37.2 | 152 |  | YT 221 | 36.9 | 152
| YT 122 | 36.6 | 155 |  | YT 222 | 36.2 | 154
Insulation Resistance
QHs -----> Masse |  | 5.9 |  | GW |  | 6.1 |  | GW
[V=3.0.KV, 1']
QHs -----> Poles |  | 4.28 |  | GW |  | 3.84 |  | GW
[V=3.0.KV, 1']
Strip -----> Strip |  | 11.6 | 11.7 | GW |  | 12.4 | 12.9 | GW
[V=3.0.KV, 30'']
Resistance [W]
|  | YT 111 | YT 121 |  |  | YT 211 | YT 221
|  | 22.31 | 22.41 |  |  | 22.34 | 22.29
|  | YT 112 | YT 122 |  |  | YT 212 | YT 222
|  | 22.59 | 22.72 |  |  | 22.66 | 22.62