## Coils-Coll.
|  |  | ELECTRICAL TESTS OF DIPOLES | [8 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
|  |  | (WITH BEAM TUBES) | [6 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F301 |  |  |  | YT111 | 21.3558 | Magnet | 6.0128 | 6012.809 | 5988.8536 | 10 | 0.876 | 21.8 |  | 0.877 | 21.8
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F301 |  |  |  | YT112 | 20.927 | Aperture1 | 3.0077 | 3007.655 | 2995.6723 | 100 | 2.2 | 18.9 |  | 2.2 | 18.9
[9 empty] | YT121 | 20.7793 | Aperture2 | 3.0051 | 3005.139 | 3005.139 | 1000 | 0.655 | 4.31 |  | 0.677 | 4.36
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 20.6673 | Apt1 Pup | 1.5047 | 1504.67 | 1498.6753
CERN PART NUMBER : |  |  | HCMBB_A00101 |  | Serial Number : |  | 000025 |  | YT211 | 21.0857 | Apt1 Plw | 1.503 | 1502.972 | 1496.9841 | 10 | 0.876 | 21.8 |  | 0.877 | 21.8
[9 empty] | YT212 | 21.0248 | Apt2 Pup | 1.5028 | 1502.778 | 1496.7908 | 100 | 2.2 | 18.9 |  | 2.2 | 18.9
CERN INSPECTION & TEST PLAN NUMBER : 9a To 9i |  |  |  |  | TEST NUMBER : |  | 9a to 9i |  | YT221 | 21.2384 | Apt2 Plw | 1.5023 | 1502.346 | 1496.3606 | 1000 | 0.66 | 4.29 |  | 0.701 | 4.4
NON CONFORMITY: |  |  | No |  | NON CONFORMITY: |  | Number |  | YT222 | 21.2078
[15 empty] | 10 | 0.881 | 43.6 |  | 0.882 | 43.6
NAME OF TEST OPERATOR: |  |  | HENRY/JARDOT | [11 empty] | 100 | 2.2 | 37.8 |  | 2.2 | 37.9
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.653 | 8.56 |  | 0.685 | 8.73
DATE OF TEST : |  |  | 21,22/10/2002 |  | TIME : |  | 10:00
[15 empty] | 10 | 0.884 | 87.3
[15 empty] | 100 | 2.21 | 75.7
[15 empty] | 1000 | 0.672 | 17.3
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 21 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 53 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1504.67 | 1502.972 |  | 1502.778 | 1502.346
Poles [20°C] |  | * | 1498.6753 | 1496.9841 |  | 1496.7908 | 1496.3606
D1 & D2 |  | * | 2995.6723 | mW |  | 3005.139 | mW
Magnet |  |  | * | 5988.8536 | mW
|  | Poles Nr. | JP 105 | JP 106 |  | JP 107 | JP 108
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  | * | 21.3558 | 20.7793 |  | 21.0857 | 21.2384
|  | * | 20.927 | 20.6673 |  | 21.0248 | 21.2078
Capacitance |  |  | ITPstep 9d/e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs | --> Dipoles |  | 200 | nF |  | 200 | nF
D1 & D2 | --> Masse |  | 200 | nF |  | 200 | nF
Magnet | --> Masse |  | * | 330 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBE)
Serial Nr: | 000025 |  |  |  |  | Date: | 21,22/10/2002
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.8 | 21.8 |  | 21.8 | 21.8
|  | (Q) | 0.876 | 0.876 |  | 0.877 | 0.877
D1 & D2 |  | (L) | 43.6 | mH |  | 43.6 | mH
|  | (Q) | 0.881 |  |  | 0.882
Magnet |  |  | (L) | 87.3 | mH
|  |  | (Q) | 0.884
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES | * | (L) | 18.9 | 18.9 |  | 18.9 | 18.9
|  | (Q) | 2.2 | 2.2 |  | 2.2 | 2.2
D1 & D2 | * | (L) | 37.8 | mH |  | 37.9 | mH
|  | (Q) | 2.2 |  |  | 2.2
Magnet | * |  | (L) | 75.7 | mH
|  |  | (Q) | 2.21
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES | * | (L) | 4.31 | 4.29 |  | 4.36 | 4.4
|  | (Q) | 0.655 | 0.66 |  | 0.677 | 0.701
D1 & D2 | * | (L) | 8.56 | mH |  | 8.73 | mH
|  | (Q) | 0.653 |  |  | 0.685
Magnet | * |  | (L) | 17.3 | mH
|  |  | (Q) | 0.672
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
Serial Nr: | 000025 |  | (WITH BEAM TUBE) |  |  | Date: | 21,22/10/2002
INSULATION RESISTANCE
Dipole D1 + QHs D1 ----> Masse * |  |  |  | 11.1 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 ----> Masse * |  |  |  | 17 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) D1 |  | * | 3.42 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) D2 |  | * | 5.7 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP -----> POLE LW |  | * | 10.9 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP -----> POLE LW |  | * | 17.6 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f/h
Dipole D1 + QHs D1 ----> Masse * |  |  |  | 0.3 | µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 ----> Masse * |  |  |  | 0.1 | µA
[V=6.0KV 5' ]
| QHs D1(8x) D1 |  | * | 0.3 | µA
[V=3.0.KV 5' ]
| QHs D2(8x) D2 |  | * | 0.2 | µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  | * | 0.1 | µA
[V=3.KV 5']
DIPOLE 2
| POLE UP -----> POLE LW |  | * | 0.1 | µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] | * | 0.87 | 0.878 |  | 0.878 | 0.882
| Dipoles D1 & D2 [4KV] * |  | 1.77 | mS | * | 1.8 | mS
| Magnet [4KV] |  | * | 2.7 | mS
| Quench Heaters |  | YT111 & 112 | YT121 & 122 |  | YT211 & 212 | YT221 & 222
| [V=1600KV] |  | 408 | 406 | µS | 394 | 413 | µS
| [C= 2µF] | Current | 42 | 41 | A | 40 | 39.5 | A
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION