## Coils-Coll.
ELECTRICAL TESTS OF DIPOLES | [11 empty] | RAW_Data [Logger] | [5 empty] | RAW_Data [Gain Phase Analyser]
AFTER COLLARING
(WITH BEAM TUBES) | [9 empty] | [DataQHs_Cl] |  | [DataPoles_Cl]
[10 empty] | QHs |  | Poles | Poles | Poles | f | Q | L |  | Q | L
[10 empty] | [W] |  | [V] | [mW] | R20[mW] | [Hz] |  | [mH] |  |  | [mH]
MAIN CONTRACTOR : |  |  | ALSTOM |  | CONTRACT NUMBER : F461 |  |  |  | YT111 | 24.2161 | Magnet | 5.9732 | 5973.217 | 6021.3881 | 10 | 0.875 | 21.9 |  | 0.876 | 21.9
SUB CONTRACTOR : |  |  | Jeumont Industrie |  | CONTRACT NUMBER : F461 |  |  |  | YT112 | 24.1659 | Aperture1 | 2.9872 | 2987.189 | 3011.2792 | 100 | 2.21 | 19 |  | 2.22 | 19
[9 empty] | YT121 | 24.0649 | Aperture2 | 2.986 | 2986.022 | 3010.1028 | 1000 | 0.665 | 4.38 |  | 0.685 | 4.45
COMPONENT NAME : |  |  | Collared Coils | [5 empty] | YT122 | 24.0557 | Apt1 Pup | 1.4952 | 1495.232 | 1507.2903
CERN PART NUMBER : |  |  | HCMBA_A00101 |  | Serial Number : |  | 000046 |  | YT211 | 22.7471 | Apt1 Plw | 1.492 | 1491.959 | 1503.9909 | 10 | 0.876 | 21.8 |  | 0.874 | 21.9
[9 empty] | YT212 | 22.6996 | Apt2 Pup | 1.4937 | 1493.702 | 1505.748 | 100 | 2.21 | 19 |  | 2.21 | 19
CERN INSPECTION & TEST PLAN NUMBER : ITP Step 9a to 9l |  |  |  |  | TEST NUMBER : |  | 9a to 9l |  | YT221 | 23.8064 | Apt2 Plw | 1.4923 | 1492.319 | 1504.3538 | 1000 | 0.669 | 4.38 |  | 0.709 | 4.5
NON CONFORMITY: |  |  | Yes/No |  | NON CONFORMITY: |  | Number |  | YT222 | 24.1586
[15 empty] | 10 | 0.878 | 43.7 |  | 0.879 | 43.7
NAME OF TEST OPERATOR: |  |  | HENRY | [11 empty] | 100 | 2.21 | 38.1 |  | 2.22 | 38.1
NAME OF RESPONSIBLE: | [14 empty] | 1000 | 0.662 | 8.71 |  | 0.695 | 8.92
DATE OF TEST : |  |  | 22/05/2003 |  | TIME :
[15 empty] | 10 | 0.88 | 87.5
[15 empty] | 100 | 2.22 | 76.2
[15 empty] | 1000 | 0.678 | 17.7
R[dc] of the Poles |  |  | ITP step 9a
I = | 1 | A | Apt 1 | [ext] |  | Apt 2 | [int]
Temp.* : | 18 | °C | Upper | Lower |  | Upper | Lower
Humidite: | 50 | % | [mW] | [mW] |  | [mW] | [mW]
Poles |  |  | 1495.232 | 1491.959 |  | 1493.702 | 1492.319
Poles [20°C] |  |  | 1507.2903 | 1503.9909 |  | 1505.748 | 1504.3538
D1 & D2 |  |  | 3011.2792 | mW |  | 3010.1028 | mW
Magnet |  |  |  | 6021.3881 | mW
|  | Poles Nr. | JP 181 | JP 184 |  | JP 182 | JP 183
R[dc] of the Quench Heaters |  |  | ITP step 9b
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Left | Right |  | Left | Right
|  |  | [W] | [W] |  | [W] | [W]
D1 & D2 |  |  | 24.2161 | 24.0649 |  | 22.7471 | 23.8064
|  |  | 24.1659 | 24.0557 |  | 22.6996 | 24.1586
Capacitance |  |  | ITP step 9d & 9e
[ Megger 1000V ; > 1'] |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
QHs -----> Dipoles |  |  | 190 | nF |  | 190 | nF
D1 & D2 -----> Ground |  |  | 190 | nF |  | 200 | nF
Magnet -----> Ground |  |  |  | 320 | nF
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000046 |  |  |  |  | Date: | 22/05/2003
INDUCTANCE L/Q[10Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 21.9 | 21.8 |  | 21.9 | 21.9
|  | (Q) | 0.875 | 0.876 |  | 0.876 | 0.874
D1 & D2 |  | (L) | 43.7 | mH |  | 43.7 | mH
|  | (Q) | 0.878 |  |  | 0.879
Magnet |  |  | (L) | 87.5 | mH
|  |  | (Q) | 0.88
INDUCTANCE L/Q[100Hz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 19 | 19 |  | 19 | 19
|  | (Q) | 2.21 | 2.21 |  | 2.21 | 2.21
D1 & D2 |  | (L) | 38.1 | mH |  | 38.1 | mH
|  | (Q) | 2.21 |  |  | 2.22
Magnet |  |  | (L) | 76.2 | mH
|  |  | (Q) | 2.22
INDUCTANCE L/Q[1.0KHz] |  |  | ITP step 9c
|  |  | Apt 1 | [ext] |  | Apt 2 | [int]
|  |  | Upper | Lower |  | Upper | Lower
|  |  | [mH] | [mH] |  | [mH] | [mH]
POLES |  | (L) | 4.38 | 4.38 |  | 4.45 | 4.5
|  | (Q) | 0.665 | 0.669 |  | 0.685 | 0.709
D1 & D2 |  | (L) | 8.71 | mH |  | 8.92 | mH
|  | (Q) | 0.662 |  |  | 0.695
Magnet |  |  | (L) | 17.7 | mH
|  |  | (Q) | 0.678
ELECTRICAL TESTS OF DIPOLES
AFTER COLLARING
(WITH BEAM TUBES)
Serial Nr: | 000046 |  |  |  |  | Date: | 22/05/2003
INSULATION RESISTANCE
Dipole D1 + QHs D1 -----> Masse |  |  |  | 43.8 | GW |  | >1000MW
[V=1.0.KV, 1']
Dipole D2 + QHs D2 -----> Masse |  |  |  | 41 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D1(8x) -----> D1 |  |  | 22.4 | GW |  | >1000MW
[V=1.0.KV, 1']
| QHs D2(8x) -----> D2 |  |  | 23.2 | GW |  | >1000MW
[V=1.0.KV, 1']
DIPOLE 1
| POLE UP ------> POLE LW |  |  | 38.2 | GW |  | >1000MW
[V=1.0.KV, 30'']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 40.4 | GW |  | >1000MW
[V=1.0KV, 30'']
DIELECTRIC TESTS |  |  | ITP step 9f to 9h
Dipole D1 + QHs D1 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
Dipole D2 + QHs D2 -----> Masse |  |  |  | 0.1 | µA |  | <15.0µA
[V=6.0KV 5' ]
| QHs D1(8x) -----> D1 |  |  | 0.4 | µA |  | <10.0µA
[V=3.0.KV 5' ]
| QHs D2(8x) -----> D2 |  |  | 0.4 | µA |  | <10.0µA
[V=3.0.KV 5' ]
DIPOLE 1
| POLE UP -----> POLE LW |  |  | 0.4 | µA |  | <10.0µA
[V=3.KV 5']
DIPOLE 2
| POLE UP ------> POLE LW |  |  | 0.4 | µA |  | <10.0µA
[V= 3KV 5']
DISCHARGE TEST |  |  | ITP step 9i & 9l
V = 100V/turn |  |  | Apt 1 | [ext] |  | Apt 2 | [int]
[C= 2µF] |  |  | Upper | Lower |  | Upper | Lower
|  |  | mS | mS |  | mS | mS
| Poles [4.0KV] |  | 0.866 | 0.864 |  | 0.868 | 0.872
| Dipoles D1 & D2 [4KV] |  | 1.78 | mS |  | 1.79 | mS
| Magnet [4KV] |  |  | 2.67 | mS
| Quench Heaters |  | YT111 & 112 | YT121 & 122 |  | YT211 & 212 | YT221 & 222
| [V=1600KV] |  | 481 | 473 | µS | 460 | 471 | µS
| [C= 2µF] | Current | 35.5 | 33 | A | 35 | 34 | A
Apt1
Apt2
Up
Lw
Dip
Magnet
LHC/MMS TRACEABILITY INFORMATION