## Alarm sheet
File name | HCMB__A001-01000346_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 1000346
Date of test Ap 1 | 2006-02-21 |  |  |  | Date of test Ap 2 | 2006-02-21
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | yellow alarm | status ok | status ok | status ok | b5 | yellow alarm | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | yellow alarm | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | yellow alarm | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | yellow alarm | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
1074 | HCMB__A001 | 1000346 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2006-02-22 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2006-02-22 | Answer to manufacturer |  |  | 2006-02-22
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -
| 2005-10-21 | P. Hagen | Check if mole positioning is within +/- 5 cm all positions

## Original data
File name |  | HCMB__A001-01000346_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 1000346 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Alstom | [9 empty] | File name for aperture 1 |  |  | 21-FEB-2006_1346_1cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 21-FEB-2006_1346_1cc2.txt |  |  |  | a even | 1
Test Operator |  | clerc |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2006-02-21 |  |  | Time of test |  |  | 15:00 - 18:07 |  |  |  | Date of test |  | 2006-02-21 |  |  | Time of test |  |  | 15:00 - 18:07
Mole name |  | Dipole 04 | [9 empty] | Mole name |  | Dipole 05 | [5 empty] | 2005-05-18
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 19 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 19
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.53 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.55 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00015674 |  |  | Dy average |  |  | 0.00055438 |  |  |  | Dx average |  | 0.00026261 |  |  | Dy average |  |  | 0.00083185
Magnetic length |  | 14.4286 |  |  | Tranfs. Function (mT/KA) |  |  | 595.3034 |  |  |  | Magnetic length |  | 14.4273 |  |  | Tranfs. Function (mT/KA) |  |  | 595.604
Phase |  | 0.12913 |  |  | Integral field (T.m) |  |  | 85.883 |  |  |  | Phase |  | -0.31072 |  |  | Integral field (T.m) |  |  | 86.1096
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.7075 | 5.9586 | 5.9543 | 5.9539 | 5.9531 | 5.9532 | 5.9528 | 5.9523 | 5.9522 | 5.9525 | 5.9516 | 5.9506 | 5.9511 | 5.9492 | 5.9512 | 5.9517 | 5.9541 | 5.9553 | 5.9569 | 3.6637 | 85.883 | C1 (T) | 3.9011 | 5.9609 | 5.9569 | 5.9562 | 5.9561 | 5.956 | 5.9558 | 5.9551 | 5.9557 | 5.9549 | 5.9542 | 5.9534 | 5.9547 | 5.9527 | 5.9543 | 5.954 | 5.957 | 5.9585 | 5.9623 | 3.7177 | 86.1096
Angle (mrad) | 3.085 | 0.081089 | 0.43846 | 0.52084 | 0.45734 | 0.14421 | -0.12198 | -0.20041 | -0.2212 | -0.27849 | -0.44155 | -0.15013 | -0.000162 | -0.41944 | -0.36222 | -0.51528 | -0.56009 | -0.41238 | -0.22354 | 0.55738 | 0.12913 | Angle (mrad) | 4.5981 | 0.2353 | 0.29316 | 0.31802 | 0.19371 | -0.018919 | -0.086236 | -0.18697 | -0.3513 | -0.306 | -0.42125 | -0.52548 | -0.34576 | -0.27631 | -0.28872 | -0.52513 | -0.52945 | -0.26429 | -0.006081 | 0.12654 | -0.31072
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.5524 | 1.4476 | 1.477 | 0.66609 | -0.017707 | 0.88218 | 0.91233 | 0.55437 | 0.96086 | 0.44291 | 0.28841 | -0.16793 | -0.17543 | -0.79127 | 0.19456 | -0.87812 | -1.676 | -0.32933 | 0.36296 | 0.61424 | 0.2859 | b2 | -0.39529 | -0.80218 | -0.7967 | -0.3681 | -0.32447 | -0.11843 | -0.28229 | -0.33592 | -0.53251 | -0.15987 | -1.0661 | -0.17397 | -0.33563 | -0.032906 | 0.025831 | -1.0225 | -1.2004 | 0.21444 | -0.18734 | 1.4874 | -0.35425
b3 | 45.6928 | -0.020201 | -0.059501 | 0.043627 | 0.11899 | -1.7595 | -0.41771 | -0.59522 | -0.46443 | -0.96327 | -1.7024 | -0.85539 | -1.2688 | -1.9668 | -2.6716 | -2.8985 | -1.3381 | 0.13705 | -0.22078 | 0.51345 | 0.6172 | b3 | 42.1164 | -0.38906 | -0.2288 | -0.86361 | -0.53444 | -1.4769 | -0.49398 | -0.69493 | -0.86628 | -1.908 | -1.8999 | -1.7814 | -1.7946 | -1.383 | -2.5358 | -3.1654 | -1.2741 | 0.65464 | 0.36108 | 0.39477 | 0.39226
b4 | 0.5124 | 0.067906 | 0.2577 | 0.33364 | -0.028789 | -0.00019948 | 0.3094 | 0.22524 | 0.22362 | 0.076022 | 0.24173 | 0.098008 | -0.10634 | 0.11752 | 0.051719 | 0.053009 | -0.273 | 0.15992 | -0.035599 | -0.12953 | 0.10452 | b4 | -0.38483 | -0.19679 | -0.15614 | -0.20272 | -0.25339 | -0.046252 | -0.0020127 | -0.349 | -0.1899 | -0.13305 | -0.3612 | -0.18879 | -0.11311 | 0.02556 | -0.013223 | -0.33668 | -0.10749 | -0.066621 | -0.19034 | -0.65493 | -0.18372
b5 | -2.385 | 0.65931 | 0.76093 | 0.95681 | 1.1728 | 1.2836 | 1.1324 | 1.1597 | 1.0367 | 1.3723 | 1.239 | 0.87899 | 0.9056 | 0.81955 | 0.88181 | 1.0096 | 0.97623 | 1.0422 | 0.99936 | -2.8133 | 0.78334 | b5 | -2.0243 | 0.62972 | 0.67126 | 0.8295 | 1.0989 | 1.3804 | 0.92663 | 1.1438 | 1.1562 | 1.1617 | 1.2616 | 1.028 | 0.82316 | 0.86366 | 0.917 | 0.67354 | 0.71431 | 0.60637 | 0.84514 | -2.5209 | 0.7174
b6 | 0.41045 | 0.042734 | -0.013876 | 0.10859 | -0.026343 | -0.057661 | 0.085097 | 0.03263 | -0.0023671 | 0.032374 | -0.034913 | 0.029714 | -0.017714 | 0.016091 | -0.016048 | 0.13484 | -0.054823 | 0.087114 | -0.0056766 | -0.13874 | 0.026512 | b6 | -0.36383 | -0.10395 | -0.040963 | -0.084241 | -0.10096 | 0.012899 | 0.025539 | -0.094247 | -0.061366 | -0.12412 | 0.018224 | -0.022163 | -0.010403 | 0.011339 | -0.0090473 | 0.17268 | 0.12861 | -0.038153 | -0.055898 | -0.15845 | -0.037013
b7 | 2.6441 | 1.1103 | 1.145 | 1.1206 | 1.1289 | 1.0753 | 1.1722 | 1.1151 | 1.1361 | 1.1279 | 1.12 | 1.1569 | 1.1226 | 1.1414 | 1.0544 | 1.0107 | 1.0861 | 1.1113 | 1.1756 | 0.86181 | 1.1585 | b7 | 2.7901 | 1.1898 | 1.2132 | 1.1111 | 1.1795 | 1.0947 | 1.1857 | 1.1429 | 1.1668 | 1.1178 | 1.1674 | 1.2101 | 1.1457 | 1.1649 | 1.0449 | 1.0447 | 1.1249 | 1.1654 | 1.1785 | 0.88677 | 1.1945
b8 | 0.12172 | 0.017248 | 0.013334 | 0.023462 | 0.037196 | 0.066731 | 0.092918 | 0.019971 | 0.05667 | 0.044619 | 0.040983 | 0.047255 | 0.0035391 | -0.0041106 | 0.05353 | 0.0065084 | 0.052972 | 0.060041 | 0.039592 | -0.10474 | 0.035545 | b8 | -0.083708 | -0.037218 | -0.0038798 | 0.0068522 | 0.033195 | 0.0019772 | -0.00026945 | -0.058047 | 0.014252 | -0.023729 | -0.028768 | 0.012564 | 0.00037423 | -0.0053493 | 0.013731 | 0.0031073 | 0.0057803 | 0.022587 | -0.0010442 | 0.0033892 | -0.0050108
b9 | 0.48335 | 0.58155 | 0.57816 | 0.61186 | 0.61373 | 0.60421 | 0.62262 | 0.60653 | 0.60664 | 0.60309 | 0.62071 | 0.61819 | 0.60017 | 0.58734 | 0.59449 | 0.55727 | 0.59103 | 0.60131 | 0.57474 | 0.5518 | 0.59331 | b9 | 0.51628 | 0.57399 | 0.58541 | 0.59239 | 0.60183 | 0.60521 | 0.62613 | 0.60121 | 0.60191 | 0.59395 | 0.5889 | 0.59085 | 0.57762 | 0.58604 | 0.58825 | 0.56149 | 0.57825 | 0.58457 | 0.57652 | 0.54079 | 0.58562
b10 | 0.064687 | -0.051627 | 0.0081117 | 0.0019975 | -0.0080197 | -0.040296 | 0.068484 | -0.026486 | 0.018833 | -0.02829 | -0.010258 | 0.024727 | -0.0097004 | -0.018878 | 0.031181 | 0.0088597 | 0.012558 | 0.048437 | 0.0024699 | -0.11758 | 0 | b10 | -0.067013 | -0.060804 | 0.019045 | 0.01589 | 0.0173 | -0.00065603 | 0.017693 | -0.06551 | -0.0069403 | -0.033705 | -0.020759 | 0.027065 | -0.0064831 | 0.0022075 | 0.029542 | 0.013695 | 0.015351 | 0.053713 | 0.029162 | -0.0030778 | 0
b11 | 0.65959 | 0.77204 | 0.77656 | 0.7874 | 0.78489 | 0.79879 | 0.78915 | 0.7883 | 0.7864 | 0.79582 | 0.79646 | 0.78677 | 0.78236 | 0.77708 | 0.79138 | 0.78383 | 0.79148 | 0.78301 | 0.77199 | 0.69011 | 0.77862 | b11 | 0.68544 | 0.76895 | 0.77031 | 0.78065 | 0.78132 | 0.79055 | 0.78095 | 0.77417 | 0.77832 | 0.78257 | 0.78001 | 0.77401 | 0.7749 | 0.77195 | 0.78237 | 0.774 | 0.77205 | 0.77069 | 0.77854 | 0.67729 | 0.77068
b12 | 0.02018 | 0.0042722 | 0.0033405 | 0.00014361 | 0.0051146 | 0.0029308 | 0.0096029 | -0.00029837 | 0.0052128 | 0.0010816 | 0.0050443 | 0.0060876 | 0.0038097 | -0.0015052 | 0.0039389 | 0.010952 | 0.008662 | 0.0063851 | 0.0055538 | -0.0090898 | 0.0045382 | b12 | -0.016455 | -0.010937 | -0.0023429 | 0.00066905 | 0.0026127 | -0.0053091 | -0.0032097 | -0.0081212 | 0.00062584 | -0.001627 | -0.0050811 | 0.0012116 | -0.0054274 | -0.0047346 | 0.0015925 | -0.0029928 | -0.0034809 | 0.002571 | 0.0050572 | 0.002163 | -0.0025077
b13 | 0.090286 | 0.058898 | 0.055812 | 0.056076 | 0.053234 | 0.05838 | 0.059905 | 0.05903 | 0.059381 | 0.059042 | 0.062344 | 0.064336 | 0.062747 | 0.065645 | 0.059071 | 0.051847 | 0.05708 | 0.054098 | 0.056906 | 0.057655 | 0.059545 | b13 | 0.097005 | 0.060324 | 0.058842 | 0.057221 | 0.058231 | 0.055182 | 0.058626 | 0.057921 | 0.056095 | 0.057913 | 0.060211 | 0.060922 | 0.060099 | 0.061948 | 0.053758 | 0.051744 | 0.055427 | 0.055179 | 0.060312 | 0.05602 | 0.059051
b14 | 0.011266 | -0.0061419 | -0.0042571 | -0.0050444 | -0.005014 | -0.0082477 | -0.0027592 | -0.0077424 | -0.0051957 | -0.0064225 | -0.0063292 | -0.0034913 | -0.0050296 | -0.0056878 | -0.0032447 | -0.0029052 | -0.002438 | -0.00021743 | -0.0034598 | -0.011032 | -0.0043351 | b14 | 0.0019811 | -0.015252 | -0.0095656 | -0.010298 | -0.010008 | -0.012189 | -0.0095518 | -0.015289 | -0.010934 | -0.011956 | -0.011692 | -0.0075687 | -0.01244 | -0.011165 | -0.009149 | -0.0093716 | -0.010058 | -0.0073075 | -0.0079382 | -0.012898 | -0.010295
b15 | 0.0034794 | 0.026781 | 0.031303 | 0.030559 | 0.032195 | 0.034908 | 0.032404 | 0.034841 | 0.034129 | 0.034043 | 0.034589 | 0.031969 | 0.032934 | 0.034306 | 0.034675 | 0.031347 | 0.033982 | 0.033409 | 0.031206 | 0.023816 | 0.03152 | b15 | 0.002327 | 0.032029 | 0.036762 | 0.038866 | 0.040806 | 0.040671 | 0.039595 | 0.039419 | 0.044073 | 0.037924 | 0.039971 | 0.036792 | 0.039329 | 0.042136 | 0.046649 | 0.042784 | 0.044548 | 0.046774 | 0.044865 | 0.031847 | 0.039182
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 30.8501 | 0.81089 | 4.3846 | 5.2084 | 4.5734 | 1.4421 | -1.2198 | -2.0041 | -2.2119 | -2.7849 | -4.4155 | -1.5013 | -0.0016192 | -4.1944 | -3.6222 | -5.1528 | -5.6009 | -4.1238 | -2.2354 | 5.5738 | 0 | a1 | 45.9811 | 2.353 | 2.9316 | 3.1802 | 1.9371 | -0.18919 | -0.86236 | -1.8697 | -3.513 | -3.06 | -4.2125 | -5.2548 | -3.4576 | -2.7631 | -2.8872 | -5.2513 | -5.2945 | -2.6429 | -0.060808 | 1.2654 | 0
a2 | -0.57666 | -0.68152 | -0.79755 | -1.0409 | -0.89354 | -1.1116 | -0.50214 | -0.14975 | 0.48219 | 0.23862 | -0.28004 | -0.26753 | -0.88415 | -1.4834 | -1.0573 | -0.29621 | -0.2834 | -1.1386 | 0.30631 | 2.617 | -0.44643 | a2 | -2.0077 | -0.18824 | -0.24484 | -0.89182 | -1.0799 | -0.74782 | 0.26087 | 0.1584 | 0.8234 | -0.69253 | 0.63115 | -0.42714 | -0.54792 | -1.2454 | -0.86545 | -1.1153 | -0.68702 | -0.97222 | -1.4412 | 3.2926 | -0.44264
a3 | -2.3861 | -0.65131 | -0.15146 | -0.61376 | -0.4012 | -0.65196 | -0.67427 | -0.44326 | -0.19994 | -0.4089 | -0.97635 | -0.18821 | -0.0055325 | 0.29339 | -0.082115 | -0.45731 | -0.088389 | 0.0075597 | 0.011725 | 0.66415 | -0.35133 | a3 | -4.1552 | -1.0037 | -1.0453 | -0.78983 | -0.36837 | -0.70987 | -0.37515 | -0.50313 | -0.22567 | -0.47736 | -0.49363 | -0.82908 | -1.0625 | -0.39965 | -0.28182 | -0.58144 | -0.60925 | -0.58281 | -0.70461 | 1.2138 | -0.6747
a4 | -1.7075 | -0.4491 | -0.43075 | -0.72187 | -0.55394 | -0.59683 | -0.38865 | -0.19519 | -0.24476 | -0.25568 | 0.025479 | 0.026901 | 0.3286 | 0.67169 | -0.059614 | -0.44107 | -0.2342 | 0.011293 | 0.13065 | 0.50187 | -0.21481 | a4 | 0.94107 | -0.023817 | -0.40774 | -0.23012 | -0.22075 | -0.13823 | -0.31767 | 0.22933 | -0.069312 | 0.68581 | 0.28244 | 0.0052389 | -0.32209 | -0.13598 | -0.22882 | 0.29389 | 0.033411 | -0.0011767 | 0.27458 | 0.81608 | 0.043303
a5 | 3.1215 | 0.06228 | 0.031434 | -0.069126 | 0.024091 | -0.278 | -0.13323 | -0.079513 | -0.057563 | 0.041401 | -0.34958 | -0.12303 | -0.088032 | 0.086894 | 0.00441 | -0.27393 | -0.045599 | 0.041001 | -0.028291 | -0.7235 | 0.013756 | a5 | 2.5992 | -0.25466 | -0.24931 | -0.091401 | 0.028415 | -0.22064 | -0.080863 | -0.10957 | -0.054759 | 0.061457 | -0.48158 | -0.29165 | -0.3097 | -0.029843 | -0.015667 | -0.23993 | -0.13789 | -0.18577 | -0.21027 | -0.76881 | -0.085646
a6 | -0.53817 | -0.065691 | 0.062211 | -0.0034631 | -0.074063 | -0.0062957 | 0.050272 | -0.023433 | 0.078953 | 0.16006 | 0.055368 | 0.078858 | -0.070153 | -0.061169 | -0.016314 | 0.02615 | -0.003874 | -0.10987 | -0.024169 | -0.077612 | -0.017136 | a6 | 0.38536 | 0.033826 | 0.020066 | -0.11085 | -0.11413 | 0.0005161 | 0.051817 | 0.065901 | 0.16614 | 0.088933 | 0.13034 | -0.010418 | 0.0055941 | -0.14903 | -0.12389 | -0.12823 | -0.15421 | -0.1298 | -0.069187 | 0.042964 | -0.0076472
a7 | 1.6329 | -0.04668 | 0.148 | 0.10181 | 0.15508 | 0.075437 | 0.10201 | 0.041286 | 0.073267 | 0.058001 | -0.0071684 | 0.010279 | 0.081013 | 0.037254 | -0.020041 | -0.087844 | -0.010761 | 0.011496 | -0.028406 | -0.095672 | 0.085873 | a7 | 1.4691 | -0.03162 | -0.0090324 | 0.024692 | 0.028545 | -0.049608 | -0.077399 | -0.086473 | -0.037937 | -0.00022992 | -0.14703 | -0.092024 | -0.017783 | 0.027621 | 0.023015 | -0.11523 | -0.10654 | -0.10135 | -0.067293 | -0.10788 | 0.0030705
a8 | -0.15398 | 0.0080718 | -0.01794 | -0.0013486 | -0.0038878 | 0.0095453 | 0.018667 | 0.010503 | 0.019014 | 0.012922 | 0.0010575 | 0.0065958 | -0.028694 | 0.0093899 | -0.0045154 | 0.00020728 | 0.023986 | -0.018597 | 0.021855 | -0.074794 | -0.0039029 | a8 | 0.096686 | 0.026484 | 0.023061 | 0.014262 | -0.0082015 | 0.031324 | 0.047777 | 0.055899 | 0.051495 | 0.065457 | 0.052746 | -0.0091389 | -0.036601 | -0.050123 | -0.054438 | -0.033227 | -0.069734 | -0.034094 | -0.039565 | -0.068702 | 0.0027911
a9 | -0.18607 | -0.02268 | -0.0066659 | -0.017833 | -0.033438 | -0.062144 | -0.046989 | -0.041803 | -0.041094 | -0.045106 | -0.048676 | -0.035046 | -0.031319 | -0.025947 | -0.045041 | -0.078779 | -0.069662 | -0.054112 | -0.027278 | 0.02858 | -0.043241 | a9 | -0.2239 | -0.034311 | -0.05593 | -0.051357 | -0.033886 | -0.072592 | -0.0691 | -0.052179 | -0.044546 | -0.072761 | -0.076635 | -0.072911 | -0.072312 | -0.044457 | -0.081135 | -0.092199 | -0.057285 | -0.05473 | -0.057734 | -0.006863 | -0.064678
a10 | -0.055764 | 0.018254 | -0.0038781 | 0.027259 | -0.0078145 | -0.0094683 | 0.015376 | -0.021612 | -0.0039597 | 0.022287 | 0.012568 | 0.026996 | -0.0017032 | -0.007604 | 0.023834 | 0.021651 | -0.014659 | -0.0092032 | 0.023872 | -0.12589 | 0 | a10 | 0.084413 | 0.034435 | 0.020512 | 0.024903 | 0.0058373 | 0.048697 | 0.029097 | 0.019913 | 0.0021172 | 0.038393 | 0.03066 | 0.048156 | 0.003514 | -0.027855 | -0.047394 | -0.036801 | -0.05833 | -0.071877 | -0.055606 | -0.10188 | 0
a11 | 0.14153 | -0.027003 | -0.016225 | -0.019978 | -0.013284 | -0.020983 | -0.017985 | -0.024289 | -0.023395 | -0.026371 | -0.038224 | -0.025475 | -0.024655 | -0.027807 | -0.030309 | -0.039326 | -0.029322 | -0.014733 | -0.022415 | 0.011802 | -0.018004 | a11 | 0.11876 | -0.051461 | -0.047775 | -0.048027 | -0.045215 | -0.050511 | -0.051134 | -0.060543 | -0.05618 | -0.051303 | -0.064958 | -0.056777 | -0.049812 | -0.045993 | -0.043787 | -0.049884 | -0.056934 | -0.05881 | -0.05085 | -0.022837 | -0.04546
a12 | -0.017427 | 0.0024587 | -0.0011308 | 3.7153e-06 | -0.0079845 | -0.0056812 | -0.0025249 | -0.0051425 | 0.00010055 | 0.0038587 | -0.00079537 | 0.003587 | -0.001264 | -0.0018437 | -0.0010755 | -0.00068136 | -0.0065738 | -0.0071055 | 0.00072354 | -0.0078923 | -0.0024317 | a12 | 0.01734 | 0.0023693 | 0.00033839 | -0.0015341 | -0.0031681 | 0.0021991 | -0.0010164 | 0.0018786 | 0.0022464 | 0.0020929 | 0.00786 | 0.0040632 | 0.0011799 | -0.0068539 | -0.0071706 | -0.010128 | -0.009666 | -0.01217 | -0.0078445 | -0.0073551 | -0.0014817
a13 | -0.028056 | -0.0081418 | -0.004144 | -0.0058579 | -0.0084748 | -0.0076685 | -0.0088675 | -0.0040863 | -0.0051573 | -0.0032475 | -0.0064224 | -0.0041406 | -0.0036288 | -0.0012561 | -0.0023906 | -0.0040608 | -0.0033608 | -0.0024954 | -0.00092739 | 0.0051087 | -0.0051284 | a13 | -0.017334 | -0.0054142 | -0.0084875 | -0.01032 | -0.0055222 | -0.0060127 | -0.0049854 | -0.0041845 | -0.0024795 | -0.0062272 | -0.0064181 | -0.0082094 | -0.0066218 | -0.0050606 | -0.0056773 | -0.0080137 | -0.0042158 | -0.0060583 | -0.0094374 | -0.0074754 | -0.0067102
a14 | -0.0011102 | -0.0086931 | -0.014463 | -0.01362 | -0.014272 | -0.015704 | -0.014648 | -0.015697 | -0.014075 | -0.013947 | -0.013283 | -0.012814 | -0.012141 | -0.01058 | -0.01263 | -0.013002 | -0.013474 | -0.010951 | -0.0092628 | -0.0044536 | -0.012303 | a14 | -0.00026806 | -0.013266 | -0.016242 | -0.015745 | -0.017287 | -0.019076 | -0.017951 | -0.016573 | -0.021717 | -0.012935 | -0.016635 | -0.015396 | -0.019736 | -0.020921 | -0.022506 | -0.017322 | -0.020936 | -0.022114 | -0.019031 | -0.009081 | -0.017181
a15 | 0.0046131 | -0.0093428 | -0.0062798 | -0.0064384 | -0.0079575 | -0.0083344 | -0.0024126 | -0.0086155 | -0.0043791 | -0.0068107 | -0.0077708 | -0.0035565 | -0.0069278 | -0.0065798 | -0.003799 | -0.0048045 | -0.0038185 | -0.0017143 | -0.0039867 | -0.020356 | -0.0058833 | a15 | -0.0037763 | -0.020684 | -0.01496 | -0.014761 | -0.015036 | -0.016491 | -0.013389 | -0.022418 | -0.016096 | -0.019062 | -0.017205 | -0.011755 | -0.017478 | -0.016454 | -0.014072 | -0.014933 | -0.014566 | -0.011686 | -0.015228 | -0.023199 | -0.015728
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00012945 | 0.00011491 | -1.7927e-05 | 0 | 1.7077e-05 | 8.515e-05 | -0.00014677 | 5.5619e-05 | -4.0935e-05 | 6.1934e-05 | 2.313e-05 | -5.1476e-05 | 2.094e-05 | 4.0651e-05 | -6.4932e-05 | -1.6809e-05 | -2.8096e-05 | -0.00010551 | 0 | 0.00029456 | 0.00015674 | Dx (m) | 0.00012555 | 0.00013895 | -3.9074e-05 | -3.1151e-05 | -3.6787e-05 | 0 | -3.423e-05 | 0.00014643 | 1.5412e-05 | 7.8345e-05 | 5.0439e-05 | -5.1397e-05 | 1.4658e-05 | 0 | -6.9702e-05 | -3.5127e-05 | -4.3023e-05 | -0.00012966 | -7.1248e-05 | 0 | 0.00026261
Dy (m) | 0.00017183 | -3.6163e-05 | 0 | -5.8924e-05 | 1.7213e-05 | 2.2396e-05 | -3.6478e-05 | 4.8322e-05 | 0 | -4.5555e-05 | -2.5713e-05 | -6.001e-05 | 0 | 1.8091e-05 | -5.3694e-05 | -4.7809e-05 | 3.0451e-05 | 1.8034e-05 | -5.2682e-05 | 0.00030552 | 0.00055438 | Dy (m) | -0.00023131 | -6.691e-05 | -4.7687e-05 | -5.6148e-05 | -1.4829e-05 | -0.0001042 | -6.5576e-05 | -3.2312e-05 | 0 | -7.8269e-05 | -6.263e-05 | -0.00010953 | 0 | 6.0825e-05 | 9.9075e-05 | 7.8539e-05 | 0.0001252 | 0.00014865 | 0.00011677 | 0.00025558 | 0.00083185
Magnetic length (m) |  | 14.4286 | [19 empty] | Magnetic length (m) |  | 14.4273
Transfer funct. (mT/KA) |  | 595.3034 | [19 empty] | Transfer funct. (mT/KA) |  | 595.604
Coil positioning (mm) |  | 0.019403 | [19 empty] | Coil positioning (mm) |  | 0.018404
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.1.5 B | [20 empty] | Soft version | DIMM V2.1.5 B
Level Average | -0.19645 | -0.11838 | -0.1025 | -0.13225 | -0.078528 | 0.13378 | -0.021422 | 0.044077 | -0.068256 | -0.22704 | -0.18804 | -0.018693 | -0.060324 | -0.064647 | 0.058603 | -0.19192 | -0.10889 | 0.016098 | -0.035411 | -0.013456 |  | Level Average | -0.11666 | -0.20824 | -0.062033 | -0.14324 | -0.13018 | -0.19423 | -0.095894 | -0.084144 | -0.091017 | -0.095481 | -0.2013 | -0.052039 | -0.14505 | -0.078113 | -0.16609 | -0.092941 | -0.12061 | -0.092193 | -0.23647 | 0.18796
Level Sigma2 | 0.005759 | 0.005559 | 0.001835 | 0.0037 | 0.01293 | 0.013749 | 0.004237 | 0.0038 | 0.008769 | 0.015638 | 0.008403 | 0.001313 | 0.00058 | 0.0024 | 0.012504 | 0.000477 | 0.004392 | 0.030474 | 0.000785 | 0.078785 |  | Level Sigma2 | 0.003169 | 0.003309 | 0.017414 | 0.000568 | 0.035412 | 0.045843 | 0.031793 | 0.10108 | 0.005516 | 0.010783 | 0.012302 | 0.00106 | 0.000413 | 0.000923 | 0.000602 | 0.009125 | 0.026675 | 0.097892 | 0.034316 | 0.000485
Position (mm) | 0 | 749 | 1499 | 2250 | 2997 | 3745 | 4496 | 5243 | 5998 | 6746 | 7498 | 8247 | 8998 | 9749 | 10497 | 11248 | 12000 | 12750 | 13499 | 14250 |  | Position (mm) | 0 | 744 | 1493 | 2244 | 2994 | 3744 | 4493 | 5235 | 5981 | 6733 | 7466 | 8214 | 8965 | 9716 | 10467 | 11218 | 11967 | 12718 | 13468 | 14218
Fist move (mm) | First move Ap1 | 940.7776 | First move Ap2 | 945.6226 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-1000346 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-1000346 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 370.7453 | 595.864 | 595.4343 | 595.3864 | 595.3058 | 595.3194 | 595.2757 | 595.2275 | 595.2172 | 595.2494 | 595.1648 | 595.0566 | 595.1121 | 594.9246 | 595.1223 | 595.1706 | 595.4096 | 595.5263 | 595.695 | 366.3663 | 8588.2984 | C1/i (mT/KA) | 390.1083 | 596.093 | 595.6853 | 595.6246 | 595.6148 | 595.5982 | 595.5827 | 595.5112 | 595.5661 | 595.4866 | 595.4205 | 595.3438 | 595.4662 | 595.2657 | 595.4276 | 595.403 | 595.7008 | 595.8493 | 596.2332 | 371.7657 | 8610.9594
Angle (mrad) | 3.085 | 0.081089 | 0.43846 | 0.52084 | 0.45734 | 0.14421 | -0.12198 | -0.20041 | -0.2212 | -0.27849 | -0.44155 | -0.15013 | -0.000162 | -0.41944 | -0.36222 | -0.51528 | -0.56009 | -0.41238 | -0.22354 | 0.55738 | 0.12913 | Angle (mrad) | 4.5981 | 0.2353 | 0.29316 | 0.31802 | 0.19371 | -0.018919 | -0.086236 | -0.18697 | -0.3513 | -0.306 | -0.42125 | -0.52548 | -0.34576 | -0.27631 | -0.28872 | -0.52513 | -0.52945 | -0.26429 | -0.006081 | 0.12654 | -0.31072
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.5524 | 1.4476 | 1.477 | 0.66609 | -0.017707 | 0.88218 | 0.91233 | 0.55437 | 0.96086 | 0.44291 | 0.28841 | -0.16793 | -0.17543 | -0.79127 | 0.19456 | -0.87812 | -1.676 | -0.32933 | 0.36296 | 0.61424 |  | b2 | -0.39529 | -0.80218 | -0.7967 | -0.3681 | -0.32447 | -0.11843 | -0.28229 | -0.33592 | -0.53251 | -0.15987 | -1.0661 | -0.17397 | -0.33563 | -0.032906 | 0.025831 | -1.0225 | -1.2004 | 0.21444 | -0.18734 | 1.4874
b3 | 45.6928 | -0.020201 | -0.059501 | 0.043627 | 0.11899 | -1.7595 | -0.41771 | -0.59522 | -0.46443 | -0.96327 | -1.7024 | -0.85539 | -1.2688 | -1.9668 | -2.6716 | -2.8985 | -1.3381 | 0.13705 | -0.22078 | 0.51345 |  | b3 | 42.1164 | -0.38906 | -0.2288 | -0.86361 | -0.53444 | -1.4769 | -0.49398 | -0.69493 | -0.86628 | -1.908 | -1.8999 | -1.7814 | -1.7946 | -1.383 | -2.5358 | -3.1654 | -1.2741 | 0.65464 | 0.36108 | 0.39477
b4 | 0.5124 | 0.067906 | 0.2577 | 0.33364 | -0.028789 | -0.00019948 | 0.3094 | 0.22524 | 0.22362 | 0.076022 | 0.24173 | 0.098008 | -0.10634 | 0.11752 | 0.051719 | 0.053009 | -0.273 | 0.15992 | -0.035599 | -0.12953 |  | b4 | -0.38483 | -0.19679 | -0.15614 | -0.20272 | -0.25339 | -0.046252 | -0.0020127 | -0.349 | -0.1899 | -0.13305 | -0.3612 | -0.18879 | -0.11311 | 0.02556 | -0.013223 | -0.33668 | -0.10749 | -0.066621 | -0.19034 | -0.65493
b5 | -2.385 | 0.65931 | 0.76093 | 0.95681 | 1.1728 | 1.2836 | 1.1324 | 1.1597 | 1.0367 | 1.3723 | 1.239 | 0.87899 | 0.9056 | 0.81955 | 0.88181 | 1.0096 | 0.97623 | 1.0422 | 0.99936 | -2.8133 |  | b5 | -2.0243 | 0.62972 | 0.67126 | 0.8295 | 1.0989 | 1.3804 | 0.92663 | 1.1438 | 1.1562 | 1.1617 | 1.2616 | 1.028 | 0.82316 | 0.86366 | 0.917 | 0.67354 | 0.71431 | 0.60637 | 0.84514 | -2.5209
b6 | 0.41045 | 0.042734 | -0.013876 | 0.10859 | -0.026343 | -0.057661 | 0.085097 | 0.03263 | -0.0023671 | 0.032374 | -0.034913 | 0.029714 | -0.017714 | 0.016091 | -0.016048 | 0.13484 | -0.054823 | 0.087114 | -0.0056766 | -0.13874 |  | b6 | -0.36383 | -0.10395 | -0.040963 | -0.084241 | -0.10096 | 0.012899 | 0.025539 | -0.094247 | -0.061366 | -0.12412 | 0.018224 | -0.022163 | -0.010403 | 0.011339 | -0.0090473 | 0.17268 | 0.12861 | -0.038153 | -0.055898 | -0.15845
b7 | 2.6441 | 1.1103 | 1.145 | 1.1206 | 1.1289 | 1.0753 | 1.1722 | 1.1151 | 1.1361 | 1.1279 | 1.12 | 1.1569 | 1.1226 | 1.1414 | 1.0544 | 1.0107 | 1.0861 | 1.1113 | 1.1756 | 0.86181 |  | b7 | 2.7901 | 1.1898 | 1.2132 | 1.1111 | 1.1795 | 1.0947 | 1.1857 | 1.1429 | 1.1668 | 1.1178 | 1.1674 | 1.2101 | 1.1457 | 1.1649 | 1.0449 | 1.0447 | 1.1249 | 1.1654 | 1.1785 | 0.88677
b8 | 0.12172 | 0.017248 | 0.013334 | 0.023462 | 0.037196 | 0.066731 | 0.092918 | 0.019971 | 0.05667 | 0.044619 | 0.040983 | 0.047255 | 0.0035391 | -0.0041106 | 0.05353 | 0.0065084 | 0.052972 | 0.060041 | 0.039592 | -0.10474 |  | b8 | -0.083708 | -0.037218 | -0.0038798 | 0.0068522 | 0.033195 | 0.0019772 | -0.00026945 | -0.058047 | 0.014252 | -0.023729 | -0.028768 | 0.012564 | 0.00037423 | -0.0053493 | 0.013731 | 0.0031073 | 0.0057803 | 0.022587 | -0.0010442 | 0.0033892
b9 | 0.48335 | 0.58155 | 0.57816 | 0.61186 | 0.61373 | 0.60421 | 0.62262 | 0.60653 | 0.60664 | 0.60309 | 0.62071 | 0.61819 | 0.60017 | 0.58734 | 0.59449 | 0.55727 | 0.59103 | 0.60131 | 0.57474 | 0.5518 |  | b9 | 0.51628 | 0.57399 | 0.58541 | 0.59239 | 0.60183 | 0.60521 | 0.62613 | 0.60121 | 0.60191 | 0.59395 | 0.5889 | 0.59085 | 0.57762 | 0.58604 | 0.58825 | 0.56149 | 0.57825 | 0.58457 | 0.57652 | 0.54079
b10 | 0.064687 | -0.051627 | 0.0081117 | 0.0019975 | -0.0080197 | -0.040296 | 0.068484 | -0.026486 | 0.018833 | -0.02829 | -0.010258 | 0.024727 | -0.0097004 | -0.018878 | 0.031181 | 0.0088597 | 0.012558 | 0.048437 | 0.0024699 | -0.11758 |  | b10 | -0.067013 | -0.060804 | 0.019045 | 0.01589 | 0.0173 | -0.00065603 | 0.017693 | -0.06551 | -0.0069403 | -0.033705 | -0.020759 | 0.027065 | -0.0064831 | 0.0022075 | 0.029542 | 0.013695 | 0.015351 | 0.053713 | 0.029162 | -0.0030778
b11 | 0.65959 | 0.77204 | 0.77656 | 0.7874 | 0.78489 | 0.79879 | 0.78915 | 0.7883 | 0.7864 | 0.79582 | 0.79646 | 0.78677 | 0.78236 | 0.77708 | 0.79138 | 0.78383 | 0.79148 | 0.78301 | 0.77199 | 0.69011 |  | b11 | 0.68544 | 0.76895 | 0.77031 | 0.78065 | 0.78132 | 0.79055 | 0.78095 | 0.77417 | 0.77832 | 0.78257 | 0.78001 | 0.77401 | 0.7749 | 0.77195 | 0.78237 | 0.774 | 0.77205 | 0.77069 | 0.77854 | 0.67729
b12 | 0.02018 | 0.0042722 | 0.0033405 | 0.00014361 | 0.0051146 | 0.0029308 | 0.0096029 | -0.00029837 | 0.0052128 | 0.0010816 | 0.0050443 | 0.0060876 | 0.0038097 | -0.0015052 | 0.0039389 | 0.010952 | 0.008662 | 0.0063851 | 0.0055538 | -0.0090898 |  | b12 | -0.016455 | -0.010937 | -0.0023429 | 0.00066905 | 0.0026127 | -0.0053091 | -0.0032097 | -0.0081212 | 0.00062584 | -0.001627 | -0.0050811 | 0.0012116 | -0.0054274 | -0.0047346 | 0.0015925 | -0.0029928 | -0.0034809 | 0.002571 | 0.0050572 | 0.002163
b13 | 0.090286 | 0.058898 | 0.055812 | 0.056076 | 0.053234 | 0.05838 | 0.059905 | 0.05903 | 0.059381 | 0.059042 | 0.062344 | 0.064336 | 0.062747 | 0.065645 | 0.059071 | 0.051847 | 0.05708 | 0.054098 | 0.056906 | 0.057655 |  | b13 | 0.097005 | 0.060324 | 0.058842 | 0.057221 | 0.058231 | 0.055182 | 0.058626 | 0.057921 | 0.056095 | 0.057913 | 0.060211 | 0.060922 | 0.060099 | 0.061948 | 0.053758 | 0.051744 | 0.055427 | 0.055179 | 0.060312 | 0.05602
b14 | 0.011266 | -0.0061419 | -0.0042571 | -0.0050444 | -0.005014 | -0.0082477 | -0.0027592 | -0.0077424 | -0.0051957 | -0.0064225 | -0.0063292 | -0.0034913 | -0.0050296 | -0.0056878 | -0.0032447 | -0.0029052 | -0.002438 | -0.00021743 | -0.0034598 | -0.011032 |  | b14 | 0.0019811 | -0.015252 | -0.0095656 | -0.010298 | -0.010008 | -0.012189 | -0.0095518 | -0.015289 | -0.010934 | -0.011956 | -0.011692 | -0.0075687 | -0.01244 | -0.011165 | -0.009149 | -0.0093716 | -0.010058 | -0.0073075 | -0.0079382 | -0.012898
b15 | 0.0034794 | 0.026781 | 0.031303 | 0.030559 | 0.032195 | 0.034908 | 0.032404 | 0.034841 | 0.034129 | 0.034043 | 0.034589 | 0.031969 | 0.032934 | 0.034306 | 0.034675 | 0.031347 | 0.033982 | 0.033409 | 0.031206 | 0.023816 |  | b15 | 0.002327 | 0.032029 | 0.036762 | 0.038866 | 0.040806 | 0.040671 | 0.039595 | 0.039419 | 0.044073 | 0.037924 | 0.039971 | 0.036792 | 0.039329 | 0.042136 | 0.046649 | 0.042784 | 0.044548 | 0.046774 | 0.044865 | 0.031847
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 30.8501 | 0.81089 | 4.3846 | 5.2084 | 4.5734 | 1.4421 | -1.2198 | -2.0041 | -2.2119 | -2.7849 | -4.4155 | -1.5013 | -0.0016192 | -4.1944 | -3.6222 | -5.1528 | -5.6009 | -4.1238 | -2.2354 | 5.5738 |  | a1 | 45.9811 | 2.353 | 2.9316 | 3.1802 | 1.9371 | -0.18919 | -0.86236 | -1.8697 | -3.513 | -3.06 | -4.2125 | -5.2548 | -3.4576 | -2.7631 | -2.8872 | -5.2513 | -5.2945 | -2.6429 | -0.060808 | 1.2654
a2 | -0.57666 | -0.68152 | -0.79755 | -1.0409 | -0.89354 | -1.1116 | -0.50214 | -0.14975 | 0.48219 | 0.23862 | -0.28004 | -0.26753 | -0.88415 | -1.4834 | -1.0573 | -0.29621 | -0.2834 | -1.1386 | 0.30631 | 2.617 |  | a2 | -2.0077 | -0.18824 | -0.24484 | -0.89182 | -1.0799 | -0.74782 | 0.26087 | 0.1584 | 0.8234 | -0.69253 | 0.63115 | -0.42714 | -0.54792 | -1.2454 | -0.86545 | -1.1153 | -0.68702 | -0.97222 | -1.4412 | 3.2926
a3 | -2.3861 | -0.65131 | -0.15146 | -0.61376 | -0.4012 | -0.65196 | -0.67427 | -0.44326 | -0.19994 | -0.4089 | -0.97635 | -0.18821 | -0.0055325 | 0.29339 | -0.082115 | -0.45731 | -0.088389 | 0.0075597 | 0.011725 | 0.66415 |  | a3 | -4.1552 | -1.0037 | -1.0453 | -0.78983 | -0.36837 | -0.70987 | -0.37515 | -0.50313 | -0.22567 | -0.47736 | -0.49363 | -0.82908 | -1.0625 | -0.39965 | -0.28182 | -0.58144 | -0.60925 | -0.58281 | -0.70461 | 1.2138
a4 | -1.7075 | -0.4491 | -0.43075 | -0.72187 | -0.55394 | -0.59683 | -0.38865 | -0.19519 | -0.24476 | -0.25568 | 0.025479 | 0.026901 | 0.3286 | 0.67169 | -0.059614 | -0.44107 | -0.2342 | 0.011293 | 0.13065 | 0.50187 |  | a4 | 0.94107 | -0.023817 | -0.40774 | -0.23012 | -0.22075 | -0.13823 | -0.31767 | 0.22933 | -0.069312 | 0.68581 | 0.28244 | 0.0052389 | -0.32209 | -0.13598 | -0.22882 | 0.29389 | 0.033411 | -0.0011767 | 0.27458 | 0.81608
a5 | 3.1215 | 0.06228 | 0.031434 | -0.069126 | 0.024091 | -0.278 | -0.13323 | -0.079513 | -0.057563 | 0.041401 | -0.34958 | -0.12303 | -0.088032 | 0.086894 | 0.00441 | -0.27393 | -0.045599 | 0.041001 | -0.028291 | -0.7235 |  | a5 | 2.5992 | -0.25466 | -0.24931 | -0.091401 | 0.028415 | -0.22064 | -0.080863 | -0.10957 | -0.054759 | 0.061457 | -0.48158 | -0.29165 | -0.3097 | -0.029843 | -0.015667 | -0.23993 | -0.13789 | -0.18577 | -0.21027 | -0.76881
a6 | -0.53817 | -0.065691 | 0.062211 | -0.0034631 | -0.074063 | -0.0062957 | 0.050272 | -0.023433 | 0.078953 | 0.16006 | 0.055368 | 0.078858 | -0.070153 | -0.061169 | -0.016314 | 0.02615 | -0.003874 | -0.10987 | -0.024169 | -0.077612 |  | a6 | 0.38536 | 0.033826 | 0.020066 | -0.11085 | -0.11413 | 0.0005161 | 0.051817 | 0.065901 | 0.16614 | 0.088933 | 0.13034 | -0.010418 | 0.0055941 | -0.14903 | -0.12389 | -0.12823 | -0.15421 | -0.1298 | -0.069187 | 0.042964
a7 | 1.6329 | -0.04668 | 0.148 | 0.10181 | 0.15508 | 0.075437 | 0.10201 | 0.041286 | 0.073267 | 0.058001 | -0.0071684 | 0.010279 | 0.081013 | 0.037254 | -0.020041 | -0.087844 | -0.010761 | 0.011496 | -0.028406 | -0.095672 |  | a7 | 1.4691 | -0.03162 | -0.0090324 | 0.024692 | 0.028545 | -0.049608 | -0.077399 | -0.086473 | -0.037937 | -0.00022992 | -0.14703 | -0.092024 | -0.017783 | 0.027621 | 0.023015 | -0.11523 | -0.10654 | -0.10135 | -0.067293 | -0.10788
a8 | -0.15398 | 0.0080718 | -0.01794 | -0.0013486 | -0.0038878 | 0.0095453 | 0.018667 | 0.010503 | 0.019014 | 0.012922 | 0.0010575 | 0.0065958 | -0.028694 | 0.0093899 | -0.0045154 | 0.00020728 | 0.023986 | -0.018597 | 0.021855 | -0.074794 |  | a8 | 0.096686 | 0.026484 | 0.023061 | 0.014262 | -0.0082015 | 0.031324 | 0.047777 | 0.055899 | 0.051495 | 0.065457 | 0.052746 | -0.0091389 | -0.036601 | -0.050123 | -0.054438 | -0.033227 | -0.069734 | -0.034094 | -0.039565 | -0.068702
a9 | -0.18607 | -0.02268 | -0.0066659 | -0.017833 | -0.033438 | -0.062144 | -0.046989 | -0.041803 | -0.041094 | -0.045106 | -0.048676 | -0.035046 | -0.031319 | -0.025947 | -0.045041 | -0.078779 | -0.069662 | -0.054112 | -0.027278 | 0.02858 |  | a9 | -0.2239 | -0.034311 | -0.05593 | -0.051357 | -0.033886 | -0.072592 | -0.0691 | -0.052179 | -0.044546 | -0.072761 | -0.076635 | -0.072911 | -0.072312 | -0.044457 | -0.081135 | -0.092199 | -0.057285 | -0.05473 | -0.057734 | -0.006863
a10 | -0.055764 | 0.018254 | -0.0038781 | 0.027259 | -0.0078145 | -0.0094683 | 0.015376 | -0.021612 | -0.0039597 | 0.022287 | 0.012568 | 0.026996 | -0.0017032 | -0.007604 | 0.023834 | 0.021651 | -0.014659 | -0.0092032 | 0.023872 | -0.12589 |  | a10 | 0.084413 | 0.034435 | 0.020512 | 0.024903 | 0.0058373 | 0.048697 | 0.029097 | 0.019913 | 0.0021172 | 0.038393 | 0.03066 | 0.048156 | 0.003514 | -0.027855 | -0.047394 | -0.036801 | -0.05833 | -0.071877 | -0.055606 | -0.10188
a11 | 0.14153 | -0.027003 | -0.016225 | -0.019978 | -0.013284 | -0.020983 | -0.017985 | -0.024289 | -0.023395 | -0.026371 | -0.038224 | -0.025475 | -0.024655 | -0.027807 | -0.030309 | -0.039326 | -0.029322 | -0.014733 | -0.022415 | 0.011802 |  | a11 | 0.11876 | -0.051461 | -0.047775 | -0.048027 | -0.045215 | -0.050511 | -0.051134 | -0.060543 | -0.05618 | -0.051303 | -0.064958 | -0.056777 | -0.049812 | -0.045993 | -0.043787 | -0.049884 | -0.056934 | -0.05881 | -0.05085 | -0.022837
a12 | -0.017427 | 0.0024587 | -0.0011308 | 3.7153e-06 | -0.0079845 | -0.0056812 | -0.0025249 | -0.0051425 | 0.00010055 | 0.0038587 | -0.00079537 | 0.003587 | -0.001264 | -0.0018437 | -0.0010755 | -0.00068136 | -0.0065738 | -0.0071055 | 0.00072354 | -0.0078923 |  | a12 | 0.01734 | 0.0023693 | 0.00033839 | -0.0015341 | -0.0031681 | 0.0021991 | -0.0010164 | 0.0018786 | 0.0022464 | 0.0020929 | 0.00786 | 0.0040632 | 0.0011799 | -0.0068539 | -0.0071706 | -0.010128 | -0.009666 | -0.01217 | -0.0078445 | -0.0073551
a13 | -0.028056 | -0.0081418 | -0.004144 | -0.0058579 | -0.0084748 | -0.0076685 | -0.0088675 | -0.0040863 | -0.0051573 | -0.0032475 | -0.0064224 | -0.0041406 | -0.0036288 | -0.0012561 | -0.0023906 | -0.0040608 | -0.0033608 | -0.0024954 | -0.00092739 | 0.0051087 |  | a13 | -0.017334 | -0.0054142 | -0.0084875 | -0.01032 | -0.0055222 | -0.0060127 | -0.0049854 | -0.0041845 | -0.0024795 | -0.0062272 | -0.0064181 | -0.0082094 | -0.0066218 | -0.0050606 | -0.0056773 | -0.0080137 | -0.0042158 | -0.0060583 | -0.0094374 | -0.0074754
a14 | -0.0011102 | -0.0086931 | -0.014463 | -0.01362 | -0.014272 | -0.015704 | -0.014648 | -0.015697 | -0.014075 | -0.013947 | -0.013283 | -0.012814 | -0.012141 | -0.01058 | -0.01263 | -0.013002 | -0.013474 | -0.010951 | -0.0092628 | -0.0044536 |  | a14 | -0.00026806 | -0.013266 | -0.016242 | -0.015745 | -0.017287 | -0.019076 | -0.017951 | -0.016573 | -0.021717 | -0.012935 | -0.016635 | -0.015396 | -0.019736 | -0.020921 | -0.022506 | -0.017322 | -0.020936 | -0.022114 | -0.019031 | -0.009081
a15 | 0.0046131 | -0.0093428 | -0.0062798 | -0.0064384 | -0.0079575 | -0.0083344 | -0.0024126 | -0.0086155 | -0.0043791 | -0.0068107 | -0.0077708 | -0.0035565 | -0.0069278 | -0.0065798 | -0.003799 | -0.0048045 | -0.0038185 | -0.0017143 | -0.0039867 | -0.020356 |  | a15 | -0.0037763 | -0.020684 | -0.01496 | -0.014761 | -0.015036 | -0.016491 | -0.013389 | -0.022418 | -0.016096 | -0.019062 | -0.017205 | -0.011755 | -0.017478 | -0.016454 | -0.014072 | -0.014933 | -0.014566 | -0.011686 | -0.015228 | -0.023199
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.1299 | -0.11495 | 0.017927 | 0.0028176 | -0.017079 | -0.08517 | 0.1467 | -0.055629 | 0.040932 | -0.061943 | -0.023129 | 0.051485 | -0.020941 | -0.040651 | 0.064925 | 0.016817 | 0.028101 | 0.1055 | 0.0039094 | -0.29486 |  | Dx (mm) | -0.12614 | -0.1389 | 0.039072 | 0.031148 | 0.036784 | -0.0080687 | 0.034221 | -0.14638 | -0.015412 | -0.07835 | -0.050459 | 0.051409 | -0.014658 | 0.0084861 | 0.069737 | 0.035143 | 0.043041 | 0.12982 | 0.071304 | 0.00089645
Dy (mm) | -0.1716 | 0.036173 | -0.0081151 | 0.058923 | -0.017214 | -0.022388 | 0.036466 | -0.048322 | -0.0073421 | 0.045557 | 0.025716 | 0.059998 | -0.0043609 | -0.01809 | 0.053686 | 0.047801 | -0.030445 | -0.017996 | 0.052681 | -0.30507 |  | Dy (mm) | 0.23121 | 0.066834 | 0.047691 | 0.056146 | 0.01483 | 0.1042 | 0.06558 | 0.03228 | 0.0035118 | 0.078265 | 0.062621 | 0.10954 | 0.0067669 | -0.060837 | -0.09908 | -0.078564 | -0.12526 | -0.14864 | -0.11676 | -0.2557
Mag. len. (m) | 14.4286 | [20 empty] | Mag. len. (m) | 14.4273
TF (mT/KA) | 595.3034 | [20 empty] | TF (mT/KA) | 595.604
Coil pos. (mm) | 0.019403 | [20 empty] | Coil pos. (mm) | 0.018404
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.017712 | [20 empty] | Twist (m2 rad) | 0.017928

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 595.3034 | 0.22757 | 0.0020898 |  | -1.2587 | 3.3602 |  |  | 595.604 | 0.24762 | 0.0018304 |  | -1.7161 | 2.8373 |  | 0 | 4
2 | 0.26381 |  | 0.22983 | 0.81871 | -0.41263 |  | -0.54516 | 0.55462 | -0.33896 |  | -0.41759 | 0.40272 | -0.48201 |  | -0.51382 | 0.64387 |  | 0 | 5
3 | 0.61636 |  | -0.94015 | 0.95199 | -0.34945 |  | -0.31616 | 0.32472 | 0.39148 |  | -1.127 | 0.97631 | -0.673 |  | -0.61199 | 0.25551 |  | 0 | 6
4 | 0.097287 |  | 0.09742 | 0.15212 | -0.23184 |  | -0.19057 | 0.35088 | -0.17975 |  | -0.16038 | 0.11104 | 0.042443 |  | -0.01865 | 0.27959 |  | 0 | 7
5 | 0.7835 |  | 1.0159 | 0.18681 | 0.014086 |  | -0.068995 | 0.12507 | 0.71867 |  | 0.9305 | 0.23049 | -0.082834 |  | -0.15764 | 0.13774 |  | 0 | 8
6 | 0.021155 |  | 0.017347 | 0.053802 | -0.016837 |  | -0.0024963 | 0.064376 | -0.030258 |  | -0.022876 | 0.073294 | -0.011257 |  | -0.024442 | 0.083233 |  | 0 | 9
7 | 1.1582 |  | 1.1169 | 0.040316 | 0.085504 |  | 0.03834 | 0.066665 | 1.1955 |  | 1.1477 | 0.048957 | 0.0044753 |  | -0.045407 | 0.055039 |  | 0 | 10
8 | 0.036214 |  | 0.036098 | 0.021971 | -0.0029959 |  | 9.4466e-05 | 0.016778 | -0.0061812 |  | -0.0039409 | 0.011102 | 0.0022567 |  | 0.001551 | 0.029104 |  | 0 | 11
9 | 0.59278 |  | 0.59788 | 0.017229 | -0.044355 |  | -0.040962 | 0.018412 | 0.58676 |  | 0.59023 | 0.014401 | -0.062857 |  | -0.059433 | 0.016773 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.77858 |  | 0.78572 | 0.00778 | -0.018016 |  | -0.024572 | 0.0071833 | 0.77088 |  | 0.77714 | 0.005609 | -0.045082 |  | -0.051968 | 0.0055876 |  | 0 | 14
12 | 0.00045809 |  | 0.0043044 | 0.002843 | -0.00020887 |  | -0.0022877 | 0.0025539 | -0.00024513 |  | -0.0023588 | 0.0026423 | -0.00017867 |  | -0.002008 | 0.0029018 |  | 0 | 15
13 | 0.0059295 |  | 0.058363 | 0.003664 | -0.00051182 |  | -0.0046067 | 0.0025494 | 0.0059154 |  | 0.057818 | 0.0025929 | -0.00065914 |  | -0.0060689 | 0.0017588 |  | 0 | 16
14 | -0.00041628 |  | -0.004799 | 0.0010645 | -0.0012669 |  | -0.013385 | 0.0018744 | -0.0010517 |  | -0.011041 | 0.0010415 | -0.0016906 |  | -0.017986 | 0.002465 |  | 0 | 17
15 | 0.003152 |  | 0.032755 | 0.0020582 | -0.00058833 |  | -0.0057516 | 0.002254 | 0.0039182 |  | 0.040777 | 0.0037918 | -0.0015728 |  | -0.015904 | 0.0027753 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.39475 | -0.67489
3 | [12 empty] | 3 | 9 | -0.587 | -0.69425
4 | [12 empty] | 4 | 16 | -1.4652 | -1.7362
5 | [12 empty] | 5 | 25 | -1.8782 | -1.725
6 | [12 empty] | 6 | 36 | -2.8327 | -2.5497
7 | [12 empty] | 7 | 49 | -2.9595 | -2.9583
8 | [12 empty] | 8 | 64 | -3.9529 | -4.0187
9 | [12 empty] | 9 | 81 | -4.028 | -4.1642
10 | [12 empty] | sigma (mm) |  | 0.019403 | 0.018404
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 370.7453 | 595.864 | 595.4343 | 595.3864 | 595.3058 | 595.3194 | 595.2757 | 595.2275 | 595.2172 | 595.2494 | 595.1648 | 595.0566 | 595.1121 | 594.9246 | 595.1223 | 595.1706 | 595.4096 | 595.5263 | 595.695 | 366.3663
2 | 0.90229 | 1.4445 | 1.4773 | 0.66182 | -0.016656 | 0.86626 | 0.91665 | 0.553 | 0.96327 | 0.4337 | 0.28082 | -0.16408 | -0.17855 | -0.8013 | 0.21445 | -0.87495 | -1.6713 | -0.33105 | 0.36313 | 0.60822 | 0.26381
3 | 45.7327 | -0.021691 | -0.059699 | 0.035955 | 0.12059 | -1.7572 | -0.42822 | -0.59135 | -0.46573 | -0.9645 | -1.7013 | -0.85599 | -1.2695 | -1.9681 | -2.6728 | -2.9024 | -1.3355 | 0.13404 | -0.21954 | 0.47969 | 0.61636
4 | 0.45926 | 0.086267 | 0.25443 | 0.33205 | -0.024174 | 0.026988 | 0.26917 | 0.24132 | 0.21373 | 0.096467 | 0.24636 | 0.085623 | -0.10178 | 0.12499 | 0.038303 | 0.045933 | -0.27913 | 0.13388 | -0.036869 | -0.27277 | 0.097287
5 | -2.3735 | 0.66006 | 0.76086 | 0.95666 | 1.173 | 1.2822 | 1.1293 | 1.1606 | 1.0366 | 1.3751 | 1.2392 | 0.87994 | 0.90558 | 0.82007 | 0.88186 | 1.0093 | 0.97672 | 1.04 | 0.99899 | -2.8183 | 0.7835
6 | 0.19033 | 0.08718 | -0.021544 | 0.10959 | -0.020481 | -0.025933 | 0.02572 | 0.05382 | -0.01897 | 0.057966 | -0.025835 | 0.00891 | -0.0095419 | 0.03223 | -0.04059 | 0.12736 | -0.06548 | 0.04566 | -0.0078268 | -0.038753 | 0.021155
7 | 2.6485 | 1.1112 | 1.145 | 1.1205 | 1.1291 | 1.0775 | 1.1668 | 1.1154 | 1.1351 | 1.1293 | 1.1205 | 1.156 | 1.1227 | 1.1412 | 1.0529 | 1.0106 | 1.0852 | 1.1088 | 1.176 | 0.85849 | 1.1582
8 | 0.1072 | 0.048319 | 0.0084821 | 0.022157 | 0.042399 | 0.091603 | 0.04913 | 0.0368 | 0.045128 | 0.061232 | 0.04715 | 0.031287 | 0.0095177 | 0.0073461 | 0.034228 | 0.00032609 | 0.046155 | 0.030646 | 0.037859 | -0.032281 | 0.036214
9 | 0.48396 | 0.57876 | 0.5781 | 0.6127 | 0.61373 | 0.60251 | 0.61759 | 0.60631 | 0.60624 | 0.6027 | 0.62075 | 0.61838 | 0.60006 | 0.58701 | 0.59409 | 0.55774 | 0.59108 | 0.59869 | 0.5754 | 0.55378 | 0.59278
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.65983 | 0.77242 | 0.77652 | 0.7874 | 0.78504 | 0.79903 | 0.78818 | 0.78845 | 0.78626 | 0.79598 | 0.79652 | 0.78671 | 0.78242 | 0.77706 | 0.79118 | 0.78369 | 0.79145 | 0.78266 | 0.772 | 0.68993 | 0.77858
12 | 0.15299 | 0.088432 | 0.02658 | -0.0021157 | 0.058593 | 0.065619 | 0.031713 | 0.02159 | 0.035238 | 0.035587 | 0.059456 | 0.035741 | 0.047484 | 0.0039456 | 0.011411 | 0.10199 | 0.07602 | 0.023881 | 0.053623 | 0.018101 | 0.0045809 |  | ten times!
13 | 0.89313 | 0.58118 | 0.5596 | 0.55474 | 0.53356 | 0.58112 | 0.59806 | 0.5928 | 0.59623 | 0.58252 | 0.61971 | 0.63885 | 0.62707 | 0.65614 | 0.58713 | 0.51409 | 0.57446 | 0.54266 | 0.56543 | 0.56206 | 0.059295 |  | ten times!
14 | 0.10242 | -0.038851 | -0.046773 | -0.054277 | -0.044484 | -0.056456 | -0.067464 | -0.058034 | -0.063196 | -0.049414 | -0.05835 | -0.050225 | -0.044368 | -0.044413 | -0.052667 | -0.035285 | -0.031287 | -0.030947 | -0.037332 | -0.0013533 | -0.0041628 |  | ten times!
15 | 0.034794 | 0.26781 | 0.31303 | 0.30559 | 0.32195 | 0.34908 | 0.32404 | 0.34841 | 0.34129 | 0.34043 | 0.34589 | 0.31969 | 0.32934 | 0.34306 | 0.34675 | 0.31347 | 0.33982 | 0.33409 | 0.31206 | 0.23816 | 0.03152 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 30.8702 | 0.8032 | 4.3862 | 5.2062 | 4.5725 | 1.4377 | -1.2174 | -2.0031 | -2.2127 | -2.7852 | -4.4163 | -1.4999 | -0.0027533 | -4.1988 | -3.6188 | -5.1501 | -5.6034 | -4.1171 | -2.2366 | 5.6302 | 0.0020898
2 | 0.38225 | -0.69024 | -0.79728 | -1.041 | -0.89411 | -1.1227 | -0.48871 | -0.15603 | 0.48275 | 0.2408 | -0.27754 | -0.26035 | -0.88481 | -1.4862 | -1.0398 | -0.27901 | -0.2879 | -1.1384 | 0.30767 | 2.6585 | -0.41263
3 | -2.3314 | -0.66085 | -0.14973 | -0.61687 | -0.40296 | -0.66093 | -0.6662 | -0.44326 | -0.19788 | -0.41231 | -0.97734 | -0.1895 | -0.0044 | 0.29858 | -0.081922 | -0.45645 | -0.088694 | 0.0078573 | 0.011966 | 0.68329 | -0.34945
4 | -1.8992 | -0.45303 | -0.42943 | -0.73509 | -0.5491 | -0.59564 | -0.39376 | -0.18305 | -0.24241 | -0.26979 | 0.01608 | 0.015982 | 0.32909 | 0.67601 | -0.07082 | -0.45134 | -0.22691 | 0.014688 | 0.11829 | 0.24973 | -0.23184
5 | 3.1628 | 0.059605 | 0.031073 | -0.071005 | 0.023586 | -0.27853 | -0.13631 | -0.079433 | -0.058519 | 0.043883 | -0.34893 | -0.12475 | -0.088487 | 0.086248 | 0.004975 | -0.27595 | -0.046058 | 0.044871 | -0.028176 | -0.74268 | 0.014086
6 | -0.4529 | -0.08176 | 0.064554 | -0.026868 | -0.06627 | 0.0044684 | 0.029904 | -0.0036044 | 0.080839 | 0.1432 | 0.045143 | 0.054174 | -0.067826 | -0.053347 | -0.035834 | 0.0096207 | 0.0079035 | -0.10324 | -0.045988 | 0.0052237 | -0.016837
7 | 1.6498 | -0.046555 | 0.14818 | 0.10124 | 0.15531 | 0.076387 | 0.099489 | 0.041924 | 0.073118 | 0.057493 | -0.0075923 | 0.0089715 | 0.080772 | 0.03738 | -0.021103 | -0.087973 | -0.010375 | 0.012749 | -0.0293 | -0.11791 | 0.085504
8 | -0.10358 | -0.0030546 | -0.015676 | -0.018291 | 0.00081525 | 0.01342 | 0.011226 | 0.023201 | 0.021901 | -0.0013218 | -0.0069838 | -0.010009 | -0.027771 | 0.013894 | -0.018158 | -0.011705 | 0.033375 | -0.010818 | 0.007657 | 0.0083885 | -0.0029959
9 | -0.17636 | -0.02058 | -0.0065942 | -0.017936 | -0.033582 | -0.063049 | -0.049505 | -0.043117 | -0.040935 | -0.043693 | -0.048382 | -0.036567 | -0.031361 | -0.026291 | -0.046746 | -0.079196 | -0.069242 | -0.053137 | -0.027397 | -0.010061 | -0.044355
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14524 | -0.02692 | -0.016194 | -0.019983 | -0.013316 | -0.021254 | -0.017972 | -0.024484 | -0.023373 | -0.026248 | -0.03832 | -0.025831 | -0.024662 | -0.027873 | -0.030401 | -0.039657 | -0.029032 | -0.014174 | -0.022606 | 0.0085023 | -0.018016
12 | -0.039181 | 0.0029419 | -0.0075865 | -0.02317 | -0.074398 | -0.052197 | -0.031487 | -0.032895 | 0.0055731 | 0.018181 | -0.020319 | 0.010128 | -0.011245 | -0.010415 | -0.032045 | -0.023826 | -0.052804 | -0.062325 | -0.0139 | 0.055865 | -0.0020887 |  | ten times!
13 | -0.26467 | -0.08736 | -0.039721 | -0.056013 | -0.087272 | -0.088325 | -0.071473 | -0.050401 | -0.047459 | -0.036844 | -0.065328 | -0.034759 | -0.0384 | -0.016637 | -0.016304 | -0.037874 | -0.031281 | -0.016148 | -0.0076031 | 0.015308 | -0.0051182 |  | ten times!
14 | -0.01112 | -0.10375 | -0.14161 | -0.15088 | -0.13927 | -0.15645 | -0.1533 | -0.14705 | -0.13721 | -0.15571 | -0.14164 | -0.14243 | -0.12143 | -0.10289 | -0.13959 | -0.1417 | -0.12534 | -0.10307 | -0.10604 | -0.034134 | -0.012669 |  | ten times!
15 | 0.046131 | -0.093428 | -0.062798 | -0.064384 | -0.079575 | -0.083344 | -0.024126 | -0.086155 | -0.043791 | -0.068107 | -0.077708 | -0.035565 | -0.069278 | -0.065798 | -0.03799 | -0.048045 | -0.038185 | -0.017143 | -0.039867 | -0.20356 | -0.0058833 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 390.1083 | 596.093 | 595.6853 | 595.6246 | 595.6148 | 595.5982 | 595.5827 | 595.5112 | 595.5661 | 595.4866 | 595.4205 | 595.3438 | 595.4662 | 595.2657 | 595.4276 | 595.403 | 595.7008 | 595.8493 | 596.2332 | 371.7657
2 | 0.1167 | -0.81643 | -0.80151 | -0.37015 | -0.3228 | -0.12853 | -0.2832 | -0.3498 | -0.53418 | -0.18186 | -1.081 | -0.17388 | -0.33957 | -0.028665 | 0.049921 | -1.004 | -1.185 | 0.21464 | -0.18069 | 1.4508 | -0.33896
3 | 42.1462 | -0.39416 | -0.23116 | -0.86478 | -0.53337 | -1.4795 | -0.49764 | -0.70264 | -0.86684 | -1.9004 | -1.9 | -1.7796 | -1.7953 | -1.3816 | -2.5316 | -3.1674 | -1.274 | 0.6562 | 0.35782 | 0.35805 | 0.39148
4 | -0.30351 | -0.18022 | -0.16511 | -0.21001 | -0.2628 | -0.049041 | -0.010722 | -0.31044 | -0.18575 | -0.1105 | -0.35332 | -0.20874 | -0.11076 | 0.024263 | -0.027905 | -0.33782 | -0.11066 | -0.078646 | -0.19874 | -0.60815 | -0.17975
5 | -2.0116 | 0.62614 | 0.67201 | 0.82844 | 1.0995 | 1.3804 | 0.92738 | 1.1403 | 1.1561 | 1.1609 | 1.2643 | 1.028 | 0.82313 | 0.8663 | 0.92079 | 0.67472 | 0.71837 | 0.6135 | 0.84869 | -2.5241 | 0.71867
6 | -0.11973 | -0.046364 | -0.057845 | -0.095967 | -0.11613 | 0.014192 | 0.0094264 | -0.036188 | -0.055066 | -0.093221 | 0.035766 | -0.047679 | -0.0045185 | 0.0072573 | -0.03557 | 0.16292 | 0.11623 | -0.086236 | -0.082782 | -0.14899 | -0.030258
7 | 2.795 | 1.1884 | 1.2137 | 1.1113 | 1.1789 | 1.096 | 1.187 | 1.1401 | 1.167 | 1.1191 | 1.1682 | 1.2095 | 1.1456 | 1.1662 | 1.0467 | 1.0457 | 1.1284 | 1.1663 | 1.1804 | 0.894 | 1.1955
8 | -0.077423 | -0.00077862 | -0.015899 | -0.0031879 | 0.02254 | 0.00071551 | -0.012485 | -0.017425 | 0.018544 | -0.0045097 | -0.017042 | -0.0054889 | 0.0041282 | -0.0064168 | -0.0017912 | -0.0027697 | -0.0025549 | -0.0092977 | -0.017217 | 0.0039868 | -0.0061812
9 | 0.52214 | 0.57074 | 0.58554 | 0.59287 | 0.60154 | 0.6079 | 0.62682 | 0.59648 | 0.60186 | 0.59414 | 0.58936 | 0.5929 | 0.57758 | 0.58693 | 0.58965 | 0.56276 | 0.58177 | 0.58653 | 0.57885 | 0.55458 | 0.58676
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.68669 | 0.76807 | 0.77038 | 0.78058 | 0.78122 | 0.79067 | 0.78098 | 0.77344 | 0.77834 | 0.7826 | 0.78016 | 0.77425 | 0.77486 | 0.77224 | 0.78275 | 0.77458 | 0.77293 | 0.77164 | 0.7789 | 0.6785 | 0.77088
12 | -0.10646 | -0.052776 | -0.042515 | -0.0099806 | 0.010429 | -0.054371 | -0.048567 | -0.022318 | 0.0123 | 0.012319 | -0.032202 | -0.01634 | -0.048372 | -0.048883 | -0.0065671 | -0.038319 | -0.047921 | -0.018499 | 0.027994 | 0.034768 | -0.0024513 |  | ten times!
13 | 0.97149 | 0.58025 | 0.58535 | 0.5679 | 0.58317 | 0.53587 | 0.57975 | 0.55801 | 0.55908 | 0.56422 | 0.58964 | 0.5993 | 0.59857 | 0.62994 | 0.55951 | 0.53036 | 0.57763 | 0.58418 | 0.62444 | 0.57805 | 0.059154 |  | ten times!
14 | 0.015038 | -0.12727 | -0.11336 | -0.11977 | -0.11428 | -0.13334 | -0.11391 | -0.11133 | -0.10422 | -0.10737 | -0.10918 | -0.10187 | -0.12063 | -0.10635 | -0.1068 | -0.096437 | -0.10135 | -0.10878 | -0.091084 | -0.08036 | -0.010517 |  | ten times!
15 | 0.02327 | 0.32029 | 0.36762 | 0.38866 | 0.40806 | 0.40671 | 0.39595 | 0.39419 | 0.44073 | 0.37924 | 0.39971 | 0.36792 | 0.39329 | 0.42136 | 0.46649 | 0.42784 | 0.44548 | 0.46774 | 0.44865 | 0.31847 | 0.039182 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 45.9716 | 2.3546 | 2.9344 | 3.183 | 1.9398 | -0.18881 | -0.8618 | -1.8677 | -3.5121 | -3.0624 | -4.2067 | -5.2524 | -3.4579 | -2.7625 | -2.8835 | -5.2537 | -5.3016 | -2.6336 | -0.05605 | 1.2876 | 0.0018304
2 | -3.215 | -0.20158 | -0.23875 | -0.88322 | -1.0774 | -0.73039 | 0.26619 | 0.15237 | 0.82335 | -0.67936 | 0.64222 | -0.39917 | -0.54832 | -1.2549 | -0.89269 | -1.1422 | -0.70271 | -0.95187 | -1.4303 | 3.3043 | -0.48201
3 | -4.1185 | -1.002 | -1.0412 | -0.78656 | -0.36627 | -0.70922 | -0.3732 | -0.49522 | -0.22574 | -0.46604 | -0.48713 | -0.82548 | -1.0632 | -0.39917 | -0.27924 | -0.58793 | -0.61188 | -0.58453 | -0.71199 | 1.1841 | -0.673
4 | 1.1283 | -0.042043 | -0.41298 | -0.24041 | -0.22483 | -0.1725 | -0.33132 | 0.21687 | -0.070466 | 0.66555 | 0.25814 | -0.01773 | -0.32447 | -0.12355 | -0.20719 | 0.30833 | 0.055862 | 0.025706 | 0.30132 | 0.66457 | 0.042443
5 | 2.6383 | -0.25124 | -0.24896 | -0.088995 | 0.030091 | -0.22104 | -0.081877 | -0.10584 | -0.053942 | 0.066363 | -0.47998 | -0.29078 | -0.30966 | -0.029268 | -0.013389 | -0.23461 | -0.1312 | -0.18249 | -0.21074 | -0.78074 | -0.082834
6 | 0.22308 | 0.0042099 | -0.00023071 | -0.13314 | -0.12067 | -0.039884 | 0.025308 | 0.048413 | 0.16449 | 0.05805 | 0.10192 | -0.055534 | 0.0027658 | -0.12411 | -0.087917 | -0.097836 | -0.10286 | -0.064016 | -0.018928 | 0.12303 | -0.011257
7 | 1.4821 | -0.029081 | -0.0093272 | 0.024351 | 0.028467 | -0.049589 | -0.078065 | -0.082332 | -0.037631 | 0.0026465 | -0.14519 | -0.092397 | -0.018005 | 0.027662 | 0.025139 | -0.11465 | -0.105 | -0.098143 | -0.066181 | -0.1075 | 0.0044753
8 | 0.027221 | 0.0061884 | 0.010951 | -0.00063686 | -0.011815 | 0.0013707 | 0.029566 | 0.043172 | 0.050177 | 0.040899 | 0.033572 | -0.037832 | -0.038939 | -0.033167 | -0.024348 | -0.010943 | -0.034487 | 0.01014 | -0.0059501 | -0.0036275 | 0.0022567
9 | -0.21006 | -0.029627 | -0.056835 | -0.05224 | -0.034136 | -0.072348 | -0.070241 | -0.049516 | -0.044516 | -0.069772 | -0.075128 | -0.075791 | -0.072262 | -0.044261 | -0.077836 | -0.090945 | -0.054938 | -0.045563 | -0.053832 | -0.0072313 | -0.062857
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.12264 | -0.050775 | -0.047711 | -0.04802 | -0.045165 | -0.050142 | -0.050975 | -0.060195 | -0.056159 | -0.051114 | -0.064496 | -0.056998 | -0.049777 | -0.046142 | -0.043361 | -0.049806 | -0.056947 | -0.05754 | -0.050106 | -0.022475 | -0.045082
12 | -0.00035289 | -0.010074 | -0.014084 | -0.03575 | -0.036343 | -0.018941 | -0.036099 | 0.0012644 | 0.020804 | -0.014509 | 0.049698 | -0.003495 | 0.0082436 | -0.041633 | -0.031314 | -0.070597 | -0.04637 | -0.058252 | -0.023985 | 0.027609 | -0.0017867 |  | ten times!
13 | -0.1771 | -0.060437 | -0.076534 | -0.095027 | -0.049224 | -0.051591 | -0.040366 | -0.056622 | -0.027061 | -0.062866 | -0.065001 | -0.069702 | -0.067786 | -0.054442 | -0.051703 | -0.081112 | -0.044903 | -0.046935 | -0.091085 | -0.099911 | -0.0065914 |  | ten times!
14 | -0.011034 | -0.17395 | -0.17204 | -0.17163 | -0.1733 | -0.22675 | -0.19713 | -0.20323 | -0.22049 | -0.16609 | -0.19411 | -0.18218 | -0.20166 | -0.18695 | -0.17891 | -0.14122 | -0.15824 | -0.15139 | -0.13823 | -0.023578 | -0.016906 |  | ten times!
15 | -0.037763 | -0.20684 | -0.1496 | -0.14761 | -0.15036 | -0.16492 | -0.13389 | -0.22418 | -0.16096 | -0.19062 | -0.17205 | -0.11755 | -0.17478 | -0.16454 | -0.14072 | -0.14933 | -0.14566 | -0.11686 | -0.15228 | -0.23199 | -0.015728 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00094167 | 0.00021985 | 0.00013939 | 3.9942e-06 | 2.684e-05 | -4.6568e-05 | -0.00012754 | -0.00014484 | -9.0747e-05 | -0.00023286 | -0.00041462 | -0.00032139 | -0.00063635 | -0.00030425 | -0.00022312 | 0.00017836 | 0.00037439 | 0.00065778 |  | 595.3034
|  | 0.00082096 | 0.00013644 | 3.4531e-05 | 1.8077e-05 | -9.794e-06 | -3.5818e-05 | -0.00015586 | -6.3689e-05 | -0.00019717 | -0.00030815 | -0.00043692 | -0.00023142 | -0.00056805 | -0.00029623 | -0.00033753 | 0.00016247 | 0.00041179 | 0.0010564 |  | 595.604
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3808.3231
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 595.2439 | -3771.5398 | 10.4174 | 3.1985 | 2.3938 | 1.0397 | 1.2682 | 0.53402 | -0.27573 | -0.44877 | 0.092189 | -1.3291 | -3.1468 | -2.2144 | -5.3644 | -2.0431 | -1.2316 | 2.7835 | 4.7441 | 7.5782 | -3845.1063
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.12587 | 3.2129 | 0.20619 | 0.56449 | 0.64649 | 0.58312 | 0.26964 | 0.0041297 | -0.074437 | -0.095402 | -0.15265 | -0.31576 | -0.024123 | 0.12559 | -0.29401 | -0.23601 | -0.38914 | -0.43448 | -0.28584 | -0.097788 | 0.68889 |  | 0
db2 | 0.22983 | 0.90229 | 1.2147 | 1.2474 | 0.43199 | -0.24649 | 0.63643 | 0.68681 | 0.32317 | 0.73343 | 0.20386 | 0.050991 | -0.39391 | -0.40838 | -1.0311 | -0.015381 | -1.1048 | -1.9011 | -0.56088 | 0.1333 | 0.60822 |  | 0
db3 | -0.93903 | 45.7327 | 0.91846 | 0.88045 | 0.97611 | 1.0607 | -0.81703 | 0.51193 | 0.34881 | 0.47443 | -0.024343 | -0.76112 | 0.08416 | -0.32934 | -1.0279 | -1.7326 | -1.9623 | -0.39531 | 1.0742 | 0.72061 | 0.47969 |  | 0
db4 | 0.09742 | 0.45926 | -0.011153 | 0.15701 | 0.23463 | -0.12159 | -0.070432 | 0.17175 | 0.1439 | 0.11631 | -0.00095309 | 0.14894 | -0.011797 | -0.1992 | 0.027574 | -0.059117 | -0.051487 | -0.37655 | 0.036457 | -0.13429 | -0.27277 |  | 0
db5 | 1.0159 | -2.3735 | -0.35583 | -0.25503 | -0.059227 | 0.15714 | 0.2663 | 0.11337 | 0.14471 | 0.020702 | 0.35919 | 0.2233 | -0.13596 | -0.11031 | -0.19582 | -0.13403 | -0.0065819 | -0.039173 | 0.024147 | -0.0169 | -2.8183 |  | 3.084e-17
db6 | 0.017347 | 0.19033 | 0.069833 | -0.038891 | 0.092248 | -0.037828 | -0.04328 | 0.0083732 | 0.036473 | -0.036316 | 0.040619 | -0.043182 | -0.0084368 | -0.026889 | 0.014883 | -0.057936 | 0.11002 | -0.082827 | 0.028314 | -0.025174 | -0.038753 |  | 0
db7 | 1.1172 | 2.6485 | -0.0056949 | 0.028105 | 0.0036216 | 0.012256 | -0.039348 | 0.049945 | -0.001511 | 0.018207 | 0.01242 | 0.003565 | 0.039151 | 0.0057976 | 0.024359 | -0.063998 | -0.10624 | -0.031677 | -0.0080399 | 0.059088 | 0.85849 |  | -4.9343e-17
db8 | 0.036098 | 0.1072 | 0.012221 | -0.027616 | -0.013941 | 0.0063013 | 0.055505 | 0.013032 | 0.00070176 | 0.0090296 | 0.025134 | 0.011052 | -0.0048106 | -0.02658 | -0.028752 | -0.0018695 | -0.035772 | 0.010057 | -0.0054522 | 0.0017607 | -0.032281 |  | -6.1679e-18
db9 | 0.59854 | 0.48396 | -0.019125 | -0.019776 | 0.014822 | 0.015848 | 0.0046265 | 0.019714 | 0.0084255 | 0.0083626 | 0.0048241 | 0.02287 | 0.020497 | 0.0021814 | -0.010872 | -0.0037882 | -0.040141 | -0.0067994 | 0.00081321 | -0.022483 | 0.55378 |  | 5.5511e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.78576 | 0.65983 | -0.013304 | -0.009197 | 0.0016804 | -0.00068026 | 0.013313 | 0.0024585 | 0.0027285 | 0.00053854 | 0.010255 | 0.010803 | 0.00098974 | -0.0033018 | -0.0086634 | 0.0054603 | -0.0020337 | 0.0057301 | -0.0030591 | -0.013717 | 0.68993 |  | 6.1679e-17
db12 | 0.0043044 | 0.015299 | 0.0045388 | -0.0016464 | -0.004516 | 0.0015549 | 0.0022575 | -0.0011331 | -0.0021454 | -0.00078057 | -0.0007457 | 0.0016412 | -0.00073031 | 0.00044401 | -0.0039098 | -0.0031633 | 0.005895 | 0.0032976 | -0.0019163 | 0.0010579 | 0.0018101 |  | 0
db13 | 0.058363 | 0.089313 | -0.00024527 | -0.002403 | -0.0028895 | -0.0050068 | -0.00025129 | 0.0014427 | 0.00091725 | 0.00126 | -0.00011109 | 0.003608 | 0.0055222 | 0.0043435 | 0.0072513 | 0.00035022 | -0.0069542 | -0.00091689 | -0.0040968 | -0.0018203 | 0.056206 |  | -1.542e-18
db14 | -0.004799 | 0.010242 | 0.00091396 | 0.0001217 | -0.00062871 | 0.00035064 | -0.00084659 | -0.0019474 | -0.0010044 | -0.0015206 | -0.0001424 | -0.0010359 | -0.0002235 | 0.00036222 | 0.0003577 | -0.00046769 | 0.0012706 | 0.0016703 | 0.0017043 | 0.0010658 | -0.00013533 |  | -3.3731e-19
db15 | 0.032755 | 0.0034794 | -0.0059735 | -0.0014517 | -0.0021959 | -0.0005591 | 0.0021537 | -0.00035061 | 0.0020867 | 0.0013742 | 0.0012884 | 0.0018349 | -0.0007858 | 0.00017956 | 0.0015516 | 0.0019206 | -0.0014074 | 0.0012278 | 0.0006549 | -0.0015483 | 0.023816 |  | 2.8912e-18
da2 | -0.54516 | 0.38225 | -0.14508 | -0.25213 | -0.49581 | -0.34895 | -0.57758 | 0.056447 | 0.38913 | 1.0279 | 0.78596 | 0.26761 | 0.28481 | -0.33965 | -0.94105 | -0.49464 | 0.26615 | 0.25726 | -0.5932 | 0.85283 | 2.6585 |  | 0
da3 | -0.31616 | -2.3314 | -0.34469 | 0.16643 | -0.30071 | -0.086798 | -0.34477 | -0.35004 | -0.1271 | 0.11828 | -0.096149 | -0.66118 | 0.12666 | 0.31176 | 0.61474 | 0.23424 | -0.14029 | 0.22747 | 0.32402 | 0.32813 | 0.68329 |  | 0
da4 | -0.19057 | -1.8992 | -0.26246 | -0.23886 | -0.54452 | -0.35853 | -0.40507 | -0.20319 | 0.0075194 | -0.051844 | -0.079223 | 0.20665 | 0.20655 | 0.51966 | 0.86658 | 0.11975 | -0.26077 | -0.036339 | 0.20526 | 0.30886 | 0.24973 |  | 0
da5 | -0.068995 | 3.1628 | 0.1286 | 0.10007 | -0.0020101 | 0.092581 | -0.20954 | -0.067315 | -0.010437 | 0.010476 | 0.11288 | -0.27994 | -0.055751 | -0.019492 | 0.15524 | 0.07397 | -0.20696 | 0.022937 | 0.11387 | 0.04082 | -0.74268 |  | 0
da6 | -0.0024963 | -0.4529 | -0.079263 | 0.067051 | -0.024372 | -0.063774 | 0.0069647 | 0.0324 | -0.0011081 | 0.083335 | 0.14569 | 0.04764 | 0.05667 | -0.06533 | -0.050851 | -0.033338 | 0.012117 | 0.0104 | -0.10074 | -0.043491 | 0.0052237 |  | 0
da7 | 0.03834 | 1.6498 | -0.084895 | 0.10984 | 0.062903 | 0.11697 | 0.038047 | 0.061149 | 0.0035845 | 0.034778 | 0.019154 | -0.045932 | -0.029368 | 0.042432 | -0.00095976 | -0.059443 | -0.12631 | -0.048715 | -0.025591 | -0.06764 | -0.11791 |  | 0
da8 | 9.4466e-05 | -0.10358 | -0.003149 | -0.01577 | -0.018385 | 0.00072078 | 0.013326 | 0.011132 | 0.023107 | 0.021807 | -0.0014163 | -0.0070783 | -0.010104 | -0.027866 | 0.013799 | -0.018253 | -0.011799 | 0.033281 | -0.010913 | 0.0075625 | 0.0083885 |  | 0
da9 | -0.040962 | -0.17636 | 0.020381 | 0.034367 | 0.023026 | 0.0073799 | -0.022087 | -0.0085435 | -0.0021558 | 2.68e-05 | -0.0027317 | -0.0074206 | 0.0043948 | 0.0096009 | 0.01467 | -0.0057848 | -0.038234 | -0.02828 | -0.012175 | 0.013565 | -0.010061 |  | -3.4694e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.024572 | 0.14524 | -0.002348 | 0.0083778 | 0.0045886 | 0.011257 | 0.0033186 | 0.0066006 | 8.8515e-05 | 0.0011995 | -0.0016757 | -0.013748 | -0.0012585 | -8.9617e-05 | -0.0033011 | -0.0058289 | -0.015085 | -0.0044595 | 0.010398 | 0.0019659 | 0.0085023 |  | 0
da12 | -0.0022877 | -0.0039181 | 0.0025819 | 0.0015291 | -2.9294e-05 | -0.0051521 | -0.002932 | -0.00086096 | -0.0010017 | 0.002845 | 0.0041058 | 0.0002558 | 0.0033005 | 0.0011632 | 0.0012462 | -0.0009168 | -9.485e-05 | -0.0029927 | -0.0039448 | 0.00089769 | 0.0055865 |  | 0
da13 | -0.0046067 | -0.026467 | -0.0041293 | 0.00063457 | -0.00099459 | -0.0041206 | -0.0042258 | -0.0025406 | -0.00043347 | -0.00013922 | 0.00092225 | -0.0019262 | 0.0011308 | 0.00076664 | 0.002943 | 0.0029763 | 0.00081931 | 0.0014786 | 0.0029918 | 0.0038464 | 0.0015308 |  | 0
da14 | -0.013385 | -0.001112 | 0.00301 | -0.00077559 | -0.0017022 | -0.00054183 | -0.0022595 | -0.0019447 | -0.00132 | -0.00033565 | -0.0021861 | -0.00077833 | -0.00085771 | 0.0012428 | 0.0030966 | -0.00057413 | -0.00078449 | 0.0008514 | 0.0030779 | 0.0027815 | -0.0034134 |  | 1.0601e-18
da15 | -0.0057516 | 0.0046131 | -0.0035912 | -0.0005282 | -0.00068681 | -0.0022059 | -0.0025828 | 0.003339 | -0.0028639 | 0.0013725 | -0.0010591 | -0.0020192 | 0.0021951 | -0.0011762 | -0.00082819 | 0.0019526 | 0.00094707 | 0.0019331 | 0.0040373 | 0.0017649 | -0.020356 |  | -6.7461e-19
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -3603.4399
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.5341 | -3449.4386 | 9.384 | 2.5381 | 1.5188 | 1.3542 | 1.0755 | 0.81523 | -0.38537 | 0.53649 | -0.79844 | -1.9084 | -3.1963 | -1.141 | -4.5077 | -1.7892 | -2.2022 | 2.7983 | 5.2919 | 11.7382 | -3757.4411
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.17161 | 4.7688 | 0.40707 | 0.46504 | 0.48991 | 0.36558 | 0.15272 | 0.085426 | -0.015167 | -0.17961 | -0.13464 | -0.24906 | -0.35363 | -0.17419 | -0.10465 | -0.11674 | -0.35377 | -0.35855 | -0.091757 | 0.166 | 0.30037 |  | 0
db2 | 0.41759 | 0.1167 | -0.39884 | -0.38392 | 0.047441 | 0.094791 | 0.28906 | 0.1344 | 0.067795 | -0.11658 | 0.23574 | -0.66338 | 0.24372 | 0.078026 | 0.38893 | 0.46751 | -0.58645 | -0.76737 | 0.63223 | 0.2369 | 1.4508 |  | 0
db3 | -1.1264 | 42.1462 | 0.73283 | 0.89584 | 0.26222 | 0.59362 | -0.35249 | 0.62935 | 0.42436 | 0.26015 | -0.77341 | -0.77297 | -0.65258 | -0.66829 | -0.25459 | -1.4046 | -2.0404 | -0.14704 | 1.7832 | 1.4848 | 0.35805 |  | 0
db4 | 0.16038 | -0.30351 | -0.019831 | -0.0047261 | -0.049626 | -0.10242 | 0.11134 | 0.14966 | -0.15005 | -0.025369 | 0.04988 | -0.19294 | -0.04836 | 0.049625 | 0.18465 | 0.13248 | -0.17743 | 0.049724 | 0.081739 | -0.038353 | -0.60815 |  | -1.542e-17
db5 | 0.92949 | -2.0116 | -0.30436 | -0.25849 | -0.10206 | 0.16895 | 0.44994 | -0.0031238 | 0.20983 | 0.22562 | 0.23036 | 0.33378 | 0.097541 | -0.10737 | -0.0642 | -0.0097052 | -0.25578 | -0.21213 | -0.317 | -0.08181 | -2.5241 |  | 0
db6 | 0.022876 | -0.11973 | -0.023488 | -0.034969 | -0.07309 | -0.093249 | 0.037068 | 0.032303 | -0.013312 | -0.032189 | -0.070345 | 0.058642 | -0.024802 | 0.018358 | 0.030134 | -0.012694 | 0.18579 | 0.13911 | -0.06336 | -0.059906 | -0.14899 |  | 0
db7 | 1.1471 | 2.795 | 0.040702 | 0.066027 | -0.036355 | 0.031245 | -0.051691 | 0.039287 | -0.0075906 | 0.019297 | -0.028579 | 0.02051 | 0.061768 | -0.0021017 | 0.018459 | -0.10099 | -0.10198 | -0.019329 | 0.0186 | 0.032715 | 0.894 |  | 4.9343e-17
db8 | 0.0039409 | -0.077423 | 0.0031623 | -0.011958 | 0.00075297 | 0.026481 | 0.0046564 | -0.0085443 | -0.013484 | 0.022484 | -0.00056877 | -0.013101 | -0.0015479 | 0.0080691 | -0.0024759 | 0.0021497 | 0.0011712 | 0.001386 | -0.0053568 | -0.013276 | 0.0039868 |  | 0
db9 | 0.5897 | 0.52214 | -0.019498 | -0.0046981 | 0.0026318 | 0.011308 | 0.017662 | 0.036586 | 0.0062436 | 0.011622 | 0.0039087 | -0.00087093 | 0.0026675 | -0.012654 | -0.0033058 | -0.00058588 | -0.027469 | -0.0084675 | -0.0036996 | -0.01138 | 0.55458 |  | 4.9343e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.77702 | 0.68669 | -0.009076 | -0.0067617 | 0.0034403 | 0.0040807 | 0.013522 | 0.0038317 | -0.0037025 | 0.0011919 | 0.0054516 | 0.0030136 | -0.0028893 | -0.0022864 | -0.0049 | 0.0056098 | -0.0025605 | -0.0042179 | -0.0055045 | 0.0017574 | 0.6785 |  | 0
db12 | 0.0023588 | -0.010646 | -0.0029187 | -0.0018927 | 0.0013608 | 0.0034017 | -0.0030782 | -0.0024978 | 0.00012702 | 0.0035888 | 0.0035907 | -0.0008614 | 0.00072485 | -0.0024784 | -0.0025295 | 0.0017021 | -0.0014731 | -0.0024333 | 0.00050891 | 0.0051582 | 0.0034768 |  | 0
db13 | 0.057818 | 0.097149 | 0.00020784 | 0.00071744 | -0.0010275 | 0.00049907 | -0.0042307 | 0.00015757 | -0.0020167 | -0.00191 | -0.0013954 | 0.0011459 | 0.0021123 | 0.0020395 | 0.0051763 | -0.0018667 | -0.0047814 | -5.4234e-05 | 0.00060009 | 0.0046267 | 0.057805 |  | 3.8549e-18
db14 | 0.011041 | 0.0015038 | -0.0016864 | -0.00029542 | -0.00093648 | -0.00038754 | -0.002293 | -0.00035006 | -9.241e-05 | 0.00061912 | 0.00030333 | 0.00012224 | 0.00085394 | -0.0010221 | 0.00040568 | 0.00036078 | 0.001397 | 0.00090609 | 0.00016289 | 0.0019323 | -0.008036 |  | 0
db15 | 0.040777 | 0.002327 | -0.0087479 | -0.0040157 | -0.0019114 | 2.8459e-05 | -0.00010628 | -0.0011824 | -0.0013583 | 0.0032952 | -0.0028537 | -0.00080607 | -0.0039853 | -0.0014484 | 0.0013591 | 0.0058716 | 0.0020067 | 0.0037708 | 0.0059962 | 0.0040873 | 0.031847 |  | 0
da2 | -0.51382 | -3.215 | 0.31223 | 0.27507 | -0.3694 | -0.56358 | -0.21657 | 0.78001 | 0.66619 | 1.3372 | -0.16554 | 1.156 | 0.11465 | -0.034506 | -0.74112 | -0.37887 | -0.62834 | -0.18889 | -0.43805 | -0.91647 | 3.3043 |  | 0
da3 | -0.61199 | -4.1185 | -0.38998 | -0.42918 | -0.17456 | 0.24572 | -0.097224 | 0.23879 | 0.11677 | 0.38626 | 0.14595 | 0.12487 | -0.21348 | -0.45117 | 0.21282 | 0.33276 | 0.024065 | 0.00011524 | 0.027466 | -0.099996 | 1.1841 |  | 0
da4 | -0.01865 | 1.1283 | -0.023393 | -0.39433 | -0.22176 | -0.20618 | -0.15385 | -0.31267 | 0.23552 | -0.051816 | 0.6842 | 0.27679 | 0.00092038 | -0.30582 | -0.1049 | -0.18854 | 0.32698 | 0.074512 | 0.044356 | 0.31997 | 0.66457 |  | 0
da5 | -0.15764 | 2.6383 | -0.093594 | -0.091322 | 0.068647 | 0.18773 | -0.063396 | 0.075765 | 0.051801 | 0.1037 | 0.22401 | -0.32234 | -0.13314 | -0.15201 | 0.12837 | 0.14425 | -0.076971 | 0.026438 | -0.024843 | -0.053095 | -0.78074 |  | 0
da6 | -0.024442 | 0.22308 | 0.028652 | 0.024211 | -0.10869 | -0.096232 | -0.015442 | 0.04975 | 0.072855 | 0.18893 | 0.082492 | 0.12636 | -0.031091 | 0.027208 | -0.099663 | -0.063475 | -0.073393 | -0.078414 | -0.039574 | 0.005514 | 0.12303 |  | 0
da7 | -0.045407 | 1.4821 | 0.016326 | 0.03608 | 0.069758 | 0.073874 | -0.0041821 | -0.032658 | -0.036925 | 0.0077759 | 0.048054 | -0.099781 | -0.04699 | 0.027402 | 0.073069 | 0.070546 | -0.069242 | -0.059597 | -0.052736 | -0.020774 | -0.1075 |  | 0
da8 | 0.001551 | 0.027221 | 0.0046374 | 0.0094003 | -0.0021878 | -0.013366 | -0.00018029 | 0.028015 | 0.041621 | 0.048626 | 0.039348 | 0.032021 | -0.039383 | -0.04049 | -0.034718 | -0.025899 | -0.012494 | -0.036038 | 0.0085891 | -0.0075011 | -0.0036275 |  | 0
da9 | -0.059433 | -0.21006 | 0.029805 | 0.0025974 | 0.0071923 | 0.025297 | -0.012915 | -0.010809 | 0.0099164 | 0.014917 | -0.010339 | -0.015695 | -0.016358 | -0.012829 | 0.015171 | -0.018403 | -0.031512 | 0.0044951 | 0.01387 | 0.0056003 | -0.0072313 |  | 6.9389e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.051968 | 0.12264 | 0.0011929 | 0.0042576 | 0.0039481 | 0.0068037 | 0.0018263 | 0.00099292 | -0.0082268 | -0.0041903 | 0.00085388 | -0.012527 | -0.0050296 | 0.0021913 | 0.0058262 | 0.0086075 | 0.0021624 | -0.004979 | -0.0055721 | 0.0018624 | -0.022475 |  | -7.7099e-19
da12 | -0.002008 | -3.5289e-05 | 0.0010006 | 0.00059958 | -0.001567 | -0.0016263 | 0.00011391 | -0.0016019 | 0.0021344 | 0.0040884 | 0.00055705 | 0.0069778 | 0.0016585 | 0.0028323 | -0.0021554 | -0.0011234 | -0.0050518 | -0.002629 | -0.0038172 | -0.00039055 | 0.0027609 |  | 0
da13 | -0.0060689 | -0.01771 | 2.515e-05 | -0.0015845 | -0.0034338 | 0.0011465 | 0.00090974 | 0.0020323 | 0.00040666 | 0.0033628 | -0.00021777 | -0.00043124 | -0.00090129 | -0.00070976 | 0.00062466 | 0.00089855 | -0.0020423 | 0.0015786 | 0.0013754 | -0.0030396 | -0.0099911 |  | 0
da14 | -0.017986 | -0.0011034 | 0.00059121 | 0.00078178 | 0.00082319 | 0.00065646 | -0.0046891 | -0.0017264 | -0.0023373 | -0.0040629 | 0.0013769 | -0.0014252 | -0.00023141 | -0.0021796 | -0.00070923 | 9.5026e-05 | 0.0038642 | 0.0021622 | 0.0028467 | 0.0041636 | -0.0023578 |  | -2.4093e-18
da15 | -0.015904 | -0.0037763 | -0.0047796 | 0.00094392 | 0.0011429 | 0.00086851 | -0.00058742 | 0.0025156 | -0.0065139 | -0.00019237 | -0.0031584 | -0.0013008 | 0.0041492 | -0.0015736 | -0.00054964 | 0.0018326 | 0.00097066 | 0.0013383 | 0.0042185 | 0.00067569 | -0.023199 |  | -7.7099e-19
Field Colinearity | 0.43985
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.085 | 0.081089 | 0.43846 | 0.52084 | 0.45734 | 0.14421 | -0.12198 | -0.20041 | -0.2212 | -0.27849 | -0.44155 | -0.15013 | -0.000162 | -0.41944 | -0.36222 | -0.51528 | -0.56009 | -0.41238 | -0.22354 | 0.55738
Twist (m2 mrad) | 23.1375 | 1.1555 | 5.9193 | 6.6407 | 5.4881 | 1.6224 | -1.2807 | -1.954 | -1.9908 | -2.2975 | -3.3116 | -1.0134 | -0.000972 | -2.202 | -1.63 | -1.9323 | -1.6803 | -0.92786 | -0.3353 | 0.20902
Twist Integral (I0) (m2 rad) | 0.017712
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.5981 | 0.2353 | 0.29316 | 0.31802 | 0.19371 | -0.018919 | -0.086236 | -0.18697 | -0.3513 | -0.306 | -0.42125 | -0.52548 | -0.34576 | -0.27631 | -0.28872 | -0.52513 | -0.52945 | -0.26429 | -0.006081 | 0.12654
Twist (m2 mrad) | 34.4856 | 3.353 | 3.9576 | 4.0547 | 2.3246 | -0.21284 | -0.90548 | -1.823 | -3.1617 | -2.5245 | -3.1594 | -3.547 | -2.0746 | -1.4506 | -1.2992 | -1.9692 | -1.5883 | -0.59466 | -0.0091215 | 0.047453
Twist Integral (I0) (m2 rad) | 0.017928
Note on B160: 4.0s
Note on B166: 3.5s
Note on B195: 3.6s
Note on B204: 4.3s
Note on B205: 4.6s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.22983 | -0.94015 | 0.09742 | 1.0159 | 1.1169 | -1.2587 | -0.54516 | -0.31616 | -0.19057 | 19.4027 | 0.5953 | -0.41759 | -1.127 | -0.16038 | 0.9305 | 1.1477 | -1.7161 | -0.51382 | -0.61199 | -0.01865 | 18.4036 | 0.5956 | 0.26381 | 0.61636 | 0.097287 | 0.7835 | 1.1582 | 0.0020898 | -0.41263 | -0.34945 | -0.23184 | 14428.635 | -0.33896 | 0.39148 | -0.17975 | 0.71867 | 1.1955 | 0.0018304 | -0.48201 | -0.673 | 0.042443 | 14427.344
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8