## Alarm sheet
File name | HCMB__A001-01000359_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 1000359
Date of test Ap 1 | 2006-04-14 |  |  |  | Date of test Ap 2 | 2006-04-14
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | yellow alarm | status ok | status ok | status ok | b3 | yellow alarm | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
1121 | HCMB__A001 | 1000359 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2006-04-18 | Analysis by |  | P. Hagen | Answer to MMS-LD |  |  | 2006-04-18 | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -
| 2005-10-21 | P. Hagen | Check if mole positioning is within +/- 5 cm all positions

## Original data
File name |  | HCMB__A001-01000359_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 1000359 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Alstom | [9 empty] | File name for aperture 1 |  |  | 14-APR-2006_1359_1cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 14-APR-2006_1359_1cc2.txt |  |  |  | a even | 1
Test Operator |  | PERROS |  |  | Test Controller |  |  | P. Hagen
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2006-04-14 |  |  | Time of test |  |  | 15:57 - 19:08 |  |  |  | Date of test |  | 2006-04-14 |  |  | Time of test |  |  | 15:57 - 19:09
Mole name |  | Dipole 04 | [9 empty] | Mole name |  | Dipole 05 | [5 empty] | 2005-05-18
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.54 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.56 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00020091 |  |  | Dy average |  |  | 0.00015561 |  |  |  | Dx average |  | 0.000159 |  |  | Dy average |  |  | 0.00030236
Magnetic length |  | 14.4275 |  |  | Tranfs. Function (mT/KA) |  |  | 595.282 |  |  |  | Magnetic length |  | 14.4274 |  |  | Tranfs. Function (mT/KA) |  |  | 595.4383
Phase |  | 0.025153 |  |  | Integral field (T.m) |  |  | 85.8493 |  |  |  | Phase |  | -0.37833 |  |  | Integral field (T.m) |  |  | 85.8533
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.6546 | 5.9568 | 5.953 | 5.952 | 5.9522 | 5.9524 | 5.9514 | 5.9524 | 5.9519 | 5.9528 | 5.9525 | 5.9514 | 5.9518 | 5.9512 | 5.9523 | 5.9513 | 5.9533 | 5.9542 | 5.9578 | 3.6756 | 85.8493 | C1 (T) | 3.6526 | 5.9588 | 5.9555 | 5.955 | 5.9537 | 5.9539 | 5.9534 | 5.9532 | 5.9536 | 5.9539 | 5.9534 | 5.9528 | 5.9521 | 5.9523 | 5.9536 | 5.9534 | 5.9553 | 5.9552 | 5.9598 | 3.6546 | 85.8533
Angle (mrad) | 3.4257 | -0.016267 | 0.20361 | 0.16141 | 0.2151 | 0.25918 | 0.27841 | 0.034444 | -0.28806 | -0.64013 | -0.57751 | -0.53144 | -0.34677 | -0.32645 | -0.26569 | -0.15207 | -0.46315 | -0.26111 | 0.047433 | 0.91616 | 0.025153 | Angle (mrad) | 3.5755 | -0.066871 | -0.090298 | 0.14303 | 0.22265 | 0.30126 | 0.23577 | 0.034052 | -0.28844 | -0.47651 | -0.30002 | -0.39716 | -0.15214 | -0.17305 | -0.13375 | -0.13763 | -0.31294 | -0.058984 | 0.058829 | -0.97976 | -0.37833
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.11772 | 0.50927 | 0.44483 | 0.30168 | 0.057356 | -0.53625 | -0.059585 | 0.041135 | -0.21684 | -0.40191 | -0.020844 | 0.38315 | -0.025043 | 0.019086 | 0.43802 | 0.062617 | -1.1334 | -1.3229 | -0.57375 | -0.34976 | -0.11324 | b2 | -3.582 | -0.40223 | 0.44794 | 0.43447 | -0.18856 | 0.2359 | 0.79635 | 1.3348 | 1.3408 | 1.6682 | 1.6051 | 0.28085 | 0.6429 | 0.53553 | 0.90632 | 0.30544 | 0.64158 | 0.38364 | 0.55843 | 0.64345 | 0.50573
b3 | 48.4365 | 1.6315 | 1.3055 | 0.97383 | 0.70173 | -0.85166 | -0.52906 | 0.68854 | 0.93541 | -0.19924 | -0.32539 | 0.26067 | 0.79322 | 0.30963 | -0.85904 | -0.7428 | 0.10383 | 1.1244 | 1.8436 | 3.0535 | 2.017 | b3 | 47.3649 | 0.98405 | 0.74845 | 0.46448 | 0.2301 | -0.75806 | 1.0869 | 1.3952 | 0.78181 | 0.081471 | -0.35913 | 1.0347 | 1.2468 | 0.73161 | -0.52041 | -0.76779 | -0.024848 | 1.0743 | 1.4574 | 0.45142 | 1.9878
b4 | 0.77692 | 0.16329 | -0.0096443 | 0.029346 | 0.076696 | 0.12819 | -0.056982 | -0.010594 | 0.17288 | 0.096816 | 0.21398 | -0.033437 | 0.22154 | 0.064871 | -0.10784 | 0.060187 | -0.092779 | -0.31668 | -0.31007 | -0.58845 | 0.020959 | b4 | -1.2974 | 0.022804 | 0.097987 | -0.032337 | 0.11983 | 0.024704 | -0.26927 | -0.020904 | -0.34418 | -0.078733 | 0.010127 | -0.20997 | -0.080368 | -0.16337 | 0.02392 | -0.18767 | -0.23381 | -0.27481 | -0.16466 | -0.28298 | -0.142
b5 | -2.7141 | 0.47177 | 0.54118 | 0.59286 | 0.50969 | 0.68685 | 0.73988 | 0.79614 | 0.67691 | 0.57723 | 0.72383 | 0.84435 | 0.67622 | 0.62044 | 0.43035 | 0.37389 | 0.58183 | 0.48035 | 0.34714 | -3.1248 | 0.36788 | b5 | -2.1893 | 0.53175 | 0.52995 | 0.72011 | 0.66606 | 0.46951 | 0.5263 | 0.53517 | 0.66189 | 0.64998 | 0.49456 | 0.55764 | 0.58835 | 0.53387 | 0.47514 | 0.59773 | 0.64589 | 0.36752 | 0.47946 | -3.1741 | 0.35052
b6 | 0.50243 | 0.017211 | -0.088083 | -0.12329 | -0.083435 | 0.10088 | -0.15876 | -0.059855 | -0.055664 | 0.03915 | -0.021542 | -0.0040483 | 0.065954 | 0.033673 | 0.034124 | -0.020675 | 0.052268 | 0.013815 | -0.025684 | 0.10287 | 0.0045788 | b6 | -0.13341 | 0.06435 | -0.0034147 | -0.028556 | 0.020217 | 0.096653 | -0.17451 | -0.086721 | -0.052112 | -0.017362 | -0.055259 | -0.112 | -0.047472 | 0.00086443 | -0.013454 | 0.066499 | 0.01236 | -0.0065333 | 0.064675 | -0.059451 | -0.020279
b7 | 2.5176 | 1.0851 | 1.1547 | 1.173 | 1.1628 | 1.0991 | 1.1067 | 1.1026 | 1.134 | 1.1257 | 1.0967 | 1.1147 | 1.1537 | 1.1163 | 1.0913 | 1.0465 | 1.1305 | 1.1671 | 1.1593 | 0.83393 | 1.1585 | b7 | 2.876 | 1.1334 | 1.1913 | 1.2187 | 1.167 | 1.0968 | 1.161 | 1.1037 | 1.1475 | 1.1318 | 1.1313 | 1.1528 | 1.121 | 1.1245 | 1.0691 | 1.0882 | 1.1678 | 1.1656 | 1.1455 | 0.93428 | 1.1887
b8 | 0.12407 | 0.063073 | -0.003259 | 0.016231 | 0.0086596 | 0.016849 | -0.072685 | 0.024665 | 0.008094 | 0.024093 | -0.050709 | -0.035005 | 0.021803 | -0.010592 | -0.014288 | 0.008047 | 0.0021614 | -0.023454 | -0.017993 | -0.043561 | 0.0007803 | b8 | -0.029442 | 0.029384 | -0.0043704 | 0.0038903 | 0.010163 | 0.028649 | -0.064025 | 0.028575 | 0.054472 | 0.062158 | 0.03162 | 0.056425 | 0.078559 | 0.051486 | 0.020247 | 0.01137 | 0.017677 | 0.020826 | 0.0040303 | -0.010986 | 0.021651
b9 | 0.42951 | 0.57455 | 0.5654 | 0.56819 | 0.57617 | 0.584 | 0.58955 | 0.589 | 0.59292 | 0.57733 | 0.57768 | 0.5841 | 0.58537 | 0.57706 | 0.56888 | 0.55572 | 0.56144 | 0.56896 | 0.55754 | 0.54335 | 0.56954 | b9 | 0.48244 | 0.55781 | 0.56984 | 0.58159 | 0.58463 | 0.57417 | 0.6204 | 0.60039 | 0.61001 | 0.59182 | 0.58755 | 0.60666 | 0.61358 | 0.60208 | 0.56433 | 0.54264 | 0.54779 | 0.57065 | 0.55358 | 0.50867 | 0.57666
b10 | 0.074971 | 0.038442 | -0.01347 | -0.015996 | 0.010017 | 0.056198 | -0.11507 | 0.00721 | -0.0065562 | 0.049568 | -0.023062 | -0.00013896 | 0.027264 | 0.0061048 | -0.0068087 | 0.013048 | 0.0025683 | -0.011437 | -0.032428 | -0.051012 | 0 | b10 | -0.007316 | 0.056656 | -0.01158 | -0.013785 | -0.0054903 | 0.030909 | -0.13119 | 0.0089248 | 0.011709 | 0.03554 | -0.042536 | 0.014206 | 0.042298 | 0.0486 | 0.029821 | 0.011359 | -0.0049562 | -0.0021186 | -0.039649 | -0.055729 | 0
b11 | 0.65027 | 0.76675 | 0.76822 | 0.7693 | 0.7739 | 0.78272 | 0.77721 | 0.77795 | 0.77943 | 0.77974 | 0.77717 | 0.77304 | 0.77129 | 0.7719 | 0.77404 | 0.76785 | 0.77464 | 0.77522 | 0.76481 | 0.66538 | 0.76621 | b11 | 0.68346 | 0.76782 | 0.76874 | 0.77748 | 0.76891 | 0.7778 | 0.7674 | 0.76431 | 0.76895 | 0.7716 | 0.76889 | 0.7655 | 0.76694 | 0.76173 | 0.77144 | 0.76106 | 0.76697 | 0.75797 | 0.75707 | 0.65219 | 0.76091
b12 | 0.018119 | 0.0057158 | 6.7021e-05 | 0.0025844 | 0.00078953 | 0.00237 | -0.011432 | 0.0028057 | -0.00016907 | 0.0028566 | -0.0028366 | 0.00026184 | 0.0026763 | 0.00088042 | -0.0015236 | -0.00023719 | -0.0024228 | -0.002787 | -0.0029761 | -0.004268 | 0.0002659 | b12 | 0.0013859 | 0.0050495 | -0.0021974 | 0.00025016 | -0.00058213 | 0.0071062 | -0.014193 | -0.0016376 | 0.0047956 | 0.0067357 | 0.0035676 | 0.005293 | 0.0067421 | 0.0032731 | 0.002321 | 0.003905 | 0.0059676 | 0.0032277 | 0.0012297 | -0.0057656 | 0.001985
b13 | 0.095276 | 0.069607 | 0.068481 | 0.070224 | 0.072464 | 0.070848 | 0.070864 | 0.070118 | 0.071665 | 0.071338 | 0.069721 | 0.069347 | 0.071258 | 0.070398 | 0.070373 | 0.067828 | 0.071952 | 0.071285 | 0.07256 | 0.064211 | 0.071158 | b13 | 0.096475 | 0.06887 | 0.070411 | 0.071787 | 0.068251 | 0.070511 | 0.069292 | 0.068161 | 0.067758 | 0.067601 | 0.070683 | 0.068848 | 0.066262 | 0.06562 | 0.069954 | 0.068972 | 0.071699 | 0.070557 | 0.071444 | 0.062946 | 0.069927
b14 | 0.011358 | -0.0028699 | -0.0042474 | -0.0047503 | -0.0040185 | -0.0020789 | -0.0080783 | -0.0040072 | -0.0047666 | -0.0023882 | -0.006327 | -0.0044625 | -0.0033843 | -0.0040334 | -0.0040711 | -0.0038262 | -0.004772 | -0.0033919 | -0.004121 | -0.0058132 | -0.0037547 | b14 | 0.00089624 | -0.0013987 | -0.0035468 | -0.004704 | -0.0043563 | -0.0025093 | -0.0086392 | -0.004381 | -0.0038115 | -0.0020775 | -0.0060164 | -0.0046349 | -0.0038665 | -0.0027811 | -0.0029922 | -0.001237 | -0.0025736 | -0.0018708 | -0.0042846 | -0.0081542 | -0.0036477
b15 | -0.0056232 | 0.023949 | 0.022698 | 0.022293 | 0.023674 | 0.026278 | 0.019498 | 0.0251 | 0.022988 | 0.024453 | 0.024061 | 0.02354 | 0.023523 | 0.024587 | 0.023441 | 0.022547 | 0.022602 | 0.023175 | 0.020272 | 0.0083937 | 0.02186 | b15 | -0.0046406 | 0.025898 | 0.026359 | 0.026435 | 0.026116 | 0.026949 | 0.018128 | 0.025838 | 0.024406 | 0.026329 | 0.023883 | 0.025324 | 0.026295 | 0.025085 | 0.026972 | 0.027113 | 0.027362 | 0.024891 | 0.024445 | 0.017574 | 0.024224
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 34.2575 | -0.16267 | 2.0361 | 1.6141 | 2.151 | 2.5918 | 2.7841 | 0.34444 | -2.8806 | -6.4013 | -5.7751 | -5.3144 | -3.4677 | -3.2645 | -2.6569 | -1.5207 | -4.6315 | -2.6112 | 0.47433 | 9.1616 | 0 | a1 | 35.755 | -0.66871 | -0.90298 | 1.4303 | 2.2265 | 3.0126 | 2.3577 | 0.34052 | -2.8844 | -4.7651 | -3.0002 | -3.9716 | -1.5214 | -1.7305 | -1.3375 | -1.3763 | -3.1294 | -0.58984 | 0.58829 | -9.7976 | 0
a2 | -4.8511 | 0.010375 | -0.32768 | -0.69708 | -0.51766 | -0.84394 | -1.3346 | -0.99657 | -0.53413 | 0.12596 | -1.186 | -0.59484 | -1.3057 | -1.4158 | -1.0148 | -0.21277 | -0.90264 | -0.6027 | -0.85309 | -1.2063 | -0.88013 | a2 | 2.6118 | 0.76185 | -0.72745 | -0.5235 | 0.20005 | 1.0735 | 0.24544 | -0.27769 | 0.41267 | 0.37843 | 0.83328 | 0.35662 | 0.25138 | 0.47609 | 1.4156 | 1.4294 | 2.0236 | 0.65691 | -0.1561 | -6.8337 | 0.32441
a3 | -1.945 | -1.0162 | -0.73492 | -0.46053 | -0.39293 | -0.28992 | -0.23642 | -0.41424 | -0.28093 | -0.38132 | -0.97548 | -0.30476 | -0.5026 | -0.19096 | -0.46143 | -0.54799 | -0.19686 | 0.05 | 0.019757 | 0.87673 | -0.41445 | a3 | -2.4794 | -0.93008 | -0.72566 | -0.56904 | -0.38239 | -0.088072 | -0.026301 | -0.066206 | -0.13609 | -0.038349 | -0.11694 | 0.63217 | -0.053832 | -0.23412 | 0.019689 | -0.024165 | 0.15112 | 0.12116 | 0.23795 | 1.0528 | -0.16146
a4 | -2.4596 | -0.56086 | -0.22359 | 0.031611 | 0.10562 | -0.044449 | 0.15207 | 0.10695 | -0.050953 | -0.21945 | 0.081206 | 0.14151 | 0.38104 | -0.0013997 | -0.31182 | -0.48338 | -0.22776 | -0.40935 | -0.014836 | 0.50687 | -0.14276 | a4 | 0.52752 | 0.1193 | -0.19185 | 0.23214 | 0.17928 | 0.023585 | 0.18808 | -0.0024184 | -0.061388 | -0.2212 | -0.20844 | 0.058632 | 0.20311 | 0.15508 | -0.16357 | -0.31685 | -0.44715 | -0.096677 | -0.17497 | -0.40598 | -0.033864
a5 | 3.263 | -0.046275 | 0.011608 | 0.033819 | -0.041668 | -0.15594 | -0.14784 | -0.091479 | -0.065613 | -0.0055302 | -0.18343 | -0.0012504 | -0.12557 | -0.034439 | -0.075413 | -0.27818 | -0.0076956 | 0.0047965 | -0.026128 | -0.67607 | 0.018184 | a5 | 3.1541 | -0.068789 | -0.1327 | -0.108 | 0.019858 | -0.0896 | 0.0019224 | 0.038992 | 0.034213 | 0.078285 | -0.095236 | 0.15593 | -0.031234 | -0.025782 | 0.06997 | -0.20272 | -0.074777 | -0.070095 | -0.033498 | -0.30626 | 0.063111
a6 | -0.60232 | -0.017976 | 0.042199 | -0.043045 | 0.038471 | 0.051874 | -0.01255 | -0.065663 | 0.034176 | 0.031598 | 0.090206 | -0.015624 | -0.042099 | -0.036724 | 0.030311 | 0.012704 | -0.031871 | -0.022546 | 0.076272 | 0.043204 | -0.011613 | a6 | 0.56773 | 0.15401 | -0.053526 | -0.025029 | -0.032624 | -0.011654 | 0.078259 | -0.048344 | 0.023752 | -0.010741 | 0.040753 | -0.084726 | -0.031623 | 0.033049 | 0.032808 | 0.0021246 | 0.0020881 | 0.042559 | 0.067164 | -0.25622 | 0.019217
a7 | 1.6651 | -0.041478 | -0.058809 | 0.0035816 | -0.017107 | 0.0069349 | -0.022314 | -0.040566 | -0.060579 | 0.0031188 | -0.027924 | -0.015123 | -0.02667 | -0.040287 | -0.065147 | -0.080049 | -0.050933 | -0.039713 | -0.062941 | -0.038273 | 0.018853 | a7 | 1.6431 | -0.039141 | -0.072589 | -0.055291 | -0.068997 | -0.04624 | -0.12083 | -0.13148 | -0.14267 | -0.11568 | -0.10071 | -0.008589 | -0.10559 | -0.12446 | -0.13538 | -0.032392 | -0.073129 | -0.1017 | -0.066375 | -0.011436 | -0.028099
a8 | -0.20834 | -0.033318 | -0.026348 | -0.016356 | -0.020329 | -0.0035377 | 0.014239 | -0.021078 | 0.020986 | 0.011709 | -0.0077009 | -0.0042414 | -0.024394 | -0.025947 | -0.026598 | 0.030176 | 0.011198 | -0.016179 | 0.0015993 | -0.080028 | -0.016299 | a8 | 0.065977 | 0.055503 | -0.0065169 | 0.029671 | -0.0020901 | -0.016715 | -0.0025143 | -0.043559 | -0.012298 | -0.050896 | -0.011347 | -0.031921 | -0.014261 | 0.0054293 | -0.022756 | 0.017075 | 0.025947 | 0.01556 | 0.048185 | -0.11549 | -0.0024857
a9 | -0.2063 | -0.021433 | -0.027745 | -0.033154 | -0.040464 | -0.017259 | -0.038198 | -0.029979 | -0.038282 | -0.057188 | -0.057383 | -0.050356 | -0.050806 | -0.046749 | -0.046543 | -0.052395 | -0.039502 | -0.035721 | -0.031474 | 0.0042123 | -0.043609 | a9 | -0.1878 | -0.049654 | -0.052089 | -0.055128 | -0.06272 | -0.031872 | -0.050582 | -0.052956 | -0.048476 | -0.057073 | -0.027224 | -0.03183 | -0.049062 | -0.060876 | -0.062349 | -0.029111 | -0.012839 | -0.024481 | -0.034956 | 0.017319 | -0.046696
a10 | -0.13776 | -0.03738 | 0.0076197 | 0.0019526 | -0.0048466 | 0.0035677 | 0.036622 | -0.019043 | 0.050178 | 0.011155 | 0.024825 | 0.022152 | -0.0094835 | 0.00034904 | 0.017185 | 0.026945 | 0.01277 | 0.0063339 | 0.016391 | -0.13391 | 0 | a10 | 0.045817 | 0.0052896 | -0.017035 | 0.0041925 | -0.02098 | 0.0044171 | 0.035689 | -0.0051021 | 0.039022 | -0.0053413 | 0.015614 | 0.01534 | 0.0065238 | 0.013191 | -0.019442 | -0.013079 | -0.023867 | 0.019764 | 0.034851 | -0.1909 | 0
a11 | 0.14023 | -0.029464 | -0.029705 | -0.026123 | -0.023456 | -0.026183 | -0.027476 | -0.032543 | -0.033357 | -0.029408 | -0.034623 | -0.028562 | -0.028481 | -0.032891 | -0.029385 | -0.034659 | -0.036407 | -0.031335 | -0.02911 | -0.00078624 | -0.023792 | a11 | 0.14877 | -0.046764 | -0.048985 | -0.046086 | -0.041463 | -0.049595 | -0.04382 | -0.038887 | -0.046746 | -0.050412 | -0.054759 | -0.041279 | -0.046729 | -0.045255 | -0.042687 | -0.057667 | -0.06217 | -0.053164 | -0.055341 | -0.022346 | -0.04131
a12 | -0.027293 | -0.0045797 | 0.0040243 | -0.00035649 | 0.00070088 | -0.00036842 | 0.003659 | -0.0038453 | 0.0064847 | 0.0024874 | 0.0051494 | 0.0031678 | -0.0014202 | -0.0024308 | 0.00059139 | -0.00012984 | 0.00081582 | -0.00090502 | 0.0021489 | -0.020616 | -0.00074328 | a12 | 0.01384 | -0.00070415 | -0.002689 | 0.0005254 | -0.0045175 | 0.00074925 | 0.0024547 | -0.0011851 | 0.0041396 | -0.003036 | 0.002335 | 0.002738 | 0.0011607 | 0.0024847 | -0.0013902 | -0.0080152 | -0.0072313 | -0.0036791 | -0.00045863 | -0.028511 | -0.0013175
a13 | -0.017441 | -0.0055505 | -0.0057358 | -0.0054784 | -0.0036516 | -0.0029819 | -0.0051137 | -0.0045656 | -0.0045543 | -0.0040509 | -0.0079892 | -0.0055027 | -0.0065774 | -0.0042297 | -0.0062584 | -0.006154 | -0.0047579 | -0.0039149 | -0.0026636 | 0.00076019 | -0.0051981 | a13 | -0.017723 | -0.0077209 | -0.0069125 | -0.0064237 | -0.0071571 | -0.0048361 | -0.0026017 | -0.0024393 | -0.0036611 | -0.0040665 | -0.0057582 | -0.0028008 | -0.0064283 | -0.0061984 | -0.007122 | -0.0079597 | -0.0045182 | -0.0058332 | -0.0035504 | 0.0019238 | -0.0054964
a14 | -0.00038569 | -0.0045524 | -0.0038542 | -0.0011183 | -0.0028319 | -0.0031365 | 8.4835e-06 | -0.0028907 | -0.00030895 | -0.0018056 | -0.0024268 | -0.0014884 | -0.0018588 | -0.0024626 | -0.0034091 | -0.0010004 | -0.0016516 | -0.0036906 | -0.0026266 | -0.00107 | -0.0021842 | a14 | 0.0024557 | -0.0052021 | -0.0066086 | -0.0033563 | -0.0048336 | -0.0070561 | -0.0048783 | -0.0085213 | -0.0063393 | -0.0082069 | -0.0061726 | -0.0050137 | -0.0056077 | -0.0054571 | -0.0084232 | -0.0071323 | -0.0071797 | -0.0065083 | -0.0065212 | -0.0055054 | -0.0059754
a15 | 0.0059375 | -0.0032383 | -0.0044625 | -0.0032644 | -0.0023395 | -0.0016461 | -0.002465 | -0.0030863 | -0.0016591 | -0.0021211 | -0.0042431 | -0.0019523 | -0.0029703 | -0.0027731 | -0.0032978 | -0.0032797 | -0.0027117 | -0.0014647 | -0.0017755 | -0.0083492 | -0.0026134 | a15 | 0.0095873 | -0.0040632 | -0.0055437 | -0.0060547 | -0.0067915 | -0.0028655 | -0.0067792 | -0.004834 | -0.0044825 | -0.0027744 | -0.0054368 | -0.0030621 | -0.0042207 | -0.0043332 | -0.0054237 | -0.0037709 | -0.0041709 | -0.0030891 | -0.0036571 | -0.011609 | -0.0042958
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00011134 | -8.8272e-05 | 3.0413e-05 | 3.5452e-05 | -2.2307e-05 | -0.00012166 | 0.00025436 | -1.7462e-05 | 1.894e-05 | -0.000107 | 5.2751e-05 | 0 | -6.0782e-05 | -1.3388e-05 | 1.6363e-05 | -2.6144e-05 | 0 | 2.5602e-05 | 7.3359e-05 | 0.00012915 | 0.00020091 | Dx (m) | 0 | -0.00012429 | 2.3112e-05 | 3.0577e-05 | 0 | -6.668e-05 | 0.0002945 | -2.0375e-05 | -2.0549e-05 | -7.8748e-05 | 9.5996e-05 | -2.9622e-05 | -9.2548e-05 | -0.00010634 | -6.7879e-05 | -2.7424e-05 | 0 | 0 | 9.4254e-05 | 0.00012635 | 0.000159
Dy (m) | 0.00038365 | 7.9513e-05 | -1.5686e-05 | 0 | 0 | -1.1816e-05 | -7.1276e-05 | 4.0882e-05 | -0.00010861 | -2.8367e-05 | -5.1958e-05 | -4.8635e-05 | 1.8661e-05 | 0 | -3.7122e-05 | -6.0835e-05 | -2.8227e-05 | -1.2855e-05 | -3.3647e-05 | 0.00034164 | 0.00015561 | Dy (m) | -0.00011253 | -1.9282e-05 | 3.9145e-05 | 0 | 4.69e-05 | -1.3912e-05 | -6.2434e-05 | 1.0311e-05 | -8.7517e-05 | 0 | -2.7681e-05 | -3.5668e-05 | -2.0111e-05 | -3.577e-05 | 3.9091e-05 | 2.7139e-05 | 5.3433e-05 | -4.378e-05 | -7.135e-05 | 0.00049974 | 0.00030236
Magnetic length (m) |  | 14.4275 | [19 empty] | Magnetic length (m) |  | 14.4274
Transfer funct. (mT/KA) |  | 595.282 | [19 empty] | Transfer funct. (mT/KA) |  | 595.4383
Coil positioning (mm) |  | 0.014292 | [19 empty] | Coil positioning (mm) |  | 0.016734
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.1.5 B | [20 empty] | Soft version | DIMM V2.1.5 B
Level Average | 0.018113 | -0.31273 | 0.12824 | -0.2775 | -0.37648 | -0.21972 | -0.39287 | -0.27234 | -0.3268 | -0.30138 | -0.27194 | -0.36434 | 0.42073 | -0.038057 | -0.26054 | -0.14301 | -0.26708 | 0.19636 | 0.11507 | 0.33895 |  | Level Average | -0.18208 | -0.13526 | -0.070919 | -0.080363 | -0.083473 | -0.093822 | -0.11827 | -0.012065 | -0.056447 | -0.24013 | -0.22805 | -0.3159 | -0.058803 | -0.087265 | -0.1672 | -0.15218 | -0.19077 | -0.067182 | -0.25272 | -0.22063
Level Sigma2 | 0.000674 | 0.038549 | 0.16626 | 0.000907 | 0.002897 | 0.001833 | 0.087449 | 0.052145 | 0.058913 | 0.004434 | 0.000375 | 0.031402 | 0.065815 | 0.001262 | 0.012195 | 0.000604 | 0.00163 | 0.010483 | 0.027275 | 0.1142 |  | Level Sigma2 | 0.68408 | 0.040063 | 0.004202 | 0.006555 | 0.002109 | 0.045819 | 0.043158 | 0.086431 | 0.16662 | 0.27972 | 0.75872 | 0.1458 | 0.00074 | 0.022705 | 0.029361 | 0.000863 | 0.006542 | 0.002873 | 0.009183 | 0.018983
Position (mm) | 0 | 751 | 1500 | 2251 | 3001 | 3751 | 4501 | 5253 | 6002 | 6753 | 7503 | 8254 | 9003 | 9754 | 10504 | 11254 | 12004 | 12754 | 13504 | 14254 |  | Position (mm) | 0 | 753 | 1502 | 2254 | 3003 | 3754 | 4503 | 5256 | 6005 | 6755 | 7504 | 8257 | 9006 | 9756 | 10507 | 11257 | 12007 | 12757 | 13506 | 14257
Fist move (mm) | First move Ap1 | 940.2921 | First move Ap2 | 936.7941 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-1000359 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-1000359 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 365.4626 | 595.6833 | 595.3049 | 595.1987 | 595.2171 | 595.2401 | 595.1386 | 595.2376 | 595.193 | 595.2823 | 595.2524 | 595.1416 | 595.1835 | 595.1175 | 595.2256 | 595.1344 | 595.3274 | 595.4201 | 595.7786 | 367.5581 | 8584.9299 | C1/i (mT/KA) | 365.2614 | 595.8846 | 595.5526 | 595.4958 | 595.3732 | 595.3942 | 595.3442 | 595.3193 | 595.356 | 595.3867 | 595.3354 | 595.2807 | 595.214 | 595.2276 | 595.3597 | 595.3363 | 595.5256 | 595.5182 | 595.9847 | 365.4557 | 8585.3272
Angle (mrad) | 3.4257 | -0.016267 | 0.20361 | 0.16141 | 0.2151 | 0.25918 | 0.27841 | 0.034444 | -0.28806 | -0.64013 | -0.57751 | -0.53144 | -0.34677 | -0.32645 | -0.26569 | -0.15207 | -0.46315 | -0.26111 | 0.047433 | 0.91616 | 0.025153 | Angle (mrad) | 3.5755 | -0.066871 | -0.090298 | 0.14303 | 0.22265 | 0.30126 | 0.23577 | 0.034052 | -0.28844 | -0.47651 | -0.30002 | -0.39716 | -0.15214 | -0.17305 | -0.13375 | -0.13763 | -0.31294 | -0.058984 | 0.058829 | -0.97976 | -0.37833
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.11772 | 0.50927 | 0.44483 | 0.30168 | 0.057356 | -0.53625 | -0.059585 | 0.041135 | -0.21684 | -0.40191 | -0.020844 | 0.38315 | -0.025043 | 0.019086 | 0.43802 | 0.062617 | -1.1334 | -1.3229 | -0.57375 | -0.34976 |  | b2 | -3.582 | -0.40223 | 0.44794 | 0.43447 | -0.18856 | 0.2359 | 0.79635 | 1.3348 | 1.3408 | 1.6682 | 1.6051 | 0.28085 | 0.6429 | 0.53553 | 0.90632 | 0.30544 | 0.64158 | 0.38364 | 0.55843 | 0.64345
b3 | 48.4365 | 1.6315 | 1.3055 | 0.97383 | 0.70173 | -0.85166 | -0.52906 | 0.68854 | 0.93541 | -0.19924 | -0.32539 | 0.26067 | 0.79322 | 0.30963 | -0.85904 | -0.7428 | 0.10383 | 1.1244 | 1.8436 | 3.0535 |  | b3 | 47.3649 | 0.98405 | 0.74845 | 0.46448 | 0.2301 | -0.75806 | 1.0869 | 1.3952 | 0.78181 | 0.081471 | -0.35913 | 1.0347 | 1.2468 | 0.73161 | -0.52041 | -0.76779 | -0.024848 | 1.0743 | 1.4574 | 0.45142
b4 | 0.77692 | 0.16329 | -0.0096443 | 0.029346 | 0.076696 | 0.12819 | -0.056982 | -0.010594 | 0.17288 | 0.096816 | 0.21398 | -0.033437 | 0.22154 | 0.064871 | -0.10784 | 0.060187 | -0.092779 | -0.31668 | -0.31007 | -0.58845 |  | b4 | -1.2974 | 0.022804 | 0.097987 | -0.032337 | 0.11983 | 0.024704 | -0.26927 | -0.020904 | -0.34418 | -0.078733 | 0.010127 | -0.20997 | -0.080368 | -0.16337 | 0.02392 | -0.18767 | -0.23381 | -0.27481 | -0.16466 | -0.28298
b5 | -2.7141 | 0.47177 | 0.54118 | 0.59286 | 0.50969 | 0.68685 | 0.73988 | 0.79614 | 0.67691 | 0.57723 | 0.72383 | 0.84435 | 0.67622 | 0.62044 | 0.43035 | 0.37389 | 0.58183 | 0.48035 | 0.34714 | -3.1248 |  | b5 | -2.1893 | 0.53175 | 0.52995 | 0.72011 | 0.66606 | 0.46951 | 0.5263 | 0.53517 | 0.66189 | 0.64998 | 0.49456 | 0.55764 | 0.58835 | 0.53387 | 0.47514 | 0.59773 | 0.64589 | 0.36752 | 0.47946 | -3.1741
b6 | 0.50243 | 0.017211 | -0.088083 | -0.12329 | -0.083435 | 0.10088 | -0.15876 | -0.059855 | -0.055664 | 0.03915 | -0.021542 | -0.0040483 | 0.065954 | 0.033673 | 0.034124 | -0.020675 | 0.052268 | 0.013815 | -0.025684 | 0.10287 |  | b6 | -0.13341 | 0.06435 | -0.0034147 | -0.028556 | 0.020217 | 0.096653 | -0.17451 | -0.086721 | -0.052112 | -0.017362 | -0.055259 | -0.112 | -0.047472 | 0.00086443 | -0.013454 | 0.066499 | 0.01236 | -0.0065333 | 0.064675 | -0.059451
b7 | 2.5176 | 1.0851 | 1.1547 | 1.173 | 1.1628 | 1.0991 | 1.1067 | 1.1026 | 1.134 | 1.1257 | 1.0967 | 1.1147 | 1.1537 | 1.1163 | 1.0913 | 1.0465 | 1.1305 | 1.1671 | 1.1593 | 0.83393 |  | b7 | 2.876 | 1.1334 | 1.1913 | 1.2187 | 1.167 | 1.0968 | 1.161 | 1.1037 | 1.1475 | 1.1318 | 1.1313 | 1.1528 | 1.121 | 1.1245 | 1.0691 | 1.0882 | 1.1678 | 1.1656 | 1.1455 | 0.93428
b8 | 0.12407 | 0.063073 | -0.003259 | 0.016231 | 0.0086596 | 0.016849 | -0.072685 | 0.024665 | 0.008094 | 0.024093 | -0.050709 | -0.035005 | 0.021803 | -0.010592 | -0.014288 | 0.008047 | 0.0021614 | -0.023454 | -0.017993 | -0.043561 |  | b8 | -0.029442 | 0.029384 | -0.0043704 | 0.0038903 | 0.010163 | 0.028649 | -0.064025 | 0.028575 | 0.054472 | 0.062158 | 0.03162 | 0.056425 | 0.078559 | 0.051486 | 0.020247 | 0.01137 | 0.017677 | 0.020826 | 0.0040303 | -0.010986
b9 | 0.42951 | 0.57455 | 0.5654 | 0.56819 | 0.57617 | 0.584 | 0.58955 | 0.589 | 0.59292 | 0.57733 | 0.57768 | 0.5841 | 0.58537 | 0.57706 | 0.56888 | 0.55572 | 0.56144 | 0.56896 | 0.55754 | 0.54335 |  | b9 | 0.48244 | 0.55781 | 0.56984 | 0.58159 | 0.58463 | 0.57417 | 0.6204 | 0.60039 | 0.61001 | 0.59182 | 0.58755 | 0.60666 | 0.61358 | 0.60208 | 0.56433 | 0.54264 | 0.54779 | 0.57065 | 0.55358 | 0.50867
b10 | 0.074971 | 0.038442 | -0.01347 | -0.015996 | 0.010017 | 0.056198 | -0.11507 | 0.00721 | -0.0065562 | 0.049568 | -0.023062 | -0.00013896 | 0.027264 | 0.0061048 | -0.0068087 | 0.013048 | 0.0025683 | -0.011437 | -0.032428 | -0.051012 |  | b10 | -0.007316 | 0.056656 | -0.01158 | -0.013785 | -0.0054903 | 0.030909 | -0.13119 | 0.0089248 | 0.011709 | 0.03554 | -0.042536 | 0.014206 | 0.042298 | 0.0486 | 0.029821 | 0.011359 | -0.0049562 | -0.0021186 | -0.039649 | -0.055729
b11 | 0.65027 | 0.76675 | 0.76822 | 0.7693 | 0.7739 | 0.78272 | 0.77721 | 0.77795 | 0.77943 | 0.77974 | 0.77717 | 0.77304 | 0.77129 | 0.7719 | 0.77404 | 0.76785 | 0.77464 | 0.77522 | 0.76481 | 0.66538 |  | b11 | 0.68346 | 0.76782 | 0.76874 | 0.77748 | 0.76891 | 0.7778 | 0.7674 | 0.76431 | 0.76895 | 0.7716 | 0.76889 | 0.7655 | 0.76694 | 0.76173 | 0.77144 | 0.76106 | 0.76697 | 0.75797 | 0.75707 | 0.65219
b12 | 0.018119 | 0.0057158 | 6.7021e-05 | 0.0025844 | 0.00078953 | 0.00237 | -0.011432 | 0.0028057 | -0.00016907 | 0.0028566 | -0.0028366 | 0.00026184 | 0.0026763 | 0.00088042 | -0.0015236 | -0.00023719 | -0.0024228 | -0.002787 | -0.0029761 | -0.004268 |  | b12 | 0.0013859 | 0.0050495 | -0.0021974 | 0.00025016 | -0.00058213 | 0.0071062 | -0.014193 | -0.0016376 | 0.0047956 | 0.0067357 | 0.0035676 | 0.005293 | 0.0067421 | 0.0032731 | 0.002321 | 0.003905 | 0.0059676 | 0.0032277 | 0.0012297 | -0.0057656
b13 | 0.095276 | 0.069607 | 0.068481 | 0.070224 | 0.072464 | 0.070848 | 0.070864 | 0.070118 | 0.071665 | 0.071338 | 0.069721 | 0.069347 | 0.071258 | 0.070398 | 0.070373 | 0.067828 | 0.071952 | 0.071285 | 0.07256 | 0.064211 |  | b13 | 0.096475 | 0.06887 | 0.070411 | 0.071787 | 0.068251 | 0.070511 | 0.069292 | 0.068161 | 0.067758 | 0.067601 | 0.070683 | 0.068848 | 0.066262 | 0.06562 | 0.069954 | 0.068972 | 0.071699 | 0.070557 | 0.071444 | 0.062946
b14 | 0.011358 | -0.0028699 | -0.0042474 | -0.0047503 | -0.0040185 | -0.0020789 | -0.0080783 | -0.0040072 | -0.0047666 | -0.0023882 | -0.006327 | -0.0044625 | -0.0033843 | -0.0040334 | -0.0040711 | -0.0038262 | -0.004772 | -0.0033919 | -0.004121 | -0.0058132 |  | b14 | 0.00089624 | -0.0013987 | -0.0035468 | -0.004704 | -0.0043563 | -0.0025093 | -0.0086392 | -0.004381 | -0.0038115 | -0.0020775 | -0.0060164 | -0.0046349 | -0.0038665 | -0.0027811 | -0.0029922 | -0.001237 | -0.0025736 | -0.0018708 | -0.0042846 | -0.0081542
b15 | -0.0056232 | 0.023949 | 0.022698 | 0.022293 | 0.023674 | 0.026278 | 0.019498 | 0.0251 | 0.022988 | 0.024453 | 0.024061 | 0.02354 | 0.023523 | 0.024587 | 0.023441 | 0.022547 | 0.022602 | 0.023175 | 0.020272 | 0.0083937 |  | b15 | -0.0046406 | 0.025898 | 0.026359 | 0.026435 | 0.026116 | 0.026949 | 0.018128 | 0.025838 | 0.024406 | 0.026329 | 0.023883 | 0.025324 | 0.026295 | 0.025085 | 0.026972 | 0.027113 | 0.027362 | 0.024891 | 0.024445 | 0.017574
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 34.2575 | -0.16267 | 2.0361 | 1.6141 | 2.151 | 2.5918 | 2.7841 | 0.34444 | -2.8806 | -6.4013 | -5.7751 | -5.3144 | -3.4677 | -3.2645 | -2.6569 | -1.5207 | -4.6315 | -2.6112 | 0.47433 | 9.1616 |  | a1 | 35.755 | -0.66871 | -0.90298 | 1.4303 | 2.2265 | 3.0126 | 2.3577 | 0.34052 | -2.8844 | -4.7651 | -3.0002 | -3.9716 | -1.5214 | -1.7305 | -1.3375 | -1.3763 | -3.1294 | -0.58984 | 0.58829 | -9.7976
a2 | -4.8511 | 0.010375 | -0.32768 | -0.69708 | -0.51766 | -0.84394 | -1.3346 | -0.99657 | -0.53413 | 0.12596 | -1.186 | -0.59484 | -1.3057 | -1.4158 | -1.0148 | -0.21277 | -0.90264 | -0.6027 | -0.85309 | -1.2063 |  | a2 | 2.6118 | 0.76185 | -0.72745 | -0.5235 | 0.20005 | 1.0735 | 0.24544 | -0.27769 | 0.41267 | 0.37843 | 0.83328 | 0.35662 | 0.25138 | 0.47609 | 1.4156 | 1.4294 | 2.0236 | 0.65691 | -0.1561 | -6.8337
a3 | -1.945 | -1.0162 | -0.73492 | -0.46053 | -0.39293 | -0.28992 | -0.23642 | -0.41424 | -0.28093 | -0.38132 | -0.97548 | -0.30476 | -0.5026 | -0.19096 | -0.46143 | -0.54799 | -0.19686 | 0.05 | 0.019757 | 0.87673 |  | a3 | -2.4794 | -0.93008 | -0.72566 | -0.56904 | -0.38239 | -0.088072 | -0.026301 | -0.066206 | -0.13609 | -0.038349 | -0.11694 | 0.63217 | -0.053832 | -0.23412 | 0.019689 | -0.024165 | 0.15112 | 0.12116 | 0.23795 | 1.0528
a4 | -2.4596 | -0.56086 | -0.22359 | 0.031611 | 0.10562 | -0.044449 | 0.15207 | 0.10695 | -0.050953 | -0.21945 | 0.081206 | 0.14151 | 0.38104 | -0.0013997 | -0.31182 | -0.48338 | -0.22776 | -0.40935 | -0.014836 | 0.50687 |  | a4 | 0.52752 | 0.1193 | -0.19185 | 0.23214 | 0.17928 | 0.023585 | 0.18808 | -0.0024184 | -0.061388 | -0.2212 | -0.20844 | 0.058632 | 0.20311 | 0.15508 | -0.16357 | -0.31685 | -0.44715 | -0.096677 | -0.17497 | -0.40598
a5 | 3.263 | -0.046275 | 0.011608 | 0.033819 | -0.041668 | -0.15594 | -0.14784 | -0.091479 | -0.065613 | -0.0055302 | -0.18343 | -0.0012504 | -0.12557 | -0.034439 | -0.075413 | -0.27818 | -0.0076956 | 0.0047965 | -0.026128 | -0.67607 |  | a5 | 3.1541 | -0.068789 | -0.1327 | -0.108 | 0.019858 | -0.0896 | 0.0019224 | 0.038992 | 0.034213 | 0.078285 | -0.095236 | 0.15593 | -0.031234 | -0.025782 | 0.06997 | -0.20272 | -0.074777 | -0.070095 | -0.033498 | -0.30626
a6 | -0.60232 | -0.017976 | 0.042199 | -0.043045 | 0.038471 | 0.051874 | -0.01255 | -0.065663 | 0.034176 | 0.031598 | 0.090206 | -0.015624 | -0.042099 | -0.036724 | 0.030311 | 0.012704 | -0.031871 | -0.022546 | 0.076272 | 0.043204 |  | a6 | 0.56773 | 0.15401 | -0.053526 | -0.025029 | -0.032624 | -0.011654 | 0.078259 | -0.048344 | 0.023752 | -0.010741 | 0.040753 | -0.084726 | -0.031623 | 0.033049 | 0.032808 | 0.0021246 | 0.0020881 | 0.042559 | 0.067164 | -0.25622
a7 | 1.6651 | -0.041478 | -0.058809 | 0.0035816 | -0.017107 | 0.0069349 | -0.022314 | -0.040566 | -0.060579 | 0.0031188 | -0.027924 | -0.015123 | -0.02667 | -0.040287 | -0.065147 | -0.080049 | -0.050933 | -0.039713 | -0.062941 | -0.038273 |  | a7 | 1.6431 | -0.039141 | -0.072589 | -0.055291 | -0.068997 | -0.04624 | -0.12083 | -0.13148 | -0.14267 | -0.11568 | -0.10071 | -0.008589 | -0.10559 | -0.12446 | -0.13538 | -0.032392 | -0.073129 | -0.1017 | -0.066375 | -0.011436
a8 | -0.20834 | -0.033318 | -0.026348 | -0.016356 | -0.020329 | -0.0035377 | 0.014239 | -0.021078 | 0.020986 | 0.011709 | -0.0077009 | -0.0042414 | -0.024394 | -0.025947 | -0.026598 | 0.030176 | 0.011198 | -0.016179 | 0.0015993 | -0.080028 |  | a8 | 0.065977 | 0.055503 | -0.0065169 | 0.029671 | -0.0020901 | -0.016715 | -0.0025143 | -0.043559 | -0.012298 | -0.050896 | -0.011347 | -0.031921 | -0.014261 | 0.0054293 | -0.022756 | 0.017075 | 0.025947 | 0.01556 | 0.048185 | -0.11549
a9 | -0.2063 | -0.021433 | -0.027745 | -0.033154 | -0.040464 | -0.017259 | -0.038198 | -0.029979 | -0.038282 | -0.057188 | -0.057383 | -0.050356 | -0.050806 | -0.046749 | -0.046543 | -0.052395 | -0.039502 | -0.035721 | -0.031474 | 0.0042123 |  | a9 | -0.1878 | -0.049654 | -0.052089 | -0.055128 | -0.06272 | -0.031872 | -0.050582 | -0.052956 | -0.048476 | -0.057073 | -0.027224 | -0.03183 | -0.049062 | -0.060876 | -0.062349 | -0.029111 | -0.012839 | -0.024481 | -0.034956 | 0.017319
a10 | -0.13776 | -0.03738 | 0.0076197 | 0.0019526 | -0.0048466 | 0.0035677 | 0.036622 | -0.019043 | 0.050178 | 0.011155 | 0.024825 | 0.022152 | -0.0094835 | 0.00034904 | 0.017185 | 0.026945 | 0.01277 | 0.0063339 | 0.016391 | -0.13391 |  | a10 | 0.045817 | 0.0052896 | -0.017035 | 0.0041925 | -0.02098 | 0.0044171 | 0.035689 | -0.0051021 | 0.039022 | -0.0053413 | 0.015614 | 0.01534 | 0.0065238 | 0.013191 | -0.019442 | -0.013079 | -0.023867 | 0.019764 | 0.034851 | -0.1909
a11 | 0.14023 | -0.029464 | -0.029705 | -0.026123 | -0.023456 | -0.026183 | -0.027476 | -0.032543 | -0.033357 | -0.029408 | -0.034623 | -0.028562 | -0.028481 | -0.032891 | -0.029385 | -0.034659 | -0.036407 | -0.031335 | -0.02911 | -0.00078624 |  | a11 | 0.14877 | -0.046764 | -0.048985 | -0.046086 | -0.041463 | -0.049595 | -0.04382 | -0.038887 | -0.046746 | -0.050412 | -0.054759 | -0.041279 | -0.046729 | -0.045255 | -0.042687 | -0.057667 | -0.06217 | -0.053164 | -0.055341 | -0.022346
a12 | -0.027293 | -0.0045797 | 0.0040243 | -0.00035649 | 0.00070088 | -0.00036842 | 0.003659 | -0.0038453 | 0.0064847 | 0.0024874 | 0.0051494 | 0.0031678 | -0.0014202 | -0.0024308 | 0.00059139 | -0.00012984 | 0.00081582 | -0.00090502 | 0.0021489 | -0.020616 |  | a12 | 0.01384 | -0.00070415 | -0.002689 | 0.0005254 | -0.0045175 | 0.00074925 | 0.0024547 | -0.0011851 | 0.0041396 | -0.003036 | 0.002335 | 0.002738 | 0.0011607 | 0.0024847 | -0.0013902 | -0.0080152 | -0.0072313 | -0.0036791 | -0.00045863 | -0.028511
a13 | -0.017441 | -0.0055505 | -0.0057358 | -0.0054784 | -0.0036516 | -0.0029819 | -0.0051137 | -0.0045656 | -0.0045543 | -0.0040509 | -0.0079892 | -0.0055027 | -0.0065774 | -0.0042297 | -0.0062584 | -0.006154 | -0.0047579 | -0.0039149 | -0.0026636 | 0.00076019 |  | a13 | -0.017723 | -0.0077209 | -0.0069125 | -0.0064237 | -0.0071571 | -0.0048361 | -0.0026017 | -0.0024393 | -0.0036611 | -0.0040665 | -0.0057582 | -0.0028008 | -0.0064283 | -0.0061984 | -0.007122 | -0.0079597 | -0.0045182 | -0.0058332 | -0.0035504 | 0.0019238
a14 | -0.00038569 | -0.0045524 | -0.0038542 | -0.0011183 | -0.0028319 | -0.0031365 | 8.4835e-06 | -0.0028907 | -0.00030895 | -0.0018056 | -0.0024268 | -0.0014884 | -0.0018588 | -0.0024626 | -0.0034091 | -0.0010004 | -0.0016516 | -0.0036906 | -0.0026266 | -0.00107 |  | a14 | 0.0024557 | -0.0052021 | -0.0066086 | -0.0033563 | -0.0048336 | -0.0070561 | -0.0048783 | -0.0085213 | -0.0063393 | -0.0082069 | -0.0061726 | -0.0050137 | -0.0056077 | -0.0054571 | -0.0084232 | -0.0071323 | -0.0071797 | -0.0065083 | -0.0065212 | -0.0055054
a15 | 0.0059375 | -0.0032383 | -0.0044625 | -0.0032644 | -0.0023395 | -0.0016461 | -0.002465 | -0.0030863 | -0.0016591 | -0.0021211 | -0.0042431 | -0.0019523 | -0.0029703 | -0.0027731 | -0.0032978 | -0.0032797 | -0.0027117 | -0.0014647 | -0.0017755 | -0.0083492 |  | a15 | 0.0095873 | -0.0040632 | -0.0055437 | -0.0060547 | -0.0067915 | -0.0028655 | -0.0067792 | -0.004834 | -0.0044825 | -0.0027744 | -0.0054368 | -0.0030621 | -0.0042207 | -0.0043332 | -0.0054237 | -0.0037709 | -0.0041709 | -0.0030891 | -0.0036571 | -0.011609
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.11308 | 0.088287 | -0.030415 | -0.035454 | 0.022306 | 0.12166 | -0.2542 | 0.017466 | -0.018949 | 0.107 | -0.052761 | -0.0021026 | 0.060781 | 0.013388 | -0.016363 | 0.026142 | 0.0043097 | -0.025601 | -0.073359 | -0.12993 |  | Dx (mm) | 0.0063095 | 0.12427 | -0.023114 | -0.030577 | -0.0096093 | 0.06667 | -0.29418 | 0.020376 | 0.020566 | 0.078735 | -0.096019 | 0.029624 | 0.092533 | 0.10634 | 0.067878 | 0.027428 | -0.0066536 | -0.0078223 | -0.09425 | -0.12806
Dy (mm) | -0.38452 | -0.079485 | 0.015686 | 0.0031109 | -0.0099702 | 0.011819 | 0.071118 | -0.040882 | 0.10863 | 0.028355 | 0.051952 | 0.048637 | -0.018658 | 0.0013392 | 0.037121 | 0.060835 | 0.028227 | 0.012855 | 0.03364 | -0.34229 |  | Dy (mm) | 0.11259 | 0.01928 | -0.039145 | 0.0073546 | -0.046903 | 0.013905 | 0.062262 | -0.010312 | 0.08752 | -0.006624 | 0.027685 | 0.035665 | 0.020098 | 0.035756 | -0.039088 | -0.027136 | -0.053442 | 0.043778 | 0.071369 | -0.50199
Mag. len. (m) | 14.4275 | [20 empty] | Mag. len. (m) | 14.4274
TF (mT/KA) | 595.282 | [20 empty] | TF (mT/KA) | 595.4383
Coil pos. (mm) | 0.014292 | [20 empty] | Coil pos. (mm) | 0.016734
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.012082 | [20 empty] | Twist (m2 rad) | 0.014691

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 595.282 | 0.18068 | -3.8333e-05 |  | -1.4835 | 3.0459 |  |  | 595.4383 | 0.20486 | -0.0017931 |  | -0.88616 | 2.2256 |  | 0 | 4
2 | -0.13108 |  | -0.11344 | 0.51374 | -0.80388 |  | -0.73237 | 0.45104 | 0.50307 |  | 0.64177 | 0.56436 | 0.30424 |  | 0.48888 | 0.70382 |  | 0 | 5
3 | 2.0194 |  | 0.39702 | 0.84822 | -0.41114 |  | -0.40565 | 0.28496 | 1.989 |  | 0.49361 | 0.73288 | -0.16101 |  | -0.12304 | 0.35903 |  | 0 | 6
4 | 0.012639 |  | 0.016224 | 0.15107 | -0.16498 |  | -0.089051 | 0.24753 | -0.14172 |  | -0.09845 | 0.13508 | -0.045731 |  | -0.041689 | 0.19781 |  | 0 | 7
5 | 0.36869 |  | 0.59196 | 0.14039 | 0.021315 |  | -0.068367 | 0.081527 | 0.3517 |  | 0.55666 | 0.088565 | 0.063102 |  | -0.029121 | 0.08738 |  | 0 | 8
6 | -0.0036747 |  | -0.01493 | 0.050871 | -0.0072738 |  | -0.0014972 | 0.040132 | -0.018462 |  | -0.016675 | 0.059213 | 0.016434 |  | 0.0053263 | 0.051626 |  | 0 | 9
7 | 1.1591 |  | 1.1227 | 0.034443 | 0.019896 |  | -0.034863 | 0.025582 | 1.1886 |  | 1.1389 | 0.03713 | -0.028335 |  | -0.085632 | 0.039494 |  | 0 | 10
8 | 0.0025311 |  | -0.0016451 | 0.019454 | -0.015852 |  | -0.013122 | 0.012989 | 0.02108 |  | 0.023379 | 0.018327 | -0.0023051 |  | -0.0040599 | 0.029966 |  | 0 | 11
9 | 0.57002 |  | 0.57421 | 0.010668 | -0.04267 |  | -0.038868 | 0.012311 | 0.57669 |  | 0.58054 | 0.022236 | -0.047013 |  | -0.043472 | 0.015092 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.76643 |  | 0.77353 | 0.0049164 | -0.023547 |  | -0.030043 | 0.0034376 | 0.76105 |  | 0.76707 | 0.0056018 | -0.041425 |  | -0.048407 | 0.0062179 |  | 0 | 14
12 | 4.1515e-05 |  | -0.00012586 | 0.0019492 | -3.3598e-05 |  | -0.00018504 | 0.0016046 | 0.00019317 |  | 0.0021045 | 0.0030142 | -0.00012901 |  | -0.0013851 | 0.0024419 |  | 0 | 15
13 | 0.0071035 |  | 0.070482 | 0.0013166 | -0.0005037 |  | -0.0048721 | 0.0013336 | 0.0069853 |  | 0.069129 | 0.0017902 | -0.00055283 |  | -0.0052386 | 0.0017697 |  | 0 | 16
14 | -0.00037593 |  | -0.0042699 | 0.00061907 | -0.00025634 |  | -0.0026597 | 0.00091837 | -0.0003633 |  | -0.0038914 | 0.0011921 | -0.00059838 |  | -0.0065067 | 0.0013817 |  | 0 | 17
15 | 0.002186 |  | 0.02326 | 0.0015812 | -0.00026134 |  | -0.0027084 | 0.0008633 | 0.0024224 |  | 0.025435 | 0.0020819 | -0.00042958 |  | -0.0045196 | 0.0012636 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.73112 | -0.46164
3 | [12 empty] | 3 | 9 | -0.71002 | -0.66756
4 | [12 empty] | 4 | 16 | -1.6431 | -1.8112
5 | [12 empty] | 5 | 25 | -2.2351 | -2.4308
6 | [12 empty] | 6 | 36 | -3.097 | -2.8952
7 | [12 empty] | 7 | 49 | -3.5172 | -3.2625
8 | [12 empty] | 8 | 64 | -4.1417 | -3.7536
9 | [12 empty] | 9 | 81 | -4.4689 | -3.9998
10 | [12 empty] | sigma (mm) |  | 0.014292 | 0.016734
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 365.4626 | 595.6833 | 595.3049 | 595.1987 | 595.2171 | 595.2401 | 595.1386 | 595.2376 | 595.193 | 595.2823 | 595.2524 | 595.1416 | 595.1835 | 595.1175 | 595.2256 | 595.1344 | 595.3274 | 595.4201 | 595.7786 | 367.5581
2 | -0.43865 | 0.50183 | 0.44815 | 0.30557 | 0.055975 | -0.52446 | -0.077385 | 0.041713 | -0.21835 | -0.40067 | -0.028826 | 0.38147 | -0.029612 | 0.018568 | 0.43435 | 0.060979 | -1.1341 | -1.3194 | -0.55776 | -0.33839 | -0.13108
3 | 48.5878 | 1.6368 | 1.3048 | 0.97403 | 0.70124 | -0.8545 | -0.52971 | 0.6878 | 0.93501 | -0.20217 | -0.32266 | 0.26188 | 0.78959 | 0.30948 | -0.86139 | -0.74826 | 0.10277 | 1.1221 | 1.8395 | 3.0094 | 2.0194
4 | 0.55391 | 0.15435 | -0.0057286 | 0.034316 | 0.074119 | 0.1081 | -0.015202 | -0.012986 | 0.17422 | 0.082247 | 0.22073 | -0.033033 | 0.21242 | 0.062906 | -0.10684 | 0.053905 | -0.09342 | -0.31378 | -0.30429 | -0.62953 | 0.012639
5 | -2.6627 | 0.47174 | 0.54059 | 0.59154 | 0.51013 | 0.68342 | 0.72774 | 0.79724 | 0.67769 | 0.57626 | 0.72487 | 0.84412 | 0.67527 | 0.62029 | 0.43085 | 0.37428 | 0.5815 | 0.48037 | 0.34734 | -3.1252 | 0.36869
6 | 0.17599 | -0.015436 | -0.076013 | -0.10861 | -0.092529 | 0.053711 | -0.060031 | -0.066067 | -0.050403 | -0.0033315 | -0.0016326 | -0.0034807 | 0.041381 | 0.028379 | 0.039573 | -0.03205 | 0.050041 | 0.02418 | 0.0035835 | 0.14573 | -0.0036747
7 | 2.5448 | 1.0839 | 1.1545 | 1.1732 | 1.1628 | 1.0982 | 1.0995 | 1.1028 | 1.135 | 1.1248 | 1.0954 | 1.1146 | 1.1533 | 1.1164 | 1.0908 | 1.0472 | 1.1306 | 1.1667 | 1.1587 | 0.84288 | 1.1591
8 | 0.13855 | 0.040003 | 0.0046286 | 0.025662 | 0.0028013 | -0.016682 | -0.0034387 | 0.0204 | 0.011424 | -0.0057403 | -0.037768 | -0.035579 | 0.0055054 | -0.014257 | -0.01072 | -0.00028946 | 0.00049799 | -0.016815 | 0.00075562 | -0.011018 | 0.0025311
9 | 0.45306 | 0.57433 | 0.56524 | 0.5679 | 0.57608 | 0.5804 | 0.57545 | 0.58935 | 0.59574 | 0.57469 | 0.57772 | 0.58467 | 0.58459 | 0.57702 | 0.56916 | 0.55641 | 0.56162 | 0.56885 | 0.55657 | 0.5641 | 0.57002
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.65573 | 0.76665 | 0.76826 | 0.76936 | 0.77388 | 0.78253 | 0.7755 | 0.77802 | 0.77988 | 0.77959 | 0.77725 | 0.77314 | 0.77121 | 0.77189 | 0.77403 | 0.76785 | 0.77466 | 0.77517 | 0.76472 | 0.66959 | 0.76643
12 | 0.15248 | 0.016893 | 0.014737 | 0.043299 | -0.0032577 | -0.037393 | 0.010269 | 0.020729 | 0.0044025 | -0.026125 | -0.0053299 | 0.0017584 | -0.0029438 | 0.0021115 | -0.0087474 | -0.017531 | -0.027365 | -0.015343 | 0.0071805 | 0.014374 | 0.00041515 |  | ten times!
13 | 0.94407 | 0.70078 | 0.68336 | 0.70093 | 0.72554 | 0.71013 | 0.69294 | 0.70262 | 0.7157 | 0.71494 | 0.69369 | 0.69284 | 0.71442 | 0.70437 | 0.70225 | 0.67858 | 0.71932 | 0.71182 | 0.72261 | 0.63913 | 0.071035 |  | ten times!
14 | 0.10002 | -0.043992 | -0.037365 | -0.041078 | -0.044342 | -0.047277 | -0.041409 | -0.042643 | -0.045563 | -0.045925 | -0.054631 | -0.044999 | -0.045161 | -0.043075 | -0.03856 | -0.044759 | -0.049153 | -0.029188 | -0.029455 | -0.025616 | -0.0037593 |  | ten times!
15 | -0.056232 | 0.23949 | 0.22698 | 0.22293 | 0.23674 | 0.26278 | 0.19498 | 0.251 | 0.22988 | 0.24453 | 0.24061 | 0.2354 | 0.23523 | 0.24587 | 0.23441 | 0.22547 | 0.22602 | 0.23175 | 0.20272 | 0.083937 | 0.02186 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 34.2924 | -0.16034 | 2.0351 | 1.6126 | 2.1517 | 2.5982 | 2.7644 | 0.34557 | -2.8798 | -6.4015 | -5.7787 | -5.3156 | -3.4631 | -3.2634 | -2.6588 | -1.5206 | -4.6294 | -2.6111 | 0.47178 | 9.1453 | -3.8333e-05
2 | -2.6341 | 0.036186 | -0.33271 | -0.69936 | -0.51581 | -0.83861 | -1.3373 | -0.99241 | -0.54671 | 0.13142 | -1.1901 | -0.59641 | -1.3004 | -1.4156 | -1.0119 | -0.20577 | -0.90288 | -0.60425 | -0.86021 | -1.0699 | -0.80388
3 | -1.8432 | -1.0052 | -0.73609 | -0.46035 | -0.39321 | -0.28923 | -0.22888 | -0.41464 | -0.28442 | -0.37766 | -0.97669 | -0.30443 | -0.50595 | -0.19097 | -0.46163 | -0.5464 | -0.19623 | 0.048869 | 0.021406 | 0.8528 | -0.41114
4 | -2.792 | -0.55108 | -0.22551 | 0.031459 | 0.10703 | -0.041895 | 0.13085 | 0.11498 | -0.068547 | -0.22316 | 0.070081 | 0.13185 | 0.3858 | -0.0014867 | -0.31587 | -0.48702 | -0.23162 | -0.41077 | -0.018034 | 0.23453 | -0.16498
5 | 3.3398 | -0.045406 | 0.012392 | 0.033483 | -0.042166 | -0.15814 | -0.14545 | -0.091862 | -0.063644 | -0.0068511 | -0.1817 | -0.0012022 | -0.12445 | -0.034308 | -0.07564 | -0.27791 | -0.0080892 | 0.0045746 | -0.024228 | -0.66407 | 0.021315
6 | -0.32711 | 0.013757 | 0.035175 | -0.044288 | 0.042697 | 0.046992 | -0.04233 | -0.049503 | -0.0097086 | 0.020214 | 0.069577 | -0.034771 | -0.03393 | -0.037062 | 0.015637 | -0.0090276 | -0.043056 | -0.0282 | 0.060878 | 0.14219 | -0.0072738
7 | 1.6944 | -0.038202 | -0.059118 | 0.003322 | -0.016884 | 0.0070301 | -0.018695 | -0.039999 | -0.060777 | 0.0023217 | -0.027007 | -0.014426 | -0.025892 | -0.040138 | -0.065108 | -0.080575 | -0.050978 | -0.03976 | -0.062644 | -0.048694 | 0.019896
8 | -0.11964 | -0.010937 | -0.030918 | -0.017741 | -0.017201 | -0.0057975 | -0.010061 | -0.0095002 | -0.009666 | 0.0068846 | -0.023249 | -0.01766 | -0.017801 | -0.026016 | -0.036894 | 0.014912 | 0.0038204 | -0.020051 | -0.0083136 | 0.0077498 | -0.015852
9 | -0.18279 | -0.018068 | -0.027511 | -0.033091 | -0.040354 | -0.01784 | -0.028937 | -0.029646 | -0.037401 | -0.058564 | -0.056055 | -0.050328 | -0.050231 | -0.046756 | -0.04626 | -0.053188 | -0.03957 | -0.035557 | -0.03026 | -0.014243 | -0.04267
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14674 | -0.028908 | -0.029626 | -0.026137 | -0.023461 | -0.026173 | -0.026348 | -0.032425 | -0.033265 | -0.029633 | -0.034352 | -0.028566 | -0.028392 | -0.032871 | -0.029342 | -0.034648 | -0.036365 | -0.031327 | -0.028943 | -0.0034647 | -0.023547
12 | -0.00040674 | -0.003283 | 0.031429 | -0.006478 | 0.012684 | -0.0070339 | -0.0081598 | -0.017655 | 0.0092844 | 0.013655 | 0.02295 | 0.0077883 | -0.0019951 | -0.024573 | -0.013249 | -0.02929 | -0.0060336 | -0.016226 | 0.0028791 | -0.050316 | -0.00033598 |  | ten times!
13 | -0.14068 | -0.054175 | -0.057745 | -0.054974 | -0.03634 | -0.026713 | -0.046728 | -0.046522 | -0.041628 | -0.038514 | -0.078357 | -0.053391 | -0.065393 | -0.042004 | -0.061855 | -0.05956 | -0.046495 | -0.039538 | -0.027049 | -0.0086771 | -0.005037 |  | ten times!
14 | -0.027192 | -0.027494 | -0.042592 | -0.012707 | -0.025946 | -0.032273 | -0.016495 | -0.020012 | -0.023914 | -0.021897 | -0.036406 | -0.024346 | -0.013487 | -0.024592 | -0.041701 | -0.020594 | -0.021673 | -0.039668 | -0.032955 | 0.0040267 | -0.0025634 |  | ten times!
15 | 0.059374 | -0.032383 | -0.044625 | -0.032644 | -0.023395 | -0.016461 | -0.02465 | -0.030863 | -0.016591 | -0.021211 | -0.042431 | -0.019523 | -0.029703 | -0.027731 | -0.032978 | -0.032797 | -0.027117 | -0.014647 | -0.017755 | -0.083492 | -0.0026134 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 365.2614 | 595.8846 | 595.5526 | 595.4958 | 595.3732 | 595.3942 | 595.3442 | 595.3193 | 595.356 | 595.3867 | 595.3354 | 595.2807 | 595.214 | 595.2276 | 595.3597 | 595.3363 | 595.5256 | 595.5182 | 595.9847 | 365.4557
2 | -3.65 | -0.41873 | 0.45332 | 0.43565 | -0.18619 | 0.2417 | 0.83377 | 1.3315 | 1.3375 | 1.6674 | 1.6007 | 0.2799 | 0.6292 | 0.52539 | 0.91038 | 0.308 | 0.64062 | 0.38525 | 0.57658 | 0.58807 | 0.50307
3 | 47.3768 | 0.98395 | 0.75018 | 0.46461 | 0.22882 | -0.7583 | 1.075 | 1.3953 | 0.78211 | 0.082823 | -0.35998 | 1.0361 | 1.2488 | 0.73566 | -0.51956 | -0.76536 | -0.020905 | 1.0732 | 1.4524 | 0.48099 | 1.989
4 | -1.2106 | 0.0069436 | 0.10209 | -0.027343 | 0.12111 | 0.017046 | -0.23281 | -0.023564 | -0.34668 | -0.090896 | 0.02068 | -0.21255 | -0.093326 | -0.17694 | 0.015688 | -0.19023 | -0.23186 | -0.27486 | -0.15459 | -0.34245 | -0.14172
5 | -2.1703 | 0.53027 | 0.53054 | 0.7198 | 0.66656 | 0.46757 | 0.51264 | 0.53584 | 0.66281 | 0.6504 | 0.49334 | 0.55772 | 0.58946 | 0.53419 | 0.47503 | 0.59718 | 0.64588 | 0.36805 | 0.48267 | -3.1386 | 0.3517
6 | -0.07452 | 0.014373 | 0.0073066 | -0.015548 | 0.025317 | 0.070618 | -0.056623 | -0.094179 | -0.064848 | -0.048542 | -0.017904 | -0.12416 | -0.084831 | -0.042912 | -0.037198 | 0.056275 | 0.016482 | -0.0048866 | 0.10111 | -0.015197 | -0.018462
7 | 2.8791 | 1.1324 | 1.1914 | 1.2188 | 1.167 | 1.0959 | 1.1531 | 1.1036 | 1.1466 | 1.1299 | 1.1324 | 1.1516 | 1.1179 | 1.1224 | 1.0689 | 1.0879 | 1.1673 | 1.166 | 1.1471 | 0.95757 | 1.1886
8 | -0.040824 | -0.0036862 | 0.0027874 | 0.012068 | 0.014191 | 0.010426 | 0.02038 | 0.023075 | 0.046572 | 0.040408 | 0.057813 | 0.047434 | 0.051376 | 0.020332 | 0.0033678 | 0.0047377 | 0.019715 | 0.022423 | 0.027409 | 0.015577 | 0.02108
9 | 0.48519 | 0.55414 | 0.57006 | 0.58139 | 0.58512 | 0.57311 | 0.60114 | 0.60033 | 0.61169 | 0.59036 | 0.58561 | 0.60673 | 0.61158 | 0.59959 | 0.56366 | 0.54267 | 0.54845 | 0.5711 | 0.55292 | 0.55563 | 0.57669
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.68446 | 0.7674 | 0.76878 | 0.77749 | 0.76905 | 0.7775 | 0.7648 | 0.76434 | 0.76912 | 0.77127 | 0.76915 | 0.76546 | 0.76656 | 0.76156 | 0.77138 | 0.76113 | 0.76725 | 0.75788 | 0.75712 | 0.66097 | 0.76105
12 | -0.0045231 | -0.010966 | -0.0085754 | 0.017662 | 0.0011777 | 0.037404 | 0.00081595 | -0.026002 | 0.035858 | 0.029977 | 0.082458 | 0.037828 | 0.023228 | -0.01809 | -0.008342 | 0.027221 | 0.064748 | 0.034371 | 0.05804 | -0.0075711 | 0.0019317 |  | ten times!
13 | 0.96682 | 0.68926 | 0.70546 | 0.71658 | 0.68392 | 0.70564 | 0.67116 | 0.68297 | 0.67393 | 0.67768 | 0.70111 | 0.68816 | 0.6645 | 0.65697 | 0.70361 | 0.69146 | 0.71979 | 0.70328 | 0.70779 | 0.64261 | 0.069853 |  | ten times!
14 | 0.018093 | -0.041136 | -0.028664 | -0.04075 | -0.038873 | -0.040217 | -0.045949 | -0.047735 | -0.045479 | -0.037696 | -0.042519 | -0.053426 | -0.059401 | -0.051055 | -0.043254 | -0.017651 | -0.022402 | -0.018218 | -0.026022 | -0.015017 | -0.003633 |  | ten times!
15 | -0.046406 | 0.25898 | 0.26359 | 0.26435 | 0.26116 | 0.26949 | 0.18128 | 0.25838 | 0.24406 | 0.26329 | 0.23883 | 0.25324 | 0.26295 | 0.25085 | 0.26972 | 0.27113 | 0.27362 | 0.24891 | 0.24445 | 0.17574 | 0.024224 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 35.7777 | -0.67383 | -0.90293 | 1.4291 | 2.2261 | 3.0082 | 2.3591 | 0.34166 | -2.8918 | -4.7662 | -2.9981 | -3.9729 | -1.5235 | -1.7346 | -1.3411 | -1.3781 | -3.1266 | -0.59053 | 0.58508 | -9.8301 | -0.0017931
2 | 1.9863 | 0.77322 | -0.72597 | -0.52595 | 0.20089 | 1.0754 | 0.23657 | -0.27584 | 0.40495 | 0.37885 | 0.83313 | 0.35008 | 0.24901 | 0.47594 | 1.4131 | 1.427 | 2.0236 | 0.65148 | -0.1657 | -6.7911 | 0.30424
3 | -2.4542 | -0.93277 | -0.72577 | -0.56774 | -0.3811 | -0.08841 | -0.013578 | -0.066235 | -0.13055 | -0.035368 | -0.12053 | 0.63318 | -0.056863 | -0.236 | 0.021814 | -0.02353 | 0.14839 | 0.12315 | 0.23711 | 1.0186 | -0.16101
4 | 0.58083 | 0.1189 | -0.18769 | 0.23012 | 0.18668 | 0.023455 | 0.1805 | -0.0013068 | -0.075184 | -0.22163 | -0.21381 | 0.052866 | 0.20101 | 0.15124 | -0.16032 | -0.31173 | -0.43915 | -0.10059 | -0.18376 | -0.79013 | -0.045731
5 | 3.1575 | -0.074782 | -0.1331 | -0.10816 | 0.020045 | -0.089766 | 0.011889 | 0.039018 | 0.03541 | 0.0785 | -0.093635 | 0.15784 | -0.030093 | -0.026824 | 0.06916 | -0.20221 | -0.074578 | -0.069913 | -0.032994 | -0.32469 | 0.063102
6 | 0.44979 | 0.14801 | -0.037659 | -0.028789 | -0.01354 | -0.015949 | 0.040202 | -0.043382 | -0.010656 | -0.0048806 | 0.026286 | -0.099147 | -0.036127 | 0.023529 | 0.0508 | 0.012861 | 0.023943 | 0.024268 | 0.036103 | -0.091205 | 0.016434
7 | 1.6443 | -0.042215 | -0.072722 | -0.05493 | -0.068809 | -0.045945 | -0.11949 | -0.13099 | -0.14453 | -0.11386 | -0.10152 | -0.0090282 | -0.10569 | -0.12546 | -0.13442 | -0.032458 | -0.072668 | -0.10202 | -0.064624 | -0.019797 | -0.028335
8 | 0.040973 | 0.053346 | 0.0034136 | 0.026864 | 0.01053 | -0.019472 | -0.027695 | -0.040138 | -0.036953 | -0.046937 | -0.020232 | -0.041659 | -0.017928 | -0.0016551 | -0.010384 | 0.024381 | 0.039683 | 0.0037134 | 0.028042 | 0.005716 | -0.0023051
9 | -0.18752 | -0.05058 | -0.052538 | -0.055006 | -0.062963 | -0.032255 | -0.040699 | -0.052852 | -0.049443 | -0.056726 | -0.025807 | -0.032339 | -0.049832 | -0.062539 | -0.061033 | -0.028758 | -0.013063 | -0.02435 | -0.031719 | -0.010434 | -0.047013
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14861 | -0.046771 | -0.049081 | -0.046077 | -0.041509 | -0.049691 | -0.042781 | -0.038882 | -0.047072 | -0.050228 | -0.054677 | -0.041453 | -0.046886 | -0.045502 | -0.042567 | -0.057456 | -0.061995 | -0.053274 | -0.055426 | -0.026581 | -0.041425
12 | 0.062518 | -0.0096416 | -0.0085616 | 0.00014069 | -0.023064 | 0.0028475 | -0.01131 | -0.0065393 | 6.7568e-05 | -0.024939 | 0.005634 | 0.010633 | 0.0064055 | 0.012938 | 0.0088114 | -0.065399 | -0.045477 | -0.058917 | -0.042941 | -0.060325 | -0.0012901 |  | ten times!
13 | -0.17812 | -0.07206 | -0.071355 | -0.064757 | -0.073489 | -0.044497 | -0.032878 | -0.02341 | -0.033063 | -0.035829 | -0.06084 | -0.025609 | -0.05972 | -0.056786 | -0.067743 | -0.078358 | -0.046599 | -0.058095 | -0.037865 | -0.017455 | -0.0055283 |  | ten times!
14 | 0.028362 | -0.051975 | -0.058644 | -0.036689 | -0.038786 | -0.072074 | -0.074502 | -0.082208 | -0.080224 | -0.078834 | -0.07147 | -0.056828 | -0.057213 | -0.058163 | -0.072518 | -0.064412 | -0.059983 | -0.074256 | -0.082418 | 0.0053522 | -0.0059838 |  | ten times!
15 | 0.095873 | -0.040632 | -0.055437 | -0.060547 | -0.067915 | -0.028655 | -0.067792 | -0.04834 | -0.044825 | -0.027744 | -0.054368 | -0.030621 | -0.042207 | -0.043332 | -0.054237 | -0.037709 | -0.041709 | -0.030891 | -0.036571 | -0.11609 | -0.0042958 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00067407 | 3.8404e-05 | -0.00014 | -0.00010909 | -7.0452e-05 | -0.00024096 | -7.4652e-05 | -0.00014957 | 4.3863e-07 | -4.979e-05 | -0.00023592 | -0.00016553 | -0.0002764 | -9.481e-05 | -0.00024802 | 7.6201e-05 | 0.00023193 | 0.00083416 |  | 595.282
|  | 0.00074959 | 0.00019202 | 9.6624e-05 | -0.00010928 | -7.4007e-05 | -0.00015798 | -0.0001998 | -0.00013816 | -8.6603e-05 | -0.00017276 | -0.00026462 | -0.00037664 | -0.0003538 | -0.00013195 | -0.00017125 | 0.00014667 | 0.00013424 | 0.0009177 |  | 595.4383
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3842.5002
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 595.2259 | -3860.1028 | 7.6841 | 1.3268 | -0.45739 | -0.14826 | 0.23814 | -1.4671 | 0.19614 | -0.55315 | 0.94712 | 0.44479 | -1.4167 | -0.71275 | -1.8216 | -0.0054601 | -1.5377 | 1.7048 | 3.2622 | 9.2851 | -3824.8977
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.14835 | 3.5776 | 0.13232 | 0.35186 | 0.30961 | 0.36352 | 0.40817 | 0.42479 | 0.18291 | -0.13963 | -0.4918 | -0.42952 | -0.38321 | -0.19796 | -0.17799 | -0.11753 | -0.0037124 | -0.31459 | -0.11276 | 0.19553 | 1.0629 |  | 0
db2 | -0.11344 | -0.43865 | 0.61528 | 0.56159 | 0.41902 | 0.16942 | -0.41102 | 0.03606 | 0.15516 | -0.1049 | -0.28723 | 0.084619 | 0.49492 | 0.083833 | 0.13201 | 0.54779 | 0.17442 | -1.0207 | -1.206 | -0.44431 | -0.33839 |  | 0
db3 | 0.39804 | 48.5878 | 1.2398 | 0.90778 | 0.57702 | 0.30423 | -1.2515 | -0.92673 | 0.29079 | 0.538 | -0.59918 | -0.71967 | -0.13514 | 0.39258 | -0.087536 | -1.2584 | -1.1453 | -0.29424 | 0.72506 | 1.4425 | 3.0094 |  | 0
db4 | 0.016224 | 0.55391 | 0.13813 | -0.021953 | 0.018092 | 0.057894 | 0.091875 | -0.031427 | -0.02921 | 0.158 | 0.066022 | 0.2045 | -0.049258 | 0.1962 | 0.046682 | -0.12306 | 0.037681 | -0.10964 | -0.33 | -0.32051 | -0.62953 |  | 0
db5 | 0.59283 | -2.6627 | -0.12022 | -0.051372 | -0.00042133 | -0.08183 | 0.091458 | 0.13579 | 0.20528 | 0.085735 | -0.015694 | 0.13291 | 0.25217 | 0.083317 | 0.028335 | -0.16111 | -0.21768 | -0.010458 | -0.11159 | -0.24462 | -3.1252 |  | 0
db6 | -0.01493 | 0.17599 | -0.00050656 | -0.061084 | -0.09368 | -0.077599 | 0.068641 | -0.045101 | -0.051137 | -0.035473 | 0.011598 | 0.013297 | 0.011449 | 0.056311 | 0.043309 | 0.054503 | -0.01712 | 0.064971 | 0.03911 | 0.018513 | 0.14573 |  | 0
db7 | 1.1233 | 2.5448 | -0.038826 | 0.031783 | 0.050503 | 0.040079 | -0.024466 | -0.023208 | -0.019884 | 0.01234 | 0.0021093 | -0.027282 | -0.0080882 | 0.030606 | -0.0063128 | -0.031883 | -0.075486 | 0.0079214 | 0.044056 | 0.036039 | 0.84288 |  | 2.4672e-17
db8 | -0.0016451 | 0.13855 | 0.041649 | 0.0062737 | 0.027308 | 0.0044464 | -0.015037 | -0.0017936 | 0.022046 | 0.013069 | -0.0040952 | -0.036123 | -0.033934 | 0.0071505 | -0.012612 | -0.009075 | 0.0013556 | 0.0021431 | -0.01517 | 0.0024007 | -0.011018 |  | 0
db9 | 0.57521 | 0.45306 | 0.00011986 | -0.0089666 | -0.006313 | 0.0018681 | 0.0061885 | 0.0012365 | 0.015136 | 0.021533 | 0.00047537 | 0.0035137 | 0.010465 | 0.010376 | 0.0028083 | -0.0050516 | -0.017803 | -0.012589 | -0.0053583 | -0.017638 | 0.5641 |  | 6.7847e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.77362 | 0.65573 | -0.0068787 | -0.0052726 | -0.0041762 | 0.00034647 | 0.0089965 | 0.0019622 | 0.0044889 | 0.0063513 | 0.006058 | 0.0037171 | -0.00039777 | -0.0023273 | -0.0016405 | 0.00050116 | -0.0056816 | 0.0011313 | 0.0016355 | -0.0088138 | 0.66959 |  | -3.084e-17
db12 | -0.00012586 | 0.015248 | 0.0018152 | 0.0015996 | 0.0044557 | -0.00019991 | -0.0036134 | 0.0011527 | 0.0021988 | 0.00056611 | -0.0024866 | -0.00040712 | 0.0003017 | -0.00016852 | 0.00033701 | -0.00074888 | -0.0016272 | -0.0026106 | -0.0014085 | 0.00084391 | 0.0014374 |  | 0
db13 | 0.070482 | 0.094407 | -0.00040468 | -0.0021464 | -0.00038962 | 0.0020715 | 0.00053055 | -0.0011885 | -0.00022027 | 0.0010875 | 0.0010114 | -0.0011132 | -0.0011984 | 0.00095944 | -4.5323e-05 | -0.00025759 | -0.0026243 | 0.00145 | 0.00069942 | 0.0017785 | 0.063913 |  | -3.8549e-18
db14 | -0.0042699 | 0.010002 | -0.0001293 | 0.00053333 | 0.00016208 | -0.0001643 | -0.00045789 | 0.00012892 | 5.5613e-06 | -0.00028643 | -0.00032263 | -0.0011932 | -0.00023007 | -0.00024621 | -3.7674e-05 | 0.00041387 | -0.00020603 | -0.00064542 | 0.0013511 | 0.0013244 | -0.0025616 |  | 4.0959e-19
db15 | 0.02326 | -0.0056232 | 0.0006887 | -0.00056197 | -0.00096703 | 0.00041402 | 0.0030183 | -0.0037622 | 0.0018398 | -0.0002714 | 0.0011934 | 0.00080125 | 0.00027978 | 0.00026344 | 0.0013275 | 0.0001807 | -0.0007131 | -0.00065794 | -8.5043e-05 | -0.0029881 | 0.0083937 |  | -4.8187e-18
da2 | -0.73237 | -2.6341 | 0.76856 | 0.39966 | 0.033018 | 0.21657 | -0.10623 | -0.60488 | -0.26003 | 0.18566 | 0.86379 | -0.45774 | 0.13597 | -0.568 | -0.68318 | -0.27955 | 0.52661 | -0.17051 | 0.12813 | -0.12784 | -1.0699 |  | 0
da3 | -0.40565 | -1.8432 | -0.59953 | -0.33044 | -0.054703 | 0.012441 | 0.11642 | 0.17677 | -0.008992 | 0.12123 | 0.02799 | -0.57104 | 0.10122 | -0.1003 | 0.21468 | -0.055977 | -0.14075 | 0.20942 | 0.45452 | 0.42705 | 0.8528 |  | 0
da4 | -0.089051 | -2.792 | -0.46202 | -0.13646 | 0.12051 | 0.19608 | 0.047156 | 0.2199 | 0.20403 | 0.020504 | -0.13411 | 0.15913 | 0.2209 | 0.47485 | 0.087565 | -0.22681 | -0.39797 | -0.14257 | -0.32172 | 0.071017 | 0.23453 |  | 0
da5 | -0.068367 | 3.3398 | 0.022961 | 0.080759 | 0.10185 | 0.026201 | -0.089777 | -0.077086 | -0.023495 | 0.0047233 | 0.061516 | -0.11333 | 0.067165 | -0.056086 | 0.034059 | -0.0072726 | -0.20954 | 0.060278 | 0.072941 | 0.044139 | -0.66407 |  | -1.0794e-17
da6 | -0.0014972 | -0.32711 | 0.015254 | 0.036672 | -0.04279 | 0.044194 | 0.048489 | -0.040833 | -0.048006 | -0.0082114 | 0.021711 | 0.071074 | -0.033274 | -0.032433 | -0.035565 | 0.017134 | -0.0075304 | -0.041559 | -0.026703 | 0.062375 | 0.14219 |  | 0
da7 | -0.034863 | 1.6944 | -0.0033394 | -0.024255 | 0.038185 | 0.017978 | 0.041893 | 0.016168 | -0.0051362 | -0.025914 | 0.037184 | 0.0078561 | 0.020437 | 0.0089705 | -0.0052754 | -0.030245 | -0.045713 | -0.016115 | -0.0048969 | -0.027781 | -0.048694 |  | 0
da8 | -0.013122 | -0.11964 | 0.0021848 | -0.017796 | -0.004619 | -0.0040793 | 0.0073241 | 0.0030602 | 0.0036214 | 0.0034556 | 0.020006 | -0.010127 | -0.0045386 | -0.0046794 | -0.012894 | -0.023772 | 0.028033 | 0.016942 | -0.0069295 | 0.0048081 | 0.0077498 |  | 0
da9 | -0.038868 | -0.18279 | 0.0208 | 0.011357 | 0.0057765 | -0.0014862 | 0.021027 | 0.0099304 | 0.0092212 | 0.0014663 | -0.019696 | -0.017187 | -0.01146 | -0.011364 | -0.0078884 | -0.0073923 | -0.01432 | -0.0007022 | 0.0033105 | 0.0086076 | -0.014243 |  | -5.7824e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.030043 | 0.14674 | 0.0011355 | 0.00041718 | 0.0039066 | 0.0065822 | 0.0038708 | 0.0036956 | -0.002382 | -0.0032218 | 0.00041089 | -0.0043088 | 0.0014773 | 0.0016509 | -0.0028272 | 0.00070169 | -0.0046044 | -0.0063214 | -0.0012836 | 0.0011004 | -0.0034647 |  | 0
da12 | -0.00018504 | -4.0674e-05 | -0.00014326 | 0.003328 | -0.00046276 | 0.0014534 | -0.00051835 | -0.00063094 | -0.0015805 | 0.0011135 | 0.0015506 | 0.00248 | 0.00096387 | -1.4469e-05 | -0.0022723 | -0.0011399 | -0.0027439 | -0.00041832 | -0.0014376 | 0.00047294 | -0.0050316 |  | 0
da13 | -0.0048721 | -0.014068 | -0.00054544 | -0.00090236 | -0.00062529 | 0.0012382 | 0.0022008 | 0.00019936 | 0.00021987 | 0.00070932 | 0.0010207 | -0.0029636 | -0.00046702 | -0.0016672 | 0.00067174 | -0.0013134 | -0.0010839 | 0.00022265 | 0.00091828 | 0.0021672 | -0.00086771 |  | 0
da14 | -0.0026597 | -0.0027192 | -8.9624e-05 | -0.0015994 | 0.001389 | 6.5147e-05 | -0.0005676 | 0.0010102 | 0.00065851 | 0.00026831 | 0.00047004 | -0.00098091 | 0.00022511 | 0.001311 | 0.00020056 | -0.0015104 | 0.00060035 | 0.00049239 | -0.0013071 | -0.00063573 | 0.00040267 |  | 1.6865e-19
da15 | -0.0027084 | 0.0059375 | -0.00052991 | -0.0017542 | -0.00055606 | 0.00036883 | 0.0010622 | 0.00024334 | -0.00037792 | 0.0010492 | 0.00058731 | -0.0015347 | 0.00075603 | -0.00026194 | -6.476e-05 | -0.00058941 | -0.00057129 | -3.3081e-06 | 0.0012437 | 0.00093286 | -0.0083492 |  | -2.5298e-19
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -3863.4003
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.3762 | -3865.0321 | 8.5388 | 2.9625 | 2.0085 | -0.050703 | 0.30201 | -0.53779 | -0.95601 | -0.3396 | 0.17604 | -0.6856 | -1.6043 | -2.7246 | -2.4962 | -0.27745 | -0.67048 | 2.509 | 2.3847 | 10.2201 | -3861.7686
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.088616 | 3.6664 | 0.021234 | -0.001677 | 0.23153 | 0.31123 | 0.38944 | 0.32452 | 0.12278 | -0.20056 | -0.388 | -0.21119 | -0.30867 | -0.063737 | -0.084843 | -0.045494 | -0.049196 | -0.22404 | 0.029563 | 0.14712 | -0.89439 |  | 0
db2 | -0.64177 | -3.65 | -1.0605 | -0.18846 | -0.20613 | -0.82796 | -0.40007 | 0.192 | 0.68973 | 0.6957 | 1.0257 | 0.95891 | -0.36188 | -0.012574 | -0.11638 | 0.26861 | -0.33378 | -0.0011595 | -0.25652 | -0.065192 | 0.58807 |  | 1.542e-17
db3 | 0.49372 | 47.3768 | 0.49035 | 0.25657 | -0.028997 | -0.26479 | -1.2519 | 0.58139 | 0.90171 | 0.2885 | -0.41079 | -0.85359 | 0.54253 | 0.7552 | 0.24205 | -1.0132 | -1.259 | -0.51451 | 0.57961 | 0.95881 | 0.48099 |  | 0
db4 | 0.09845 | -1.2106 | 0.10539 | 0.20054 | 0.071107 | 0.21956 | 0.1155 | -0.13436 | 0.074885 | -0.24823 | 0.0075535 | 0.11913 | -0.1141 | 0.0051242 | -0.078491 | 0.11414 | -0.091785 | -0.13341 | -0.17641 | -0.05614 | -0.34245 |  | 0
db5 | 0.55727 | -2.1703 | -0.02639 | -0.026124 | 0.16314 | 0.1099 | -0.089097 | -0.044028 | -0.020826 | 0.10615 | 0.093738 | -0.063328 | 0.0010602 | 0.032795 | -0.022475 | -0.081633 | 0.040513 | 0.089216 | -0.18861 | -0.073998 | -3.1386 |  | -1.542e-17
db6 | 0.016675 | -0.07452 | 0.031048 | 0.023982 | 0.0011274 | 0.041992 | 0.087294 | -0.039948 | -0.077504 | -0.048173 | -0.031867 | -0.0012286 | -0.10749 | -0.068156 | -0.026237 | -0.020523 | 0.07295 | 0.033157 | 0.011789 | 0.11778 | -0.015197 |  | 0
db7 | 1.1398 | 2.8791 | -0.0065413 | 0.052469 | 0.079947 | 0.028138 | -0.042993 | 0.014248 | -0.035285 | 0.0076681 | -0.0089838 | -0.0064556 | 0.012702 | -0.021011 | -0.016537 | -0.070012 | -0.050966 | 0.028386 | 0.027063 | 0.008162 | 0.95757 |  | -7.4015e-17
db8 | -0.023379 | -0.040824 | -0.027066 | -0.020592 | -0.011311 | -0.009188 | -0.012953 | -0.002999 | -0.00030462 | 0.023192 | 0.017029 | 0.034434 | 0.024054 | 0.027997 | -0.0030473 | -0.020012 | -0.018642 | -0.0036648 | -0.00095669 | 0.0040297 | 0.015577 |  | 1.3492e-18
db9 | 0.5822 | 0.48519 | -0.026398 | -0.01048 | 0.0008515 | 0.0045864 | -0.0074238 | 0.020606 | 0.01979 | 0.031158 | 0.0098223 | 0.0050777 | 0.026196 | 0.03104 | 0.019056 | -0.016879 | -0.037868 | -0.032084 | -0.0094314 | -0.027617 | 0.55563 |  | 6.7847e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.76726 | 0.68446 | 0.00033322 | 0.0017103 | 0.01042 | 0.0019782 | 0.010435 | -0.0022704 | -0.0027312 | 0.0020503 | 0.0041996 | 0.0020843 | -0.001609 | -0.00051325 | -0.0055042 | 0.0043066 | -0.0059382 | 0.00017788 | -0.0091845 | -0.0099449 | 0.66097 |  | 3.084e-17
db12 | -0.0021045 | -0.00045231 | -0.0032011 | -0.002962 | -0.00033826 | -0.0019867 | 0.0016359 | -0.0020229 | -0.0047047 | 0.0014813 | 0.00089316 | 0.0061413 | 0.0016783 | 0.00021832 | -0.0039135 | -0.0029387 | 0.00061757 | 0.0043703 | 0.0013326 | 0.0036995 | -0.00075711 |  | 0
db13 | 0.069129 | 0.096682 | -0.00020343 | 0.0014172 | 0.0025293 | -0.00073709 | 0.0014347 | -0.0020133 | -0.00083266 | -0.0017357 | -0.0013617 | 0.00098136 | -0.00031305 | -0.0026796 | -0.0034327 | 0.0012315 | 1.6608e-05 | 0.0028497 | 0.0011989 | 0.00165 | 0.064261 |  | -2.313e-18
db14 | 0.0038914 | 0.0018093 | -0.00022225 | 0.001025 | -0.00018362 | 4.049e-06 | -0.00013033 | -0.00070355 | -0.00088215 | -0.00065656 | 0.00012181 | -0.00036051 | -0.0014513 | -0.0020487 | -0.0012141 | -0.00043398 | 0.0021262 | 0.0016512 | 0.0020696 | 0.0012892 | -0.0015017 |  | 0
db15 | 0.025435 | -0.0046406 | 0.00046346 | 0.0009242 | 0.0010003 | 0.0006807 | 0.0015139 | -0.0073068 | 0.0004032 | -0.0010293 | 0.00089434 | -0.0015523 | -0.00011093 | 0.00085979 | -0.00034968 | 0.0015372 | 0.0016783 | 0.0019269 | -0.00054359 | -0.00098983 | 0.017574 |  | -2.1202e-18
da2 | 0.48888 | 1.9863 | 0.28434 | -1.2149 | -1.0148 | -0.28799 | 0.58656 | -0.25231 | -0.76472 | -0.083929 | -0.11002 | 0.34425 | -0.1388 | -0.23986 | -0.012938 | 0.92421 | 0.93816 | 1.5347 | 0.16261 | -0.65458 | -6.7911 |  | 0
da3 | -0.12304 | -2.4542 | -0.80973 | -0.60272 | -0.4447 | -0.25805 | 0.034634 | 0.10947 | 0.056809 | -0.0075049 | 0.087676 | 0.002518 | 0.75622 | 0.066181 | -0.11296 | 0.14486 | 0.099514 | 0.27143 | 0.24619 | 0.36016 | 1.0186 |  | 0
da4 | -0.041689 | 0.58083 | 0.16059 | -0.146 | 0.27181 | 0.22837 | 0.065144 | 0.22219 | 0.040383 | -0.033495 | -0.17995 | -0.17212 | 0.094555 | 0.2427 | 0.19292 | -0.11863 | -0.27004 | -0.39746 | -0.058902 | -0.14207 | -0.79013 |  | 0
da5 | -0.029121 | 3.1575 | -0.045661 | -0.10398 | -0.079038 | 0.049166 | -0.060645 | 0.041011 | 0.06814 | 0.064532 | 0.10762 | -0.064513 | 0.18697 | -0.00097125 | 0.0022971 | 0.098282 | -0.17308 | -0.045457 | -0.040792 | -0.0038723 | -0.32469 |  | 0
da6 | 0.0053263 | 0.44979 | 0.14268 | -0.042985 | -0.034116 | -0.018867 | -0.021275 | 0.034876 | -0.048708 | -0.015983 | -0.010207 | 0.020959 | -0.10447 | -0.041453 | 0.018202 | 0.045474 | 0.0075346 | 0.018617 | 0.018942 | 0.030776 | -0.091205 |  | 0
da7 | -0.085632 | 1.6443 | 0.043418 | 0.012911 | 0.030703 | 0.016823 | 0.039687 | -0.033857 | -0.045358 | -0.058896 | -0.02823 | -0.015887 | 0.076604 | -0.02006 | -0.039828 | -0.048787 | 0.053175 | 0.012964 | -0.01639 | 0.021009 | -0.019797 |  | 0
da8 | -0.0040599 | 0.040973 | 0.057406 | 0.0074736 | 0.030924 | 0.01459 | -0.015412 | -0.023635 | -0.036078 | -0.032893 | -0.042877 | -0.016172 | -0.037599 | -0.013868 | 0.0024049 | -0.0063244 | 0.028441 | 0.043743 | 0.0077733 | 0.032102 | 0.005716 |  | 0
da9 | -0.043472 | -0.18752 | -0.0071076 | -0.0090654 | -0.011534 | -0.019491 | 0.011217 | 0.002773 | -0.0093801 | -0.0059708 | -0.013253 | 0.017666 | 0.011133 | -0.0063595 | -0.019066 | -0.017561 | 0.014714 | 0.030409 | 0.019123 | 0.011753 | -0.010434 |  | -3.2767e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.048407 | 0.14861 | 0.0016363 | -0.00067356 | 0.0023304 | 0.0068984 | -0.0012841 | 0.0056262 | 0.0095247 | 0.0013349 | -0.0018211 | -0.0062702 | 0.0069536 | 0.0015211 | 0.0029051 | 0.0058403 | -0.0090487 | -0.013587 | -0.0048672 | -0.007019 | -0.026581 |  | -4.6259e-18
da12 | -0.0013851 | 0.0062518 | 0.0004209 | 0.0005289 | 0.0013991 | -0.00092134 | 0.0016698 | 0.00025411 | 0.00073114 | 0.0013918 | -0.0011088 | 0.0019485 | 0.0024484 | 0.0020256 | 0.0026788 | 0.0022662 | -0.0051548 | -0.0031627 | -0.0045066 | -0.002909 | -0.0060325 |  | 0
da13 | -0.0052386 | -0.017812 | -0.0019674 | -0.0018968 | -0.0012371 | -0.0021103 | 0.00078894 | 0.0019509 | 0.0028976 | 0.0019324 | 0.0016557 | -0.00084542 | 0.0026778 | -0.00073343 | -0.00043996 | -0.0015356 | -0.0025972 | 0.0005787 | -0.00057084 | 0.0014521 | -0.0017455 |  | 4.0959e-19
da14 | -0.0065067 | 0.0028362 | 0.0013091 | 0.00064223 | 0.0028378 | 0.0026281 | -0.0007007 | -0.00094357 | -0.0017142 | -0.0015157 | -0.0013767 | -0.00064039 | 0.0008239 | 0.00078535 | 0.00069037 | -0.00074517 | 6.5409e-05 | 0.00050832 | -0.00091893 | -0.0017352 | 0.00053522 |  | 8.4327e-19
da15 | -0.0045196 | 0.0095873 | 0.00045638 | -0.0010241 | -0.0015351 | -0.0022719 | 0.0016541 | -0.0022596 | -0.00031437 | 3.7105e-05 | 0.0017452 | -0.00091718 | 0.0014575 | 0.00029893 | 0.00018639 | -0.00090406 | 0.00074876 | 0.00034875 | 0.0014306 | 0.00086255 | -0.011609 |  | -2.8912e-19
Field Colinearity | 0.40348
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.4257 | -0.016267 | 0.20361 | 0.16141 | 0.2151 | 0.25918 | 0.27841 | 0.034444 | -0.28806 | -0.64013 | -0.57751 | -0.53144 | -0.34677 | -0.32645 | -0.26569 | -0.15207 | -0.46315 | -0.26111 | 0.047433 | 0.91616
Twist (m2 mrad) | 25.693 | -0.2318 | 2.7487 | 2.0579 | 2.5812 | 2.9158 | 2.9233 | 0.33583 | -2.5925 | -5.2811 | -4.3313 | -3.5872 | -2.0806 | -1.7139 | -1.1956 | -0.57027 | -1.3895 | -0.58751 | 0.07115 | 0.34356
Twist Integral (I0) (m2 rad) | 0.012082
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.5755 | -0.066871 | -0.090298 | 0.14303 | 0.22265 | 0.30126 | 0.23577 | 0.034052 | -0.28844 | -0.47651 | -0.30002 | -0.39716 | -0.15214 | -0.17305 | -0.13375 | -0.13763 | -0.31294 | -0.058984 | 0.058829 | -0.97976
Twist (m2 mrad) | 26.8161 | -0.95291 | -1.219 | 1.8236 | 2.6718 | 3.3892 | 2.4756 | 0.33201 | -2.596 | -3.9312 | -2.2501 | -2.6809 | -0.91284 | -0.90851 | -0.60189 | -0.51611 | -0.93883 | -0.13271 | 0.088244 | -0.36741
Twist Integral (I0) (m2 rad) | 0.014691
Note on B158: 3.7s
Note on B193: 3.8s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.11344 | 0.39702 | 0.016224 | 0.59196 | 1.1227 | -1.4835 | -0.73237 | -0.40565 | -0.089051 | 14.2919 | 0.59528 | 0.64177 | 0.49361 | -0.09845 | 0.55666 | 1.1389 | -0.88616 | 0.48888 | -0.12304 | -0.041689 | 16.7338 | 0.59544 | -0.13108 | 2.0194 | 0.012639 | 0.36869 | 1.1591 | -3.8333e-05 | -0.80388 | -0.41114 | -0.16498 | 14427.515 | 0.50307 | 1.989 | -0.14172 | 0.3517 | 1.1886 | -0.0017931 | 0.30424 | -0.16101 | -0.045731 | 14427.371
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8