## Alarm sheet
File name | HCMB__A001-01000387_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 1000387
Date of test Ap 1 | 2006-03-07 |  |  |  | Date of test Ap 2 | 2006-03-07
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | yellow alarm | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | yellow alarm | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | yellow alarm | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | yellow alarm | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | yellow alarm
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
1087 | HCMB__A001 | 1000387 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2006-03-08 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2006-03-08 | Answer to manufacturer |  |  | 2006-03-08
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -
| 2005-10-21 | P. Hagen | Check if mole positioning is within +/- 5 cm all positions

## Original data
File name |  | HCMB__A001-01000387_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 1000387 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Alstom | [9 empty] | File name for aperture 1 |  |  | 07-MAR-2006_1387_1cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 07-MAR-2006_1387_1cc2.txt |  |  |  | a even | 1
Test Operator |  | clerc |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2006-03-07 |  |  | Time of test |  |  | 19:33 - 22:31 |  |  |  | Date of test |  | 2006-03-07 |  |  | Time of test |  |  | 19:34 - 22:32
Mole name |  | Dipole 04 | [9 empty] | Mole name |  | Dipole 05 | [5 empty] | 2005-05-18
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 17 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 17
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.54 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.56 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00021482 |  |  | Dy average |  |  | 0.00061148 |  |  |  | Dx average |  | 0.00030918 |  |  | Dy average |  |  | 0.00083104
Magnetic length |  | 14.4295 |  |  | Tranfs. Function (mT/KA) |  |  | 594.6023 |  |  |  | Magnetic length |  | 14.4294 |  |  | Tranfs. Function (mT/KA) |  |  | 594.9648
Phase |  | -0.089879 |  |  | Integral field (T.m) |  |  | 85.7867 |  |  |  | Phase |  | -0.32877 |  |  | Integral field (T.m) |  |  | 85.9343
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.7285 | 5.9535 | 5.9486 | 5.9478 | 5.9468 | 5.9476 | 5.9445 | 5.9438 | 5.9439 | 5.9446 | 5.944 | 5.9441 | 5.9445 | 5.9444 | 5.9449 | 5.9447 | 5.9445 | 5.9455 | 5.9506 | 3.6405 | 85.7867 | C1 (T) | 3.7599 | 5.9564 | 5.9527 | 5.9514 | 5.9512 | 5.9516 | 5.9478 | 5.947 | 5.9482 | 5.9483 | 5.9476 | 5.948 | 5.9483 | 5.9476 | 5.948 | 5.9483 | 5.9484 | 5.9489 | 5.9539 | 3.7404 | 85.9343
Angle (mrad) | 4.2978 | 0.31212 | 0.13037 | 0.12748 | 0.34447 | -0.16328 | -0.31382 | -0.26341 | -0.21533 | -0.34783 | -0.42761 | -0.11935 | -0.32497 | -0.24155 | -0.27392 | -0.199 | -0.28216 | -0.37366 | -0.21587 | 0.24678 | -0.089879 | Angle (mrad) | 3.9538 | 0.3977 | 0.18183 | 0.20609 | 0.2934 | 0.11628 | -0.15069 | -0.32914 | -0.072847 | -0.34844 | -0.2561 | -0.090639 | -0.027247 | -0.13837 | -0.42545 | -0.37988 | -0.31276 | -0.22708 | -0.17397 | -1.2132 | -0.32877
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.58677 | -0.93177 | -0.87044 | -1.4915 | -0.45451 | -0.94075 | 0.051528 | 0.055186 | 0.049586 | 0.44312 | -0.46663 | 0.17532 | 0.086571 | 0.18295 | -0.59777 | -0.71373 | -1.014 | -1.4584 | -0.88453 | -0.29964 | -0.44688 | b2 | 1.1248 | -0.84075 | -0.22273 | -0.19432 | 0.044791 | -0.46381 | 0.010976 | -0.12391 | 0.4201 | 0.66664 | -0.53699 | 0.57976 | -0.80914 | 0.25572 | 0.49806 | -0.012114 | 0.2014 | 0.4919 | 0.7458 | 1.0355 | 0.10759
b3 | 45.3533 | -0.061473 | -2.4052 | -2.8089 | -3.0307 | -2.6444 | -1.6869 | -1.4659 | -1.4335 | -2.6953 | -2.1474 | -2.6317 | -2.8648 | -2.9878 | -3.3936 | -2.5237 | -1.2814 | -1.7356 | 0.61398 | 1.1846 | -0.41657 | b3 | 44.4476 | -0.47308 | -1.0425 | -1.1989 | -1.2758 | -2.6986 | -0.86054 | -1.2102 | -0.92563 | -2.0159 | -2.8268 | -1.8536 | -1.4864 | -1.7227 | -2.6356 | -3.0888 | -1.8927 | -0.78953 | 0.76301 | 1.5793 | 0.096111
b4 | 1.3419 | -0.30248 | -0.38055 | -0.11079 | -0.033499 | -0.011122 | 0.44497 | 0.36361 | 0.42024 | 0.35828 | 0.074071 | 0.45252 | 0.20329 | 0.08892 | 0.27046 | -0.023245 | -0.067088 | -0.26633 | -0.094674 | 0.094183 | 0.11873 | b4 | 0.078873 | -0.29005 | -0.12864 | -0.011552 | 0.057206 | -0.12095 | 0.1924 | 0.17228 | 0.054922 | 0.23801 | -0.64346 | -0.41732 | -0.40165 | -0.30564 | -0.46909 | -0.18082 | -0.064815 | -0.17863 | -0.30679 | -0.35456 | -0.1546
b5 | -2.2012 | 0.97353 | 1.0268 | 1.0343 | 1.066 | 1.0708 | 0.89729 | 1.0327 | 1.0838 | 0.9661 | 0.61881 | 1.134 | 0.93361 | 0.89762 | 0.99039 | 0.92726 | 0.71793 | 0.81481 | 0.58694 | -2.9079 | 0.70751 | b5 | -1.9418 | 0.73656 | 0.88964 | 0.95034 | 0.87724 | 0.91601 | 0.84632 | 1.0497 | 0.99681 | 1.1634 | 0.81022 | 0.97603 | 0.83249 | 0.82997 | 0.57228 | 0.8583 | 0.74223 | 0.51532 | 0.43691 | -3.095 | 0.61403
b6 | 0.74871 | -0.016864 | -0.12491 | -0.05161 | -0.15181 | 0.16595 | 0.19051 | 0.14516 | 0.069814 | 0.10209 | -0.022444 | -0.19847 | -0.073235 | -0.091124 | -0.045925 | -0.012184 | 0.04312 | -0.059628 | -0.054066 | 0.23281 | 0.022156 | b6 | -0.084038 | 0.0028105 | 0.082951 | 0.036034 | -0.010804 | 0.021471 | 0.14715 | 0.10439 | 0.0081481 | -0.081028 | -0.015234 | -0.16107 | -0.027401 | -0.15987 | -0.14599 | 0.061845 | -0.032523 | -0.091513 | -0.13481 | -0.10013 | -0.026554
b7 | 2.7786 | 1.1589 | 1.1229 | 1.0897 | 1.1434 | 1.1125 | 1.2566 | 1.2295 | 1.1782 | 1.1897 | 1.2298 | 1.1598 | 1.1624 | 1.1986 | 1.0631 | 1.1828 | 1.1944 | 1.1182 | 1.2179 | 1.0072 | 1.2146 | b7 | 2.9166 | 1.2615 | 1.1857 | 1.1671 | 1.1827 | 1.1567 | 1.2235 | 1.1959 | 1.2249 | 1.2246 | 1.1916 | 1.1888 | 1.1976 | 1.2138 | 1.1234 | 1.1595 | 1.1928 | 1.2281 | 1.2445 | 0.89905 | 1.2446
b8 | 0.14895 | -0.037836 | 0.01322 | 0.00067375 | -0.022544 | 0.073653 | 0.081395 | 0.038501 | -0.019419 | -0.015898 | -0.08218 | -0.083161 | -0.0198 | -0.0099849 | 0.044936 | 0.043661 | 0.024553 | -0.036302 | 0.010213 | 0.03117 | 0.0060378 | b8 | -0.017696 | -0.084071 | -0.013183 | 0.014839 | -0.01928 | 0.0031851 | 0.010388 | 0.0046241 | -0.011763 | -0.0647 | -0.01322 | -0.036702 | -0.0097895 | 0.01036 | 0.02066 | -0.012248 | 0.047312 | -0.012595 | 0.031315 | 0.022049 | -0.0068665
b9 | 0.46023 | 0.56096 | 0.57339 | 0.58324 | 0.61003 | 0.622 | 0.59913 | 0.60214 | 0.62272 | 0.59476 | 0.62316 | 0.60886 | 0.59002 | 0.58772 | 0.58119 | 0.58361 | 0.58201 | 0.56799 | 0.58124 | 0.58858 | 0.58854 | b9 | 0.56052 | 0.58754 | 0.57174 | 0.58635 | 0.60883 | 0.59357 | 0.58052 | 0.56484 | 0.57874 | 0.59271 | 0.5839 | 0.59402 | 0.58416 | 0.58494 | 0.58708 | 0.55766 | 0.5694 | 0.57107 | 0.57584 | 0.56939 | 0.58072
b10 | 0.13415 | -0.13699 | -0.0066074 | 0.036292 | -0.044097 | 0.097794 | 0.10533 | 0.040701 | -0.046352 | -0.085605 | -0.12026 | -0.12698 | 0.017539 | -0.022029 | 0.043231 | 0.068768 | 0.060106 | -0.0054981 | 0.031847 | 0.014311 | 0 | b10 | 0.015824 | -0.088967 | 0.043004 | 0.045058 | -4.4423e-05 | 0.049616 | 0.066543 | 0.026741 | -0.028626 | -0.087359 | -0.0040986 | -0.087369 | 0.02605 | -0.014163 | 0.031926 | 0.019906 | 0.011506 | -0.032462 | 0.034297 | -0.034266 | 0
b11 | 0.64215 | 0.76602 | 0.77571 | 0.78543 | 0.77957 | 0.7775 | 0.76007 | 0.75853 | 0.76716 | 0.77192 | 0.76339 | 0.77199 | 0.7697 | 0.77012 | 0.77294 | 0.76671 | 0.76495 | 0.76576 | 0.76003 | 0.67495 | 0.76216 | b11 | 0.68015 | 0.75521 | 0.75738 | 0.76574 | 0.77955 | 0.77172 | 0.74246 | 0.74051 | 0.7521 | 0.75332 | 0.75746 | 0.7549 | 0.75511 | 0.74947 | 0.75602 | 0.75779 | 0.76003 | 0.75267 | 0.75392 | 0.6613 | 0.7508
b12 | 0.021554 | -0.011177 | 0.0039754 | 0.0022973 | -0.0019258 | 0.01302 | 0.0085071 | 0.0016714 | -0.0078919 | -0.0078427 | -0.0079768 | -0.015257 | -0.0018124 | -0.0036422 | -0.00053307 | 0.0083897 | 0.0067976 | -0.0027305 | 0.0060616 | 0.011843 | 0.00055621 | b12 | -0.0023017 | -0.012657 | 0.0019241 | 0.0061068 | -0.0055746 | 0.0063633 | -0.0013904 | -0.0023224 | -0.005989 | -0.011536 | -0.0013862 | -0.00567 | 0.002129 | 0.0027338 | 0.0083046 | 0.0031663 | 0.0055607 | -0.002887 | 0.0069263 | 0.00031835 | -0.00038707
b13 | 0.095105 | 0.059548 | 0.051008 | 0.053775 | 0.054958 | 0.051986 | 0.058858 | 0.057214 | 0.05652 | 0.061772 | 0.062318 | 0.059972 | 0.058317 | 0.062072 | 0.058554 | 0.057015 | 0.063058 | 0.058367 | 0.064473 | 0.058625 | 0.05953 | b13 | 0.092103 | 0.05768 | 0.053175 | 0.054048 | 0.0555 | 0.049761 | 0.058203 | 0.05601 | 0.05801 | 0.056843 | 0.05694 | 0.057204 | 0.056052 | 0.057373 | 0.055883 | 0.056438 | 0.060093 | 0.061995 | 0.063649 | 0.059245 | 0.058166
b14 | 0.0075026 | -0.012768 | -0.0072194 | -0.0045507 | -0.009945 | 0.0010465 | -0.00097964 | -0.0036408 | -0.010108 | -0.011752 | -0.01271 | -0.014509 | -0.0072887 | -0.009224 | -0.0052357 | -0.0021946 | -0.0038946 | -0.0099791 | -0.0064002 | -0.011069 | -0.0069351 | b14 | 0.0046119 | -0.015445 | -0.010186 | -0.010099 | -0.015326 | -0.0089889 | -0.0074314 | -0.011519 | -0.015564 | -0.018947 | -0.012903 | -0.018717 | -0.010923 | -0.012229 | -0.0081672 | -0.0099583 | -0.010417 | -0.01418 | -0.010084 | -0.011687 | -0.011709
b15 | 0.0094241 | 0.028517 | 0.035084 | 0.039445 | 0.036121 | 0.037642 | 0.034054 | 0.038732 | 0.036564 | 0.028024 | 0.02888 | 0.027396 | 0.037604 | 0.033781 | 0.038427 | 0.03699 | 0.037797 | 0.03921 | 0.035334 | 0.030736 | 0.03401 | b15 | 0.0020461 | 0.030662 | 0.043035 | 0.046198 | 0.04157 | 0.045946 | 0.039557 | 0.041379 | 0.039466 | 0.031974 | 0.040156 | 0.033809 | 0.038221 | 0.033262 | 0.034878 | 0.040745 | 0.044181 | 0.038247 | 0.040558 | 0.027078 | 0.037494
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 42.9782 | 3.1212 | 1.3037 | 1.2748 | 3.4447 | -1.6328 | -3.1382 | -2.6341 | -2.1533 | -3.4783 | -4.2761 | -1.1935 | -3.2497 | -2.4155 | -2.7392 | -1.99 | -2.8216 | -3.7366 | -2.1587 | 2.4678 | 0 | a1 | 39.5386 | 3.977 | 1.8183 | 2.0609 | 2.934 | 1.1628 | -1.5069 | -3.2914 | -0.72847 | -3.4844 | -2.561 | -0.90639 | -0.27247 | -1.3837 | -4.2545 | -3.7988 | -3.1276 | -2.2708 | -1.7397 | -12.1317 | 0
a2 | -6.9347 | 2.1041 | 0.47761 | -0.90953 | -1.9517 | 0.46807 | -0.51679 | -1.282 | -0.46082 | -0.18673 | -1.2811 | -0.91773 | -1.075 | -1.8595 | -0.3614 | -0.30699 | -0.39579 | -0.12237 | -0.0065613 | 0.15858 | -0.6669 | a2 | 0.50465 | 0.89072 | 0.42122 | -0.6059 | -1.1606 | 0.70542 | 0.36771 | 0.36857 | 0.49251 | 0.056536 | -0.29806 | 1.2868 | 0.43653 | 0.30833 | -0.11471 | -0.17021 | 0.47815 | -0.30942 | -0.76912 | 3.296 | 0.24791
a3 | -2.3541 | -0.79856 | -0.28843 | -0.45557 | -0.15304 | -1.04 | -0.77727 | -0.40267 | -0.019294 | 0.2109 | -0.23553 | -0.085622 | -0.24028 | -0.20125 | -0.1794 | -0.032071 | -0.0014666 | -0.19611 | -0.37074 | 0.8711 | -0.3228 | a3 | -2.6754 | -0.77277 | -0.67276 | -0.31761 | -0.51803 | -0.22871 | -0.57463 | -0.70535 | 0.044518 | -0.28024 | -0.68283 | -0.25947 | -0.64042 | -0.41922 | -0.26652 | -0.45874 | -0.20678 | -0.64859 | -0.68023 | 0.18825 | -0.51199
a4 | -2.4536 | -0.48026 | 0.14599 | 0.37186 | 0.24551 | -0.3507 | -0.2419 | 0.0013159 | 0.089323 | -0.53404 | -0.41114 | -0.024922 | -0.1853 | 0.19107 | -0.21681 | 0.065436 | 0.34791 | 0.32304 | 0.093582 | -0.41559 | -0.12283 | a4 | -0.37997 | -0.43831 | -0.23106 | -0.21835 | -0.064152 | -0.43476 | 0.20831 | 0.27319 | 0.13028 | -0.40615 | -0.49207 | 0.0060546 | -0.10635 | 0.029952 | -0.2343 | -0.18721 | -0.40749 | -0.05406 | 0.076382 | 0.67635 | -0.12282
a5 | 2.9223 | -0.2906 | -0.12061 | -0.30795 | -0.12807 | -0.20543 | -0.13396 | -0.10069 | 0.026115 | 0.10965 | -0.087106 | -0.13656 | -0.023967 | 0.17312 | 0.099194 | -0.17012 | -0.092541 | -0.069126 | -0.25458 | -1.0016 | -0.025707 | a5 | 3.193 | -0.24105 | 0.045996 | 0.056135 | 0.052983 | -0.30976 | -0.17216 | -0.039149 | -0.087164 | 0.11026 | -0.49874 | -0.38821 | -0.25902 | -0.14381 | 0.064469 | -0.18066 | -0.16919 | -0.20622 | -0.23465 | -0.52385 | -0.04732
a6 | -0.80101 | 0.1516 | 0.13616 | 0.056696 | 0.021703 | -0.010822 | -0.090915 | -0.088335 | 0.012056 | 0.028332 | 0.1054 | 0.17339 | 0.11137 | 0.10241 | 0.16715 | 0.15491 | 0.064687 | 0.09441 | 0.078226 | -0.36737 | 0.028138 | a6 | 0.054789 | 0.051833 | -0.043336 | -0.068482 | 0.081654 | 0.18143 | -0.033884 | -0.0070937 | 0.093466 | 0.088111 | -0.09036 | 0.099863 | 0.038926 | 0.081438 | 0.038533 | 0.19268 | 0.1303 | 0.10003 | 0.077112 | -0.17572 | 0.048615
a7 | 1.5601 | -0.062316 | -0.13242 | -0.013411 | -0.062983 | -0.15066 | -0.16001 | -0.072774 | -0.075163 | -0.11842 | 0.022651 | 0.012907 | -0.042844 | 0.03101 | 0.0085843 | -0.03272 | -0.062619 | -0.08209 | -0.072539 | -0.17438 | -0.010104 | a7 | 1.6127 | 0.022487 | 0.081754 | 0.042068 | 0.011534 | -0.20385 | -0.2258 | -0.2298 | -0.15033 | -0.1253 | 0.079556 | 0.018899 | 0.028632 | 0.016202 | -0.02327 | 0.016013 | -0.092843 | -0.092551 | -0.074747 | -0.058648 | 0.0042112
a8 | -0.16972 | 0.084705 | 0.07569 | 0.080346 | 0.033036 | 0.018854 | -0.0058028 | -0.10418 | -0.039427 | 0.0024501 | 0.0054682 | 0.0068754 | -0.07746 | -0.055068 | -0.010014 | 0.0091047 | 0.021167 | -0.014865 | 0.030826 | -0.071277 | -0.0045783 | a8 | 0.09016 | 0.020475 | -0.030104 | -0.035371 | 0.040155 | 0.042002 | 0.028302 | -0.014162 | -0.019839 | 0.026643 | -0.044311 | -0.004432 | -0.031334 | 0.0029093 | 0.0011815 | -0.0045041 | -0.039172 | -0.020452 | 0.018716 | -0.10472 | -0.0037429
a9 | -0.2134 | -0.015248 | -0.048496 | -0.057178 | -0.047507 | -0.034835 | -0.022453 | -0.023431 | -0.01257 | -0.023869 | -0.011994 | -0.030414 | -0.026492 | -0.031362 | -0.030174 | -0.001532 | -0.0088718 | -0.0006065 | -0.032034 | -0.014703 | -0.031286 | a9 | -0.16947 | -0.036386 | -0.036584 | -0.016611 | -0.019356 | -0.074909 | -0.06712 | -0.06254 | -0.067765 | -0.086148 | -0.042758 | -0.04667 | -0.057702 | -0.060872 | -0.048508 | -0.027958 | -0.013005 | -0.054386 | -0.048645 | -0.030674 | -0.051621
a10 | -0.058454 | 0.019516 | 0.036233 | -0.0083737 | 0.0019973 | 0.0062557 | 0.013058 | -0.045944 | -0.027778 | 0.076952 | 0.024547 | 0.047021 | -0.020361 | 0.01095 | 0.02408 | 0.014848 | -0.024505 | -0.041017 | 0.0043729 | -0.12287 | 0 | a10 | 0.089137 | 0.043242 | -0.023286 | -0.036637 | -0.024251 | -0.036055 | 0.011445 | -0.010541 | -0.012141 | 0.01731 | -0.023268 | 0.016186 | 0.0076518 | 0.045106 | 0.048953 | 0.017877 | -0.035497 | -0.014563 | -0.0059384 | -0.066659 | 0
a11 | 0.13814 | -0.02913 | -0.038963 | -0.02942 | -0.027151 | -0.042319 | -0.044692 | -0.038382 | -0.036424 | -0.027976 | -0.032519 | -0.025185 | -0.027533 | -0.01884 | -0.033839 | -0.034707 | -0.036867 | -0.033459 | -0.040478 | -0.027058 | -0.027435 | a11 | 0.13279 | -0.051782 | -0.040054 | -0.037417 | -0.044459 | -0.052453 | -0.054718 | -0.057256 | -0.050632 | -0.05135 | -0.051553 | -0.048924 | -0.041707 | -0.042542 | -0.046023 | -0.0553 | -0.058078 | -0.051618 | -0.050052 | -0.025362 | -0.042467
a12 | -0.014951 | 0.003043 | -0.0017218 | -0.0089896 | -0.0054999 | 0.00050689 | -0.0020128 | -0.0038729 | -0.003301 | 0.0053776 | 0.0010398 | 0.0039877 | -0.0014505 | 0.00059863 | 0.0088461 | -0.0014138 | -0.0077038 | -0.0073928 | 0.0013266 | -0.0034899 | -0.0015667 | a12 | 0.006514 | 0.0040938 | -0.0021905 | -0.0082335 | -0.0094012 | -0.0042021 | 0.0013707 | -0.00050238 | 0.0012176 | -0.0021492 | -0.0049007 | 0.0035922 | 0.0028689 | 0.0072386 | 0.0066754 | 0.0074029 | -0.00064597 | -0.00069822 | -0.00043847 | -0.0078993 | 1.2425e-05
a13 | -0.02782 | -0.0038065 | -0.0021293 | -0.0050294 | -0.0026333 | -0.0031253 | -0.0048194 | -0.003446 | -0.00035567 | 0.00069528 | 0.0014877 | -0.001522 | -0.00070872 | 0.00095499 | -0.00019154 | -0.0015879 | 0.0002086 | 0.0022854 | -0.0011497 | 0.0039067 | -0.0020756 | a13 | -0.021094 | -0.0073088 | -0.0073856 | -0.0095128 | -0.0062713 | -0.0033423 | -0.0045996 | -0.0018147 | -0.0011605 | 0.001522 | -0.0051277 | -0.0029837 | -0.0063993 | -0.0046565 | -0.0068562 | -0.0018643 | -0.0032747 | -0.0035644 | -0.0056957 | -0.0049802 | -0.0050241
a14 | 0.00029341 | -0.010137 | -0.016182 | -0.016489 | -0.012699 | -0.014255 | -0.010935 | -0.015999 | -0.012884 | -0.010132 | -0.010778 | -0.010629 | -0.018573 | -0.014846 | -0.015629 | -0.01603 | -0.016375 | -0.019334 | -0.01277 | -0.0076571 | -0.013471 | a14 | 0.00044838 | -0.014496 | -0.021862 | -0.02369 | -0.01993 | -0.023247 | -0.015123 | -0.017458 | -0.018103 | -0.01576 | -0.021894 | -0.015812 | -0.019209 | -0.014332 | -0.016661 | -0.021138 | -0.02317 | -0.018025 | -0.018815 | -0.01027 | -0.017907
a15 | 0.0022129 | -0.020185 | -0.0092982 | -0.0062748 | -0.012721 | -0.00043859 | -0.00030522 | -0.0045897 | -0.01159 | -0.0097815 | -0.015212 | -0.015881 | -0.008866 | -0.010148 | -0.0065802 | -0.0026776 | -0.0052222 | -0.011124 | -0.0074685 | -0.011366 | -0.0085214 | a15 | -0.00056456 | -0.026226 | -0.016852 | -0.016552 | -0.02061 | -0.015369 | -0.012135 | -0.015831 | -0.020799 | -0.025342 | -0.019404 | -0.026512 | -0.017759 | -0.018611 | -0.01516 | -0.014574 | -0.016405 | -0.022433 | -0.015616 | -0.026256 | -0.018331
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00030717 | 0.00030547 | 1.8436e-05 | -7.9114e-05 | 9.6199e-05 | -0.00021264 | -0.0002332 | -9.6161e-05 | 9.9567e-05 | 0.00019436 | 0.00026976 | 0.00028299 | -4.0293e-05 | 4.9192e-05 | -9.2547e-05 | -0.00015074 | -0.00013588 | 0 | -7.0523e-05 | -4.7909e-05 | 0.00021482 | Dx (m) | -8.0025e-05 | 0.0002061 | -9.901e-05 | -0.00010373 | 0 | -0.00011417 | -0.00014958 | -6.2879e-05 | 6.2582e-05 | 0.00019904 | 0 | 0.00019832 | -5.7518e-05 | 3.7802e-05 | -6.4785e-05 | -4.1499e-05 | -3.1609e-05 | 7.0733e-05 | -7.7896e-05 | 8.1342e-05 | 0.00030918
Dy (m) | 0.00022212 | -3.1843e-05 | -7.8483e-05 | 1.5184e-05 | 0 | -2.5305e-05 | -4.2912e-05 | 9.8107e-05 | 6.6324e-05 | -0.00016258 | -4.3227e-05 | -9.4643e-05 | 4.3525e-05 | -2.2971e-05 | -5.7024e-05 | -3.9762e-05 | 4.7985e-05 | 9.1396e-05 | -1.3543e-05 | 0.00030774 | 0.00061148 | Dy (m) | -0.00020705 | -8.3395e-05 | 4.7039e-05 | 7.6315e-05 | 5.2708e-05 | 7.171e-05 | -3.7223e-05 | 1.9335e-05 | 3.1663e-05 | -2.5493e-05 | 5.2598e-05 | -2.3661e-05 | -2.0408e-05 | -0.00010014 | -0.00011405 | -4.3136e-05 | 7.6993e-05 | 3.7747e-05 | 0 | 0.00017439 | 0.00083104
Magnetic length (m) |  | 14.4295 | [19 empty] | Magnetic length (m) |  | 14.4294
Transfer funct. (mT/KA) |  | 594.6023 | [19 empty] | Transfer funct. (mT/KA) |  | 594.9648
Coil positioning (mm) |  | 0.022379 | [19 empty] | Coil positioning (mm) |  | 0.020256
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.1.5 B | [20 empty] | Soft version | DIMM V2.1.5 B
Level Average | -0.059186 | -0.054059 | -0.082663 | -0.066875 | -0.056555 | 0.12626 | -0.12708 | -0.047365 | -0.10494 | 0.07633 | -0.068286 | -0.26868 | -0.058089 | -0.045359 | -0.048292 | -0.07594 | -0.044938 | 0.067732 | 0.050683 | 0.151 |  | Level Average | -0.010008 | -0.13575 | 0.16229 | 0.0733 | -0.098428 | -0.14058 | -0.063977 | -0.083325 | -0.11763 | -0.11807 | 0.19415 | -0.15672 | -0.082764 | -0.14271 | -0.11482 | -0.12525 | 0.008918 | -0.20435 | -0.012088 | 0.091308
Level Sigma2 | 0.009568 | 0.013561 | 0.085663 | 0.000735 | 0.002943 | 0.088157 | 0.00049 | 0.006746 | 0.002616 | 0.010133 | 0.009498 | 0.012685 | 0.01101 | 0.005967 | 0.003182 | 0.002298 | 0.025202 | 0.044995 | 0.024281 | 0.009284 |  | Level Sigma2 | 0.010809 | 0.000253 | 0.006511 | 0.047081 | 0.032042 | 0.081641 | 0.000342 | 0.037442 | 0.020161 | 0.036907 | 0.027689 | 0.011869 | 0.015869 | 0.000348 | 0.073935 | 0.066618 | 0.005505 | 0.005128 | 0.033365 | 0.38017
Position (mm) | 0 | 749 | 1499 | 2243 | 2992 | 3748 | 4499 | 5244 | 5998 | 6749 | 7497 | 8248 | 8998 | 9749 | 10499 | 11249 | 12000 | 12750 | 13500 | 14250 |  | Position (mm) | 0 | 750 | 1500 | 2247 | 2999 | 3749 | 4499 | 5250 | 6000 | 6751 | 7480 | 8232 | 8983 | 9732 | 10483 | 11232 | 11983 | 12732 | 13484 | 14234
Fist move (mm) | First move Ap1 | 941.4996 | First move Ap2 | 954.476 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-1000387 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-1000387 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 372.8531 | 595.3508 | 594.8583 | 594.7796 | 594.6842 | 594.7616 | 594.4516 | 594.3787 | 594.3949 | 594.4553 | 594.3992 | 594.4122 | 594.4492 | 594.4413 | 594.4871 | 594.4725 | 594.4506 | 594.5527 | 595.0611 | 364.0458 | 8578.6732 | C1/i (mT/KA) | 375.9884 | 595.6374 | 595.2715 | 595.1414 | 595.122 | 595.1593 | 594.7833 | 594.6999 | 594.8233 | 594.8272 | 594.7597 | 594.7974 | 594.8264 | 594.762 | 594.8036 | 594.8273 | 594.8382 | 594.8946 | 595.3913 | 374.0374 | 8593.4286
Angle (mrad) | 4.2978 | 0.31212 | 0.13037 | 0.12748 | 0.34447 | -0.16328 | -0.31382 | -0.26341 | -0.21533 | -0.34783 | -0.42761 | -0.11935 | -0.32497 | -0.24155 | -0.27392 | -0.199 | -0.28216 | -0.37366 | -0.21587 | 0.24678 | -0.089879 | Angle (mrad) | 3.9538 | 0.3977 | 0.18183 | 0.20609 | 0.2934 | 0.11628 | -0.15069 | -0.32914 | -0.072847 | -0.34844 | -0.2561 | -0.090639 | -0.027247 | -0.13837 | -0.42545 | -0.37988 | -0.31276 | -0.22708 | -0.17397 | -1.2132 | -0.32877
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.58677 | -0.93177 | -0.87044 | -1.4915 | -0.45451 | -0.94075 | 0.051528 | 0.055186 | 0.049586 | 0.44312 | -0.46663 | 0.17532 | 0.086571 | 0.18295 | -0.59777 | -0.71373 | -1.014 | -1.4584 | -0.88453 | -0.29964 |  | b2 | 1.1248 | -0.84075 | -0.22273 | -0.19432 | 0.044791 | -0.46381 | 0.010976 | -0.12391 | 0.4201 | 0.66664 | -0.53699 | 0.57976 | -0.80914 | 0.25572 | 0.49806 | -0.012114 | 0.2014 | 0.4919 | 0.7458 | 1.0355
b3 | 45.3533 | -0.061473 | -2.4052 | -2.8089 | -3.0307 | -2.6444 | -1.6869 | -1.4659 | -1.4335 | -2.6953 | -2.1474 | -2.6317 | -2.8648 | -2.9878 | -3.3936 | -2.5237 | -1.2814 | -1.7356 | 0.61398 | 1.1846 |  | b3 | 44.4476 | -0.47308 | -1.0425 | -1.1989 | -1.2758 | -2.6986 | -0.86054 | -1.2102 | -0.92563 | -2.0159 | -2.8268 | -1.8536 | -1.4864 | -1.7227 | -2.6356 | -3.0888 | -1.8927 | -0.78953 | 0.76301 | 1.5793
b4 | 1.3419 | -0.30248 | -0.38055 | -0.11079 | -0.033499 | -0.011122 | 0.44497 | 0.36361 | 0.42024 | 0.35828 | 0.074071 | 0.45252 | 0.20329 | 0.08892 | 0.27046 | -0.023245 | -0.067088 | -0.26633 | -0.094674 | 0.094183 |  | b4 | 0.078873 | -0.29005 | -0.12864 | -0.011552 | 0.057206 | -0.12095 | 0.1924 | 0.17228 | 0.054922 | 0.23801 | -0.64346 | -0.41732 | -0.40165 | -0.30564 | -0.46909 | -0.18082 | -0.064815 | -0.17863 | -0.30679 | -0.35456
b5 | -2.2012 | 0.97353 | 1.0268 | 1.0343 | 1.066 | 1.0708 | 0.89729 | 1.0327 | 1.0838 | 0.9661 | 0.61881 | 1.134 | 0.93361 | 0.89762 | 0.99039 | 0.92726 | 0.71793 | 0.81481 | 0.58694 | -2.9079 |  | b5 | -1.9418 | 0.73656 | 0.88964 | 0.95034 | 0.87724 | 0.91601 | 0.84632 | 1.0497 | 0.99681 | 1.1634 | 0.81022 | 0.97603 | 0.83249 | 0.82997 | 0.57228 | 0.8583 | 0.74223 | 0.51532 | 0.43691 | -3.095
b6 | 0.74871 | -0.016864 | -0.12491 | -0.05161 | -0.15181 | 0.16595 | 0.19051 | 0.14516 | 0.069814 | 0.10209 | -0.022444 | -0.19847 | -0.073235 | -0.091124 | -0.045925 | -0.012184 | 0.04312 | -0.059628 | -0.054066 | 0.23281 |  | b6 | -0.084038 | 0.0028105 | 0.082951 | 0.036034 | -0.010804 | 0.021471 | 0.14715 | 0.10439 | 0.0081481 | -0.081028 | -0.015234 | -0.16107 | -0.027401 | -0.15987 | -0.14599 | 0.061845 | -0.032523 | -0.091513 | -0.13481 | -0.10013
b7 | 2.7786 | 1.1589 | 1.1229 | 1.0897 | 1.1434 | 1.1125 | 1.2566 | 1.2295 | 1.1782 | 1.1897 | 1.2298 | 1.1598 | 1.1624 | 1.1986 | 1.0631 | 1.1828 | 1.1944 | 1.1182 | 1.2179 | 1.0072 |  | b7 | 2.9166 | 1.2615 | 1.1857 | 1.1671 | 1.1827 | 1.1567 | 1.2235 | 1.1959 | 1.2249 | 1.2246 | 1.1916 | 1.1888 | 1.1976 | 1.2138 | 1.1234 | 1.1595 | 1.1928 | 1.2281 | 1.2445 | 0.89905
b8 | 0.14895 | -0.037836 | 0.01322 | 0.00067375 | -0.022544 | 0.073653 | 0.081395 | 0.038501 | -0.019419 | -0.015898 | -0.08218 | -0.083161 | -0.0198 | -0.0099849 | 0.044936 | 0.043661 | 0.024553 | -0.036302 | 0.010213 | 0.03117 |  | b8 | -0.017696 | -0.084071 | -0.013183 | 0.014839 | -0.01928 | 0.0031851 | 0.010388 | 0.0046241 | -0.011763 | -0.0647 | -0.01322 | -0.036702 | -0.0097895 | 0.01036 | 0.02066 | -0.012248 | 0.047312 | -0.012595 | 0.031315 | 0.022049
b9 | 0.46023 | 0.56096 | 0.57339 | 0.58324 | 0.61003 | 0.622 | 0.59913 | 0.60214 | 0.62272 | 0.59476 | 0.62316 | 0.60886 | 0.59002 | 0.58772 | 0.58119 | 0.58361 | 0.58201 | 0.56799 | 0.58124 | 0.58858 |  | b9 | 0.56052 | 0.58754 | 0.57174 | 0.58635 | 0.60883 | 0.59357 | 0.58052 | 0.56484 | 0.57874 | 0.59271 | 0.5839 | 0.59402 | 0.58416 | 0.58494 | 0.58708 | 0.55766 | 0.5694 | 0.57107 | 0.57584 | 0.56939
b10 | 0.13415 | -0.13699 | -0.0066074 | 0.036292 | -0.044097 | 0.097794 | 0.10533 | 0.040701 | -0.046352 | -0.085605 | -0.12026 | -0.12698 | 0.017539 | -0.022029 | 0.043231 | 0.068768 | 0.060106 | -0.0054981 | 0.031847 | 0.014311 |  | b10 | 0.015824 | -0.088967 | 0.043004 | 0.045058 | -4.4423e-05 | 0.049616 | 0.066543 | 0.026741 | -0.028626 | -0.087359 | -0.0040986 | -0.087369 | 0.02605 | -0.014163 | 0.031926 | 0.019906 | 0.011506 | -0.032462 | 0.034297 | -0.034266
b11 | 0.64215 | 0.76602 | 0.77571 | 0.78543 | 0.77957 | 0.7775 | 0.76007 | 0.75853 | 0.76716 | 0.77192 | 0.76339 | 0.77199 | 0.7697 | 0.77012 | 0.77294 | 0.76671 | 0.76495 | 0.76576 | 0.76003 | 0.67495 |  | b11 | 0.68015 | 0.75521 | 0.75738 | 0.76574 | 0.77955 | 0.77172 | 0.74246 | 0.74051 | 0.7521 | 0.75332 | 0.75746 | 0.7549 | 0.75511 | 0.74947 | 0.75602 | 0.75779 | 0.76003 | 0.75267 | 0.75392 | 0.6613
b12 | 0.021554 | -0.011177 | 0.0039754 | 0.0022973 | -0.0019258 | 0.01302 | 0.0085071 | 0.0016714 | -0.0078919 | -0.0078427 | -0.0079768 | -0.015257 | -0.0018124 | -0.0036422 | -0.00053307 | 0.0083897 | 0.0067976 | -0.0027305 | 0.0060616 | 0.011843 |  | b12 | -0.0023017 | -0.012657 | 0.0019241 | 0.0061068 | -0.0055746 | 0.0063633 | -0.0013904 | -0.0023224 | -0.005989 | -0.011536 | -0.0013862 | -0.00567 | 0.002129 | 0.0027338 | 0.0083046 | 0.0031663 | 0.0055607 | -0.002887 | 0.0069263 | 0.00031835
b13 | 0.095105 | 0.059548 | 0.051008 | 0.053775 | 0.054958 | 0.051986 | 0.058858 | 0.057214 | 0.05652 | 0.061772 | 0.062318 | 0.059972 | 0.058317 | 0.062072 | 0.058554 | 0.057015 | 0.063058 | 0.058367 | 0.064473 | 0.058625 |  | b13 | 0.092103 | 0.05768 | 0.053175 | 0.054048 | 0.0555 | 0.049761 | 0.058203 | 0.05601 | 0.05801 | 0.056843 | 0.05694 | 0.057204 | 0.056052 | 0.057373 | 0.055883 | 0.056438 | 0.060093 | 0.061995 | 0.063649 | 0.059245
b14 | 0.0075026 | -0.012768 | -0.0072194 | -0.0045507 | -0.009945 | 0.0010465 | -0.00097964 | -0.0036408 | -0.010108 | -0.011752 | -0.01271 | -0.014509 | -0.0072887 | -0.009224 | -0.0052357 | -0.0021946 | -0.0038946 | -0.0099791 | -0.0064002 | -0.011069 |  | b14 | 0.0046119 | -0.015445 | -0.010186 | -0.010099 | -0.015326 | -0.0089889 | -0.0074314 | -0.011519 | -0.015564 | -0.018947 | -0.012903 | -0.018717 | -0.010923 | -0.012229 | -0.0081672 | -0.0099583 | -0.010417 | -0.01418 | -0.010084 | -0.011687
b15 | 0.0094241 | 0.028517 | 0.035084 | 0.039445 | 0.036121 | 0.037642 | 0.034054 | 0.038732 | 0.036564 | 0.028024 | 0.02888 | 0.027396 | 0.037604 | 0.033781 | 0.038427 | 0.03699 | 0.037797 | 0.03921 | 0.035334 | 0.030736 |  | b15 | 0.0020461 | 0.030662 | 0.043035 | 0.046198 | 0.04157 | 0.045946 | 0.039557 | 0.041379 | 0.039466 | 0.031974 | 0.040156 | 0.033809 | 0.038221 | 0.033262 | 0.034878 | 0.040745 | 0.044181 | 0.038247 | 0.040558 | 0.027078
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 42.9782 | 3.1212 | 1.3037 | 1.2748 | 3.4447 | -1.6328 | -3.1382 | -2.6341 | -2.1533 | -3.4783 | -4.2761 | -1.1935 | -3.2497 | -2.4155 | -2.7392 | -1.99 | -2.8216 | -3.7366 | -2.1587 | 2.4678 |  | a1 | 39.5386 | 3.977 | 1.8183 | 2.0609 | 2.934 | 1.1628 | -1.5069 | -3.2914 | -0.72847 | -3.4844 | -2.561 | -0.90639 | -0.27247 | -1.3837 | -4.2545 | -3.7988 | -3.1276 | -2.2708 | -1.7397 | -12.1317
a2 | -6.9347 | 2.1041 | 0.47761 | -0.90953 | -1.9517 | 0.46807 | -0.51679 | -1.282 | -0.46082 | -0.18673 | -1.2811 | -0.91773 | -1.075 | -1.8595 | -0.3614 | -0.30699 | -0.39579 | -0.12237 | -0.0065613 | 0.15858 |  | a2 | 0.50465 | 0.89072 | 0.42122 | -0.6059 | -1.1606 | 0.70542 | 0.36771 | 0.36857 | 0.49251 | 0.056536 | -0.29806 | 1.2868 | 0.43653 | 0.30833 | -0.11471 | -0.17021 | 0.47815 | -0.30942 | -0.76912 | 3.296
a3 | -2.3541 | -0.79856 | -0.28843 | -0.45557 | -0.15304 | -1.04 | -0.77727 | -0.40267 | -0.019294 | 0.2109 | -0.23553 | -0.085622 | -0.24028 | -0.20125 | -0.1794 | -0.032071 | -0.0014666 | -0.19611 | -0.37074 | 0.8711 |  | a3 | -2.6754 | -0.77277 | -0.67276 | -0.31761 | -0.51803 | -0.22871 | -0.57463 | -0.70535 | 0.044518 | -0.28024 | -0.68283 | -0.25947 | -0.64042 | -0.41922 | -0.26652 | -0.45874 | -0.20678 | -0.64859 | -0.68023 | 0.18825
a4 | -2.4536 | -0.48026 | 0.14599 | 0.37186 | 0.24551 | -0.3507 | -0.2419 | 0.0013159 | 0.089323 | -0.53404 | -0.41114 | -0.024922 | -0.1853 | 0.19107 | -0.21681 | 0.065436 | 0.34791 | 0.32304 | 0.093582 | -0.41559 |  | a4 | -0.37997 | -0.43831 | -0.23106 | -0.21835 | -0.064152 | -0.43476 | 0.20831 | 0.27319 | 0.13028 | -0.40615 | -0.49207 | 0.0060546 | -0.10635 | 0.029952 | -0.2343 | -0.18721 | -0.40749 | -0.05406 | 0.076382 | 0.67635
a5 | 2.9223 | -0.2906 | -0.12061 | -0.30795 | -0.12807 | -0.20543 | -0.13396 | -0.10069 | 0.026115 | 0.10965 | -0.087106 | -0.13656 | -0.023967 | 0.17312 | 0.099194 | -0.17012 | -0.092541 | -0.069126 | -0.25458 | -1.0016 |  | a5 | 3.193 | -0.24105 | 0.045996 | 0.056135 | 0.052983 | -0.30976 | -0.17216 | -0.039149 | -0.087164 | 0.11026 | -0.49874 | -0.38821 | -0.25902 | -0.14381 | 0.064469 | -0.18066 | -0.16919 | -0.20622 | -0.23465 | -0.52385
a6 | -0.80101 | 0.1516 | 0.13616 | 0.056696 | 0.021703 | -0.010822 | -0.090915 | -0.088335 | 0.012056 | 0.028332 | 0.1054 | 0.17339 | 0.11137 | 0.10241 | 0.16715 | 0.15491 | 0.064687 | 0.09441 | 0.078226 | -0.36737 |  | a6 | 0.054789 | 0.051833 | -0.043336 | -0.068482 | 0.081654 | 0.18143 | -0.033884 | -0.0070937 | 0.093466 | 0.088111 | -0.09036 | 0.099863 | 0.038926 | 0.081438 | 0.038533 | 0.19268 | 0.1303 | 0.10003 | 0.077112 | -0.17572
a7 | 1.5601 | -0.062316 | -0.13242 | -0.013411 | -0.062983 | -0.15066 | -0.16001 | -0.072774 | -0.075163 | -0.11842 | 0.022651 | 0.012907 | -0.042844 | 0.03101 | 0.0085843 | -0.03272 | -0.062619 | -0.08209 | -0.072539 | -0.17438 |  | a7 | 1.6127 | 0.022487 | 0.081754 | 0.042068 | 0.011534 | -0.20385 | -0.2258 | -0.2298 | -0.15033 | -0.1253 | 0.079556 | 0.018899 | 0.028632 | 0.016202 | -0.02327 | 0.016013 | -0.092843 | -0.092551 | -0.074747 | -0.058648
a8 | -0.16972 | 0.084705 | 0.07569 | 0.080346 | 0.033036 | 0.018854 | -0.0058028 | -0.10418 | -0.039427 | 0.0024501 | 0.0054682 | 0.0068754 | -0.07746 | -0.055068 | -0.010014 | 0.0091047 | 0.021167 | -0.014865 | 0.030826 | -0.071277 |  | a8 | 0.09016 | 0.020475 | -0.030104 | -0.035371 | 0.040155 | 0.042002 | 0.028302 | -0.014162 | -0.019839 | 0.026643 | -0.044311 | -0.004432 | -0.031334 | 0.0029093 | 0.0011815 | -0.0045041 | -0.039172 | -0.020452 | 0.018716 | -0.10472
a9 | -0.2134 | -0.015248 | -0.048496 | -0.057178 | -0.047507 | -0.034835 | -0.022453 | -0.023431 | -0.01257 | -0.023869 | -0.011994 | -0.030414 | -0.026492 | -0.031362 | -0.030174 | -0.001532 | -0.0088718 | -0.0006065 | -0.032034 | -0.014703 |  | a9 | -0.16947 | -0.036386 | -0.036584 | -0.016611 | -0.019356 | -0.074909 | -0.06712 | -0.06254 | -0.067765 | -0.086148 | -0.042758 | -0.04667 | -0.057702 | -0.060872 | -0.048508 | -0.027958 | -0.013005 | -0.054386 | -0.048645 | -0.030674
a10 | -0.058454 | 0.019516 | 0.036233 | -0.0083737 | 0.0019973 | 0.0062557 | 0.013058 | -0.045944 | -0.027778 | 0.076952 | 0.024547 | 0.047021 | -0.020361 | 0.01095 | 0.02408 | 0.014848 | -0.024505 | -0.041017 | 0.0043729 | -0.12287 |  | a10 | 0.089137 | 0.043242 | -0.023286 | -0.036637 | -0.024251 | -0.036055 | 0.011445 | -0.010541 | -0.012141 | 0.01731 | -0.023268 | 0.016186 | 0.0076518 | 0.045106 | 0.048953 | 0.017877 | -0.035497 | -0.014563 | -0.0059384 | -0.066659
a11 | 0.13814 | -0.02913 | -0.038963 | -0.02942 | -0.027151 | -0.042319 | -0.044692 | -0.038382 | -0.036424 | -0.027976 | -0.032519 | -0.025185 | -0.027533 | -0.01884 | -0.033839 | -0.034707 | -0.036867 | -0.033459 | -0.040478 | -0.027058 |  | a11 | 0.13279 | -0.051782 | -0.040054 | -0.037417 | -0.044459 | -0.052453 | -0.054718 | -0.057256 | -0.050632 | -0.05135 | -0.051553 | -0.048924 | -0.041707 | -0.042542 | -0.046023 | -0.0553 | -0.058078 | -0.051618 | -0.050052 | -0.025362
a12 | -0.014951 | 0.003043 | -0.0017218 | -0.0089896 | -0.0054999 | 0.00050689 | -0.0020128 | -0.0038729 | -0.003301 | 0.0053776 | 0.0010398 | 0.0039877 | -0.0014505 | 0.00059863 | 0.0088461 | -0.0014138 | -0.0077038 | -0.0073928 | 0.0013266 | -0.0034899 |  | a12 | 0.006514 | 0.0040938 | -0.0021905 | -0.0082335 | -0.0094012 | -0.0042021 | 0.0013707 | -0.00050238 | 0.0012176 | -0.0021492 | -0.0049007 | 0.0035922 | 0.0028689 | 0.0072386 | 0.0066754 | 0.0074029 | -0.00064597 | -0.00069822 | -0.00043847 | -0.0078993
a13 | -0.02782 | -0.0038065 | -0.0021293 | -0.0050294 | -0.0026333 | -0.0031253 | -0.0048194 | -0.003446 | -0.00035567 | 0.00069528 | 0.0014877 | -0.001522 | -0.00070872 | 0.00095499 | -0.00019154 | -0.0015879 | 0.0002086 | 0.0022854 | -0.0011497 | 0.0039067 |  | a13 | -0.021094 | -0.0073088 | -0.0073856 | -0.0095128 | -0.0062713 | -0.0033423 | -0.0045996 | -0.0018147 | -0.0011605 | 0.001522 | -0.0051277 | -0.0029837 | -0.0063993 | -0.0046565 | -0.0068562 | -0.0018643 | -0.0032747 | -0.0035644 | -0.0056957 | -0.0049802
a14 | 0.00029341 | -0.010137 | -0.016182 | -0.016489 | -0.012699 | -0.014255 | -0.010935 | -0.015999 | -0.012884 | -0.010132 | -0.010778 | -0.010629 | -0.018573 | -0.014846 | -0.015629 | -0.01603 | -0.016375 | -0.019334 | -0.01277 | -0.0076571 |  | a14 | 0.00044838 | -0.014496 | -0.021862 | -0.02369 | -0.01993 | -0.023247 | -0.015123 | -0.017458 | -0.018103 | -0.01576 | -0.021894 | -0.015812 | -0.019209 | -0.014332 | -0.016661 | -0.021138 | -0.02317 | -0.018025 | -0.018815 | -0.01027
a15 | 0.0022129 | -0.020185 | -0.0092982 | -0.0062748 | -0.012721 | -0.00043859 | -0.00030522 | -0.0045897 | -0.01159 | -0.0097815 | -0.015212 | -0.015881 | -0.008866 | -0.010148 | -0.0065802 | -0.0026776 | -0.0052222 | -0.011124 | -0.0074685 | -0.011366 |  | a15 | -0.00056456 | -0.026226 | -0.016852 | -0.016552 | -0.02061 | -0.015369 | -0.012135 | -0.015831 | -0.020799 | -0.025342 | -0.019404 | -0.026512 | -0.017759 | -0.018611 | -0.01516 | -0.014574 | -0.016405 | -0.022433 | -0.015616 | -0.026256
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.30761 | -0.30522 | -0.018423 | 0.079119 | -0.096196 | 0.21245 | 0.23307 | 0.096181 | -0.099567 | -0.19441 | -0.26964 | -0.2827 | 0.040294 | -0.04919 | 0.092587 | 0.15068 | 0.13589 | -0.0082115 | 0.070514 | 0.048374 |  | Dx (mm) | 0.08 | -0.20597 | 0.099012 | 0.10376 | 0.0029098 | 0.11417 | 0.14962 | 0.062884 | -0.062573 | -0.19888 | -0.0056184 | -0.19828 | 0.05752 | -0.037811 | 0.064847 | 0.04151 | 0.031618 | -0.070731 | 0.077882 | -0.081396
Dy (mm) | -0.22092 | 0.031705 | 0.078482 | -0.015161 | 0.0010052 | 0.025242 | 0.042912 | -0.098102 | -0.066282 | 0.16243 | 0.043177 | 0.094322 | -0.043528 | 0.022968 | 0.057016 | 0.039744 | -0.04791 | -0.091416 | 0.013537 | -0.30754 |  | Dy (mm) | 0.20717 | 0.083217 | -0.047031 | -0.076266 | -0.052719 | -0.071665 | 0.037233 | -0.019338 | -0.031656 | 0.025506 | -0.052603 | 0.023599 | 0.020404 | 0.10017 | 0.11402 | 0.043133 | -0.076981 | -0.037743 | -0.0082198 | -0.17448
Mag. len. (m) | 14.4295 | [20 empty] | Mag. len. (m) | 14.4294
TF (mT/KA) | 594.6023 | [20 empty] | TF (mT/KA) | 594.9648
Coil pos. (mm) | 0.022379 | [20 empty] | Coil pos. (mm) | 0.020256
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.015423 | [20 empty] | Twist (m2 rad) | 0.020988

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 594.6023 | 0.26783 | 0.0041496 |  | -1.5818 | 2.309 |  |  | 594.9648 | 0.26126 | 0.0012726 |  | -0.96417 | 2.4538 |  | 0 | 4
2 | -0.50133 |  | -0.48962 | 0.5767 | -0.61807 |  | -0.47004 | 0.94106 | 0.092443 |  | 0.039902 | 0.4791 | 0.21488 |  | 0.13307 | 0.61106 |  | 0 | 5
3 | -0.41669 |  | -2.0674 | 1.053 | -0.31934 |  | -0.29563 | 0.31051 | 0.094714 |  | -1.5131 | 0.952 | -0.51213 |  | -0.46065 | 0.2314 |  | 0 | 6
4 | 0.12505 |  | 0.079939 | 0.26911 | -0.14259 |  | -0.03466 | 0.29995 | -0.14881 |  | -0.15504 | 0.24548 | -0.12598 |  | -0.14148 | 0.23416 |  | 0 | 7
5 | 0.70644 |  | 0.93021 | 0.15563 | -0.018946 |  | -0.093066 | 0.13233 | 0.61326 |  | 0.83208 | 0.18222 | -0.046441 |  | -0.14331 | 0.17043 |  | 0 | 8
6 | 0.012797 |  | -0.0056315 | 0.098597 | 0.028836 |  | 0.065749 | 0.073845 | -0.024193 |  | -0.021721 | 0.081147 | 0.04306 |  | 0.0575 | 0.080862 |  | 0 | 9
7 | 1.2127 |  | 1.165 | 0.0513 | -0.008143 |  | -0.058519 | 0.058013 | 1.2443 |  | 1.1971 | 0.033215 | 0.0044112 |  | -0.049742 | 0.10316 |  | 0 | 10
8 | 0.0071593 |  | 0.0033322 | 0.0203 | -0.0029892 |  | -0.00052464 | 0.046582 | -0.007618 |  | -0.0078262 | 0.020974 | -0.0035837 |  | -0.0032441 | 0.028379 |  | 0 | 11
9 | 0.58378 |  | 0.58778 | 0.018578 | -0.028428 |  | -0.023513 | 0.016649 | 0.57946 |  | 0.57999 | 0.011723 | -0.051015 |  | -0.047166 | 0.021 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.76154 |  | 0.76871 | 0.0073044 | -0.026954 |  | -0.032997 | 0.0069157 | 0.7507 |  | 0.75623 | 0.0092037 | -0.042383 |  | -0.049114 | 0.0060009 |  | 0 | 14
12 | 4.8351e-05 |  | 2.7712e-05 | 0.0030197 | -9.7972e-05 |  | -0.001563 | 0.0032965 | -5.2089e-05 |  | -0.00033699 | 0.0038609 | -4.8347e-06 |  | 0.00018211 | 0.003127 |  | 0 | 15
13 | 0.0058941 |  | 0.05768 | 0.0037951 | -0.00019321 |  | -0.0011913 | 0.0021147 | 0.0058081 |  | 0.056831 | 0.0031571 | -0.00049854 |  | -0.0043161 | 0.0024091 |  | 0 | 16
14 | -0.00071711 |  | -0.0075089 | 0.0010539 | -0.0014098 |  | -0.015139 | 0.0019122 | -0.0011861 |  | -0.012632 | 0.0010247 | -0.001794 |  | -0.018917 | 0.0017929 |  | 0 | 17
15 | 0.003401 |  | 0.034978 | 0.0040646 | -0.00085215 |  | -0.008798 | 0.0052895 | 0.0037494 |  | 0.039102 | 0.0045987 | -0.0018331 |  | -0.018677 | 0.0042006 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.30559 | -0.61421
3 | [12 empty] | 3 | 9 | -0.55896 | -0.75641
4 | [12 empty] | 4 | 16 | -1.2584 | -1.4281
5 | [12 empty] | 5 | 25 | -1.9414 | -1.736
6 | [12 empty] | 6 | 36 | -2.4613 | -2.5133
7 | [12 empty] | 7 | 49 | -2.9086 | -2.8381
8 | [12 empty] | 8 | 64 | -3.4818 | -3.7133
9 | [12 empty] | 9 | 81 | -4.0406 | -4.1547
10 | [12 empty] | sigma (mm) |  | 0.022379 | 0.020256
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 372.8531 | 595.3508 | 594.8583 | 594.7796 | 594.6842 | 594.7616 | 594.4516 | 594.3787 | 594.3949 | 594.4553 | 594.3992 | 594.4122 | 594.4492 | 594.4413 | 594.4871 | 594.4725 | 594.4506 | 594.5527 | 595.0611 | 364.0458
2 | -0.99338 | -0.93696 | -0.87832 | -1.4645 | -0.48882 | -0.87774 | 0.093857 | 0.07642 | 0.032946 | 0.3855 | -0.53595 | 0.086842 | 0.10138 | 0.16512 | -0.56201 | -0.66914 | -0.99353 | -1.4579 | -0.89021 | -0.3379 | -0.50133
3 | 45.3761 | -0.080452 | -2.4044 | -2.8083 | -3.0312 | -2.6455 | -1.707 | -1.4721 | -1.4271 | -2.6983 | -2.147 | -2.6095 | -2.8648 | -2.9863 | -3.4002 | -2.5227 | -1.2827 | -1.7412 | 0.61538 | 1.2064 | -0.41669
4 | 1.3493 | -0.23473 | -0.37833 | -0.12894 | -0.0094017 | -0.065868 | 0.39441 | 0.34257 | 0.44522 | 0.40667 | 0.11245 | 0.52493 | 0.19468 | 0.10024 | 0.25021 | -0.057711 | -0.089 | -0.26327 | -0.10522 | 0.19976 | 0.12505
5 | -2.2169 | 0.97343 | 1.0293 | 1.0353 | 1.0617 | 1.0603 | 0.88308 | 1.0311 | 1.0856 | 0.97329 | 0.61837 | 1.1224 | 0.93305 | 0.89699 | 0.99444 | 0.92961 | 0.7153 | 0.81212 | 0.58837 | -2.878 | 0.70644
6 | 0.3254 | 0.10729 | -0.12128 | -0.081969 | -0.11301 | 0.081189 | 0.084719 | 0.10595 | 0.11297 | 0.17693 | 0.094941 | -0.082319 | -0.089107 | -0.070064 | -0.080492 | -0.075544 | -0.013106 | -0.053738 | -0.084724 | 0.23454 | 0.012797
7 | 2.7751 | 1.1553 | 1.1254 | 1.0892 | 1.1425 | 1.1063 | 1.2487 | 1.2322 | 1.1784 | 1.1886 | 1.2208 | 1.1503 | 1.1641 | 1.1978 | 1.0611 | 1.1802 | 1.1926 | 1.1186 | 1.2178 | 1.0156 | 1.2127
8 | 0.10451 | 0.042508 | 0.016399 | -0.020634 | 0.0050487 | 0.011053 | 0.015231 | 0.012329 | 0.01015 | 0.036691 | -0.0033507 | -0.0035118 | -0.030445 | 0.0032808 | 0.018804 | 0.00225 | -0.012464 | -0.034081 | -0.0092779 | 0.019899 | 0.0071593
9 | 0.44522 | 0.53915 | 0.57483 | 0.58179 | 0.60779 | 0.61109 | 0.58643 | 0.60245 | 0.62125 | 0.59256 | 0.60655 | 0.5922 | 0.59011 | 0.58728 | 0.57979 | 0.57844 | 0.57831 | 0.56995 | 0.58008 | 0.60822 | 0.58378
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.64 | 0.76387 | 0.77567 | 0.7854 | 0.77944 | 0.77572 | 0.75873 | 0.75867 | 0.76679 | 0.7715 | 0.76202 | 0.76944 | 0.76979 | 0.77001 | 0.7733 | 0.76585 | 0.76459 | 0.76619 | 0.75976 | 0.67527 | 0.76154
12 | 0.052413 | 0.015677 | 0.045208 | -0.0065211 | 0.018042 | 0.051681 | -0.01322 | -0.019743 | -0.039029 | 0.0071419 | 0.0393 | -0.033912 | -0.034493 | -0.014714 | -0.043676 | 0.02281 | 0.0074202 | -0.025397 | 0.028415 | 0.089927 | 0.00048351 |  | ten times!
13 | 0.9329 | 0.56323 | 0.49935 | 0.54242 | 0.54216 | 0.51541 | 0.58673 | 0.58682 | 0.56403 | 0.58766 | 0.59342 | 0.56068 | 0.5916 | 0.61464 | 0.58243 | 0.5678 | 0.64063 | 0.59656 | 0.64686 | 0.60835 | 0.058941 |  | ten times!
14 | 0.047126 | -0.061266 | -0.072881 | -0.070425 | -0.07094 | -0.055484 | -0.075266 | -0.063379 | -0.064775 | -0.085737 | -0.068373 | -0.093646 | -0.082187 | -0.080475 | -0.084747 | -0.068722 | -0.079184 | -0.088765 | -0.085353 | -0.094146 | -0.0071711 |  | ten times!
15 | 0.094241 | 0.28517 | 0.35084 | 0.39445 | 0.36121 | 0.37642 | 0.34054 | 0.38732 | 0.36564 | 0.28024 | 0.2888 | 0.27396 | 0.37604 | 0.33781 | 0.38427 | 0.3699 | 0.37797 | 0.3921 | 0.35334 | 0.30736 | 0.03401 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 43.1113 | 3.1607 | 1.3083 | 1.2777 | 3.4337 | -1.6373 | -3.1312 | -2.6265 | -2.1558 | -3.4847 | -4.2953 | -1.2097 | -3.247 | -2.4211 | -2.7352 | -1.9856 | -2.8213 | -3.7445 | -2.158 | 2.4619 | 0.0041496
2 | -5.6707 | 2.0757 | 0.4992 | -0.9103 | -1.9531 | 0.50191 | -0.48696 | -1.2944 | -0.47222 | -0.1304 | -1.2776 | -0.89137 | -1.0886 | -1.8526 | -0.33668 | -0.29462 | -0.40299 | -0.14123 | -0.0044636 | 0.19648 | -0.61807
3 | -2.1686 | -0.82274 | -0.28268 | -0.46106 | -0.14887 | -1.0268 | -0.77069 | -0.3964 | -0.012809 | 0.18231 | -0.25566 | -0.094398 | -0.2374 | -0.19995 | -0.17857 | -0.033648 | -0.010377 | -0.19994 | -0.37167 | 0.87976 | -0.31934
4 | -2.7795 | -0.50839 | 0.12651 | 0.38128 | 0.24236 | -0.34679 | -0.24362 | 0.027432 | 0.10684 | -0.56595 | -0.42296 | -0.059174 | -0.17551 | 0.18822 | -0.23226 | 0.062796 | 0.35896 | 0.34043 | 0.095937 | -0.61461 | -0.14259
5 | 3.0434 | -0.27683 | -0.11699 | -0.30949 | -0.12741 | -0.20599 | -0.13013 | -0.093998 | 0.027829 | 0.10639 | -0.078462 | -0.11663 | -0.026225 | 0.17522 | 0.095412 | -0.17685 | -0.094518 | -0.070501 | -0.25599 | -0.97528 | -0.018946
6 | -0.75374 | 0.13192 | 0.1042 | 0.062902 | 0.019159 | -0.0094364 | -0.096784 | -0.043295 | 0.036976 | -0.047997 | 0.088815 | 0.13607 | 0.12984 | 0.093236 | 0.14548 | 0.14006 | 0.087887 | 0.13025 | 0.074212 | -0.25507 | 0.028836
7 | 1.5951 | -0.051176 | -0.13227 | -0.016024 | -0.061665 | -0.15308 | -0.16089 | -0.067093 | -0.077309 | -0.11717 | 0.024719 | 0.016937 | -0.041914 | 0.029989 | 0.0079111 | -0.033999 | -0.063319 | -0.083507 | -0.073491 | -0.16902 | -0.008143
8 | -0.090977 | 0.074146 | 0.054092 | 0.086636 | 0.030597 | 0.014949 | -0.015439 | -0.075322 | -0.020593 | -0.045195 | -0.0087155 | -0.024196 | -0.064871 | -0.062146 | -0.024293 | -0.0017019 | 0.034856 | 0.0095672 | 0.028187 | 0.014239 | -0.0029892
9 | -0.18819 | -0.0097953 | -0.047868 | -0.056536 | -0.047382 | -0.036845 | -0.026457 | -0.018977 | -0.015661 | -0.008587 | -0.0057406 | -0.017036 | -0.025653 | -0.030809 | -0.03266 | -0.0041634 | -0.0055843 | -0.0010509 | -0.032426 | -0.0092268 | -0.028428
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.1442 | -0.028299 | -0.039185 | -0.028938 | -0.027492 | -0.042601 | -0.044625 | -0.038034 | -0.036975 | -0.026475 | -0.032115 | -0.023525 | -0.027547 | -0.018767 | -0.034349 | -0.034785 | -0.035978 | -0.03366 | -0.040592 | -0.024593 | -0.026954
12 | 0.059213 | 0.0089019 | -0.045753 | -0.081332 | -0.057177 | 0.00049299 | -0.030028 | 0.0032304 | -0.0068163 | -0.016094 | -0.0057641 | -0.0030897 | 0.0036155 | -0.0037454 | 0.065021 | -0.028445 | -0.055913 | -0.036132 | 0.0076782 | 0.091033 | -0.00097972 |  | ten times!
13 | -0.26622 | -0.058628 | -0.01924 | -0.040846 | -0.035599 | -0.0082955 | -0.028383 | -0.025424 | -0.01849 | 0.006486 | -0.0031514 | -0.027733 | -0.0037904 | 0.0055856 | 0.011433 | 0.0032586 | 0.017675 | 0.014664 | -0.0039485 | 0.015869 | -0.0019321 |  | ten times!
14 | 0.014474 | -0.15955 | -0.18591 | -0.15588 | -0.13737 | -0.14961 | -0.1208 | -0.12506 | -0.11839 | -0.15447 | -0.15183 | -0.16454 | -0.16931 | -0.15896 | -0.16931 | -0.16908 | -0.14299 | -0.16457 | -0.1273 | 0.0058 | -0.014098 |  | ten times!
15 | 0.022129 | -0.20185 | -0.092982 | -0.062748 | -0.12721 | -0.0043859 | -0.0030522 | -0.045897 | -0.1159 | -0.097815 | -0.15212 | -0.15881 | -0.08866 | -0.10148 | -0.065802 | -0.026776 | -0.052222 | -0.11124 | -0.074685 | -0.11366 | -0.0085215 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 375.9884 | 595.6374 | 595.2715 | 595.1414 | 595.122 | 595.1593 | 594.7833 | 594.6999 | 594.8233 | 594.8272 | 594.7597 | 594.7974 | 594.8264 | 594.762 | 594.8036 | 594.8273 | 594.8382 | 594.8946 | 595.3913 | 374.0374
2 | 0.64129 | -0.85978 | -0.20687 | -0.17684 | 0.048441 | -0.42564 | 0.023606 | -0.11335 | 0.41312 | 0.61863 | -0.53463 | 0.5358 | -0.80062 | 0.24312 | 0.51459 | 0.00064212 | 0.21032 | 0.48821 | 0.73947 | 1.0467 | 0.092443
3 | 44.4326 | -0.49006 | -1.0383 | -1.1957 | -1.2752 | -2.6906 | -0.86425 | -1.213 | -0.92575 | -2.0094 | -2.8229 | -1.8682 | -1.4827 | -1.7242 | -2.635 | -3.0889 | -1.8868 | -0.7914 | 0.76712 | 1.5534 | 0.094714
4 | 0.27107 | -0.25907 | -0.14988 | -0.035761 | 0.055948 | -0.14033 | 0.1611 | 0.15693 | 0.070247 | 0.29311 | -0.63622 | -0.37394 | -0.41416 | -0.30165 | -0.4761 | -0.19104 | -0.067272 | -0.16822 | -0.31434 | -0.39233 | -0.14881
5 | -1.9365 | 0.738 | 0.88782 | 0.95078 | 0.87598 | 0.91146 | 0.83947 | 1.0478 | 0.99609 | 1.1593 | 0.81159 | 0.96733 | 0.83319 | 0.83059 | 0.57635 | 0.85999 | 0.73958 | 0.5123 | 0.43981 | -3.0884 | 0.61326
6 | -0.048469 | 0.095178 | 0.04016 | -0.0078377 | -0.012234 | -0.01998 | 0.079576 | 0.079415 | 0.036878 | 0.0038055 | -0.014348 | -0.077713 | -0.051507 | -0.1431 | -0.17264 | 0.045102 | -0.043311 | -0.059621 | -0.1688 | -0.070689 | -0.024193
7 | 2.9249 | 1.2551 | 1.1868 | 1.1675 | 1.1818 | 1.1553 | 1.2233 | 1.1959 | 1.2248 | 1.2196 | 1.1926 | 1.1858 | 1.1975 | 1.2141 | 1.1229 | 1.1596 | 1.1934 | 1.2281 | 1.2434 | 0.90732 | 1.2443
8 | -0.055321 | -0.028547 | -0.039013 | -0.013195 | -0.019633 | -0.026179 | -0.031662 | -0.011522 | 0.0062883 | -0.010262 | -0.010617 | 0.018207 | -0.026156 | 0.017898 | 0.0001412 | -0.023709 | 0.039311 | 0.0073795 | 0.010399 | 0.046378 | -0.007618
9 | 0.56963 | 0.57974 | 0.57006 | 0.58535 | 0.60951 | 0.59194 | 0.57547 | 0.56406 | 0.578 | 0.58374 | 0.58453 | 0.58505 | 0.58345 | 0.58704 | 0.58894 | 0.55763 | 0.57066 | 0.57015 | 0.57446 | 0.57407 | 0.57946
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.68114 | 0.75375 | 0.75732 | 0.76574 | 0.77989 | 0.77144 | 0.74262 | 0.74062 | 0.75183 | 0.7518 | 0.75762 | 0.75422 | 0.75507 | 0.75 | 0.75617 | 0.75791 | 0.75995 | 0.75256 | 0.75357 | 0.66221 | 0.7507
12 | -0.10588 | -0.046998 | -0.015472 | 0.026603 | -0.054552 | 0.025222 | -0.076583 | -0.047838 | -0.034007 | -0.035285 | -0.0096994 | 0.022869 | -0.0023898 | 0.039358 | 0.051947 | 0.014558 | 0.043975 | 0.0030325 | 0.034602 | 0.043357 | -0.00052089 |  | ten times!
13 | 0.91892 | 0.54325 | 0.54732 | 0.56231 | 0.56338 | 0.5182 | 0.58623 | 0.56822 | 0.57704 | 0.53654 | 0.57765 | 0.54081 | 0.56232 | 0.55921 | 0.54835 | 0.56057 | 0.61709 | 0.61748 | 0.64368 | 0.59888 | 0.058081 |  | ten times!
14 | 0.043807 | -0.12042 | -0.13043 | -0.13007 | -0.14531 | -0.12402 | -0.12678 | -0.1341 | -0.12988 | -0.14242 | -0.11877 | -0.13712 | -0.13032 | -0.12729 | -0.11453 | -0.11869 | -0.10527 | -0.11255 | -0.1258 | -0.060995 | -0.011861 |  | ten times!
15 | 0.020461 | 0.30662 | 0.43034 | 0.46198 | 0.4157 | 0.45946 | 0.39557 | 0.41379 | 0.39466 | 0.31974 | 0.40156 | 0.33809 | 0.38221 | 0.33262 | 0.34878 | 0.40745 | 0.44181 | 0.38247 | 0.40558 | 0.27078 | 0.037494 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 39.5226 | 3.9919 | 1.8152 | 2.0637 | 2.9343 | 1.1561 | -1.5102 | -3.2929 | -0.72587 | -3.4847 | -2.5628 | -0.89219 | -0.27298 | -1.3845 | -4.2574 | -3.7984 | -3.1276 | -2.271 | -1.7358 | -12.1053 | 0.0012726
2 | -0.55351 | 0.87663 | 0.42329 | -0.61278 | -1.1683 | 0.68574 | 0.3816 | 0.37104 | 0.48939 | 0.056028 | -0.31601 | 1.2858 | 0.44443 | 0.32676 | -0.077321 | -0.15229 | 0.46177 | -0.31832 | -0.76215 | 3.3302 | 0.21488
3 | -2.6729 | -0.78444 | -0.66979 | -0.31376 | -0.51746 | -0.22148 | -0.58139 | -0.70779 | 0.046263 | -0.29557 | -0.68929 | -0.25752 | -0.6379 | -0.41361 | -0.2544 | -0.45599 | -0.20539 | -0.65045 | -0.68173 | 0.18704 | -0.51213
4 | -0.34542 | -0.46442 | -0.22228 | -0.20267 | -0.053307 | -0.41099 | 0.20695 | 0.27854 | 0.13642 | -0.40797 | -0.4827 | -0.017477 | -0.10684 | 0.0091115 | -0.25064 | -0.19415 | -0.39279 | -0.052915 | 0.081527 | 0.53926 | -0.12598
5 | 3.1968 | -0.23798 | 0.048405 | 0.059033 | 0.052745 | -0.3154 | -0.17228 | -0.038424 | -0.085368 | 0.11602 | -0.49912 | -0.38126 | -0.25952 | -0.13819 | 0.06863 | -0.18379 | -0.17114 | -0.20515 | -0.23674 | -0.53319 | -0.046441
6 | -0.20401 | 0.016416 | -0.026511 | -0.038608 | 0.10365 | 0.2189 | -0.038039 | 0.0061688 | 0.10383 | 0.068292 | -0.068079 | 0.091284 | 0.029721 | 0.038744 | -0.0061431 | 0.17479 | 0.16374 | 0.11408 | 0.082777 | -0.12204 | 0.04306
7 | 1.6112 | 0.027105 | 0.082726 | 0.044045 | 0.011067 | -0.20573 | -0.2277 | -0.2294 | -0.151 | -0.12244 | 0.079167 | 0.018893 | 0.029456 | 0.01582 | -0.024272 | 0.016308 | -0.090833 | -0.093342 | -0.075241 | -0.060574 | 0.0044112
8 | 0.041893 | -0.00606 | -0.015745 | -0.013516 | 0.055286 | 0.066045 | 0.022856 | -0.0071716 | -0.013213 | 0.011466 | -0.02997 | -0.015384 | -0.035381 | -0.025746 | -0.02884 | -0.015277 | -0.018351 | -0.012119 | 0.022726 | -0.059146 | -0.0035837
9 | -0.17499 | -0.027751 | -0.034292 | -0.012779 | -0.019319 | -0.070848 | -0.069338 | -0.061915 | -0.068646 | -0.083146 | -0.042941 | -0.043879 | -0.058216 | -0.059218 | -0.052115 | -0.028806 | -0.011942 | -0.05558 | -0.048251 | -0.036712 | -0.051015
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.13276 | -0.050555 | -0.039855 | -0.036563 | -0.044632 | -0.051847 | -0.054817 | -0.057264 | -0.050705 | -0.051436 | -0.051618 | -0.048377 | -0.041842 | -0.042543 | -0.046916 | -0.055588 | -0.057788 | -0.05172 | -0.049993 | -0.025742 | -0.042383
12 | -0.05764 | -0.0035704 | 0.00090952 | -0.046271 | -0.07323 | -0.014155 | 0.0032679 | 0.0034273 | 0.024626 | -0.02959 | -0.028067 | 0.022217 | 0.023214 | 0.030577 | 0.024914 | 0.057391 | 0.026925 | 0.0077548 | 0.0024396 | -0.0088852 | -4.8347e-05 |  | ten times!
13 | -0.21852 | -0.086092 | -0.060967 | -0.082221 | -0.068448 | -0.018054 | -0.026577 | -0.011455 | -0.024036 | -0.0050524 | -0.057408 | -0.050434 | -0.053839 | -0.041341 | -0.053178 | -0.0086489 | -0.033277 | -0.049486 | -0.046385 | -0.071788 | -0.0049854 |  | ten times!
14 | 0.0013649 | -0.21046 | -0.18821 | -0.19374 | -0.18076 | -0.1909 | -0.14841 | -0.15979 | -0.18146 | -0.20582 | -0.20244 | -0.20798 | -0.1901 | -0.17655 | -0.19127 | -0.22087 | -0.19942 | -0.18143 | -0.17539 | -0.081389 | -0.01794 |  | ten times!
15 | -0.0056456 | -0.26226 | -0.16852 | -0.16552 | -0.2061 | -0.15369 | -0.12135 | -0.15831 | -0.20799 | -0.25342 | -0.19404 | -0.26512 | -0.17759 | -0.18611 | -0.1516 | -0.14574 | -0.16405 | -0.22433 | -0.15616 | -0.26256 | -0.018331 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.0012589 | 0.00043059 | 0.00029823 | 0.00013779 | 0.00026796 | -0.0002534 | -0.000376 | -0.00034876 | -0.00024718 | -0.00034153 | -0.00031966 | -0.00025744 | -0.00027072 | -0.0001937 | -0.00021825 | -0.00025508 | -8.337e-05 | 0.00077165 |  | 594.6023
|  | 0.0011305 | 0.00051555 | 0.00029688 | 0.00026427 | 0.00032697 | -0.000305 | -0.00044518 | -0.00023777 | -0.00023122 | -0.00034467 | -0.00028131 | -0.00023256 | -0.0003408 | -0.00027088 | -0.00023105 | -0.00021273 | -0.00011793 | 0.00071691 |  | 594.9648
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3802.6437
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 594.5268 | -3728.5738 | 13.8596 | 5.5757 | 4.2519 | 2.6473 | 3.9491 | -1.2651 | -2.4913 | -2.2188 | -1.2028 | -2.1465 | -1.9278 | -1.3054 | -1.4383 | -0.66797 | -0.91354 | -1.2819 | 0.43543 | 8.9868 | -3876.7135
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.15818 | 4.4693 | 0.47425 | 0.28901 | 0.28595 | 0.50155 | -0.0055463 | -0.15494 | -0.10447 | -0.057398 | -0.19029 | -0.27134 | 0.03721 | -0.16651 | -0.083931 | -0.11534 | -0.040382 | -0.12395 | -0.21627 | -0.057615 | 0.40438 |  | -1.3974e-17
db2 | -0.48962 | -0.99338 | -0.44734 | -0.3887 | -0.97491 | 0.00079032 | -0.38813 | 0.58347 | 0.56604 | 0.52256 | 0.87512 | -0.046333 | 0.57646 | 0.591 | 0.65473 | -0.072396 | -0.17952 | -0.50391 | -0.96832 | -0.4006 | -0.3379 |  | 0
db3 | -2.0658 | 45.3761 | 1.987 | -0.337 | -0.74093 | -0.96378 | -0.57809 | 0.36042 | 0.59536 | 0.6403 | -0.6309 | -0.079584 | -0.54213 | -0.79744 | -0.91886 | -1.3328 | -0.45524 | 0.7847 | 0.32624 | 2.6828 | 1.2064 |  | 0
db4 | 0.079939 | 1.3493 | -0.31467 | -0.45826 | -0.20888 | -0.089341 | -0.14581 | 0.31447 | 0.26263 | 0.36528 | 0.32673 | 0.032508 | 0.44499 | 0.11474 | 0.020306 | 0.17027 | -0.13765 | -0.16894 | -0.34321 | -0.18516 | 0.19976 |  | 0
db5 | 0.93182 | -2.2169 | 0.043221 | 0.099077 | 0.10506 | 0.13149 | 0.13011 | -0.047126 | 0.10092 | 0.15542 | 0.043081 | -0.31183 | 0.19215 | 0.0028459 | -0.033215 | 0.064235 | -0.00060096 | -0.21491 | -0.11808 | -0.34184 | -2.878 |  | 0
db6 | -0.0056315 | 0.3254 | 0.11292 | -0.11565 | -0.076337 | -0.10738 | 0.08682 | 0.090351 | 0.11158 | 0.11861 | 0.18256 | 0.10057 | -0.076688 | -0.083476 | -0.064433 | -0.07486 | -0.069913 | -0.0074741 | -0.048107 | -0.079092 | 0.23454 |  | 0
db7 | 1.1671 | 2.7751 | -0.0097015 | -0.039597 | -0.075785 | -0.022509 | -0.058717 | 0.083671 | 0.067153 | 0.013448 | 0.023602 | 0.055787 | -0.014658 | -0.00089133 | 0.03283 | -0.10387 | 0.015214 | 0.027633 | -0.046413 | 0.0528 | 1.0156 |  | -1.2336e-17
db8 | 0.0033322 | 0.10451 | 0.039176 | 0.013067 | -0.023966 | 0.0017165 | 0.0077204 | 0.011899 | 0.0089964 | 0.0068182 | 0.033359 | -0.006683 | -0.006844 | -0.033778 | -5.1439e-05 | 0.015472 | -0.0010822 | -0.015797 | -0.037413 | -0.01261 | 0.019899 |  | 0
db9 | 0.59301 | 0.44522 | -0.048629 | -0.012949 | -0.005992 | 0.020007 | 0.023307 | -0.0013554 | 0.014672 | 0.033467 | 0.0047842 | 0.018771 | 0.0044195 | 0.0023303 | -0.00050201 | -0.0079865 | -0.0093455 | -0.0094753 | -0.017826 | -0.0076983 | 0.60822 |  | -6.1679e-18
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.7693 | 0.64 | -0.0048341 | 0.0069573 | 0.016689 | 0.010735 | 0.0070102 | -0.0099795 | -0.010035 | -0.0019173 | 0.002788 | -0.0066855 | 0.0007317 | 0.0010839 | 0.0013056 | 0.0045876 | -0.0028557 | -0.0041155 | -0.0025225 | -0.0089439 | 0.67527 |  | -1.8504e-17
db12 | 2.7712e-05 | 0.0052413 | 0.00154 | 0.0044931 | -0.00067982 | 0.0017764 | 0.0051404 | -0.0013497 | -0.0020021 | -0.0039307 | 0.00068648 | 0.0039023 | -0.0034189 | -0.003477 | -0.0014991 | -0.0043953 | 0.0022532 | 0.00071431 | -0.0025674 | 0.0028138 | 0.0089927 |  | 0
db13 | 0.05768 | 0.09329 | -0.0013575 | -0.0077449 | -0.0034384 | -0.0034637 | -0.006139 | 0.00099318 | 0.0010015 | -0.0012772 | 0.0010859 | 0.0016615 | -0.0016118 | 0.0014795 | 0.003784 | 0.00056275 | -0.00089981 | 0.0063824 | 0.0019757 | 0.0070059 | 0.060835 |  | 1.1179e-17
db14 | -0.0075089 | 0.0047126 | 0.0013824 | 0.00022083 | 0.00046645 | 0.0004149 | 0.0019605 | -1.7647e-05 | 0.001171 | 0.0010314 | -0.0010648 | 0.00067166 | -0.0018557 | -0.00070984 | -0.0005386 | -0.00096576 | 0.00063667 | -0.00040945 | -0.0013676 | -0.0010264 | -0.0094146 |  | -8.1917e-19
db15 | 0.034978 | 0.0094241 | -0.0064605 | 0.00010638 | 0.0044669 | 0.0011434 | 0.0026637 | -0.00092426 | 0.0037541 | 0.0015866 | -0.0069538 | -0.0060975 | -0.0075824 | 0.0026265 | -0.001197 | 0.0034494 | 0.0020116 | 0.0028193 | 0.0042318 | 0.00035585 | 0.030736 |  | 3.8549e-18
da2 | -0.47004 | -5.6707 | 2.5457 | 0.96924 | -0.44026 | -1.483 | 0.97195 | -0.016921 | -0.82432 | -0.0021796 | 0.33964 | -0.80761 | -0.42133 | -0.61852 | -1.3826 | 0.13336 | 0.17542 | 0.067047 | 0.32881 | 0.46558 | 0.19648 |  | 0
da3 | -0.29563 | -2.1686 | -0.52711 | 0.012948 | -0.16543 | 0.14676 | -0.73115 | -0.47506 | -0.10077 | 0.28282 | 0.47794 | 0.039973 | 0.20123 | 0.058227 | 0.095676 | 0.11706 | 0.26198 | 0.28525 | 0.095688 | -0.076044 | 0.87976 |  | 0
da4 | -0.03466 | -2.7795 | -0.47373 | 0.16117 | 0.41594 | 0.27702 | -0.31213 | -0.20896 | 0.062092 | 0.1415 | -0.53129 | -0.3883 | -0.024514 | -0.14085 | 0.22288 | -0.1976 | 0.097456 | 0.39362 | 0.37509 | 0.1306 | -0.61461 |  | 0
da5 | -0.093066 | 3.0434 | -0.18377 | -0.023924 | -0.21643 | -0.034348 | -0.11292 | -0.037069 | -0.00093255 | 0.12089 | 0.19945 | 0.014603 | -0.023569 | 0.066841 | 0.26828 | 0.18848 | -0.083781 | -0.0014527 | 0.022565 | -0.16293 | -0.97528 |  | 0
da6 | 0.065749 | -0.75374 | 0.06617 | 0.038446 | -0.002847 | -0.04659 | -0.075185 | -0.16253 | -0.10904 | -0.028773 | -0.11375 | 0.023066 | 0.07032 | 0.064087 | 0.027487 | 0.079727 | 0.074314 | 0.022138 | 0.064499 | 0.0084634 | -0.25507 |  | 0
da7 | -0.058519 | 1.5951 | 0.0073428 | -0.073751 | 0.042495 | -0.003146 | -0.094557 | -0.10237 | -0.008574 | -0.01879 | -0.058646 | 0.083238 | 0.075456 | 0.016605 | 0.088508 | 0.06643 | 0.02452 | -0.0047994 | -0.024987 | -0.014971 | -0.16902 |  | 0
da8 | -0.00052464 | -0.090977 | 0.07467 | 0.054617 | 0.08716 | 0.031121 | 0.015473 | -0.014914 | -0.074798 | -0.020068 | -0.04467 | -0.0081908 | -0.023671 | -0.064347 | -0.061622 | -0.023769 | -0.0011772 | 0.03538 | 0.010092 | 0.028711 | 0.014239 |  | 0
da9 | -0.023513 | -0.18819 | 0.013718 | -0.024356 | -0.033023 | -0.023869 | -0.013332 | -0.0029442 | 0.0045354 | 0.007852 | 0.014926 | 0.017772 | 0.006477 | -0.0021403 | -0.0072964 | -0.0091466 | 0.01935 | 0.017929 | 0.022462 | -0.0089129 | -0.0092268 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.032997 | 0.1442 | 0.0046973 | -0.0061884 | 0.0040591 | 0.0055047 | -0.0096043 | -0.011628 | -0.0050376 | -0.0039783 | 0.0065217 | 0.00088212 | 0.009472 | 0.0054502 | 0.01423 | -0.0013522 | -0.0017881 | -0.0029816 | -0.00066335 | -0.0075948 | -0.024593 |  | 1.7347e-18
da12 | -0.001563 | 0.0059213 | 0.0024532 | -0.0030122 | -0.0065702 | -0.0041546 | 0.0016123 | -0.0014397 | 0.0018861 | 0.00088142 | -4.6307e-05 | 0.00098663 | 0.0012541 | 0.0019246 | 0.0011885 | 0.0080651 | -0.0012814 | -0.0040283 | -0.0020501 | 0.0023309 | 0.0091033 |  | 0
da13 | -0.0011913 | -0.026622 | -0.0046716 | -0.00073273 | -0.0028933 | -0.0023686 | 0.00036171 | -0.0016471 | -0.0013512 | -0.00065776 | 0.0018399 | 0.00087612 | -0.001582 | 0.00081222 | 0.0017498 | 0.0023345 | 0.0015171 | 0.0029588 | 0.0026577 | 0.00079641 | 0.0015869 |  | 0
da14 | -0.015139 | 0.0014474 | -0.00081622 | -0.0034525 | -0.00044897 | 0.0014015 | 0.00017793 | 0.0030587 | 0.0026322 | 0.0033 | -0.00030856 | -4.4826e-05 | -0.0013156 | -0.0017926 | -0.00075739 | -0.0017925 | -0.0017698 | 0.00083921 | -0.0013187 | 0.0024083 | 0.00058 |  | 1.2529e-18
da15 | -0.008798 | 0.0022129 | -0.011387 | -0.00050018 | 0.0025232 | -0.0039231 | 0.0083594 | 0.0084928 | 0.0042084 | -0.0027919 | -0.0009835 | -0.0064144 | -0.0070831 | -6.7966e-05 | -0.0013504 | 0.0022178 | 0.0061205 | 0.0035758 | -0.0023258 | 0.0013295 | -0.011366 |  | 9.6374e-19
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -3696.1611
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 594.8961 | -3679.7632 | 12.4615 | 6.3109 | 4.1239 | 3.7978 | 4.4248 | -1.8956 | -3.2975 | -1.2232 | -1.1577 | -2.2923 | -1.6586 | -1.1711 | -2.2536 | -1.5544 | -1.156 | -0.97275 | -0.024689 | 8.3247 | -3712.5589
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.096417 | 4.0487 | 0.49561 | 0.27794 | 0.30279 | 0.38985 | 0.21203 | -0.054601 | -0.23287 | 0.02383 | -0.25206 | -0.15986 | 0.0071981 | 0.069119 | -0.042038 | -0.32933 | -0.28342 | -0.21634 | -0.13068 | -0.077163 | -1.1141 |  | 0
db2 | -0.039902 | 0.64129 | -0.89968 | -0.24677 | -0.21674 | 0.008539 | -0.46554 | -0.016296 | -0.15325 | 0.37322 | 0.57873 | -0.57453 | 0.4959 | -0.84052 | 0.20322 | 0.47469 | -0.03926 | 0.17042 | 0.44831 | 0.69956 | 1.0467 |  | 0
db3 | -1.513 | 44.4326 | 1.023 | 0.47473 | 0.31735 | 0.23786 | -1.1776 | 0.64882 | 0.30005 | 0.58732 | -0.49632 | -1.3098 | -0.35511 | 0.030364 | -0.21109 | -1.1219 | -1.5758 | -0.37373 | 0.72167 | 2.2802 | 1.5534 |  | 0
db4 | 0.15504 | 0.27107 | -0.10403 | 0.0051605 | 0.11928 | 0.21098 | 0.014701 | 0.31614 | 0.31196 | 0.22528 | 0.44815 | -0.48118 | -0.21891 | -0.25912 | -0.14661 | -0.32106 | -0.035999 | 0.087764 | -0.013187 | -0.15931 | -0.39233 |  | 0
db5 | 0.83332 | -1.9365 | -0.094082 | 0.055739 | 0.1187 | 0.0439 | 0.079379 | 0.0073876 | 0.21572 | 0.16401 | 0.32725 | -0.020489 | 0.13524 | 0.0011089 | -0.001494 | -0.25573 | 0.027909 | -0.092503 | -0.31978 | -0.39227 | -3.0884 |  | 0
db6 | 0.021721 | -0.048469 | 0.1169 | 0.061881 | 0.013883 | 0.0094874 | 0.0017407 | 0.1013 | 0.10114 | 0.058599 | 0.025526 | 0.0073726 | -0.055992 | -0.029786 | -0.12138 | -0.15092 | 0.066823 | -0.02159 | -0.0379 | -0.14708 | -0.070689 |  | 0
db7 | 1.1979 | 2.9249 | 0.058004 | -0.010276 | -0.029548 | -0.015236 | -0.041805 | 0.026218 | -0.0011906 | 0.027733 | 0.022534 | -0.0045294 | -0.011284 | 0.00045702 | 0.016978 | -0.074209 | -0.037484 | -0.0036599 | 0.030983 | 0.046315 | 0.90732 |  | 3.7007e-17
db8 | 0.0078262 | -0.055321 | -0.020721 | -0.031187 | -0.0053688 | -0.011807 | -0.018353 | -0.023836 | -0.0036957 | 0.014114 | -0.0024356 | -0.0027912 | 0.026033 | -0.018329 | 0.025725 | 0.0079674 | -0.015883 | 0.047137 | 0.015206 | 0.018225 | 0.046378 |  | 0
db9 | 0.58183 | 0.56963 | -0.00024798 | -0.0099231 | 0.0053649 | 0.029519 | 0.011955 | -0.0045171 | -0.015929 | -0.0019881 | 0.0037536 | 0.004546 | 0.0050626 | 0.0034609 | 0.0070566 | 0.0089479 | -0.022361 | -0.009329 | -0.0098378 | -0.0055321 | 0.57407 |  | 0
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.75641 | 0.68114 | -0.0024792 | 0.0010964 | 0.009512 | 0.023659 | 0.015215 | -0.013603 | -0.015611 | -0.0043963 | -0.0044304 | 0.0013926 | -0.0020028 | -0.0011614 | -0.006223 | -5.9808e-05 | 0.0016868 | 0.0037261 | -0.0036692 | -0.002652 | 0.66221 |  | 3.084e-17
db12 | 0.00033699 | -0.010588 | -0.0043628 | -0.0012102 | 0.0029973 | -0.0051182 | 0.0028592 | -0.0073214 | -0.0044468 | -0.0030638 | -0.0031915 | -0.00063294 | 0.0026239 | 9.8012e-05 | 0.0042728 | 0.0055317 | 0.0017928 | 0.0047345 | 0.00064025 | 0.0037972 | 0.0043357 |  | 0
db13 | 0.056831 | 0.091892 | -0.0025066 | -0.0020989 | -0.00060038 | -0.00049366 | -0.0050118 | 0.0017914 | -9.0038e-06 | 0.00087261 | -0.0031774 | 0.00093376 | -0.0027503 | -0.00059901 | -0.00090998 | -0.0019965 | -0.00077462 | 0.0048776 | 0.0049165 | 0.0075362 | 0.059888 |  | 0
db14 | 0.012632 | 0.0043807 | 0.00059041 | -0.00041055 | -0.00037484 | -0.0018992 | 0.00023025 | -4.5562e-05 | -0.00077782 | -0.00035632 | -0.0016103 | 0.00075502 | -0.0010797 | -0.00040022 | -9.6576e-05 | 0.0011787 | 0.00076322 | 0.0021047 | 0.0013768 | 5.2062e-05 | -0.0060995 |  | -7.7099e-19
db15 | 0.039102 | 0.0020461 | -0.0084406 | 0.0039321 | 0.0070951 | 0.0024679 | 0.0068434 | 0.00045485 | 0.002277 | 0.00036394 | -0.0071288 | 0.0010531 | -0.0052932 | -0.000881 | -0.0058408 | -0.004224 | 0.0016425 | 0.0050784 | -0.00085567 | 0.001456 | 0.027078 |  | -2.6985e-18
da2 | 0.13307 | -0.55351 | 0.74355 | 0.29021 | -0.74585 | -1.3014 | 0.55266 | 0.24852 | 0.23796 | 0.35632 | -0.077047 | -0.44908 | 1.1528 | 0.31136 | 0.19369 | -0.2104 | -0.28537 | 0.3287 | -0.45139 | -0.89523 | 3.3302 |  | 0
da3 | -0.46065 | -2.6729 | -0.32379 | -0.20914 | 0.14689 | -0.056813 | 0.23917 | -0.12074 | -0.24714 | 0.50691 | 0.16508 | -0.22864 | 0.20313 | -0.17725 | 0.047036 | 0.20625 | 0.0046617 | 0.25526 | -0.1898 | -0.22108 | 0.18704 |  | 0
da4 | -0.14148 | -0.34542 | -0.32294 | -0.080805 | -0.061191 | 0.088171 | -0.26951 | 0.34843 | 0.42002 | 0.27789 | -0.26649 | -0.34123 | 0.124 | 0.034637 | 0.15059 | -0.10916 | -0.052677 | -0.25131 | 0.088562 | 0.223 | 0.53926 |  | 0
da5 | -0.14331 | 3.1968 | -0.094671 | 0.19171 | 0.20234 | 0.19605 | -0.17209 | -0.028971 | 0.10488 | 0.057941 | 0.25933 | -0.35582 | -0.23796 | -0.11621 | 0.0051154 | 0.21194 | -0.040486 | -0.027835 | -0.061845 | -0.093435 | -0.53319 |  | 0
da6 | 0.0575 | -0.20401 | -0.041084 | -0.084012 | -0.096108 | 0.046148 | 0.1614 | -0.095539 | -0.051331 | 0.046331 | 0.010792 | -0.12558 | 0.033783 | -0.02778 | -0.018757 | -0.063643 | 0.11729 | 0.10624 | 0.056576 | 0.025276 | -0.12204 |  | 0
da7 | -0.049742 | 1.6112 | 0.076847 | 0.13247 | 0.093787 | 0.060809 | -0.15598 | -0.17796 | -0.17965 | -0.10125 | -0.072698 | 0.12891 | 0.068636 | 0.079198 | 0.065562 | 0.025471 | 0.06605 | -0.041091 | -0.0436 | -0.025499 | -0.060574 |  | 0
da8 | -0.0032441 | 0.041893 | -0.0028159 | -0.012501 | -0.010272 | 0.05853 | 0.069289 | 0.0261 | -0.0039274 | -0.0099692 | 0.01471 | -0.026726 | -0.01214 | -0.032137 | -0.022502 | -0.025596 | -0.012033 | -0.015107 | -0.0088747 | 0.025971 | -0.059146 |  | 0
da9 | -0.047166 | -0.17499 | 0.019414 | 0.012874 | 0.034386 | 0.027846 | -0.023682 | -0.022172 | -0.01475 | -0.021481 | -0.03598 | 0.0042247 | 0.0032867 | -0.011051 | -0.012052 | -0.0049497 | 0.01836 | 0.035224 | -0.0084139 | -0.001085 | -0.036712 |  | 2.8912e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.049114 | 0.13276 | -0.0014402 | 0.0092593 | 0.012551 | 0.0044829 | -0.0027328 | -0.005703 | -0.0081498 | -0.0015908 | -0.0023217 | -0.002504 | 0.00073779 | 0.0072721 | 0.0065718 | 0.0021983 | -0.0064733 | -0.0086734 | -0.0026058 | -0.00087844 | -0.025742 |  | -1.1565e-18
da12 | 0.00018211 | -0.005764 | -0.00053915 | -9.1157e-05 | -0.0048092 | -0.0075051 | -0.0015976 | 0.00014468 | 0.00016062 | 0.0022805 | -0.0031411 | -0.0029888 | 0.0020395 | 0.0021393 | 0.0028756 | 0.0023093 | 0.005557 | 0.0025104 | 0.00059337 | 6.1847e-05 | -0.00088852 |  | -4.2163e-20
da13 | -0.0043161 | -0.021852 | -0.0042931 | -0.0017806 | -0.003906 | -0.0025287 | 0.0025108 | 0.0016584 | 0.0031706 | 0.0019126 | 0.0038109 | -0.0014247 | -0.00072731 | -0.0010678 | 0.00018197 | -0.0010017 | 0.0034512 | 0.00098841 | -0.00063253 | -0.00032239 | -0.0071788 |  | -1.6865e-19
da14 | -0.018917 | 0.00013649 | -0.0021296 | 9.5313e-05 | -0.00045705 | 0.00084044 | -0.00017321 | 0.004076 | 0.0029379 | 0.00077039 | -0.0016652 | -0.0013272 | -0.0018816 | -9.2883e-05 | 0.0012617 | -0.00021014 | -0.0031702 | -0.0010257 | 0.00077335 | 0.0013776 | -0.0081389 |  | 0
da15 | -0.018677 | -0.00056456 | -0.0075491 | 0.0018249 | 0.0021257 | -0.0019331 | 0.0033086 | 0.0065425 | 0.0028459 | -0.0021218 | -0.0066651 | -0.0007268 | -0.0078349 | 0.00091839 | 6.616e-05 | 0.0035173 | 0.0041033 | 0.0022727 | -0.003756 | 0.0030614 | -0.026256 |  | 0
Field Colinearity | 0.23889
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.2978 | 0.31212 | 0.13037 | 0.12748 | 0.34447 | -0.16328 | -0.31382 | -0.26341 | -0.21533 | -0.34783 | -0.42761 | -0.11935 | -0.32497 | -0.24155 | -0.27392 | -0.199 | -0.28216 | -0.37366 | -0.21587 | 0.24678
Twist (m2 mrad) | 32.2334 | 4.4477 | 1.76 | 1.6253 | 4.1336 | -1.8369 | -3.2951 | -2.5682 | -1.938 | -2.8696 | -3.2071 | -0.80561 | -1.9498 | -1.2681 | -1.2326 | -0.74626 | -0.84649 | -0.84074 | -0.32381 | 0.092543
Twist Integral (I0) (m2 rad) | 0.015423
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.9538 | 0.3977 | 0.18183 | 0.20609 | 0.2934 | 0.11628 | -0.15069 | -0.32914 | -0.072847 | -0.34844 | -0.2561 | -0.090639 | -0.027247 | -0.13837 | -0.42545 | -0.37988 | -0.31276 | -0.22708 | -0.17397 | -1.2132
Twist (m2 mrad) | 29.6538 | 5.6673 | 2.4547 | 2.6276 | 3.5208 | 1.3082 | -1.5823 | -3.2091 | -0.65562 | -2.8746 | -1.9207 | -0.61181 | -0.16348 | -0.72645 | -1.9145 | -1.4246 | -0.93829 | -0.51092 | -0.26095 | -0.45494
Twist Integral (I0) (m2 rad) | 0.020988
Note on B160: 3.6s
Note on B203: 3.6s
Note on B204: 4.9s
Note on B205: 4.2s
Note on V219: 3.8s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.48962 | -2.0674 | 0.079939 | 0.93021 | 1.165 | -1.5818 | -0.47004 | -0.29563 | -0.03466 | 22.3791 | 0.5946 | 0.039902 | -1.5131 | -0.15504 | 0.83208 | 1.1971 | -0.96417 | 0.13307 | -0.46065 | -0.14148 | 20.2556 | 0.59496 | -0.50133 | -0.41669 | 0.12505 | 0.70644 | 1.2127 | 0.0041496 | -0.61807 | -0.31934 | -0.14259 | 14429.462 | 0.092443 | 0.094714 | -0.14881 | 0.61326 | 1.2443 | 0.0012726 | 0.21488 | -0.51213 | -0.12598 | 14429.441
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8