## Alarm sheet
File name | HCMB__A001-03000004_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000004
Date of test Ap 1 | 2003-09-23 |  |  |  | Date of test Ap 2 | 0
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length
| Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field
Angle |  | status ok | status ok | status ok | Angle
b2 | status ok | status ok | status ok | status ok | b2
b3 | status ok | status ok | status ok | status ok | b3
b4 | status ok | status ok | status ok | status ok | b4
b5 | status ok | status ok | status ok | status ok | b5
b6 | status ok | status ok | status ok | status ok | b6
b7 | status ok | status ok | status ok | status ok | b7
b8 | status ok | status ok | status ok | status ok | b8
b9 | status ok | status ok | status ok | status ok | b9
b10 | status ok | status ok | status ok | status ok | b10
b11 | status ok | status ok | status ok | status ok | b11
b12 | status ok | status ok | status ok | status ok | b12
b13 | status ok | status ok | status ok | status ok | b13
b14 | status ok | status ok | status ok | status ok | b14
b15 | status ok | status ok | status ok | yellow alarm | b15
a2 | status ok | status ok | status ok | status ok | a2
a3 | status ok | status ok | status ok | status ok | a3
a4 | status ok | status ok | status ok | status ok | a4
a5 | status ok | status ok | status ok | status ok | a5
a6 | status ok | status ok | status ok | status ok | a6
a7 | status ok | status ok | status ok | status ok | a7
a8 | status ok | status ok | status ok | status ok | a8
a9 | status ok | status ok | status ok | status ok | a9
a10 | status ok | status ok | status ok | status ok | a10
a11 | status ok | status ok | status ok | status ok | a11
a12 | status ok | status ok | status ok | status ok | a12
a13 | status ok | status ok | status ok | status ok | a13
a14 | status ok | status ok | status ok | status ok | a14
a15 | status ok | status ok | status ok | status ok | a15
Coil Positioning | status ok |  |  |  | Coil Positioning
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
13 | HCMB__A001 | 3000004 |  | 1 |  |  | 3
| COMMENTS: | New measurement of Ap. 1 for recuperation of the coils.
| HOLDING POINT | NA | Data at CERN |  | 2003-09-24 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-09-24 | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| #DIV/0! | 0.2 | #DIV/0! | #DIV/0! | 0.8 | #DIV/0! |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.95 | 0.15 | 0.15 | 0.95 | Up
| Pre-stress (estimate, Mpa) | [6 empty] | 0.95 | 0.15 | 0.15 | 0.95 | Down
| Inner average |  |  | Outer average |  |  |  | Seen from conection side
| Used | Nominal | Differ. | Used | Nominal | Differ.
| 75 | 75 | 0 | 75 | 75 | 0
| Peak pre-stress during collaring |  |  |  |  | 110
| Collared coil template - version 11.02.02 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans Assembly Resume et Original Data
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
Remarks: Magnet went back to CC; Aperture 1 has been measured again to be re-used. Measurement has been carried out in inverse direction.
1. In the sheet "Summary Data" are the inversed values due to the change of measurement direction.
2. In the sheet "Difference" is the difference between the actual and the old measurement.
3. The OK for re-using the coils in aperture 1 has been given on this assumption.

## Original data
File name |  | HCMB__A001-03000004_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000004 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | noell0106cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | noell0106cct2 |  |  |  | a even | 1
Test Operator |  | Oswald |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-09-23 |  |  | Time of test |  |  | 18:45 - 21:30 |  |  |  | Date of test |  |  |  |  | Time of test
Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03 |  |  |  | Mole name |  |  |  |  | Calibration date
Coil length |  | 0.7499 |  |  | Coil temperature |  |  | 24 |  |  |  | Coil length |  |  |  |  | Coil temperature
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  |  |  |  | Analysis tools
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  |  |  |  | N. measure
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  |  |  |  | I. Gain
Dx average |  | 1.538e-05 |  |  | Dy average |  |  | 0.00099699 |  |  |  | Dx average |  |  |  |  | Dy average
Magnetic length |  | 14.4536 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1699 |  |  |  | Magnetic length |  |  |  |  | Tranfs. Function (mT/KA)
Phase |  | -4.2564 |  |  | Integral field (T.m) |  |  | 0.073243 |  |  |  | Phase |  |  |  |  | Integral field (T.m)
File | Aperture 1 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003254 | 0.00507 | 0.005067 | 0.005067 | 0.005067 | 0.005066 | 0.005066 | 0.005066 | 0.005066 | 0.005066 | 0.005067 | 0.005068 | 0.005068 | 0.005068 | 0.005067 | 0.005067 | 0.005067 | 0.005069 | 0.005072 | 0.003202 | 0.073243
Angle (mrad) | -0.099546 | 0.23467 | 1.3441 | 0.4451 | 0.60781 | 0.10637 | 1.1611 | 1.1878 | 0.99933 | -1.3519 | -1.1837 | -0.72536 | -0.50854 | -0.61581 | -1.4089 | -0.54443 | -0.23074 | -0.40447 | -0.88429 | 2.8449 | -4.2564
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.61228 | 0.26123 | 0.17983 | -0.39754 | -0.35476 | -0.30233 | -0.68789 | -0.85599 | -1.3076 | 0.075306 | 0.20629 | -0.18024 | -0.93733 | -0.32723 | -0.87499 | -1.0811 | -0.13843 | -0.21802 | -0.24961 | -1.6023 | -0.40513
b3 | 3.5976 | 2.9654 | 3.0348 | 3.8505 | 2.9273 | 3.0487 | 3.0865 | 3.6645 | 2.7525 | 1.7583 | 3.6542 | 4.0203 | 3.8954 | 3.4294 | 2.7906 | 2.6606 | 3.3726 | 2.998 | 2.7058 | 36.7211 | 4.2609
b4 | -0.25762 | -0.18554 | -0.19841 | -0.11498 | 0.099571 | -0.072748 | -0.043905 | 0.24054 | -0.13565 | 0.028537 | -0.077574 | -0.23096 | -0.40055 | -0.54428 | -0.57645 | -0.6337 | -0.54797 | -0.36256 | -0.4275 | 0.12756 | -0.22151
b5 | -3.1618 | 0.33904 | 0.41239 | 0.61694 | 0.66986 | 0.63276 | 0.60255 | 0.55783 | 0.63186 | 0.9867 | 0.7116 | 0.39782 | 0.23431 | 0.41603 | 0.27844 | 0.35143 | 0.46082 | 0.55283 | 0.73782 | -4.4719 | 0.24567
b6 | 0.09546 | 0.0016916 | 0.083609 | -0.03418 | 0.078989 | 0.11406 | 0.095541 | 0.035054 | 0.013583 | -0.053335 | 0.030061 | -0.028989 | 0.058664 | -0.18109 | -0.12218 | -0.062197 | -0.059627 | -0.10827 | -0.048137 | -0.090619 | -0.0094855
b7 | 0.31256 | 0.61086 | 0.62535 | 0.6209 | 0.64214 | 0.6371 | 0.61406 | 0.60297 | 0.67882 | 0.63241 | 0.69993 | 0.65525 | 0.62428 | 0.69328 | 0.63729 | 0.6736 | 0.63522 | 0.55182 | 0.50155 | 2.4423 | 0.67866
b8 | -0.027824 | -0.038811 | -0.040119 | -0.0015956 | -0.019176 | -0.0098399 | -0.036308 | 0.01898 | -0.013614 | 0.024371 | 0.018141 | -0.023688 | 0.043298 | 0.0030711 | 0.0093713 | 0.008766 | -0.0037483 | -0.00050821 | 0.018004 | 0.0069615 | -0.0029511
b9 | 0.28165 | 0.27429 | 0.27198 | 0.26428 | 0.24988 | 0.26971 | 0.27251 | 0.27132 | 0.25819 | 0.23087 | 0.2535 | 0.26864 | 0.26534 | 0.27078 | 0.27053 | 0.25441 | 0.2687 | 0.22162 | 0.26169 | 0.27708 | 0.26223
b10 | -0.11573 | -0.026064 | -0.0015148 | 0.027121 | 0.031401 | 0.01275 | -0.027044 | 0.0034461 | 0.01514 | 0.012947 | 0.080474 | -0.022028 | 0.0016925 | 0.00043813 | -0.0051953 | -0.02934 | 0.0072847 | 0.0020567 | 0.0044406 | -0.021624 | 0
b11 | 0.633 | 0.72558 | 0.72248 | 0.7215 | 0.72219 | 0.7194 | 0.71626 | 0.70998 | 0.71011 | 0.70459 | 0.70241 | 0.7114 | 0.70904 | 0.7064 | 0.7128 | 0.71686 | 0.71567 | 0.71495 | 0.72852 | 0.62908 | 0.70946
b12 | -0.015958 | -0.0054906 | 0.0006949 | 0.00059032 | 0.0003058 | -0.0019419 | -0.009814 | -0.0086527 | 0.00069365 | 0.0096657 | 0.0099905 | -0.0072405 | -0.0098966 | -0.0081686 | -0.0066007 | -0.0075736 | 0.0019961 | 0.0034753 | 0.0041688 | -0.0029853 | -0.0023828
b13 | 0.065118 | 0.087562 | 0.086072 | 0.083528 | 0.080966 | 0.081022 | 0.082267 | 0.080771 | 0.082173 | 0.076688 | 0.080182 | 0.077397 | 0.079592 | 0.079577 | 0.076612 | 0.078832 | 0.079671 | 0.08186 | 0.079155 | 0.11474 | 0.081366
b14 | -0.010908 | -0.0047432 | -0.0026854 | -0.00086858 | -0.0016031 | -0.0008232 | -0.0043946 | -0.004072 | -0.0016382 | -0.0042281 | 0.0050723 | -0.0071882 | -0.0075832 | -0.0072816 | -0.0058302 | -0.0056655 | -0.0030762 | -0.0037608 | -0.0023491 | 0.0040402 | -0.0034844
b15 | 0.03793 | 0.051854 | 0.053627 | 0.0522 | 0.054522 | 0.055141 | 0.054531 | 0.054242 | 0.052675 | 0.054937 | 0.049258 | 0.0523 | 0.050849 | 0.054205 | 0.052909 | 0.049942 | 0.051018 | 0.052691 | 0.053526 | 0.010663 | 0.050925
b16
b17
a1 | -0.99546 | 2.3467 | 13.4408 | 4.451 | 6.0782 | 1.0637 | 11.6106 | 11.8779 | 9.9933 | -13.519 | -11.8367 | -7.2536 | -5.0854 | -6.1581 | -14.0892 | -5.4443 | -2.3074 | -4.0447 | -8.8429 | 28.4486 | 0
a2 | -6.5762 | -1.5064 | -1.4897 | -1.4237 | -1.9905 | -0.31357 | 0.41306 | 0.44496 | -0.22845 | 0.79618 | -0.82024 | -2.3318 | -2.3739 | -1.6626 | -1.7053 | -1.0065 | -0.75543 | -0.85302 | 0.31454 | 11.1397 | -0.70971
a3 | 0.46842 | 0.77671 | 0.62109 | 0.34609 | 0.6595 | 0.92563 | 0.97094 | 0.88546 | 1.0231 | 1.5279 | 0.58403 | 0.26049 | 0.72446 | 1.021 | 1.0542 | 1.3658 | 0.44859 | 0.66998 | 0.00018324 | -5.0625 | 0.569
a4 | -0.2529 | -0.37019 | -0.18838 | -0.44573 | -0.033498 | 0.34425 | 0.5638 | 0.63481 | 0.64857 | -0.30641 | -0.39344 | -0.3711 | -0.58807 | 0.029746 | -0.23182 | -0.50349 | -0.32857 | -0.22012 | -0.049633 | 2.4739 | -0.021247
a5 | -0.2789 | 0.15646 | 0.077778 | 0.0052918 | 0.01524 | 0.051088 | 0.087831 | 0.17934 | 0.33301 | 0.47421 | 0.021576 | 0.034328 | 0.15512 | 0.10389 | 0.21578 | 0.26622 | 0.10661 | -0.056529 | -0.19339 | 1.6833 | 0.15138
a6 | -0.029962 | -0.031727 | 0.017146 | 0.0078036 | 0.022369 | 0.13755 | 0.13391 | 0.215 | 0.15096 | 0.2023 | -0.057159 | -0.098367 | -0.024142 | 0.057803 | 0.0097351 | 0.0023651 | 0.054715 | 0.068225 | 0.096366 | 0.63356 | 0.069821
a7 | 0.045163 | 0.02153 | 0.0038596 | 0.039157 | 0.083091 | 0.086342 | -0.015934 | -0.018329 | -0.030109 | -0.083235 | -0.034897 | -0.044856 | -0.020625 | 0.040244 | 0.074639 | 0.057862 | 0.010841 | 0.024334 | -0.042224 | 1.4431 | 0.056686
a8 | -0.0016172 | 0.016525 | 0.010722 | 0.010652 | 0.012819 | 0.019743 | 0.0049552 | 0.054298 | -0.04023 | -0.02121 | -0.0064897 | -0.017605 | 0.029467 | -0.027078 | -0.033676 | 0.016029 | 0.0037736 | -0.050325 | 0.0072637 | 0.14014 | 0.0040018
a9 | -0.023181 | -0.00016388 | 0.018582 | 0.0075641 | 0.025404 | 0.022369 | 0.011696 | 0.021545 | 0.010114 | 0.030158 | 0.0082048 | -0.013447 | -0.025497 | 0.011282 | -0.011335 | 0.0082944 | -0.0047283 | 0.0047737 | 0.01592 | -0.32055 | -0.0039737
a10 | -0.05875 | 0.0031417 | -0.0012503 | 0.011246 | -0.0043156 | -0.0089808 | 0.0024792 | 0.0047792 | 0.0051859 | 0.02178 | 0.019659 | 0.0012964 | 0.017971 | -0.018222 | -0.017508 | -0.017452 | -0.021758 | -0.0067323 | 0.00537 | 0.064962 | 0
a11 | -0.013667 | -0.022173 | -0.010064 | -0.021014 | -0.017967 | -0.019583 | -0.013795 | -0.013169 | -0.017447 | -0.039736 | -0.041576 | -0.040672 | -0.040498 | -0.03657 | -0.039973 | -0.035752 | -0.037422 | -0.044052 | -0.050137 | 0.17576 | -0.02278
a12 | -0.01363 | -0.0070649 | -0.005581 | 1.1112e-05 | -0.0096127 | -0.008603 | -0.0066662 | -0.0045138 | -0.00047258 | 0.0024482 | -0.0073747 | -0.012056 | -0.0071499 | -0.0099509 | -0.011583 | -0.0083723 | -0.0067926 | -0.0032333 | -0.0060816 | 0.011093 | -0.0059354
a13 | 0.0014697 | 0.0019016 | -0.0012665 | -0.00047472 | -0.0011517 | -0.0034106 | -0.00048429 | -0.00011371 | -0.0027047 | 0.00078091 | -0.0073822 | -0.0052062 | -0.0027404 | -0.0032847 | -0.0023501 | -0.0023479 | -0.0057795 | -0.0042612 | -0.0019786 | -0.021038 | -0.0028307
a14 | -0.016181 | -0.027939 | -0.028396 | -0.026968 | -0.029953 | -0.031223 | -0.030046 | -0.030266 | -0.028869 | -0.033861 | -0.025697 | -0.030115 | -0.028052 | -0.029959 | -0.029983 | -0.026679 | -0.026164 | -0.027322 | -0.028402 | 0.0026422 | -0.027426
a15 | -0.016446 | -0.0070944 | -0.0047724 | -0.00014776 | 0.00095496 | -0.0021373 | -0.0070212 | -0.0032514 | -0.0012634 | -0.0010926 | 0.0039314 | -0.0097727 | -0.006506 | -0.0067516 | -0.0074799 | -0.0088734 | -0.0063192 | -0.0059557 | -0.0041376 | 0.0018468 | -0.004517
a16
a17
Dx (m) | 0.00030706 | 6.1243e-05 | 0 | -6.3076e-05 | -7.4115e-05 | -3.068e-05 | 6.4293e-05 | 0 | -3.592e-05 | -2.8172e-05 | -0.00019146 | 5.2651e-05 | 0 | 0 | 1.0009e-05 | 6.7333e-05 | -1.9945e-05 | 0 | 0 | 0 | 1.538e-05
Dy (m) | 0.00016537 | 0 | 0 | -2.8333e-05 | 0 | 2.0388e-05 | 0 | -1.1592e-05 | -1.3297e-05 | -5.4149e-05 | -5.8853e-05 | 0 | -4.3186e-05 | 4.3671e-05 | 4.2312e-05 | 4.477e-05 | 5.0638e-05 | 1.5647e-05 | -1.3182e-05 | -0.00017788 | 0.00099699
Magnetic length (m) |  | 14.4536
Transfer funct. (mT/KA) |  | 596.1699
Coil positioning (mm) |  | 0.024428
Current (A) |  | 8.5

## Summary Data
File | HCMB__A001-3000004 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 376.7059 | 596.7059 | 596.3529 | 596.1176 | 596.1176 | 596.1176 | 596.2353 | 596.2353 | 596.2353 | 596.1176 | 596 | 596 | 596 | 596 | 596 | 596.1176 | 596.1176 | 596.1176 | 596.4706 | 382.8235 | 0.073243
Angle (mrad) | 2.8449 | -0.88429 | -0.40447 | -0.23074 | -0.54443 | -1.4089 | -0.61581 | -0.50854 | -0.72536 | -1.1837 | -1.3519 | 0.99933 | 1.1878 | 1.1611 | 0.10637 | 0.60781 | 0.4451 | 1.3441 | 0.23467 | -0.099546 | -4.2564
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -1.6023 | -0.24961 | -0.21802 | -0.13843 | -1.0811 | -0.87499 | -0.32723 | -0.93733 | -0.18024 | 0.20629 | 0.075306 | -1.3076 | -0.85599 | -0.68789 | -0.30233 | -0.35476 | -0.39754 | 0.17983 | 0.26123 | 0.61228 | -0.40513
b3 | 36.7211 | 2.7058 | 2.998 | 3.3726 | 2.6606 | 2.7906 | 3.4294 | 3.8954 | 4.0203 | 3.6542 | 1.7583 | 2.7525 | 3.6645 | 3.0865 | 3.0487 | 2.9273 | 3.8505 | 3.0348 | 2.9654 | 3.5976 | 4.2609
b4 | 0.12756 | -0.4275 | -0.36256 | -0.54797 | -0.6337 | -0.57645 | -0.54428 | -0.40055 | -0.23096 | -0.077574 | 0.028537 | -0.13565 | 0.24054 | -0.043905 | -0.072748 | 0.099571 | -0.11498 | -0.19841 | -0.18554 | -0.25762 | -0.22151
b5 | -4.4719 | 0.73782 | 0.55283 | 0.46082 | 0.35143 | 0.27844 | 0.41603 | 0.23431 | 0.39782 | 0.7116 | 0.9867 | 0.63186 | 0.55783 | 0.60255 | 0.63276 | 0.66986 | 0.61694 | 0.41239 | 0.33904 | -3.1618 | 0.24567
b6 | -0.090619 | -0.048137 | -0.10827 | -0.059627 | -0.062197 | -0.12218 | -0.18109 | 0.058664 | -0.028989 | 0.030061 | -0.053335 | 0.013583 | 0.035054 | 0.095541 | 0.11406 | 0.078989 | -0.03418 | 0.083609 | 0.0016916 | 0.09546 | -0.0094855
b7 | 2.4423 | 0.50155 | 0.55182 | 0.63522 | 0.6736 | 0.63729 | 0.69328 | 0.62428 | 0.65525 | 0.69993 | 0.63241 | 0.67882 | 0.60297 | 0.61406 | 0.6371 | 0.64214 | 0.6209 | 0.62535 | 0.61086 | 0.31256 | 0.67866
b8 | 0.0069615 | 0.018004 | -0.00050821 | -0.0037483 | 0.008766 | 0.0093713 | 0.0030711 | 0.043298 | -0.023688 | 0.018141 | 0.024371 | -0.013614 | 0.01898 | -0.036308 | -0.0098399 | -0.019176 | -0.0015956 | -0.040119 | -0.038811 | -0.027824 | -0.0029511
b9 | 0.27708 | 0.26169 | 0.22162 | 0.2687 | 0.25441 | 0.27053 | 0.27078 | 0.26534 | 0.26864 | 0.2535 | 0.23087 | 0.25819 | 0.27132 | 0.27251 | 0.26971 | 0.24988 | 0.26428 | 0.27198 | 0.27429 | 0.28165 | 0.26223
b10 | -0.021624 | 0.0044406 | 0.0020567 | 0.0072847 | -0.02934 | -0.0051953 | 0.00043813 | 0.0016925 | -0.022028 | 0.080474 | 0.012947 | 0.01514 | 0.0034461 | -0.027044 | 0.01275 | 0.031401 | 0.027121 | -0.0015148 | -0.026064 | -0.11573 | 0
b11 | 0.62908 | 0.72852 | 0.71495 | 0.71567 | 0.71686 | 0.7128 | 0.7064 | 0.70904 | 0.7114 | 0.70241 | 0.70459 | 0.71011 | 0.70998 | 0.71626 | 0.7194 | 0.72219 | 0.7215 | 0.72248 | 0.72558 | 0.633 | 0.70946
b12 | -0.0029853 | 0.0041688 | 0.0034753 | 0.0019961 | -0.0075736 | -0.0066007 | -0.0081686 | -0.0098966 | -0.0072405 | 0.0099905 | 0.0096657 | 0.00069365 | -0.0086527 | -0.009814 | -0.0019419 | 0.0003058 | 0.00059032 | 0.0006949 | -0.0054906 | -0.015958 | -0.0023828
b13 | 0.11474 | 0.079155 | 0.08186 | 0.079671 | 0.078832 | 0.076612 | 0.079577 | 0.079592 | 0.077397 | 0.080182 | 0.076688 | 0.082173 | 0.080771 | 0.082267 | 0.081022 | 0.080966 | 0.083528 | 0.086072 | 0.087562 | 0.065118 | 0.081366
b14 | 0.0040402 | -0.0023491 | -0.0037608 | -0.0030762 | -0.0056655 | -0.0058302 | -0.0072816 | -0.0075832 | -0.0071882 | 0.0050723 | -0.0042281 | -0.0016382 | -0.004072 | -0.0043946 | -0.0008232 | -0.0016031 | -0.00086858 | -0.0026854 | -0.0047432 | -0.010908 | -0.0034844
b15 | 0.010663 | 0.053526 | 0.052691 | 0.051018 | 0.049942 | 0.052909 | 0.054205 | 0.050849 | 0.0523 | 0.049258 | 0.054937 | 0.052675 | 0.054242 | 0.054531 | 0.055141 | 0.054522 | 0.0522 | 0.053627 | 0.051854 | 0.03793 | 0.050925
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 28.4486 | -8.8429 | -4.0447 | -2.3074 | -5.4443 | -14.0892 | -6.1581 | -5.0854 | -7.2536 | -11.8367 | -13.519 | 9.9933 | 11.8779 | 11.6106 | 1.0637 | 6.0782 | 4.451 | 13.4408 | 2.3467 | -0.99546 | 0
a2 | 11.1397 | 0.31454 | -0.85302 | -0.75543 | -1.0065 | -1.7053 | -1.6626 | -2.3739 | -2.3318 | -0.82024 | 0.79618 | -0.22845 | 0.44496 | 0.41306 | -0.31357 | -1.9905 | -1.4237 | -1.4897 | -1.5064 | -6.5762 | -0.70971
a3 | -5.0625 | 0.00018324 | 0.66998 | 0.44859 | 1.3658 | 1.0542 | 1.021 | 0.72446 | 0.26049 | 0.58403 | 1.5279 | 1.0231 | 0.88546 | 0.97094 | 0.92563 | 0.6595 | 0.34609 | 0.62109 | 0.77671 | 0.46842 | 0.569
a4 | 2.4739 | -0.049633 | -0.22012 | -0.32857 | -0.50349 | -0.23182 | 0.029746 | -0.58807 | -0.3711 | -0.39344 | -0.30641 | 0.64857 | 0.63481 | 0.5638 | 0.34425 | -0.033498 | -0.44573 | -0.18838 | -0.37019 | -0.2529 | -0.021247
a5 | 1.6833 | -0.19339 | -0.056529 | 0.10661 | 0.26622 | 0.21578 | 0.10389 | 0.15512 | 0.034328 | 0.021576 | 0.47421 | 0.33301 | 0.17934 | 0.087831 | 0.051088 | 0.01524 | 0.0052918 | 0.077778 | 0.15646 | -0.2789 | 0.15138
a6 | 0.63356 | 0.096366 | 0.068225 | 0.054715 | 0.0023651 | 0.0097351 | 0.057803 | -0.024142 | -0.098367 | -0.057159 | 0.2023 | 0.15096 | 0.215 | 0.13391 | 0.13755 | 0.022369 | 0.0078036 | 0.017146 | -0.031727 | -0.029962 | 0.069821
a7 | 1.4431 | -0.042224 | 0.024334 | 0.010841 | 0.057862 | 0.074639 | 0.040244 | -0.020625 | -0.044856 | -0.034897 | -0.083235 | -0.030109 | -0.018329 | -0.015934 | 0.086342 | 0.083091 | 0.039157 | 0.0038596 | 0.02153 | 0.045163 | 0.056686
a8 | 0.14014 | 0.0072637 | -0.050325 | 0.0037736 | 0.016029 | -0.033676 | -0.027078 | 0.029467 | -0.017605 | -0.0064897 | -0.02121 | -0.04023 | 0.054298 | 0.0049552 | 0.019743 | 0.012819 | 0.010652 | 0.010722 | 0.016525 | -0.0016172 | 0.0040018
a9 | -0.32055 | 0.01592 | 0.0047737 | -0.0047283 | 0.0082944 | -0.011335 | 0.011282 | -0.025497 | -0.013447 | 0.0082048 | 0.030158 | 0.010114 | 0.021545 | 0.011696 | 0.022369 | 0.025404 | 0.0075641 | 0.018582 | -0.00016388 | -0.023181 | -0.0039737
a10 | 0.064962 | 0.00537 | -0.0067323 | -0.021758 | -0.017452 | -0.017508 | -0.018222 | 0.017971 | 0.0012964 | 0.019659 | 0.02178 | 0.0051859 | 0.0047792 | 0.0024792 | -0.0089808 | -0.0043156 | 0.011246 | -0.0012503 | 0.0031417 | -0.05875 | 0
a11 | 0.17576 | -0.050137 | -0.044052 | -0.037422 | -0.035752 | -0.039973 | -0.03657 | -0.040498 | -0.040672 | -0.041576 | -0.039736 | -0.017447 | -0.013169 | -0.013795 | -0.019583 | -0.017967 | -0.021014 | -0.010064 | -0.022173 | -0.013667 | -0.02278
a12 | 0.011093 | -0.0060816 | -0.0032333 | -0.0067926 | -0.0083723 | -0.011583 | -0.0099509 | -0.0071499 | -0.012056 | -0.0073747 | 0.0024482 | -0.00047258 | -0.0045138 | -0.0066662 | -0.008603 | -0.0096127 | 1.1112e-05 | -0.005581 | -0.0070649 | -0.01363 | -0.0059354
a13 | -0.021038 | -0.0019786 | -0.0042612 | -0.0057795 | -0.0023479 | -0.0023501 | -0.0032847 | -0.0027404 | -0.0052062 | -0.0073822 | 0.00078091 | -0.0027047 | -0.00011371 | -0.00048429 | -0.0034106 | -0.0011517 | -0.00047472 | -0.0012665 | 0.0019016 | 0.0014697 | -0.0028307
a14 | 0.0026422 | -0.028402 | -0.027322 | -0.026164 | -0.026679 | -0.029983 | -0.029959 | -0.028052 | -0.030115 | -0.025697 | -0.033861 | -0.028869 | -0.030266 | -0.030046 | -0.031223 | -0.029953 | -0.026968 | -0.028396 | -0.027939 | -0.016181 | -0.027426
a15 | 0.0018468 | -0.0041376 | -0.0059557 | -0.0063192 | -0.0088734 | -0.0074799 | -0.0067516 | -0.006506 | -0.0097727 | 0.0039314 | -0.0010926 | -0.0012634 | -0.0032514 | -0.0070212 | -0.0021373 | 0.00095496 | -0.00014776 | -0.0047724 | -0.0070944 | -0.016446 | -0.004517
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0 | 0 | 0 | 0.019945 | -0.067333 | -0.010009 | 0 | 0 | -0.052651 | 0.19146 | 0.028172 | 0.03592 | 0 | -0.064293 | 0.03068 | 0.074115 | 0.063076 | 0 | -0.061243 | -0.30706
Dy (mm) | 0.17788 | 0.013182 | -0.015647 | -0.050638 | -0.04477 | -0.042312 | -0.043671 | 0.043186 | 0 | 0.058853 | 0.054149 | 0.013297 | 0.011592 | 0 | -0.020388 | 0 | 0.028334 | 0 | 0 | -0.16537
Mag. len. (m) | 14.4536
TF (mT/KA) | 596.1699
Coil pos. (mm) | 0.024428
Current (A) | 8.5

## Difference
File | HCMB__A001-3000004 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 14.1176 | -0.11765 | -0.11765 | -0.11765 | 0 | -0.11765 | -0.11765 | -0.11765 | 0 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -0.11765 | -8.7059 | 2.4e-05
Angle (mrad) | -1.2218 | -0.12217 | 0.22696 | 0.006999 | -0.10503 | -0.8055 | -0.4114 | -0.83139 | -1.0944 | -0.20242 | -0.37771 | 1.5142 | 1.0735 | 1.1971 | 0.21243 | 0.44325 | -0.79901 | -0.035997 | -0.023235 | 1.1251 | -7.6864
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b2 | 1.4929 | -0.10486 | -0.10417 | -0.14625 | -0.079264 | -0.086163 | -0.098036 | -0.1155 | -0.081096 | -0.064969 | -0.20339 | -0.071678 | -0.10725 | -0.11768 | -0.028042 | -0.05998 | -0.091112 | -0.078188 | -0.064986 | -0.079876 | -0.046206
b3 | 2.8394 | 0.45306 | 0.50458 | 0.37365 | 0.41381 | 0.40411 | 0.50292 | 0.39188 | 0.36038 | 0.34301 | 0.55824 | 0.45215 | 0.45654 | 0.41175 | 0.42707 | 0.4176 | 0.38847 | 0.32583 | 0.33983 | 0.30116 | 0.53087
b4 | 0.23981 | -0.073373 | -0.078726 | -0.056258 | -0.051095 | -0.046095 | -0.033586 | -0.063499 | -0.070625 | -0.042326 | -0.098229 | -0.034589 | -0.049691 | -0.05552 | -0.028826 | -0.050364 | -0.070172 | -0.057157 | -0.057568 | -0.0021901 | -0.044876
b5 | 0.84263 | -0.11768 | -0.11073 | -0.10211 | -0.0912 | -0.089985 | -0.093323 | -0.078927 | -0.080298 | -0.091484 | -0.10484 | -0.10864 | -0.097143 | -0.10109 | -0.10634 | -0.098296 | -0.085284 | -0.093039 | -0.07591 | -0.14598 | -0.071244
b6 | -0.24703 | 0.0022635 | -0.0075293 | 0.013237 | 0.0066961 | 0.0053389 | -0.00469 | 0.0053881 | -0.015516 | 0.032799 | 0.032573 | -0.0087812 | 0.0069848 | -0.015562 | 0.023206 | -0.00080542 | 0.0021141 | 0.001136 | 0.0059293 | 0.0071523 | -0.0033325
b7 | 0.095988 | 0.011748 | 0.012002 | 0.0076837 | 0.0022043 | 0.012334 | -0.0060306 | 0.0025893 | 0.0004575 | 0.0038063 | 0.0019725 | 0.0047533 | 0.0014782 | 0.0061488 | 0.0076862 | 0.007422 | 0.0078576 | -0.0033179 | 0.007099 | -0.002755 | 0.0099379
b8 | -0.020273 | -0.0015561 | -0.01807 | 0.0013709 | 0.003266 | -0.0069054 | -0.015787 | 0.0020762 | -0.0076173 | 0.010034 | -0.0038524 | -0.0086783 | -0.0010153 | -0.0081865 | 0.0083693 | 0.00044924 | 0.00016055 | -0.0037683 | 0.0034277 | 0.0014704 | -0.0029618
b9 | 0.0094647 | -0.012065 | -0.010148 | -0.010103 | -0.0091186 | -0.014467 | -0.011523 | -0.011361 | -0.012448 | -0.0077942 | -0.0093714 | -0.0099023 | -0.010831 | -0.01038 | -0.008633 | -0.010625 | -0.011357 | -0.011506 | -0.012472 | -0.01419 | -0.010257
b10 | -0.08675 | -0.0013135 | -0.034222 | 0.013058 | 0.014477 | 0.0027734 | -0.032286 | 0.016936 | -0.0071652 | 0.05933 | 0.010106 | -0.020244 | -0.010847 | -0.026804 | 0.031277 | 0.00051652 | 0.0038782 | -0.0010532 | 0.022737 | 0.01591 | 0
b11 | 0.012985 | -0.022979 | -0.022402 | -0.020637 | -0.01967 | -0.021595 | -0.020674 | -0.019778 | -0.019784 | -0.02107 | -0.020191 | -0.019215 | -0.017146 | -0.020424 | -0.020184 | -0.019992 | -0.01994 | -0.019452 | -0.019601 | -0.028526 | -0.019537
b12 | -0.015064 | -0.0027169 | -0.0049908 | 0.0005791 | -0.00071534 | -0.0026863 | -0.0069751 | -0.00029629 | -0.0049292 | 0.0065845 | 0.0011063 | -0.0047826 | -0.0033195 | -0.0045235 | 0.0037859 | -0.0021982 | -0.00085535 | -0.0013113 | 0.0017229 | 0.00071203 | -0.0018185
b13 | 0.0013568 | -0.0037963 | -0.0049838 | -0.0034322 | -0.0047571 | -0.0045294 | -0.0048944 | -0.0038063 | -0.0055794 | -0.0042027 | -0.0048679 | -0.0056452 | -0.00494 | -0.0053588 | -0.0054843 | -0.0051712 | -0.0047067 | -0.0045686 | -0.0042168 | -0.0069466 | -0.0045498
b14 | 0.0033216 | -0.0098285 | -0.012186 | -0.0084354 | -0.0085468 | -0.0096657 | -0.01286 | -0.0095075 | -0.012658 | -0.0039713 | -0.0086561 | -0.010108 | -0.010013 | -0.011179 | -0.0070636 | -0.0075597 | -0.008333 | -0.0097426 | -0.0080695 | -0.0094134 | -0.0089408
b15 | 0.017202 | 0.020508 | 0.023115 | 0.021291 | 0.019822 | 0.021331 | 0.022639 | 0.018309 | 0.021447 | 0.020582 | 0.021282 | 0.022349 | 0.022634 | 0.023676 | 0.022838 | 0.025045 | 0.021266 | 0.018759 | 0.020515 | 0.01985 | 0.021289
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | -12.2185 | -1.2217 | 2.2696 | 0.06999 | -1.0503 | -8.055 | -4.114 | -8.3139 | -10.9439 | -2.0242 | -3.7771 | 15.1422 | 10.735 | 11.9708 | 2.1242 | 4.4325 | -7.9901 | -0.35997 | -0.23235 | 11.2512 | 0
a2 | 10.5929 | -0.026855 | -0.017919 | -0.03553 | 0.007811 | -0.10866 | -0.013515 | 0.067441 | 0.021331 | 0.017287 | -0.013586 | 0.02492 | 0.0054414 | -0.060126 | -0.03547 | -0.032444 | -0.035152 | 0.080241 | 0.064288 | -0.59129 | 0.32846
a3 | -1.858 | -0.02714 | -0.02859 | -0.016558 | -0.056142 | -0.052727 | -0.1348 | -0.073428 | -0.064183 | -0.083794 | -0.053093 | -0.079642 | -0.060181 | -0.059324 | -0.076463 | -0.094133 | -0.085498 | -0.045617 | -0.011893 | -0.033009 | -0.12383
a4 | 2.2463 | -0.0073517 | -0.016443 | 0.024835 | 0.024994 | 0.05092 | -0.00013106 | 0.023273 | 0.026273 | -0.0036724 | 0.054629 | -0.0005669 | -0.023031 | -0.013873 | -0.031028 | -0.016716 | 0.0037798 | -0.0008785 | 0.000695 | -0.12122 | 0.074979
a5 | -0.37902 | -0.034829 | -0.0098174 | -0.019397 | -0.015393 | -0.012543 | -0.029628 | -0.015879 | -0.020258 | -0.033453 | -0.074273 | -0.03971 | -0.02747 | -0.029218 | -0.032749 | -0.031329 | -0.025975 | -0.015156 | -0.019132 | 0.036924 | -0.033734
a6 | 0.47668 | 0.012108 | 0.0043033 | 0.0036606 | 0.0023692 | -0.0044093 | -0.0033703 | 0.014632 | -0.0035723 | 0.0059815 | 0.014481 | 0.0091185 | -0.001249 | -0.011876 | -0.0050775 | -0.00069954 | 0.015209 | 0.026259 | 0.0016521 | -0.017401 | 0.019349
a7 | -0.2745 | -0.02593 | -0.02999 | -0.031538 | -0.036688 | -0.036842 | -0.040886 | -0.035588 | -0.036408 | -0.031727 | -0.030664 | -0.022118 | -0.024133 | -0.021248 | -0.029744 | -0.029241 | -0.035575 | -0.037453 | -0.025817 | -0.020556 | -0.036787
a8 | 0.16482 | 0.00019812 | -0.0057279 | -0.0025403 | -0.0022937 | -0.0093376 | 0.00086819 | 0.0040692 | -0.0030245 | -0.00080855 | -0.0050778 | 0.0020029 | 0.00059619 | -0.0052146 | -0.0026377 | 0.00045228 | 0.0051121 | 0.015673 | -0.0060146 | -0.012559 | 0.0042345
a9 | -0.025038 | -0.017207 | -0.0081203 | -0.014998 | -0.015683 | -0.016937 | -0.017902 | -0.016714 | -0.015068 | -0.0097853 | -0.014902 | -0.014698 | -0.013929 | -0.013362 | -0.013084 | -0.01604 | -0.012847 | -0.015294 | -0.016869 | -0.013899 | -0.015305
a10 | 0.055361 | 0.011981 | -0.018085 | -0.0072729 | -0.0025666 | -0.026995 | -0.0079093 | 0.021992 | -0.011212 | -0.0020473 | -0.0027904 | 0.0077475 | 0.0044938 | -0.021956 | -0.010537 | -0.013189 | 0.026883 | 0.047205 | -0.011164 | -0.031423 | 0
a11 | -0.059967 | -0.060651 | -0.058959 | -0.05957 | -0.060886 | -0.065274 | -0.062828 | -0.066766 | -0.0683 | -0.060337 | -0.060812 | -0.048166 | -0.051699 | -0.049351 | -0.058907 | -0.056057 | -0.065396 | -0.059779 | -0.060354 | -0.051618 | -0.059141
a12 | 0.022333 | -0.0010528 | -0.0071247 | -0.0040078 | -0.0038092 | -0.0068332 | -0.0042811 | 0.00042472 | -0.0047968 | -0.0019932 | -0.0043955 | -0.0038745 | -0.0019871 | -0.0056323 | -0.0048974 | -0.0041398 | 2.6773e-06 | 0.0032036 | -0.0055482 | -0.0092804 | -0.0027378
a13 | -0.010205 | -0.0044861 | -0.006771 | -0.0075592 | -0.0063266 | -0.0069123 | -0.0063561 | -0.006774 | -0.0070317 | -0.0096561 | -0.0063937 | -0.005981 | -0.0044647 | -0.0050218 | -0.0084783 | -0.0060133 | -0.0068509 | -0.0075411 | -0.0077221 | -0.0053466 | -0.0067726
a14 | -0.00035536 | -0.018019 | -0.020186 | -0.017817 | -0.017635 | -0.019509 | -0.019336 | -0.015157 | -0.018435 | -0.017006 | -0.020181 | -0.018569 | -0.018151 | -0.020383 | -0.01971 | -0.019628 | -0.017148 | -0.015885 | -0.018281 | -0.013231 | -0.017617
a15 | -0.011713 | -0.011204 | -0.012698 | -0.010899 | -0.014322 | -0.013078 | -0.015485 | -0.012301 | -0.014595 | -0.0034522 | -0.0087273 | -0.010837 | -0.011183 | -0.014095 | -0.0092164 | -0.0092845 | -0.0098037 | -0.014193 | -0.012047 | -0.018278 | -0.011741
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.16557 | -0.012804 | -0.084149 | 0.034266 | 0.034866 | 0.0076615 | -0.07553 | 0.035852 | -0.019235 | 0.14048 | 0.028172 | -0.046175 | -0.033356 | -0.064293 | 0.07296 | 0.0025113 | 0.012113 | 0 | 0.048157 | 0.034725
Dy (mm) | 0.21492 | 0.028317 | -0.040142 | -0.017626 | -0.013871 | -0.064882 | -0.016825 | 0.043186 | -0.030344 | 0.0091653 | -0.0032343 | 0.013297 | 0.011592 | -0.056285 | -0.020388 | -0.016639 | 0.067237 | 0.11042 | -0.03825 | -0.1144
Mag. len. (m) | 0.00698
TF (mT/KA) | -0.10458
Coil pos. (mm) | -0.00070645
Current (A) | 0 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19
|  | -6.375 | -5.625 | -4.875 | -4.125 | -3.375 | -2.625 | -1.875 | -1.125 | -0.375 | 0.375 | 1.125 | 1.875 | 2.625 | 3.375 | 4.125 | 4.875 | 5.625 | 6.375

## Summary Data (2)
File | HCMB__A001-3000004 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 382.8235 | 596.4706 | 596.1176 | 596.1176 | 596.1176 | 596 | 596 | 596 | 596 | 596 | 596.1176 | 596.2353 | 596.2353 | 596.2353 | 596.1176 | 596.1176 | 596.1176 | 596.3529 | 596.7059 | 376.7059 | 0.073243
Angle (mrad) | -0.099546 | 0.23467 | 1.3441 | 0.4451 | 0.60781 | 0.10637 | 1.1611 | 1.1878 | 0.99933 | -1.3519 | -1.1837 | -0.72536 | -0.50854 | -0.61581 | -1.4089 | -0.54443 | -0.23074 | -0.40447 | -0.88429 | 2.8449 | -4.2564
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.61228 | 0.26123 | 0.17983 | -0.39754 | -0.35476 | -0.30233 | -0.68789 | -0.85599 | -1.3076 | 0.075306 | 0.20629 | -0.18024 | -0.93733 | -0.32723 | -0.87499 | -1.0811 | -0.13843 | -0.21802 | -0.24961 | -1.6023 | -0.40513
b3 | 3.5976 | 2.9654 | 3.0348 | 3.8505 | 2.9273 | 3.0487 | 3.0865 | 3.6645 | 2.7525 | 1.7583 | 3.6542 | 4.0203 | 3.8954 | 3.4294 | 2.7906 | 2.6606 | 3.3726 | 2.998 | 2.7058 | 36.7211 | 4.2609
b4 | -0.25762 | -0.18554 | -0.19841 | -0.11498 | 0.099571 | -0.072748 | -0.043905 | 0.24054 | -0.13565 | 0.028537 | -0.077574 | -0.23096 | -0.40055 | -0.54428 | -0.57645 | -0.6337 | -0.54797 | -0.36256 | -0.4275 | 0.12756 | -0.22151
b5 | -3.1618 | 0.33904 | 0.41239 | 0.61694 | 0.66986 | 0.63276 | 0.60255 | 0.55783 | 0.63186 | 0.9867 | 0.7116 | 0.39782 | 0.23431 | 0.41603 | 0.27844 | 0.35143 | 0.46082 | 0.55283 | 0.73782 | -4.4719 | 0.24567
b6 | 0.09546 | 0.0016916 | 0.083609 | -0.03418 | 0.078989 | 0.11406 | 0.095541 | 0.035054 | 0.013583 | -0.053335 | 0.030061 | -0.028989 | 0.058664 | -0.18109 | -0.12218 | -0.062197 | -0.059627 | -0.10827 | -0.048137 | -0.090619 | -0.0094855
b7 | 0.31256 | 0.61086 | 0.62535 | 0.6209 | 0.64214 | 0.6371 | 0.61406 | 0.60297 | 0.67882 | 0.63241 | 0.69993 | 0.65525 | 0.62428 | 0.69328 | 0.63729 | 0.6736 | 0.63522 | 0.55182 | 0.50155 | 2.4423 | 0.67866
b8 | -0.027824 | -0.038811 | -0.040119 | -0.0015956 | -0.019176 | -0.0098399 | -0.036308 | 0.01898 | -0.013614 | 0.024371 | 0.018141 | -0.023688 | 0.043298 | 0.0030711 | 0.0093713 | 0.008766 | -0.0037483 | -0.00050821 | 0.018004 | 0.0069615 | -0.0029511
b9 | 0.28165 | 0.27429 | 0.27198 | 0.26428 | 0.24988 | 0.26971 | 0.27251 | 0.27132 | 0.25819 | 0.23087 | 0.2535 | 0.26864 | 0.26534 | 0.27078 | 0.27053 | 0.25441 | 0.2687 | 0.22162 | 0.26169 | 0.27708 | 0.26223
b10 | -0.11573 | -0.026064 | -0.0015148 | 0.027121 | 0.031401 | 0.01275 | -0.027044 | 0.0034461 | 0.01514 | 0.012947 | 0.080474 | -0.022028 | 0.0016925 | 0.00043813 | -0.0051953 | -0.02934 | 0.0072847 | 0.0020567 | 0.0044406 | -0.021624 | 0
b11 | 0.633 | 0.72558 | 0.72248 | 0.7215 | 0.72219 | 0.7194 | 0.71626 | 0.70998 | 0.71011 | 0.70459 | 0.70241 | 0.7114 | 0.70904 | 0.7064 | 0.7128 | 0.71686 | 0.71567 | 0.71495 | 0.72852 | 0.62908 | 0.70946
b12 | -0.015958 | -0.0054906 | 0.0006949 | 0.00059032 | 0.0003058 | -0.0019419 | -0.009814 | -0.0086527 | 0.00069365 | 0.0096657 | 0.0099905 | -0.0072405 | -0.0098966 | -0.0081686 | -0.0066007 | -0.0075736 | 0.0019961 | 0.0034753 | 0.0041688 | -0.0029853 | -0.0023828
b13 | 0.065118 | 0.087562 | 0.086072 | 0.083528 | 0.080966 | 0.081022 | 0.082267 | 0.080771 | 0.082173 | 0.076688 | 0.080182 | 0.077397 | 0.079592 | 0.079577 | 0.076612 | 0.078832 | 0.079671 | 0.08186 | 0.079155 | 0.11474 | 0.081366
b14 | -0.010908 | -0.0047432 | -0.0026854 | -0.00086858 | -0.0016031 | -0.0008232 | -0.0043946 | -0.004072 | -0.0016382 | -0.0042281 | 0.0050723 | -0.0071882 | -0.0075832 | -0.0072816 | -0.0058302 | -0.0056655 | -0.0030762 | -0.0037608 | -0.0023491 | 0.0040402 | -0.0034844
b15 | 0.03793 | 0.051854 | 0.053627 | 0.0522 | 0.054522 | 0.055141 | 0.054531 | 0.054242 | 0.052675 | 0.054937 | 0.049258 | 0.0523 | 0.050849 | 0.054205 | 0.052909 | 0.049942 | 0.051018 | 0.052691 | 0.053526 | 0.010663 | 0.050925
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | -0.99546 | 2.3467 | 13.4408 | 4.451 | 6.0782 | 1.0637 | 11.6106 | 11.8779 | 9.9933 | -13.519 | -11.8367 | -7.2536 | -5.0854 | -6.1581 | -14.0892 | -5.4443 | -2.3074 | -4.0447 | -8.8429 | 28.4486 | 0
a2 | -6.5762 | -1.5064 | -1.4897 | -1.4237 | -1.9905 | -0.31357 | 0.41306 | 0.44496 | -0.22845 | 0.79618 | -0.82024 | -2.3318 | -2.3739 | -1.6626 | -1.7053 | -1.0065 | -0.75543 | -0.85302 | 0.31454 | 11.1397 | -0.70971
a3 | 0.46842 | 0.77671 | 0.62109 | 0.34609 | 0.6595 | 0.92563 | 0.97094 | 0.88546 | 1.0231 | 1.5279 | 0.58403 | 0.26049 | 0.72446 | 1.021 | 1.0542 | 1.3658 | 0.44859 | 0.66998 | 0.00018324 | -5.0625 | 0.569
a4 | -0.2529 | -0.37019 | -0.18838 | -0.44573 | -0.033498 | 0.34425 | 0.5638 | 0.63481 | 0.64857 | -0.30641 | -0.39344 | -0.3711 | -0.58807 | 0.029746 | -0.23182 | -0.50349 | -0.32857 | -0.22012 | -0.049633 | 2.4739 | -0.021247
a5 | -0.2789 | 0.15646 | 0.077778 | 0.0052918 | 0.01524 | 0.051088 | 0.087831 | 0.17934 | 0.33301 | 0.47421 | 0.021576 | 0.034328 | 0.15512 | 0.10389 | 0.21578 | 0.26622 | 0.10661 | -0.056529 | -0.19339 | 1.6833 | 0.15138
a6 | -0.029962 | -0.031727 | 0.017146 | 0.0078036 | 0.022369 | 0.13755 | 0.13391 | 0.215 | 0.15096 | 0.2023 | -0.057159 | -0.098367 | -0.024142 | 0.057803 | 0.0097351 | 0.0023651 | 0.054715 | 0.068225 | 0.096366 | 0.63356 | 0.069821
a7 | 0.045163 | 0.02153 | 0.0038596 | 0.039157 | 0.083091 | 0.086342 | -0.015934 | -0.018329 | -0.030109 | -0.083235 | -0.034897 | -0.044856 | -0.020625 | 0.040244 | 0.074639 | 0.057862 | 0.010841 | 0.024334 | -0.042224 | 1.4431 | 0.056686
a8 | -0.0016172 | 0.016525 | 0.010722 | 0.010652 | 0.012819 | 0.019743 | 0.0049552 | 0.054298 | -0.04023 | -0.02121 | -0.0064897 | -0.017605 | 0.029467 | -0.027078 | -0.033676 | 0.016029 | 0.0037736 | -0.050325 | 0.0072637 | 0.14014 | 0.0040018
a9 | -0.023181 | -0.00016388 | 0.018582 | 0.0075641 | 0.025404 | 0.022369 | 0.011696 | 0.021545 | 0.010114 | 0.030158 | 0.0082048 | -0.013447 | -0.025497 | 0.011282 | -0.011335 | 0.0082944 | -0.0047283 | 0.0047737 | 0.01592 | -0.32055 | -0.0039737
a10 | -0.05875 | 0.0031417 | -0.0012503 | 0.011246 | -0.0043156 | -0.0089808 | 0.0024792 | 0.0047792 | 0.0051859 | 0.02178 | 0.019659 | 0.0012964 | 0.017971 | -0.018222 | -0.017508 | -0.017452 | -0.021758 | -0.0067323 | 0.00537 | 0.064962 | 0
a11 | -0.013667 | -0.022173 | -0.010064 | -0.021014 | -0.017967 | -0.019583 | -0.013795 | -0.013169 | -0.017447 | -0.039736 | -0.041576 | -0.040672 | -0.040498 | -0.03657 | -0.039973 | -0.035752 | -0.037422 | -0.044052 | -0.050137 | 0.17576 | -0.02278
a12 | -0.01363 | -0.0070649 | -0.005581 | 1.1112e-05 | -0.0096127 | -0.008603 | -0.0066662 | -0.0045138 | -0.00047258 | 0.0024482 | -0.0073747 | -0.012056 | -0.0071499 | -0.0099509 | -0.011583 | -0.0083723 | -0.0067926 | -0.0032333 | -0.0060816 | 0.011093 | -0.0059354
a13 | 0.0014697 | 0.0019016 | -0.0012665 | -0.00047472 | -0.0011517 | -0.0034106 | -0.00048429 | -0.00011371 | -0.0027047 | 0.00078091 | -0.0073822 | -0.0052062 | -0.0027404 | -0.0032847 | -0.0023501 | -0.0023479 | -0.0057795 | -0.0042612 | -0.0019786 | -0.021038 | -0.0028307
a14 | -0.016181 | -0.027939 | -0.028396 | -0.026968 | -0.029953 | -0.031223 | -0.030046 | -0.030266 | -0.028869 | -0.033861 | -0.025697 | -0.030115 | -0.028052 | -0.029959 | -0.029983 | -0.026679 | -0.026164 | -0.027322 | -0.028402 | 0.0026422 | -0.027426
a15 | -0.016446 | -0.0070944 | -0.0047724 | -0.00014776 | 0.00095496 | -0.0021373 | -0.0070212 | -0.0032514 | -0.0012634 | -0.0010926 | 0.0039314 | -0.0097727 | -0.006506 | -0.0067516 | -0.0074799 | -0.0088734 | -0.0063192 | -0.0059557 | -0.0041376 | 0.0018468 | -0.004517
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.30706 | -0.061243 | 0 | 0.063076 | 0.074115 | 0.03068 | -0.064293 | 0 | 0.03592 | 0.028172 | 0.19146 | -0.052651 | 0 | 0 | -0.010009 | -0.067333 | 0.019945 | 0 | 0 | 0
Dy (mm) | -0.16537 | 0 | 0 | 0.028334 | 0 | -0.020388 | 0 | 0.011592 | 0.013297 | 0.054149 | 0.058853 | 0 | 0.043186 | -0.043671 | -0.042312 | -0.04477 | -0.050638 | -0.015647 | 0.013182 | 0.17788
Mag. len. (m) | 14.4536
TF (mT/KA) | 596.1699
Coil pos. (mm) | 0.024428
Current (A) | 8.5

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -0.40513 | 0.56744 | -0.40398 | 0.45874 | -0.70971 | 3.1763 | -0.91696 | 0.97931 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
3 | 4.2609 | 7.5215 | 3.1447 | 0.55554 | 0.569 | 1.3507 | 0.76991 | 0.37652 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
4 | -0.22151 | 0.25107 | -0.23494 | 0.24943 | -0.021247 | 0.68851 | -0.10123 | 0.395 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
5 | 0.24567 | 1.3673 | 0.53269 | 0.18955 | 0.15138 | 0.39304 | 0.11259 | 0.1501 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
6 | -0.0094855 | 0.082982 | -0.013311 | 0.080661 | 0.069821 | 0.15591 | 0.053599 | 0.087682 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
7 | 0.67866 | 0.41752 | 0.62972 | 0.047288 | 0.056686 | 0.32375 | 0.0083853 | 0.048656 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
8 | -0.0029511 | 0.022866 | -0.0035536 | 0.022053 | 0.0040018 | 0.040303 | -0.00069139 | 0.026278 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
9 | 0.26223 | 0.014961 | 0.26042 | 0.014911 | -0.0039737 | 0.074661 | 0.0074478 | 0.014506 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
10 | 0 | 0.036545 | 0.00057572 | 0.0013812 | 0 | 0.023665 | -5.7027e-05 | 0.0011736 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
11 | 0.70946 | 0.026765 | 0.7149 | 0.0074674 | -0.02278 | 0.047549 | -0.030133 | 0.012829 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
12 | -0.0023828 | 0.0067889 | -0.0024587 | 0.0047254 | -0.0059354 | 0.0056738 | -0.0062222 | 0.0033074 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
13 | 0.081366 | 0.008961 | 0.08068 | 0.0029772 | -0.0028307 | 0.0048484 | -0.002164 | 0.0022552 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
14 | -0.0034844 | 0.0037451 | -0.0038661 | 0.0017408 | -0.027426 | 0.0077 | -0.029012 | 0.0025674 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
15 | 0.050925 | 0.0099601 | 0.052801 | 0.0017238 | -0.004517 | 0.0046757 | -0.0043161 | 0.0036944 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.40009 | #REF!
3 | [12 empty] | 3 | 9 | -0.7823 | #REF!
4 | [12 empty] | 4 | 16 | -1.1587 | #REF!
5 | [12 empty] | 5 | 25 | -1.7798 | #REF!
6 | [12 empty] | 6 | 36 | -2.4758 | #REF!
7 | [12 empty] | 7 | 49 | -3.0372 | #REF!
8 | [12 empty] | 8 | 64 | -3.7267 | #REF!
9 | [12 empty] | 9 | 81 | -4.2194 | #REF!
10 | [12 empty] | sigma (mm) |  | 0.024428 | #REF!
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.61258 | 0.15315 | 0.036755 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.35886 | 0.089715 | 0.021532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.20767 | 0.051917 | 0.01246 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.11871 | 0.029678 | 0.0071228 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.067037 | 0.016759 | 0.0040222 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.037395 | 0.0093488 | 0.0022437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.020607 | 0.0051516 | 0.0012364 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.011217 | 0.0028042 | 0.00067302 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0060316 | 0.0015079 | 0.0003619 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0032039 | 0.00080097 | 0.00019223 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0016811 | 0.00042029 | 0.00010087 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0008714 | 0.00021785 | 5.2284e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.7082 | -0.24961 | -0.21926 | -0.14901 | -1.0673 | -0.87695 | -0.33248 | -0.93365 | -0.15533 | 0.12802 | 0.079212 | -1.3176 | -0.85478 | -0.66454 | -0.31556 | -0.38028 | -0.42496 | 0.17983 | 0.28259 | 0.73313 | -0.40513
3 | 36.7987 | 2.7057 | 2.9986 | 3.3775 | 2.6571 | 2.7913 | 3.4291 | 3.8909 | 4.0182 | 3.6528 | 1.7552 | 2.7548 | 3.6658 | 3.086 | 3.0478 | 2.926 | 3.8496 | 3.0348 | 2.9634 | 3.591 | 4.2609
4 | 0.19802 | -0.4281 | -0.36235 | -0.5514 | -0.63093 | -0.57795 | -0.54535 | -0.39898 | -0.22603 | -0.10933 | 0.028039 | -0.13995 | 0.24103 | -0.03479 | -0.077561 | 0.08789 | -0.1241 | -0.19841 | -0.18066 | -0.4752 | -0.22151
5 | -4.4387 | 0.7382 | 0.55251 | 0.46035 | 0.35016 | 0.27796 | 0.41529 | 0.234 | 0.39737 | 0.70892 | 0.99036 | 0.6323 | 0.55856 | 0.60436 | 0.63091 | 0.66814 | 0.61763 | 0.41239 | 0.33907 | -3.1517 | 0.24567
6 | -2.2853e-05 | -0.048333 | -0.1084 | -0.064292 | -0.047103 | -0.12104 | -0.18171 | 0.05835 | -0.016813 | -0.017961 | -0.061214 | 0.004836 | 0.034979 | 0.10948 | 0.10654 | 0.062192 | -0.047611 | 0.083609 | 0.014895 | 0.1267 | -0.0094855
7 | 2.4526 | 0.50159 | 0.55214 | 0.63517 | 0.67355 | 0.63791 | 0.69377 | 0.62481 | 0.65474 | 0.69834 | 0.63166 | 0.6788 | 0.60323 | 0.6131 | 0.63706 | 0.64273 | 0.62106 | 0.62535 | 0.60988 | 0.30915 | 0.67866
8 | -0.01987 | 0.018103 | -0.00054336 | -0.0061576 | 0.016653 | 0.010871 | 0.0028392 | 0.04278 | -0.017032 | -0.0044723 | 0.022079 | -0.017915 | 0.019098 | -0.028063 | -0.013948 | -0.027891 | -0.0093392 | -0.040119 | -0.030906 | 0.014678 | -0.0029511
9 | 0.28319 | 0.26173 | 0.22168 | 0.26921 | 0.25378 | 0.27089 | 0.2712 | 0.26575 | 0.26803 | 0.24596 | 0.2313 | 0.25794 | 0.27135 | 0.27159 | 0.2696 | 0.24865 | 0.26354 | 0.27198 | 0.27345 | 0.26798 | 0.26223
10 | -0.0032336 | 0.0040519 | 0.0024622 | 2.9328e-06 | -5.6662e-06 | -3.8014e-06 | 0.0013776 | 0.00066369 | 5.0659e-06 | -7.4715e-05 | 4.5859e-06 | -1.3017e-07 | 0.0033563 | 4.509e-05 | 2.1964e-06 | -8.4758e-05 | 3.1807e-08 | -0.0015148 | 7.5365e-05 | -6.9111e-05 | 0
11 | 0.63036 | 0.72847 | 0.71499 | 0.71586 | 0.71677 | 0.71308 | 0.70668 | 0.70884 | 0.71116 | 0.70089 | 0.7045 | 0.71009 | 0.70994 | 0.71585 | 0.71956 | 0.72218 | 0.72148 | 0.72248 | 0.72536 | 0.63129 | 0.70946
12 | -0.056268 | 0.041504 | 0.035223 | 0.01081 | -0.037526 | -0.059892 | -0.080674 | -0.099802 | -0.04364 | -0.011529 | 0.081705 | -0.014152 | -0.086537 | -0.060804 | -0.036475 | -0.039301 | -0.031382 | 0.006949 | -0.017053 | -0.020157 | -0.023828 | ten times!
13 | 1.151 | 0.78869 | 0.82187 | 0.80731 | 0.79453 | 0.77538 | 0.80578 | 0.78666 | 0.77108 | 0.78283 | 0.75377 | 0.81924 | 0.80502 | 0.82051 | 0.81528 | 0.81057 | 0.82985 | 0.86072 | 0.8734 | 0.64603 | 0.81366 | ten times!
14 | 0.043108 | -0.02394 | -0.03684 | -0.036506 | -0.02569 | -0.051335 | -0.070388 | -0.078145 | -0.049205 | -0.025038 | -0.055513 | -0.032102 | -0.041031 | -0.015073 | -0.021805 | -0.049308 | -0.035836 | -0.026854 | -0.021279 | 0.0092334 | -0.034844 | ten times!
15 | 0.10663 | 0.53526 | 0.52691 | 0.51018 | 0.49942 | 0.52909 | 0.54205 | 0.50849 | 0.523 | 0.49258 | 0.54937 | 0.52675 | 0.54242 | 0.54531 | 0.55141 | 0.54522 | 0.522 | 0.53627 | 0.51854 | 0.3793 | 0.50925 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.3712 | 0.31034 | -0.8475 | -0.73639 | -0.98168 | -1.6902 | -1.645 | -2.3937 | -2.3302 | -0.85869 | 0.77991 | -0.23708 | 0.43996 | 0.42041 | -0.3096 | -1.9963 | -1.4391 | -1.4897 | -1.5008 | -6.4893 | -0.70971
3 | -5.0665 | 0.0011777 | 0.66898 | 0.44485 | 1.3548 | 1.0494 | 1.0168 | 0.72751 | 0.25704 | 0.59813 | 1.5292 | 1.0194 | 0.88497 | 0.97734 | 0.9235 | 0.65994 | 0.35163 | 0.62109 | 0.77271 | 0.44719 | 0.569
4 | 2.661 | -0.051922 | -0.21808 | -0.32358 | -0.49557 | -0.22854 | 0.034021 | -0.59045 | -0.37067 | -0.40427 | -0.32212 | 0.64378 | 0.63329 | 0.56513 | 0.34692 | -0.033763 | -0.44993 | -0.18838 | -0.36793 | -0.39607 | -0.021247
5 | 1.688 | -0.19321 | -0.057027 | 0.1054 | 0.26545 | 0.21429 | 0.10157 | 0.15438 | 0.032804 | 0.024274 | 0.47339 | 0.33136 | 0.17922 | 0.090363 | 0.050531 | 0.014752 | 0.0054318 | 0.077778 | 0.15589 | -0.27696 | 0.15138
6 | 0.48023 | 0.094032 | 0.071272 | 0.065991 | 0.014384 | 0.019516 | 0.068489 | -0.033657 | -0.099201 | -0.06934 | 0.19104 | 0.14816 | 0.21254 | 0.13355 | 0.1412 | 0.020196 | 0.00072292 | 0.017146 | -0.031262 | -0.0068244 | 0.069821
7 | 1.4426 | -0.042321 | 0.024331 | 0.010732 | 0.058468 | 0.074664 | 0.040299 | -0.021395 | -0.045237 | -0.034825 | -0.083533 | -0.02944 | -0.018419 | -0.015803 | 0.08601 | 0.082699 | 0.038899 | 0.0038596 | 0.021947 | 0.043064 | 0.056686
8 | 0.11695 | 0.0056403 | -0.048693 | 0.010221 | 0.021652 | -0.028342 | -0.021513 | 0.024075 | -0.017938 | -0.01425 | -0.027493 | -0.042017 | 0.052818 | 0.005309 | 0.022007 | 0.011933 | 0.0069039 | 0.010722 | 0.01652 | 0.016952 | 0.0040018
9 | -0.31851 | 0.015889 | 0.0047907 | -0.0043033 | 0.0069769 | -0.011544 | 0.011292 | -0.025535 | -0.013411 | 0.0037047 | 0.029462 | 0.0099084 | 0.021523 | 0.01178 | 0.022652 | 0.025573 | 0.0067817 | 0.018582 | -6.2015e-05 | -0.042864 | -0.0039737
10 | -0.00086094 | -0.00027916 | -0.00015169 | -1.7429e-06 | 1.0878e-05 | -2.4512e-06 | -7.5187e-05 | -4.0683e-05 | 3.6768e-05 | 2.4604e-05 | -4.7287e-06 | 1.498e-07 | -6.2012e-05 | 0.0019575 | 2.2086e-07 | -0.0035322 | 7.7517e-07 | -0.0012503 | 0.0023429 | 0.00035875 | 0
11 | 0.1761 | -0.050172 | -0.044017 | -0.037269 | -0.036336 | -0.040229 | -0.036801 | -0.040221 | -0.041083 | -0.041043 | -0.040119 | -0.017441 | -0.013104 | -0.014072 | -0.019438 | -0.017506 | -0.021025 | -0.010064 | -0.022453 | -0.018083 | -0.02278 | ten times!
12 | -0.033146 | -0.068181 | -0.023291 | -0.038634 | -0.059926 | -0.093109 | -0.074978 | -0.095762 | -0.12249 | -0.097081 | -0.0049856 | -0.011753 | -0.051747 | -0.066881 | -0.073631 | -0.095525 | -0.016383 | -0.05581 | -0.069827 | -0.057102 | -0.059354 | ten times!
13 | -0.21587 | -0.019549 | -0.043062 | -0.054995 | -0.039156 | -0.027683 | -0.035279 | -0.0249 | -0.064187 | -0.038482 | 0.016855 | -0.01895 | -0.00077618 | -0.019615 | -0.026909 | 0.0054596 | 0.0084492 | -0.012665 | 0.0059311 | -0.03709 | -0.028307 | ten times!
14 | 0.010802 | -0.28983 | -0.26643 | -0.23933 | -0.2533 | -0.28201 | -0.28009 | -0.2986 | -0.30539 | -0.28704 | -0.36285 | -0.29408 | -0.30784 | -0.30418 | -0.30243 | -0.30012 | -0.28179 | -0.28396 | -0.28297 | -0.15174 | -0.27426 | ten times!
15 | 0.018468 | -0.041376 | -0.059557 | -0.063192 | -0.088734 | -0.074799 | -0.067516 | -0.06506 | -0.097727 | 0.039314 | -0.010926 | -0.012634 | -0.032514 | -0.070212 | -0.021373 | 0.0095496 | -0.0014776 | -0.047724 | -0.070944 | -0.16446 | -0.04517 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
6 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
7 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
8 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
9 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
10 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
11 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
12 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
13 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
14 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
15 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
16 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
17 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
6 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
7 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
8 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
9 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
10 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
11 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
12 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
13 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
14 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
15 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
16 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
17 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00089898 | 0.00030697 | -8.7706e-05 | -8.7706e-05 | -8.7706e-05 | 0.00010963 | 0.00010963 | 0.00010963 | -8.7706e-05 | -0.00028504 | -0.00028504 | -0.00028504 | -0.00028504 | -0.00028504 | -8.7706e-05 | -8.7706e-05 | -8.7706e-05 | 0.00050431 |  | 596.1699
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF!
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3629.3665
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.934 | -3680.6789 | 9.8678 | 3.9471 | 0 | 0 | 0 | 1.9736 | 1.9736 | 1.9736 | 0 | -1.9736 | -1.9736 | -1.9736 | -1.9736 | -1.9736 | 0 | 0 | 0 | 5.9207 | -3578.0541
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.098439 | 2.9433 | -0.78585 | -0.30603 | -0.1323 | -0.44599 | -1.3105 | -0.51737 | -0.4101 | -0.62692 | -1.0852 | -1.2535 | 1.0978 | 1.2862 | 1.2595 | 0.20481 | 0.70625 | 0.54354 | 1.4425 | 0.33311 | -0.0011074
db2 | -0.39947 | -1.6023 | 0.14986 | 0.18145 | 0.26104 | -0.68167 | -0.47552 | 0.072235 | -0.53787 | 0.21923 | 0.60576 | 0.47478 | -0.90814 | -0.45652 | -0.28842 | 0.097138 | 0.044711 | 0.0019346 | 0.5793 | 0.6607 | 0.61228
db3 | 1.8053 | 36.7211 | -0.43952 | -0.14729 | 0.22732 | -0.48467 | -0.35474 | 0.28406 | 0.75011 | 0.87502 | 0.50894 | -1.387 | -0.39284 | 0.51922 | -0.058831 | -0.096631 | -0.21796 | 0.70522 | -0.11048 | -0.17991 | 3.5976
db4 | -0.23245 | 0.12756 | -0.19505 | -0.13011 | -0.31552 | -0.40125 | -0.344 | -0.31183 | -0.1681 | 0.0014939 | 0.15488 | 0.26099 | 0.096801 | 0.473 | 0.18855 | 0.1597 | 0.33202 | 0.11747 | 0.034046 | 0.046908 | -0.25762
db5 | 0.44833 | -4.4719 | 0.20499 | 0.019993 | -0.072018 | -0.18141 | -0.2544 | -0.11681 | -0.29852 | -0.13501 | 0.17877 | 0.45386 | 0.099023 | 0.024992 | 0.06972 | 0.099928 | 0.13703 | 0.084101 | -0.12044 | -0.19379 | -3.1618
db6 | -0.010375 | -0.090619 | -0.037762 | -0.097895 | -0.049252 | -0.051822 | -0.1118 | -0.17071 | 0.069039 | -0.018614 | 0.040436 | -0.04296 | 0.023958 | 0.045429 | 0.10592 | 0.12443 | 0.089364 | -0.023805 | 0.093983 | 0.012066 | 0.09546
db7 | 0.73782 | 2.4423 | -0.12827 | -0.078006 | 0.005395 | 0.043778 | 0.0074614 | 0.06346 | -0.0055403 | 0.025425 | 0.070102 | 0.0025892 | 0.048993 | -0.026854 | -0.015762 | 0.0072742 | 0.012318 | -0.0089267 | -0.0044698 | -0.018967 | 0.31256
db8 | -0.0024114 | 0.0069615 | 0.020415 | 0.0019032 | -0.0013369 | 0.011177 | 0.011783 | 0.0054825 | 0.045709 | -0.021276 | 0.020552 | 0.026782 | -0.011203 | 0.021392 | -0.033896 | -0.0074285 | -0.016764 | 0.00081576 | -0.037708 | -0.0364 | -0.027824
db9 | 0.22836 | 0.27708 | 0.00067443 | -0.039392 | 0.0076856 | -0.0066022 | 0.009515 | 0.0097676 | 0.0043225 | 0.0076285 | -0.0075124 | -0.030144 | -0.002827 | 0.01031 | 0.011497 | 0.0086989 | -0.011134 | 0.0032617 | 0.010971 | 0.01328 | 0.28165
db10 | 0.0048892 | -0.021624 | -0.00044857 | -0.0028325 | 0.0023955 | -0.03423 | -0.010084 | -0.0044511 | -0.0031967 | -0.026917 | 0.075585 | 0.0080574 | 0.010251 | -0.0014431 | -0.031933 | 0.0078611 | 0.026511 | 0.022232 | -0.006404 | -0.030953 | -0.11573
db11 | 0.71916 | 0.62908 | 0.013512 | -5.4906e-05 | 0.00065629 | 0.0018509 | -0.0022076 | -0.0086077 | -0.0059643 | -0.0036067 | -0.012599 | -0.010414 | -0.0049017 | -0.0050326 | 0.0012525 | 0.0043953 | 0.0071835 | 0.0064949 | 0.0074697 | 0.010573 | 0.633
db12 | -0.0018777 | -0.0029853 | 0.0060465 | 0.0053529 | 0.0038738 | -0.005696 | -0.004723 | -0.006291 | -0.008019 | -0.0053628 | 0.011868 | 0.011543 | 0.0025713 | -0.006775 | -0.0079363 | -6.4232e-05 | 0.0021835 | 0.002468 | 0.0025726 | -0.0036129 | -0.015958
db13 | 0.080774 | 0.11474 | -0.0016185 | 0.0010863 | -0.0011027 | -0.001942 | -0.0041615 | -0.0011964 | -0.0011816 | -0.0033763 | -0.00059139 | -0.0040858 | 0.001399 | -2.9044e-06 | 0.0014932 | 0.00024788 | 0.00019274 | 0.002754 | 0.0052982 | 0.006788 | 0.065118
db14 | -0.0034844 | 0.0040402 | 0.0011353 | -0.0002764 | 0.00040819 | -0.0021811 | -0.0023459 | -0.0037972 | -0.0040988 | -0.0037038 | 0.0085567 | -0.00074368 | 0.0018462 | -0.00058767 | -0.00091022 | 0.0026612 | 0.0018813 | 0.0026158 | 0.00079896 | -0.0012588 | -0.010908
db15 | 0.052801 | 0.010663 | 0.0007244 | -0.00011022 | -0.0017834 | -0.0028597 | 0.00010705 | 0.0014033 | -0.0019522 | -0.00050176 | -0.0035439 | 0.0021356 | -0.00012673 | 0.001441 | 0.0017298 | 0.0023398 | 0.0017201 | -0.00060101 | 0.00082579 | -0.00094791 | 0.03793
da2 | -0.91625 | 11.1397 | 1.2308 | 0.063228 | 0.16081 | -0.090264 | -0.78909 | -0.74637 | -1.4576 | -1.4155 | 0.096008 | 1.7124 | 0.68779 | 1.3612 | 1.3293 | 0.60268 | -1.0743 | -0.50749 | -0.57341 | -0.59014 | -6.5762
da3 | 0.77029 | -5.0625 | -0.7701 | -0.10031 | -0.3217 | 0.59547 | 0.28386 | 0.25074 | -0.045828 | -0.5098 | -0.18626 | 0.75764 | 0.25286 | 0.11518 | 0.20066 | 0.15534 | -0.11079 | -0.4242 | -0.1492 | 0.0064257 | 0.46842
da4 | -0.10051 | 2.4739 | 0.050881 | -0.1196 | -0.22805 | -0.40297 | -0.13131 | 0.13026 | -0.48755 | -0.27059 | -0.29293 | -0.20589 | 0.74908 | 0.73533 | 0.66431 | 0.44477 | 0.067017 | -0.34522 | -0.087862 | -0.26967 | -0.2529
da5 | 0.11299 | 1.6833 | -0.30638 | -0.16952 | -0.0063837 | 0.15323 | 0.10279 | -0.0090993 | 0.042132 | -0.078664 | -0.091416 | 0.36122 | 0.22002 | 0.066345 | -0.025161 | -0.061904 | -0.097752 | -0.1077 | -0.035213 | 0.043468 | -0.2789
da6 | 0.053603 | 0.63356 | 0.042763 | 0.014621 | 0.0011114 | -0.051238 | -0.043868 | 0.0041993 | -0.077745 | -0.15197 | -0.11076 | 0.1487 | 0.09736 | 0.1614 | 0.080305 | 0.083948 | -0.031234 | -0.0458 | -0.036458 | -0.08533 | -0.029962
da7 | 0.0084273 | 1.4431 | -0.050651 | 0.015907 | 0.0024137 | 0.049435 | 0.066212 | 0.031816 | -0.029052 | -0.053283 | -0.043324 | -0.091662 | -0.038537 | -0.026756 | -0.024361 | 0.077915 | 0.074663 | 0.030729 | -0.0045677 | 0.013103 | 0.045163
da8 | -0.00057591 | 0.14014 | 0.0078396 | -0.049749 | 0.0043495 | 0.016605 | -0.0331 | -0.026502 | 0.030043 | -0.017029 | -0.0059138 | -0.020634 | -0.039654 | 0.054874 | 0.0055311 | 0.020319 | 0.013395 | 0.011228 | 0.011298 | 0.017101 | -0.0016172
da9 | 0.0078186 | -0.32055 | 0.0081015 | -0.0030449 | -0.012547 | 0.00047586 | -0.019153 | 0.0034634 | -0.033315 | -0.021266 | 0.00038624 | 0.022339 | 0.002295 | 0.013726 | 0.0038772 | 0.01455 | 0.017585 | -0.00025452 | 0.010763 | -0.0079825 | -0.023181
da10 | -0.00018392 | 0.064962 | 0.0055539 | -0.0065483 | -0.021575 | -0.017268 | -0.017324 | -0.018038 | 0.018155 | 0.0014804 | 0.019843 | 0.021964 | 0.0053698 | 0.0049631 | 0.0026631 | -0.0087969 | -0.0041317 | 0.01143 | -0.0010664 | 0.0033256 | -0.05875
da11 | -0.030089 | 0.17576 | -0.020048 | -0.013963 | -0.0073329 | -0.0056628 | -0.0098846 | -0.006481 | -0.010409 | -0.010584 | -0.011487 | -0.0096471 | 0.012642 | 0.01692 | 0.016294 | 0.010506 | 0.012121 | 0.0090751 | 0.020025 | 0.0079156 | -0.013667
da12 | -0.0062583 | 0.011093 | 0.0001767 | 0.003025 | -0.00053435 | -0.002114 | -0.0053242 | -0.0036927 | -0.00089166 | -0.0057977 | -0.0011165 | 0.0087065 | 0.0057857 | 0.0017445 | -0.00040789 | -0.0023447 | -0.0033545 | 0.0062694 | 0.00067729 | -0.00080665 | -0.01363
da13 | -0.0023475 | -0.021038 | 0.00036891 | -0.0019138 | -0.003432 | -4.3438e-07 | -2.6514e-06 | -0.00093725 | -0.00039296 | -0.0028587 | -0.0050348 | 0.0031284 | -0.00035721 | 0.0022338 | 0.0018632 | -0.0010631 | 0.0011958 | 0.0018728 | 0.001081 | 0.004249 | 0.0014697
da14 | -0.028883 | 0.0026422 | 0.00048067 | 0.001561 | 0.0027189 | 0.0022039 | -0.0011002 | -0.0010759 | 0.00083152 | -0.0012321 | 0.0031863 | -0.0049778 | 1.4032e-05 | -0.0013831 | -0.0011633 | -0.00234 | -0.0010702 | 0.0019146 | 0.00048748 | 0.00094414 | -0.016181
da15 | -0.0043161 | 0.0018468 | 0.0001785 | -0.0016395 | -0.0020031 | -0.0045573 | -0.0031638 | -0.0024355 | -0.0021898 | -0.0054566 | 0.0082475 | 0.0032235 | 0.0030527 | 0.0010647 | -0.002705 | 0.0021788 | 0.0052711 | 0.0041684 | -0.00045629 | -0.0027783 | -0.016446
Check Computation for Aperture2
Magnetic Length
mbh dB/B | #REF!
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db6 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db7 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db8 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db9 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db10 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db11 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db12 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db13 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db14 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
db15 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da6 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da7 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da8 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da9 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da10 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da11 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da12 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da13 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da14 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
da15 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Field Colinearity | #REF!
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 40.6 | 4.6 | 3.5 | 7 | -3.47 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.34 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -2.45 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.73 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.61 | 0.3 | 3.5 | 7 | 0.62 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.33 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.266 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 24.5 | 56.7 | 8.4 | 72.8 | -12.22 | 5.28 | -20.97 | 14.03
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.41 | 3.09 | -2.16 | 4.84 | -0.64 | 0.64 | -1.28 | 1.28 | -7 | 2.1 | -11.55 | 6.65 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.415 | 1.045 | 0.1 | 1.36 | -0.216 | 0.216 | -0.432 | 0.432 | 1.56 | 3.66 | 0.51 | 4.71 | -0.08 | 1.32 | -0.78 | 2.02
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.225 | 0.435 | 0.12 | 0.54 | -0.08 | 0.08 | -0.16 | 0.16 | 0.091 | 0.441 | -0.084 | 0.616 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 24 | 4.6 | 3.5 | 7 | -6.652 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.43 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -5.57 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.54 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 1.62 | 0.3 | 3.5 | 7 | -0.0495 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.28 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.147 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 7.9 | 40.1 | -8.2 | 56.2 | -15.402 | 2.098 | -24.152 | 10.848
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.32 | 3.18 | -2.07 | 4.93 | -0.64 | 0.64 | -1.28 | 1.28 | -10.12 | -1.02 | -14.67 | 3.53 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.225 | 0.855 | -0.09 | 1.17 | -0.216 | 0.216 | -0.432 | 0.432 | 0.57 | 2.67 | -0.48 | 3.72 | -0.7495 | 0.6505 | -1.4495 | 1.3505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.175 | 0.385 | 0.07 | 0.49 | -0.08 | 0.08 | -0.16 | 0.16 | -0.028 | 0.322 | -0.203 | 0.497 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 34.5 | 4.6 | 3.5 | 7 | 2.43 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.74 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.44 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.7 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.54 | 0.3 | 3.5 | 7 | 0.389 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.26 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.282 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 18.4 | 50.6 | 2.3 | 66.7 | -6.32 | 11.18 | -15.07 | 19.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.01 | 2.49 | -2.76 | 4.24 | -0.64 | 0.64 | -1.28 | 1.28 | -8.99 | 0.11 | -13.54 | 4.66 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.385 | 1.015 | 0.07 | 1.33 | -0.216 | 0.216 | -0.432 | 0.432 | 1.49 | 3.59 | 0.44 | 4.64 | -0.311 | 1.089 | -1.011 | 1.789
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.155 | 0.365 | 0.05 | 0.47 | -0.08 | 0.08 | -0.16 | 0.16 | 0.107 | 0.457 | -0.068 | 0.632 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.39947 | 3.1453 | -0.23245 | 0.53283 | 0.62982 | -0.98439 | -0.91625 | 0.77029 | -0.10051 | 24.4277 | 0.59617 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | -0.40513 | 4.2609 | -0.22151 | 0.24567 | 0.67866 | 0 | -0.70971 | 0.569 | -0.021247 | 14453.583 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.15 | 0.15 | 0.95 | 0.95 | 0.15 | 0.15 | 0.95 | 0.95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0