## Alarm sheet
File name | HCMB__A001-03003036_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000036
Date of test Ap 1 | 2003-05-19 |  |  |  | Date of test Ap 2 | 2003-05-19
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | yellow alarm | status ok | status ok | Main field | status ok | yellow alarm | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | yellow alarm | status ok | status ok | b5 | status ok | yellow alarm | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
103 | HCMB__A001 | 3000036 |  | 2 |  |  | 3
| COMMENTS: | Main field component shows no increase towards position 1 (conn.side). Asked for re-measurement.
| HOLDING POINT | HOLD | Data at CERN |  | 2003-05-20 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-05-20 | Answer to manufacturer |  |  | 2003-05-22
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 11.07.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)

## Original data
File name |  | HCMB__A001-03003036_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000036 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 2.0 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | noell36cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | noell36cct2 |  |  |  | a even | 1
Test Operator |  | Ortega |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-05-19 |  |  | Time of test |  |  | 6.45 - 9.30 |  |  |  | Date of test |  | 2003-05-19 |  |  | Time of test |  |  | 9.45 - 12.45
Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 1999-07-20 |  |  |  | Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 1999-07-20
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00015231 |  |  | Dy average |  |  | 0.00031929 |  |  |  | Dx average |  | 0.00014078 |  |  | Dy average |  |  | 0.00038493
Magnetic length |  | 14.4543 |  |  | Tranfs. Function (mT/KA) |  |  | 596.5098 |  |  |  | Magnetic length |  | 14.4557 |  |  | Tranfs. Function (mT/KA) |  |  | 596.4248
Phase |  | -3.984 |  |  | Integral field (T.m) |  |  | 0.073288 |  |  |  | Phase |  | -4.0092 |  |  | Integral field (T.m) |  |  | 0.073285
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003238 | 0.00507 | 0.00507 | 0.005071 | 0.00507 | 0.005071 | 0.005071 | 0.005069 | 0.005069 | 0.00507 | 0.00507 | 0.00507 | 0.00507 | 0.00507 | 0.00507 | 0.005071 | 0.005071 | 0.005071 | 0.005072 | 0.003227 | 0.073288 | C1 (T) | 0.003238 | 0.005069 | 0.005068 | 0.005069 | 0.005068 | 0.005069 | 0.00507 | 0.005069 | 0.005069 | 0.005069 | 0.00507 | 0.00507 | 0.00507 | 0.005069 | 0.00507 | 0.00507 | 0.005071 | 0.005071 | 0.005072 | 0.003236 | 0.073285
Angle (mrad) | 0.20803 | -0.58246 | -0.77114 | -0.64817 | -1.1548 | -0.49339 | -0.24637 | 0.62591 | 0.40608 | 0.74751 | 0.21455 | 1.2814 | -0.063452 | 0.99196 | -0.19456 | -0.15825 | -0.63238 | 0.14096 | -0.77971 | 1.8674 | -3.984 | Angle (mrad) | 1.2897 | -1.1381 | -0.54045 | -1.2396 | -0.27434 | -1.1003 | 0.63396 | -0.34105 | 1.2188 | -0.43007 | 0.33579 | -0.48339 | 1.355 | 0.113 | 0.75181 | -0.68339 | -0.27848 | -0.25123 | -0.006847 | 2.3982 | -4.0092
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.617 | 0.90426 | 0.94342 | 1.4821 | 1.0794 | 0.38531 | 0.62073 | 0.22469 | 0.21156 | 0.60257 | -0.26692 | -0.41555 | 0.77901 | 1.0303 | 1.0185 | 0.89769 | 0.35801 | 0.33253 | -0.12837 | 0.68155 | 0.59794 | b2 | -3.2418 | 0.13026 | -0.23816 | -0.087807 | -0.84279 | 0.024661 | -0.49 | 0.20789 | 0.16629 | -0.36166 | -0.0018967 | 0.016656 | -0.18681 | -0.41155 | -0.30611 | -0.40551 | -0.013469 | 0.012913 | -0.41777 | 0.82187 | -0.24645
b3 | 39.0575 | 0.002401 | 0.84752 | -0.15513 | 2.1343 | 1.8604 | 2.0059 | 0.90355 | 1.5186 | 1.3765 | 1.2416 | 1.164 | 1.4841 | 0.51666 | 0.81239 | 1.5955 | 1.489 | 1.5207 | 2.2298 | 2.0915 | 2.533 | b3 | 37.9901 | -0.49039 | 0.42834 | 0.99692 | 1.3272 | 0.43904 | 1.2605 | 1.29 | 1.2302 | 0.95193 | -0.16976 | -0.17875 | 0.59975 | -0.38882 | 0.16803 | 1.119 | 1.2632 | 1.1963 | 1.9604 | 0.21266 | 1.9403
b4 | 0.45086 | -0.01611 | -0.14341 | 0.0096284 | -0.26239 | -0.079258 | 0.033279 | -0.26405 | -0.23353 | 0.039515 | 0.00027024 | -0.037186 | 0.13748 | -0.095713 | -0.11493 | 0.15231 | 0.042437 | 0.13855 | 0.081257 | -0.037381 | -0.018028 | b4 | -0.44996 | 0.11627 | -0.096371 | 0.10044 | -0.12286 | 0.29673 | 0.16109 | 0.22395 | 0.022556 | 0.1186 | 0.049329 | -0.019834 | 0.028299 | 0.16971 | 0.10578 | -0.14119 | -0.07147 | -0.15442 | -0.11318 | -0.077885 | 0.017438
b5 | -2.5709 | -0.032392 | -0.48231 | 0.24051 | 0.23004 | 0.39181 | 0.36945 | 0.22605 | 0.18739 | 0.31671 | 0.5774 | 0.22063 | 0.01253 | 0.028438 | 0.28467 | 0.24649 | 0.2026 | 0.42283 | 0.41255 | -1.9497 | 0.050484 | b5 | -2.4207 | 0.074911 | -0.40574 | 0.30119 | 0.40905 | 0.3887 | 0.35596 | 0.27377 | 0.41437 | 0.42787 | 1.0036 | 0.65786 | 0.50707 | 0.4353 | 0.428 | 0.43858 | 0.46277 | 0.62533 | 0.6686 | -1.8287 | 0.24664
b6 | 0.36437 | 0.039021 | -0.014761 | -0.19199 | -0.05341 | -0.026261 | -0.010385 | -0.032475 | 0.02276 | 0.064477 | 0.051568 | -0.035728 | 0.033483 | -0.070203 | -0.019911 | -0.073201 | -0.095713 | 0.045653 | 0.013366 | 0.051704 | -0.0045715 | b6 | -0.021278 | 0.067634 | 0.030284 | 0.052949 | 0.072885 | -0.0023354 | 0.033315 | -0.006572 | -0.071935 | 0.0795 | -0.031134 | -0.015302 | -0.1788 | 0.014662 | 0.011427 | -0.087649 | -0.18462 | -0.053895 | 0.0014 | -0.061887 | -0.016672
b7 | 3.0103 | 0.91134 | 0.86657 | 1.0589 | 0.99108 | 0.99509 | 0.94702 | 0.9281 | 0.95801 | 0.97262 | 0.9366 | 0.87042 | 0.81654 | 0.90693 | 0.89155 | 0.87723 | 0.92153 | 1.0137 | 1.0058 | 0.96177 | 1.0067 | b7 | 3.0925 | 0.92056 | 0.97905 | 1.0906 | 1.106 | 1.1411 | 1.0767 | 1.0661 | 1.0661 | 1.0453 | 1.061 | 1.0803 | 1.1148 | 1.1413 | 1.1722 | 1.0962 | 1.1178 | 1.1867 | 1.1173 | 0.89408 | 1.1477
b8 | 0.073097 | -0.015505 | -0.011544 | -0.010037 | 0.025379 | 0.031164 | 0.029772 | 0.022847 | 0.030966 | 0.053118 | 0.029886 | 0.015105 | 0.021578 | 0.015009 | 0.020802 | -0.0026579 | -0.0090368 | 0.051282 | -0.016672 | -0.07916 | 0.01441 | b8 | -0.074395 | 0.032877 | 0.0056066 | 0.0040251 | 0.013519 | 0.082362 | 0.054511 | 0.017697 | 0.026066 | 0.033758 | 0.0049403 | -0.026628 | -0.044356 | 0.018182 | 0.0051449 | -0.025828 | -0.0024833 | 0.014841 | 0.046878 | 0.0016765 | 0.011136
b9 | 0.52584 | 0.52884 | 0.48311 | 0.50763 | 0.52673 | 0.50969 | 0.49713 | 0.50911 | 0.53475 | 0.51804 | 0.50933 | 0.4938 | 0.50106 | 0.50069 | 0.5037 | 0.50005 | 0.50011 | 0.52245 | 0.53484 | 0.5613 | 0.51227 | b9 | 0.54721 | 0.52161 | 0.49188 | 0.51906 | 0.52569 | 0.49149 | 0.49862 | 0.50565 | 0.50792 | 0.49876 | 0.46772 | 0.45698 | 0.48883 | 0.49131 | 0.49598 | 0.49126 | 0.50643 | 0.51771 | 0.52784 | 0.5706 | 0.50415
b10 | 0.089889 | -0.042505 | -0.030291 | -0.052037 | 0.014354 | 0.031333 | 0.031426 | -0.019015 | -0.0083019 | 0.067674 | 0.055798 | 0.012903 | -0.023039 | 0.0062609 | -0.0030668 | -0.018852 | -0.05301 | 0.034515 | -0.021426 | -0.063053 | 0 | b10 | 0.050164 | 0.039716 | 0.0095029 | 0.029593 | 0.04732 | 0.069205 | 0.061391 | 0.0056563 | -0.010742 | 0.027095 | -0.014497 | -0.043879 | -0.070537 | -0.0015881 | 0.008157 | -0.046072 | -0.10022 | -0.014503 | 0.020986 | -0.076087 | 0
b11 | 0.68057 | 0.74494 | 0.74443 | 0.75209 | 0.75302 | 0.75957 | 0.75813 | 0.75702 | 0.75239 | 0.75213 | 0.76398 | 0.76091 | 0.75677 | 0.75422 | 0.75879 | 0.7573 | 0.75198 | 0.74974 | 0.74943 | 0.6847 | 0.74953 | b11 | 0.66219 | 0.73185 | 0.72957 | 0.74303 | 0.74386 | 0.74849 | 0.74694 | 0.7453 | 0.75124 | 0.74822 | 0.76403 | 0.75165 | 0.74739 | 0.75133 | 0.75038 | 0.74792 | 0.74292 | 0.74528 | 0.74546 | 0.68679 | 0.74162
b12 | 0.017478 | -0.0069328 | -0.0024533 | -0.00024506 | 0.002769 | 0.0039898 | 0.0041203 | -0.0021252 | 0.0028995 | 0.0042072 | 0.0018615 | 0.0014938 | -0.0011458 | 0.0019637 | 0.00021017 | -0.0025839 | -0.0047473 | 0.0023014 | -0.0062875 | -0.012245 | 0.00013817 | b12 | -0.00019328 | 0.002799 | -0.00092166 | -0.00049589 | 0.0062716 | 0.0095764 | 0.0078907 | -0.00037581 | 0.0050817 | 0.0039143 | 0.0048341 | -0.0029792 | -0.007315 | -0.0014301 | -0.00096404 | -0.0046077 | -0.0091901 | 6.4932e-06 | 0.0041104 | -0.0062753 | 0.0006263
b13 | 0.10449 | 0.075827 | 0.077207 | 0.070281 | 0.07298 | 0.073238 | 0.072824 | 0.071695 | 0.070016 | 0.07033 | 0.068945 | 0.072249 | 0.071396 | 0.071736 | 0.069525 | 0.07368 | 0.075208 | 0.07292 | 0.075449 | 0.065135 | 0.073343 | b13 | 0.10446 | 0.071831 | 0.076659 | 0.072193 | 0.072787 | 0.074017 | 0.076138 | 0.076425 | 0.073465 | 0.072999 | 0.072639 | 0.074674 | 0.074858 | 0.073736 | 0.071943 | 0.075111 | 0.072437 | 0.074432 | 0.074478 | 0.060545 | 0.074503
b14 | 0.0062452 | -0.0049606 | -0.0036154 | -0.0025839 | -0.0023886 | -0.0021537 | -0.0026887 | -0.0036374 | -0.0035086 | 1.2532e-05 | -0.00060799 | -0.0023194 | -0.0044911 | -0.0032204 | -0.0043319 | -0.0045987 | -0.005729 | -0.0023563 | -0.0056995 | -0.0062915 | -0.0030555 | b14 | 0.0075407 | -0.0025121 | -0.0037532 | -0.0028814 | -0.00099381 | 0.00056019 | 0.00028777 | -0.0035074 | -0.0035914 | -0.0032365 | -0.0042907 | -0.006647 | -0.0074714 | -0.0035614 | -0.0034561 | -0.0063898 | -0.0084495 | -0.0035097 | -0.0025695 | -0.0064187 | -0.0033852
b15 | -0.0039762 | 0.02652 | 0.025748 | 0.026882 | 0.025384 | 0.025364 | 0.027616 | 0.027963 | 0.028348 | 0.027674 | 0.026859 | 0.026684 | 0.02814 | 0.028966 | 0.029307 | 0.026327 | 0.023487 | 0.025638 | 0.025262 | 0.010631 | 0.025234 | b15 | 0.00045834 | 0.03113 | 0.025508 | 0.026261 | 0.026333 | 0.027901 | 0.027836 | 0.026746 | 0.029506 | 0.029106 | 0.028347 | 0.025583 | 0.025183 | 0.026715 | 0.027979 | 0.026823 | 0.024443 | 0.025045 | 0.028069 | 0.012742 | 0.025779
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 2.0803 | -5.8246 | -7.7114 | -6.4817 | -11.5483 | -4.9339 | -2.4637 | 6.2591 | 4.0608 | 7.4751 | 2.1455 | 12.8138 | -0.63452 | 9.9197 | -1.9456 | -1.5825 | -6.3238 | 1.4096 | -7.7971 | 18.6742 | 0 | a1 | 12.8974 | -11.381 | -5.4045 | -12.3958 | -2.7434 | -11.0029 | 6.3396 | -3.4105 | 12.1881 | -4.3007 | 3.3579 | -4.8339 | 13.5505 | 1.13 | 7.5181 | -6.8339 | -2.7848 | -2.5123 | -0.068468 | 23.9822 | 0
a2 | -1.9131 | -0.21539 | 0.87467 | 1.6588 | -0.26894 | -0.072117 | 0.72271 | 0.32151 | -0.066958 | 0.72813 | 0.24462 | -0.20722 | 0.45433 | -0.0032723 | -0.30911 | -0.30166 | 0.1217 | -0.49615 | -0.60306 | 4.699 | 0.22571 | a2 | 0.52486 | 2.0212 | 2.387 | 3.2688 | 2.0496 | 0.5877 | 1.5036 | 0.91947 | 2.0171 | 1.1189 | 1.0842 | 1.9361 | 0.87404 | 1.2463 | 1.8415 | 2.4367 | 1.0283 | 2.0417 | 1.6222 | 2.1785 | 1.645
a3 | -2.2351 | 0.50757 | 0.18065 | 0.67523 | 0.26848 | -0.053094 | 0.33103 | 0.1494 | 0.26011 | 0.18399 | 0.75121 | 0.41407 | -0.51884 | -0.37238 | -0.2056 | 0.15198 | 0.37539 | 0.80224 | 0.10804 | 0.064941 | 0.13612 | a3 | -1.446 | 1.0639 | 1.0175 | 0.66687 | 0.83576 | 0.68344 | 0.62615 | 0.4681 | 0.60822 | 0.84189 | 0.77301 | 0.63629 | 0.43148 | 0.56019 | 0.25153 | 0.43945 | 0.29401 | 0.49848 | 0.45748 | 0.66157 | 0.55258
a4 | -0.0020932 | 0.33719 | 0.28195 | -0.18662 | 0.030451 | 0.26222 | 0.2352 | 0.06374 | 0.36534 | 0.12417 | 0.2179 | 0.32013 | 0.35874 | 1.0349 | 0.74827 | 0.76941 | 0.3227 | 0.53574 | 0.57858 | 0.54052 | 0.34981 | a4 | 0.98242 | -0.39188 | -0.20129 | -0.0237 | 0.26465 | 0.3466 | -0.24501 | -0.2504 | -0.24491 | 0.47859 | 0.75322 | 0.58685 | 0.41178 | 0.38251 | -0.072863 | 0.059465 | -0.0061009 | 0.027701 | -0.26415 | -0.19168 | 0.10977
a5 | 2.031 | -0.017144 | -0.14295 | 0.28706 | 0.027878 | -0.023358 | 0.15149 | 0.1476 | 0.2052 | 0.1362 | 0.26884 | 0.16676 | -0.019895 | 0.068852 | -0.15466 | 0.021992 | -0.020193 | 0.16064 | 0.035817 | -0.29066 | 0.12515 | a5 | 2.2141 | 0.010036 | 0.1427 | 0.18978 | 0.006066 | -0.094844 | 0.10534 | 0.093784 | 0.14045 | -0.059575 | 0.084484 | 0.12819 | 0.056784 | -0.079027 | -0.1925 | -0.073611 | -0.096848 | 0.11778 | 0.014592 | -0.34616 | 0.087496
a6 | -0.16389 | 0.017225 | 0.10778 | -0.054259 | 0.018949 | 0.079879 | 0.11275 | 0.015334 | 0.066219 | 0.077304 | 0.14469 | 0.15291 | 0.12986 | 0.16016 | 0.22346 | 0.26153 | 0.18245 | 0.17148 | 0.17729 | 0.30478 | 0.11073 | a6 | -0.3358 | -0.13931 | -0.13669 | -0.030492 | -0.03786 | 0.0060733 | -0.076808 | -0.23223 | -0.20268 | -0.10768 | -0.05208 | -0.0042892 | -0.074101 | -0.036857 | -0.020187 | -0.039529 | -0.081806 | 0.044653 | -0.046794 | -0.31599 | -0.087396
a7 | 1.2976 | -0.05124 | -0.074891 | -0.072674 | -0.062333 | -0.047746 | -0.028042 | -0.077799 | -0.061119 | -0.080751 | -0.086232 | -0.023673 | -0.085508 | 0.023618 | -0.04601 | -0.020809 | -0.043885 | -0.028969 | -0.092348 | 0.022659 | -0.0060846 | a7 | 1.3218 | -0.042934 | -0.038035 | -0.016546 | 0.022799 | 0.012263 | 0.059187 | 0.0019931 | 0.039124 | -0.030774 | -0.033818 | -0.0079518 | -0.0044386 | -0.0071457 | -0.070566 | 0.007061 | -0.023549 | -0.058831 | -0.033844 | 0.017148 | 0.032636
a8 | -0.11334 | -0.046288 | -0.044494 | -0.026131 | -0.03915 | -0.02268 | -0.02086 | -0.046624 | -0.016124 | -0.010016 | 0.044163 | 0.037685 | 0.019164 | 0.035193 | 0.037794 | 0.066778 | 0.062201 | 0.043275 | -0.0070406 | 0.048099 | 0.0013016 | a8 | -0.077414 | -0.028592 | -0.038676 | -0.0073662 | -0.013337 | -0.037521 | -0.0094614 | -0.084104 | -0.071935 | -0.057949 | -0.057713 | -0.057196 | -0.096666 | -0.063065 | -0.092843 | -0.062005 | -0.057155 | 0.0032479 | -0.010971 | 0.024826 | -0.045489
a9 | -0.22297 | 0.01323 | -0.01794 | -0.004312 | 0.0039716 | -0.013963 | -0.014597 | -0.010153 | -0.012632 | -0.016879 | -0.00050739 | -0.011493 | -0.036135 | -0.016735 | -0.048037 | -0.0075616 | -0.035453 | -0.023187 | -0.037754 | -0.048624 | -0.024047 | a9 | -0.24901 | 0.028174 | 0.0068229 | -0.0047456 | 0.0068328 | -0.03882 | -0.026981 | -0.013747 | -0.023934 | -0.029049 | -0.0077133 | -0.020763 | -0.019985 | -0.034084 | -0.047108 | -0.032193 | -0.047754 | -0.024759 | -0.019219 | -0.067862 | -0.028603
a10 | -0.040511 | -0.021642 | 0.018922 | -0.019957 | 0.022719 | 0.057246 | 0.031548 | -0.046239 | -0.062655 | -0.021819 | 0.027741 | 0.032677 | -0.051887 | -0.035109 | -0.019183 | 0.040555 | 0.040851 | 0.043747 | 0.0084657 | -0.031684 | 0 | a10 | -0.10902 | -0.044009 | 0.021502 | 0.0068296 | 0.043852 | 0.01663 | 0.014047 | 0.0024128 | -0.051492 | -0.004795 | 0.028425 | 0.040151 | -0.012258 | -0.003954 | -0.024523 | 0.0088142 | -0.0096338 | 0.034181 | 0.011957 | -0.01334 | 0
a11 | 0.13376 | -0.043252 | -0.045618 | -0.046499 | -0.048607 | -0.034949 | -0.037239 | -0.030651 | -0.028864 | -0.030948 | -0.03768 | -0.028619 | -0.041161 | -0.027111 | -0.042642 | -0.033453 | -0.046287 | -0.036619 | -0.048783 | -0.019515 | -0.031959 | a11 | 0.13328 | -0.057455 | -0.04115 | -0.043857 | -0.035242 | -0.048197 | -0.028455 | -0.030905 | -0.023276 | -0.04222 | -0.02777 | -0.039088 | -0.028176 | -0.037798 | -0.040972 | -0.04723 | -0.042833 | -0.039953 | -0.036879 | -0.010338 | -0.031794
a12 | -0.012017 | -0.0044622 | 0.0036909 | -0.0052802 | -0.00095529 | 0.0064138 | 0.0064092 | -0.0068318 | -0.005817 | -0.00056309 | 0.0058989 | 0.0059453 | -0.0062248 | -0.0056667 | 0.0015904 | 0.0091251 | 0.0087184 | 0.0047546 | 0.0040753 | -0.0012824 | 0.00064013 | a12 | -0.033515 | -0.010074 | -0.0020888 | -0.0012059 | 0.0017953 | -0.00068421 | 0.0028008 | -0.003802 | -0.0093826 | -0.0050862 | -0.000495 | 0.0028584 | -0.0061809 | -0.0052003 | -0.0065256 | -0.002843 | -0.010314 | -0.0015163 | -0.0036548 | -0.01527 | -0.0048116
a13 | -0.028359 | -0.002753 | -0.0037314 | -0.0016547 | -0.0019443 | -0.0036647 | -0.0021582 | 0.0011655 | 0.00020549 | -0.00064293 | 0.0014133 | 0.00095881 | -0.0026565 | -0.0023105 | -0.0046374 | -0.0032275 | -0.004364 | -0.0033753 | -0.0027621 | -0.0074019 | -0.0030592 | a13 | -0.029787 | -0.00062635 | 0.00033546 | -0.0024879 | -0.0026051 | -0.005327 | -0.0030413 | -0.00073596 | -0.0014037 | -0.0047497 | -0.0047523 | -0.0050184 | -0.0025754 | -0.0052455 | -0.0075164 | -0.0060651 | -0.003249 | -0.0034462 | -0.0024245 | -0.0029878 | -0.0042467
a14 | 0.0021054 | -0.0065164 | -0.0044964 | -0.0092307 | -0.0078176 | -0.0057398 | -0.0067491 | -0.010214 | -0.0088713 | -0.0091439 | -0.0069614 | -0.0069826 | -0.010764 | -0.0081074 | -0.0083721 | -0.0046606 | -0.00468 | -0.0054043 | -0.0067404 | -0.0021518 | -0.006821 | a14 | 0.00040936 | -0.012977 | -0.006898 | -0.0070593 | -0.0065224 | -0.0080315 | -0.0069916 | -0.0086674 | -0.010944 | -0.0077451 | -0.0067398 | -0.0067077 | -0.010259 | -0.0089343 | -0.011608 | -0.01078 | -0.010768 | -0.0083679 | -0.010813 | -0.0032366 | -0.0084361
a15 | 0.0066545 | -0.0067268 | -0.0047481 | -0.0037946 | -0.0020505 | -0.0012496 | -0.00096189 | -0.003844 | -0.0045747 | -0.00047926 | -0.0006215 | -0.0014528 | -0.0066661 | -0.0037567 | -0.004678 | -0.0045735 | -0.0046525 | -0.0019243 | -0.0042345 | -0.0062109 | -0.0031488 | a15 | 0.004882 | -0.004219 | -0.0034436 | -0.0035125 | -0.001694 | -0.00068938 | -0.00011247 | -0.0029121 | -0.0047481 | -0.0033207 | -0.0033286 | -0.0050282 | -0.0081029 | -0.004852 | -0.004859 | -0.0060989 | -0.0078435 | -0.0036256 | -0.0020466 | -0.0055846 | -0.0036771
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00019689 | 9.381e-05 | 7.1547e-05 | 0.00011437 | -2.8975e-05 | -6.4051e-05 | -6.681e-05 | 3.8412e-05 | 1.333e-05 | -0.00015474 | -0.00012079 | -2.6035e-05 | 4.5262e-05 | -1.6925e-05 | 0 | 4.6219e-05 | 0.000125 | -7.3221e-05 | 4.9646e-05 | 0.00015425 | 0.00015231 | Dx (m) | -6.9331e-05 | -9.9599e-05 | -1.926e-05 | -6.655e-05 | -0.00010316 | -0.00015418 | -0.00013833 | -1.2652e-05 | 2.0695e-05 | -6.1983e-05 | 3.4519e-05 | 0.00010368 | 0.00015918 | 0 | -2.1445e-05 | 0.00010558 | 0.00022734 | 3.7163e-05 | -4.6401e-05 | 0.00018776 | 0.00014078
Dy (m) | 0.00014031 | 5.4868e-05 | -3.8835e-05 | 5.2192e-05 | -5.3164e-05 | -0.00013107 | -7.4034e-05 | 0.00010538 | 0.00014204 | 4.2956e-05 | -6.7706e-05 | -7.3979e-05 | 0.000119 | 7.8518e-05 | 4.323e-05 | -8.8984e-05 | -8.4708e-05 | -0.00010277 | -1.5979e-05 | 8.3183e-05 | 0.00031929 | Dy (m) | 0.00029281 | 9.4422e-05 | -5.119e-05 | -1.9549e-05 | -0.00010515 | -4.7731e-05 | -3.7281e-05 | 0 | 0.00011711 | 0 | -6.1984e-05 | -8.543e-05 | 3.3964e-05 | 0 | 5.438e-05 | -1.3355e-05 | 3.5407e-05 | -7.5979e-05 | -2.9564e-05 | 3.6109e-05 | 0.00038493
Magnetic length (m) |  | 14.4543 | [19 empty] | Magnetic length (m) |  | 14.4557
Transfer funct. (mT/KA) |  | 596.5098 | [19 empty] | Transfer funct. (mT/KA) |  | 596.4248
Coil positioning (mm) |  | 0.018006 | [19 empty] | Coil positioning (mm) |  | 0.01784
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMB__A001-3000036 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000036 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 380.9412 | 596.4706 | 596.4706 | 596.5882 | 596.4706 | 596.5882 | 596.5882 | 596.3529 | 596.3529 | 596.4706 | 596.4706 | 596.4706 | 596.4706 | 596.4706 | 596.4706 | 596.5882 | 596.5882 | 596.5882 | 596.7059 | 379.6471 | 0.073288 | C1/i (mT/KA) | 380.9412 | 596.3529 | 596.2353 | 596.3529 | 596.2353 | 596.3529 | 596.4706 | 596.3529 | 596.3529 | 596.3529 | 596.4706 | 596.4706 | 596.4706 | 596.3529 | 596.4706 | 596.4706 | 596.5882 | 596.5882 | 596.7059 | 380.7059 | 0.073285
Angle (mrad) | 0.20803 | -0.58246 | -0.77114 | -0.64817 | -1.1548 | -0.49339 | -0.24637 | 0.62591 | 0.40608 | 0.74751 | 0.21455 | 1.2814 | -0.063452 | 0.99196 | -0.19456 | -0.15825 | -0.63238 | 0.14096 | -0.77971 | 1.8674 | -3.984 | Angle (mrad) | 1.2897 | -1.1381 | -0.54045 | -1.2396 | -0.27434 | -1.1003 | 0.63396 | -0.34105 | 1.2188 | -0.43007 | 0.33579 | -0.48339 | 1.355 | 0.113 | 0.75181 | -0.68339 | -0.27848 | -0.25123 | -0.006847 | 2.3982 | -4.0092
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.617 | 0.90426 | 0.94342 | 1.4821 | 1.0794 | 0.38531 | 0.62073 | 0.22469 | 0.21156 | 0.60257 | -0.26692 | -0.41555 | 0.77901 | 1.0303 | 1.0185 | 0.89769 | 0.35801 | 0.33253 | -0.12837 | 0.68155 | 0.59794 | b2 | -3.2418 | 0.13026 | -0.23816 | -0.087807 | -0.84279 | 0.024661 | -0.49 | 0.20789 | 0.16629 | -0.36166 | -0.0018967 | 0.016656 | -0.18681 | -0.41155 | -0.30611 | -0.40551 | -0.013469 | 0.012913 | -0.41777 | 0.82187 | -0.24645
b3 | 39.0575 | 0.002401 | 0.84752 | -0.15513 | 2.1343 | 1.8604 | 2.0059 | 0.90355 | 1.5186 | 1.3765 | 1.2416 | 1.164 | 1.4841 | 0.51666 | 0.81239 | 1.5955 | 1.489 | 1.5207 | 2.2298 | 2.0915 | 2.533 | b3 | 37.9901 | -0.49039 | 0.42834 | 0.99692 | 1.3272 | 0.43904 | 1.2605 | 1.29 | 1.2302 | 0.95193 | -0.16976 | -0.17875 | 0.59975 | -0.38882 | 0.16803 | 1.119 | 1.2632 | 1.1963 | 1.9604 | 0.21266 | 1.9403
b4 | 0.45086 | -0.01611 | -0.14341 | 0.0096284 | -0.26239 | -0.079258 | 0.033279 | -0.26405 | -0.23353 | 0.039515 | 0.00027024 | -0.037186 | 0.13748 | -0.095713 | -0.11493 | 0.15231 | 0.042437 | 0.13855 | 0.081257 | -0.037381 | -0.018028 | b4 | -0.44996 | 0.11627 | -0.096371 | 0.10044 | -0.12286 | 0.29673 | 0.16109 | 0.22395 | 0.022556 | 0.1186 | 0.049329 | -0.019834 | 0.028299 | 0.16971 | 0.10578 | -0.14119 | -0.07147 | -0.15442 | -0.11318 | -0.077885 | 0.017438
b5 | -2.5709 | -0.032392 | -0.48231 | 0.24051 | 0.23004 | 0.39181 | 0.36945 | 0.22605 | 0.18739 | 0.31671 | 0.5774 | 0.22063 | 0.01253 | 0.028438 | 0.28467 | 0.24649 | 0.2026 | 0.42283 | 0.41255 | -1.9497 | 0.050484 | b5 | -2.4207 | 0.074911 | -0.40574 | 0.30119 | 0.40905 | 0.3887 | 0.35596 | 0.27377 | 0.41437 | 0.42787 | 1.0036 | 0.65786 | 0.50707 | 0.4353 | 0.428 | 0.43858 | 0.46277 | 0.62533 | 0.6686 | -1.8287 | 0.24664
b6 | 0.36437 | 0.039021 | -0.014761 | -0.19199 | -0.05341 | -0.026261 | -0.010385 | -0.032475 | 0.02276 | 0.064477 | 0.051568 | -0.035728 | 0.033483 | -0.070203 | -0.019911 | -0.073201 | -0.095713 | 0.045653 | 0.013366 | 0.051704 | -0.0045715 | b6 | -0.021278 | 0.067634 | 0.030284 | 0.052949 | 0.072885 | -0.0023354 | 0.033315 | -0.006572 | -0.071935 | 0.0795 | -0.031134 | -0.015302 | -0.1788 | 0.014662 | 0.011427 | -0.087649 | -0.18462 | -0.053895 | 0.0014 | -0.061887 | -0.016672
b7 | 3.0103 | 0.91134 | 0.86657 | 1.0589 | 0.99108 | 0.99509 | 0.94702 | 0.9281 | 0.95801 | 0.97262 | 0.9366 | 0.87042 | 0.81654 | 0.90693 | 0.89155 | 0.87723 | 0.92153 | 1.0137 | 1.0058 | 0.96177 | 1.0067 | b7 | 3.0925 | 0.92056 | 0.97905 | 1.0906 | 1.106 | 1.1411 | 1.0767 | 1.0661 | 1.0661 | 1.0453 | 1.061 | 1.0803 | 1.1148 | 1.1413 | 1.1722 | 1.0962 | 1.1178 | 1.1867 | 1.1173 | 0.89408 | 1.1477
b8 | 0.073097 | -0.015505 | -0.011544 | -0.010037 | 0.025379 | 0.031164 | 0.029772 | 0.022847 | 0.030966 | 0.053118 | 0.029886 | 0.015105 | 0.021578 | 0.015009 | 0.020802 | -0.0026579 | -0.0090368 | 0.051282 | -0.016672 | -0.07916 | 0.01441 | b8 | -0.074395 | 0.032877 | 0.0056066 | 0.0040251 | 0.013519 | 0.082362 | 0.054511 | 0.017697 | 0.026066 | 0.033758 | 0.0049403 | -0.026628 | -0.044356 | 0.018182 | 0.0051449 | -0.025828 | -0.0024833 | 0.014841 | 0.046878 | 0.0016765 | 0.011136
b9 | 0.52584 | 0.52884 | 0.48311 | 0.50763 | 0.52673 | 0.50969 | 0.49713 | 0.50911 | 0.53475 | 0.51804 | 0.50933 | 0.4938 | 0.50106 | 0.50069 | 0.5037 | 0.50005 | 0.50011 | 0.52245 | 0.53484 | 0.5613 | 0.51227 | b9 | 0.54721 | 0.52161 | 0.49188 | 0.51906 | 0.52569 | 0.49149 | 0.49862 | 0.50565 | 0.50792 | 0.49876 | 0.46772 | 0.45698 | 0.48883 | 0.49131 | 0.49598 | 0.49126 | 0.50643 | 0.51771 | 0.52784 | 0.5706 | 0.50415
b10 | 0.089889 | -0.042505 | -0.030291 | -0.052037 | 0.014354 | 0.031333 | 0.031426 | -0.019015 | -0.0083019 | 0.067674 | 0.055798 | 0.012903 | -0.023039 | 0.0062609 | -0.0030668 | -0.018852 | -0.05301 | 0.034515 | -0.021426 | -0.063053 | 0 | b10 | 0.050164 | 0.039716 | 0.0095029 | 0.029593 | 0.04732 | 0.069205 | 0.061391 | 0.0056563 | -0.010742 | 0.027095 | -0.014497 | -0.043879 | -0.070537 | -0.0015881 | 0.008157 | -0.046072 | -0.10022 | -0.014503 | 0.020986 | -0.076087 | 0
b11 | 0.68057 | 0.74494 | 0.74443 | 0.75209 | 0.75302 | 0.75957 | 0.75813 | 0.75702 | 0.75239 | 0.75213 | 0.76398 | 0.76091 | 0.75677 | 0.75422 | 0.75879 | 0.7573 | 0.75198 | 0.74974 | 0.74943 | 0.6847 | 0.74953 | b11 | 0.66219 | 0.73185 | 0.72957 | 0.74303 | 0.74386 | 0.74849 | 0.74694 | 0.7453 | 0.75124 | 0.74822 | 0.76403 | 0.75165 | 0.74739 | 0.75133 | 0.75038 | 0.74792 | 0.74292 | 0.74528 | 0.74546 | 0.68679 | 0.74162
b12 | 0.017478 | -0.0069328 | -0.0024533 | -0.00024506 | 0.002769 | 0.0039898 | 0.0041203 | -0.0021252 | 0.0028995 | 0.0042072 | 0.0018615 | 0.0014938 | -0.0011458 | 0.0019637 | 0.00021017 | -0.0025839 | -0.0047473 | 0.0023014 | -0.0062875 | -0.012245 | 0.00013817 | b12 | -0.00019328 | 0.002799 | -0.00092166 | -0.00049589 | 0.0062716 | 0.0095764 | 0.0078907 | -0.00037581 | 0.0050817 | 0.0039143 | 0.0048341 | -0.0029792 | -0.007315 | -0.0014301 | -0.00096404 | -0.0046077 | -0.0091901 | 6.4932e-06 | 0.0041104 | -0.0062753 | 0.0006263
b13 | 0.10449 | 0.075827 | 0.077207 | 0.070281 | 0.07298 | 0.073238 | 0.072824 | 0.071695 | 0.070016 | 0.07033 | 0.068945 | 0.072249 | 0.071396 | 0.071736 | 0.069525 | 0.07368 | 0.075208 | 0.07292 | 0.075449 | 0.065135 | 0.073343 | b13 | 0.10446 | 0.071831 | 0.076659 | 0.072193 | 0.072787 | 0.074017 | 0.076138 | 0.076425 | 0.073465 | 0.072999 | 0.072639 | 0.074674 | 0.074858 | 0.073736 | 0.071943 | 0.075111 | 0.072437 | 0.074432 | 0.074478 | 0.060545 | 0.074503
b14 | 0.0062452 | -0.0049606 | -0.0036154 | -0.0025839 | -0.0023886 | -0.0021537 | -0.0026887 | -0.0036374 | -0.0035086 | 1.2532e-05 | -0.00060799 | -0.0023194 | -0.0044911 | -0.0032204 | -0.0043319 | -0.0045987 | -0.005729 | -0.0023563 | -0.0056995 | -0.0062915 | -0.0030555 | b14 | 0.0075407 | -0.0025121 | -0.0037532 | -0.0028814 | -0.00099381 | 0.00056019 | 0.00028777 | -0.0035074 | -0.0035914 | -0.0032365 | -0.0042907 | -0.006647 | -0.0074714 | -0.0035614 | -0.0034561 | -0.0063898 | -0.0084495 | -0.0035097 | -0.0025695 | -0.0064187 | -0.0033852
b15 | -0.0039762 | 0.02652 | 0.025748 | 0.026882 | 0.025384 | 0.025364 | 0.027616 | 0.027963 | 0.028348 | 0.027674 | 0.026859 | 0.026684 | 0.02814 | 0.028966 | 0.029307 | 0.026327 | 0.023487 | 0.025638 | 0.025262 | 0.010631 | 0.025234 | b15 | 0.00045834 | 0.03113 | 0.025508 | 0.026261 | 0.026333 | 0.027901 | 0.027836 | 0.026746 | 0.029506 | 0.029106 | 0.028347 | 0.025583 | 0.025183 | 0.026715 | 0.027979 | 0.026823 | 0.024443 | 0.025045 | 0.028069 | 0.012742 | 0.025779
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 2.0803 | -5.8246 | -7.7114 | -6.4817 | -11.5483 | -4.9339 | -2.4637 | 6.2591 | 4.0608 | 7.4751 | 2.1455 | 12.8138 | -0.63452 | 9.9197 | -1.9456 | -1.5825 | -6.3238 | 1.4096 | -7.7971 | 18.6742 | 0 | a1 | 12.8974 | -11.381 | -5.4045 | -12.3958 | -2.7434 | -11.0029 | 6.3396 | -3.4105 | 12.1881 | -4.3007 | 3.3579 | -4.8339 | 13.5505 | 1.13 | 7.5181 | -6.8339 | -2.7848 | -2.5123 | -0.068468 | 23.9822 | 0
a2 | -1.9131 | -0.21539 | 0.87467 | 1.6588 | -0.26894 | -0.072117 | 0.72271 | 0.32151 | -0.066958 | 0.72813 | 0.24462 | -0.20722 | 0.45433 | -0.0032723 | -0.30911 | -0.30166 | 0.1217 | -0.49615 | -0.60306 | 4.699 | 0.22571 | a2 | 0.52486 | 2.0212 | 2.387 | 3.2688 | 2.0496 | 0.5877 | 1.5036 | 0.91947 | 2.0171 | 1.1189 | 1.0842 | 1.9361 | 0.87404 | 1.2463 | 1.8415 | 2.4367 | 1.0283 | 2.0417 | 1.6222 | 2.1785 | 1.645
a3 | -2.2351 | 0.50757 | 0.18065 | 0.67523 | 0.26848 | -0.053094 | 0.33103 | 0.1494 | 0.26011 | 0.18399 | 0.75121 | 0.41407 | -0.51884 | -0.37238 | -0.2056 | 0.15198 | 0.37539 | 0.80224 | 0.10804 | 0.064941 | 0.13612 | a3 | -1.446 | 1.0639 | 1.0175 | 0.66687 | 0.83576 | 0.68344 | 0.62615 | 0.4681 | 0.60822 | 0.84189 | 0.77301 | 0.63629 | 0.43148 | 0.56019 | 0.25153 | 0.43945 | 0.29401 | 0.49848 | 0.45748 | 0.66157 | 0.55258
a4 | -0.0020932 | 0.33719 | 0.28195 | -0.18662 | 0.030451 | 0.26222 | 0.2352 | 0.06374 | 0.36534 | 0.12417 | 0.2179 | 0.32013 | 0.35874 | 1.0349 | 0.74827 | 0.76941 | 0.3227 | 0.53574 | 0.57858 | 0.54052 | 0.34981 | a4 | 0.98242 | -0.39188 | -0.20129 | -0.0237 | 0.26465 | 0.3466 | -0.24501 | -0.2504 | -0.24491 | 0.47859 | 0.75322 | 0.58685 | 0.41178 | 0.38251 | -0.072863 | 0.059465 | -0.0061009 | 0.027701 | -0.26415 | -0.19168 | 0.10977
a5 | 2.031 | -0.017144 | -0.14295 | 0.28706 | 0.027878 | -0.023358 | 0.15149 | 0.1476 | 0.2052 | 0.1362 | 0.26884 | 0.16676 | -0.019895 | 0.068852 | -0.15466 | 0.021992 | -0.020193 | 0.16064 | 0.035817 | -0.29066 | 0.12515 | a5 | 2.2141 | 0.010036 | 0.1427 | 0.18978 | 0.006066 | -0.094844 | 0.10534 | 0.093784 | 0.14045 | -0.059575 | 0.084484 | 0.12819 | 0.056784 | -0.079027 | -0.1925 | -0.073611 | -0.096848 | 0.11778 | 0.014592 | -0.34616 | 0.087496
a6 | -0.16389 | 0.017225 | 0.10778 | -0.054259 | 0.018949 | 0.079879 | 0.11275 | 0.015334 | 0.066219 | 0.077304 | 0.14469 | 0.15291 | 0.12986 | 0.16016 | 0.22346 | 0.26153 | 0.18245 | 0.17148 | 0.17729 | 0.30478 | 0.11073 | a6 | -0.3358 | -0.13931 | -0.13669 | -0.030492 | -0.03786 | 0.0060733 | -0.076808 | -0.23223 | -0.20268 | -0.10768 | -0.05208 | -0.0042892 | -0.074101 | -0.036857 | -0.020187 | -0.039529 | -0.081806 | 0.044653 | -0.046794 | -0.31599 | -0.087396
a7 | 1.2976 | -0.05124 | -0.074891 | -0.072674 | -0.062333 | -0.047746 | -0.028042 | -0.077799 | -0.061119 | -0.080751 | -0.086232 | -0.023673 | -0.085508 | 0.023618 | -0.04601 | -0.020809 | -0.043885 | -0.028969 | -0.092348 | 0.022659 | -0.0060846 | a7 | 1.3218 | -0.042934 | -0.038035 | -0.016546 | 0.022799 | 0.012263 | 0.059187 | 0.0019931 | 0.039124 | -0.030774 | -0.033818 | -0.0079518 | -0.0044386 | -0.0071457 | -0.070566 | 0.007061 | -0.023549 | -0.058831 | -0.033844 | 0.017148 | 0.032636
a8 | -0.11334 | -0.046288 | -0.044494 | -0.026131 | -0.03915 | -0.02268 | -0.02086 | -0.046624 | -0.016124 | -0.010016 | 0.044163 | 0.037685 | 0.019164 | 0.035193 | 0.037794 | 0.066778 | 0.062201 | 0.043275 | -0.0070406 | 0.048099 | 0.0013016 | a8 | -0.077414 | -0.028592 | -0.038676 | -0.0073662 | -0.013337 | -0.037521 | -0.0094614 | -0.084104 | -0.071935 | -0.057949 | -0.057713 | -0.057196 | -0.096666 | -0.063065 | -0.092843 | -0.062005 | -0.057155 | 0.0032479 | -0.010971 | 0.024826 | -0.045489
a9 | -0.22297 | 0.01323 | -0.01794 | -0.004312 | 0.0039716 | -0.013963 | -0.014597 | -0.010153 | -0.012632 | -0.016879 | -0.00050739 | -0.011493 | -0.036135 | -0.016735 | -0.048037 | -0.0075616 | -0.035453 | -0.023187 | -0.037754 | -0.048624 | -0.024047 | a9 | -0.24901 | 0.028174 | 0.0068229 | -0.0047456 | 0.0068328 | -0.03882 | -0.026981 | -0.013747 | -0.023934 | -0.029049 | -0.0077133 | -0.020763 | -0.019985 | -0.034084 | -0.047108 | -0.032193 | -0.047754 | -0.024759 | -0.019219 | -0.067862 | -0.028603
a10 | -0.040511 | -0.021642 | 0.018922 | -0.019957 | 0.022719 | 0.057246 | 0.031548 | -0.046239 | -0.062655 | -0.021819 | 0.027741 | 0.032677 | -0.051887 | -0.035109 | -0.019183 | 0.040555 | 0.040851 | 0.043747 | 0.0084657 | -0.031684 | 0 | a10 | -0.10902 | -0.044009 | 0.021502 | 0.0068296 | 0.043852 | 0.01663 | 0.014047 | 0.0024128 | -0.051492 | -0.004795 | 0.028425 | 0.040151 | -0.012258 | -0.003954 | -0.024523 | 0.0088142 | -0.0096338 | 0.034181 | 0.011957 | -0.01334 | 0
a11 | 0.13376 | -0.043252 | -0.045618 | -0.046499 | -0.048607 | -0.034949 | -0.037239 | -0.030651 | -0.028864 | -0.030948 | -0.03768 | -0.028619 | -0.041161 | -0.027111 | -0.042642 | -0.033453 | -0.046287 | -0.036619 | -0.048783 | -0.019515 | -0.031959 | a11 | 0.13328 | -0.057455 | -0.04115 | -0.043857 | -0.035242 | -0.048197 | -0.028455 | -0.030905 | -0.023276 | -0.04222 | -0.02777 | -0.039088 | -0.028176 | -0.037798 | -0.040972 | -0.04723 | -0.042833 | -0.039953 | -0.036879 | -0.010338 | -0.031794
a12 | -0.012017 | -0.0044622 | 0.0036909 | -0.0052802 | -0.00095529 | 0.0064138 | 0.0064092 | -0.0068318 | -0.005817 | -0.00056309 | 0.0058989 | 0.0059453 | -0.0062248 | -0.0056667 | 0.0015904 | 0.0091251 | 0.0087184 | 0.0047546 | 0.0040753 | -0.0012824 | 0.00064013 | a12 | -0.033515 | -0.010074 | -0.0020888 | -0.0012059 | 0.0017953 | -0.00068421 | 0.0028008 | -0.003802 | -0.0093826 | -0.0050862 | -0.000495 | 0.0028584 | -0.0061809 | -0.0052003 | -0.0065256 | -0.002843 | -0.010314 | -0.0015163 | -0.0036548 | -0.01527 | -0.0048116
a13 | -0.028359 | -0.002753 | -0.0037314 | -0.0016547 | -0.0019443 | -0.0036647 | -0.0021582 | 0.0011655 | 0.00020549 | -0.00064293 | 0.0014133 | 0.00095881 | -0.0026565 | -0.0023105 | -0.0046374 | -0.0032275 | -0.004364 | -0.0033753 | -0.0027621 | -0.0074019 | -0.0030592 | a13 | -0.029787 | -0.00062635 | 0.00033546 | -0.0024879 | -0.0026051 | -0.005327 | -0.0030413 | -0.00073596 | -0.0014037 | -0.0047497 | -0.0047523 | -0.0050184 | -0.0025754 | -0.0052455 | -0.0075164 | -0.0060651 | -0.003249 | -0.0034462 | -0.0024245 | -0.0029878 | -0.0042467
a14 | 0.0021054 | -0.0065164 | -0.0044964 | -0.0092307 | -0.0078176 | -0.0057398 | -0.0067491 | -0.010214 | -0.0088713 | -0.0091439 | -0.0069614 | -0.0069826 | -0.010764 | -0.0081074 | -0.0083721 | -0.0046606 | -0.00468 | -0.0054043 | -0.0067404 | -0.0021518 | -0.006821 | a14 | 0.00040936 | -0.012977 | -0.006898 | -0.0070593 | -0.0065224 | -0.0080315 | -0.0069916 | -0.0086674 | -0.010944 | -0.0077451 | -0.0067398 | -0.0067077 | -0.010259 | -0.0089343 | -0.011608 | -0.01078 | -0.010768 | -0.0083679 | -0.010813 | -0.0032366 | -0.0084361
a15 | 0.0066545 | -0.0067268 | -0.0047481 | -0.0037946 | -0.0020505 | -0.0012496 | -0.00096189 | -0.003844 | -0.0045747 | -0.00047926 | -0.0006215 | -0.0014528 | -0.0066661 | -0.0037567 | -0.004678 | -0.0045735 | -0.0046525 | -0.0019243 | -0.0042345 | -0.0062109 | -0.0031488 | a15 | 0.004882 | -0.004219 | -0.0034436 | -0.0035125 | -0.001694 | -0.00068938 | -0.00011247 | -0.0029121 | -0.0047481 | -0.0033207 | -0.0033286 | -0.0050282 | -0.0081029 | -0.004852 | -0.004859 | -0.0060989 | -0.0078435 | -0.0036256 | -0.0020466 | -0.0055846 | -0.0036771
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.19689 | -0.09381 | -0.071547 | -0.11437 | 0.028975 | 0.064051 | 0.06681 | -0.038412 | -0.01333 | 0.15474 | 0.12079 | 0.026035 | -0.045262 | 0.016925 | 0 | -0.046219 | -0.125 | 0.073221 | -0.049646 | -0.15425 |  | Dx (mm) | 0.069331 | 0.099599 | 0.01926 | 0.06655 | 0.10316 | 0.15418 | 0.13833 | 0.012652 | -0.020695 | 0.061983 | -0.034519 | -0.10368 | -0.15918 | 0 | 0.021445 | -0.10558 | -0.22734 | -0.037163 | 0.046401 | -0.18776
Dy (mm) | -0.14031 | -0.054868 | 0.038835 | -0.052192 | 0.053164 | 0.13107 | 0.074034 | -0.10538 | -0.14204 | -0.042956 | 0.067706 | 0.073979 | -0.119 | -0.078518 | -0.04323 | 0.088984 | 0.084708 | 0.10277 | 0.015979 | -0.083183 |  | Dy (mm) | -0.29281 | -0.094422 | 0.05119 | 0.019549 | 0.10515 | 0.047731 | 0.037281 | 0 | -0.11711 | 0 | 0.061984 | 0.08543 | -0.033964 | 0 | -0.05438 | 0.013355 | -0.035407 | 0.075979 | 0.029564 | -0.036109
Mag. len. (m) | 14.4223 | [20 empty] | Mag. len. (m) | 14.4237
TF (mT/KA) | 596.5098 | [20 empty] | TF (mT/KA) | 596.4248
Coil pos. (mm) | 0.018006 | [20 empty] | Coil pos. (mm) | 0.01784
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5
Twist (m2 rad) | -0.014939 | [20 empty] | Twist (m2 rad) | -0.019707

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.59794 | 0.53879 | 0.55841 | 0.51163 | 0.22571 | 1.2612 | 0.13968 | 0.57559 | -0.24645 | 0.77713 | -0.1772 | 0.27949 | 1.645 | 0.70582 | 1.6629 | 0.68098 |  | 0
3 | 2.533 | 8.4694 | 1.2529 | 0.67243 | 0.13612 | 0.64442 | 0.22232 | 0.35551 | 1.9403 | 8.3665 | 0.72299 | 0.70022 | 0.55258 | 0.50901 | 0.62129 | 0.22468 |  | 0
4 | -0.018028 | 0.1653 | -0.036115 | 0.13324 | 0.34981 | 0.29186 | 0.35324 | 0.29192 | 0.017438 | 0.16851 | 0.039449 | 0.12976 | 0.10977 | 0.38505 | 0.087578 | 0.33684 |  | 0
5 | 0.050484 | 0.79883 | 0.21431 | 0.23206 | 0.12515 | 0.46534 | 0.072246 | 0.12446 | 0.24664 | 0.8315 | 0.41343 | 0.28191 | 0.087496 | 0.51045 | 0.027227 | 0.10727 |  | 0
6 | -0.0045715 | 0.10523 | -0.019001 | 0.053336 | 0.11073 | 0.10883 | 0.11149 | 0.075335 | -0.016672 | 0.073471 | -0.015217 | 0.048126 | -0.087396 | 0.10352 | -0.075095 | 0.061108 |  | 0
7 | 1.0067 | 0.46699 | 0.93702 | 0.061529 | -0.0060846 | 0.303 | -0.053531 | 0.030024 | 1.1477 | 0.45661 | 1.0869 | 0.063708 | 0.032636 | 0.29975 | -0.013188 | 0.033145 |  | 0
8 | 0.01441 | 0.032717 | 0.016191 | 0.01159 | 0.0013016 | 0.046812 | 0.0025351 | 0.038102 | 0.011136 | 0.034876 | 0.013689 | 0.01917 | -0.045489 | 0.03396 | -0.049542 | 0.026737 |  | 0
9 | 0.51227 | 0.018202 | 0.51027 | 0.014917 | -0.024047 | 0.048716 | -0.01688 | 0.014912 | 0.50415 | 0.025654 | 0.49879 | 0.019405 | -0.028603 | 0.055137 | -0.019744 | 0.020347 |  | 0
10 | 0 | 0.041892 | -0.0001106 | 0.00046728 | 0 | 0.037412 | 3.678e-06 | 2.8307e-05 | 0 | 0.047307 | -9.07e-05 | 0.00037754 | 0 | 0.035657 | -0.0002494 | 0.0013744 |  | 0
11 | 0.74953 | 0.022615 | 0.75441 | 0.0053942 | -0.031959 | 0.039094 | -0.038353 | 0.0072994 | 0.74162 | 0.023562 | 0.74622 | 0.0074703 | -0.031794 | 0.039406 | -0.03854 | 0.0084035 |  | 0
12 | 0.00013817 | 0.0059146 | 0.00011267 | 0.0024224 | 0.00064013 | 0.0061156 | 0.00080307 | 0.0024203 | 0.0006263 | 0.0050758 | 0.00076285 | 0.0026817 | -0.0048116 | 0.0080811 | -0.0040111 | 0.0020374 |  | 0
13 | 0.073343 | 0.007738 | 0.072514 | 0.0020819 | -0.0030592 | 0.0062145 | -0.0020154 | 0.0020134 | 0.074503 | 0.0077322 | 0.073767 | 0.0014598 | -0.0042467 | 0.0062272 | -0.0034019 | 0.0022781 |  | 0
14 | -0.0030555 | 0.0027097 | -0.0031827 | 0.0010715 | -0.006821 | 0.0029538 | -0.0074536 | 0.0010911 | -0.0033852 | 0.003462 | -0.0037761 | 0.00065425 | -0.0084361 | 0.0030581 | -0.0092764 | 0.0014005 |  | 0
15 | 0.025234 | 0.0077306 | 0.026787 | 0.0014978 | -0.0031488 | 0.0030045 | -0.0033883 | 0.0019598 | 0.025779 | 0.0068315 | 0.02714 | 0.0017473 | -0.0036771 | 0.0028526 | -0.0039132 | 0.0021282 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.61127 | -0.8295
3 | [12 empty] | 3 | 9 | -0.71553 | -0.92471
4 | [12 empty] | 4 | 16 | -1.6234 | -1.5651
5 | [12 empty] | 5 | 25 | -1.7722 | -1.7493
6 | [12 empty] | 6 | 36 | -2.7585 | -2.9145
7 | [12 empty] | 7 | 49 | -3.147 | -3.0802
8 | [12 empty] | 8 | 64 | -3.8626 | -3.7881
9 | [12 empty] | 9 | 81 | -4.2054 | -3.9185
10 | [12 empty] | sigma (mm) |  | 0.018006 | 0.01784
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | 0.74916 | 0.90101 | 0.95138 | 1.4758 | 1.0738 | 0.37048 | 0.60785 | 0.22693 | 0.20959 | 0.57659 | -0.27858 | -0.41552 | 0.79418 | 1.0327 | 1.0195 | 0.90796 | 0.38365 | 0.32913 | -0.11515 | 0.71887 | 0.59794
3 | 39.0419 | -0.0011306 | 0.84764 | -0.15322 | 2.136 | 1.8673 | 2.0086 | 0.90058 | 1.5089 | 1.3745 | 1.2442 | 1.1683 | 1.4777 | 0.50261 | 0.80668 | 1.6088 | 1.4947 | 1.5286 | 2.2321 | 2.0825 | 2.533
4 | 0.50291 | -0.016604 | -0.15283 | 0.012575 | -0.26361 | -0.085883 | 0.03011 | -0.26567 | -0.2398 | 0.026607 | -0.011857 | -0.035635 | 0.13817 | -0.097098 | -0.11336 | 0.15545 | 0.047993 | 0.13515 | 0.08621 | -0.10245 | -0.018028
5 | -2.5852 | -0.031593 | -0.48139 | 0.23488 | 0.23079 | 0.39538 | 0.37211 | 0.22521 | 0.18472 | 0.3128 | 0.57844 | 0.22423 | 0.0084304 | 0.025089 | 0.28183 | 0.25234 | 0.20363 | 0.42703 | 0.41358 | -1.9548 | 0.050484
6 | 0.090918 | 0.070187 | 0.0060952 | -0.14791 | -0.064714 | -0.050965 | -0.033449 | -0.017 | 0.030331 | 0.012582 | 0.0095797 | -0.044344 | 0.050118 | -0.076275 | -0.019209 | -0.059545 | -0.05637 | 0.018405 | 0.030468 | 0.1034 | -0.0045715
7 | 3.0109 | 0.91179 | 0.86552 | 1.059 | 0.98992 | 0.99304 | 0.94557 | 0.93048 | 0.95912 | 0.96941 | 0.93634 | 0.87141 | 0.816 | 0.90569 | 0.89088 | 0.87963 | 0.92323 | 1.014 | 1.0054 | 0.95509 | 1.0067
8 | 0.039098 | 0.0074998 | 0.0043936 | 0.017389 | 0.018297 | 0.01494 | 0.013634 | 0.032553 | 0.035164 | 0.015735 | 0.00091931 | 0.0086553 | 0.034273 | 0.01164 | 0.021779 | 0.0079016 | 0.018968 | 0.032159 | -0.0044604 | -0.036512 | 0.01441
9 | 0.51948 | 0.52735 | 0.48235 | 0.50503 | 0.52715 | 0.5126 | 0.49725 | 0.5113 | 0.53941 | 0.51299 | 0.50675 | 0.4949 | 0.50378 | 0.50209 | 0.50414 | 0.5015 | 0.49843 | 0.52349 | 0.53435 | 0.55754 | 0.51227
10 | 2.7036e-05 | -1.1845e-06 | -2.2778e-06 | -1.3438e-05 | -2.3535e-07 | 2.0361e-05 | 1.0314e-05 | -1.0524e-05 | 9.3448e-06 | -7.373e-06 | 1.4959e-05 | 4.4108e-06 | -9.4333e-06 | 4.1158e-06 | -0.0019824 | -1.3853e-05 | -2.4379e-05 | 9.6895e-06 | 1.1384e-06 | 2.9927e-05 | 0
11 | 0.67943 | 0.74467 | 0.74441 | 0.75225 | 0.75293 | 0.75995 | 0.75826 | 0.75743 | 0.75295 | 0.75173 | 0.76409 | 0.76117 | 0.75722 | 0.75449 | 0.75875 | 0.75774 | 0.75207 | 0.74994 | 0.74927 | 0.68355 | 0.74953
12 | 0.057642 | -0.01805 | 0.013437 | 0.054896 | 0.012034 | 0.0033948 | 0.0057317 | -0.0026791 | 0.035377 | -0.034555 | -0.039493 | 0.002161 | 0.013584 | 0.012347 | 0.0035168 | -0.0038284 | 0.016276 | -0.017124 | -0.036746 | -0.047182 | 0.0013817 | ten times!
13 | 1.0333 | 0.75745 | 0.76875 | 0.70424 | 0.72715 | 0.72768 | 0.72579 | 0.72411 | 0.70944 | 0.70629 | 0.68641 | 0.719 | 0.72221 | 0.72264 | 0.69802 | 0.732 | 0.74358 | 0.72628 | 0.7515 | 0.64529 | 0.73343 | ten times!
14 | 0.06121 | -0.026079 | -0.022501 | 0.0011113 | -0.030841 | -0.036265 | -0.042668 | -0.024192 | -0.026623 | -0.034971 | -0.033143 | -0.0298 | -0.02789 | -0.033812 | -0.041654 | -0.039318 | -0.036358 | -0.040651 | -0.047224 | -0.045156 | -0.030555 | ten times!
15 | -0.039762 | 0.2652 | 0.25748 | 0.26882 | 0.25384 | 0.25364 | 0.27616 | 0.27963 | 0.28348 | 0.27674 | 0.26859 | 0.26684 | 0.2814 | 0.28966 | 0.29307 | 0.26327 | 0.23487 | 0.25638 | 0.25262 | 0.10631 | 0.25234 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.2166 | -0.20977 | 0.87232 | 1.6669 | -0.28321 | -0.1004 | 0.70264 | 0.33339 | -0.041173 | 0.73174 | 0.22405 | -0.21862 | 0.47234 | 0.0022418 | -0.30498 | -0.31754 | 0.11239 | -0.52145 | -0.60663 | 4.7206 | 0.22571
3 | -2.2238 | 0.51299 | 0.18519 | 0.67155 | 0.27079 | -0.054225 | 0.32783 | 0.14492 | 0.25511 | 0.1809 | 0.74656 | 0.41309 | -0.51308 | -0.3768 | -0.20648 | 0.15587 | 0.38187 | 0.79281 | 0.11288 | 0.079106 | 0.13612
4 | -0.18106 | 0.33639 | 0.28395 | -0.17594 | 0.027383 | 0.25048 | 0.22638 | 0.070679 | 0.37224 | 0.12241 | 0.20106 | 0.31527 | 0.35888 | 1.0351 | 0.75117 | 0.76449 | 0.31807 | 0.52275 | 0.57745 | 0.49181 | 0.34981
5 | 2.0555 | -0.016039 | -0.14051 | 0.28229 | 0.028552 | -0.023851 | 0.1495 | 0.14676 | 0.20641 | 0.13349 | 0.26267 | 0.16636 | -0.016994 | 0.066434 | -0.15491 | 0.027463 | -0.0111 | 0.15556 | 0.038343 | -0.27556 | 0.12515
6 | -0.10499 | 0.033177 | 0.094013 | -0.037686 | 0.00098983 | 0.034927 | 0.088671 | 0.048797 | 0.11396 | 0.096461 | 0.12598 | 0.1304 | 0.16279 | 0.18515 | 0.23706 | 0.23364 | 0.15296 | 0.13546 | 0.17 | 0.33425 | 0.11073
7 | 1.311 | -0.053378 | -0.076017 | -0.07412 | -0.062421 | -0.04883 | -0.028376 | -0.077545 | -0.059396 | -0.079174 | -0.089262 | -0.024538 | -0.084093 | 0.023858 | -0.045639 | -0.019441 | -0.040368 | -0.032444 | -0.092382 | 0.023003 | -0.0060846
8 | -0.057964 | -0.032049 | -0.053927 | -0.013895 | -0.052382 | -0.053696 | -0.037721 | -0.021561 | 0.019541 | 0.0016857 | 0.027964 | 0.020635 | 0.046454 | 0.053826 | 0.048041 | 0.045674 | 0.04018 | 0.018807 | -0.011944 | 0.066541 | 0.0013016
9 | -0.21207 | 0.01092 | -0.0166 | -0.0069582 | 0.0032191 | -0.018078 | -0.016945 | -0.012154 | -0.013698 | -0.013552 | -0.0042814 | -0.012449 | -0.03883 | -0.01616 | -0.048107 | -0.0056812 | -0.030373 | -0.02676 | -0.037351 | -0.053988 | -0.024047
10 | 0.00016911 | 1.4139e-05 | -3.2608e-06 | 5.0049e-06 | -2.0177e-06 | 1.6121e-06 | -4.5277e-06 | -6.9191e-06 | -1.5479e-05 | 3.4682e-06 | -8.8329e-06 | 3.4922e-06 | -9.1719e-06 | -3.9767e-06 | 0.00011221 | 5.4137e-06 | -2.1976e-05 | 2.4221e-07 | -3.2147e-06 | 4.8449e-05 | 0
11 | 0.13687 | -0.043769 | -0.045385 | -0.046898 | -0.048684 | -0.035553 | -0.037713 | -0.030966 | -0.028647 | -0.030775 | -0.038222 | -0.028791 | -0.041431 | -0.026949 | -0.042637 | -0.033032 | -0.045322 | -0.036997 | -0.048587 | -0.020302 | -0.031959 | ten times!
12 | 0.022739 | -0.017077 | 0.01386 | -0.028245 | -0.036543 | -0.0019627 | 0.027052 | -0.014673 | 0.012051 | 0.016397 | 0.024833 | 0.021548 | -0.0031245 | -0.016632 | 0.037119 | 0.043918 | 0.038364 | -0.0036081 | 0.031275 | 0.017362 | 0.0064013 | ten times!
13 | -0.28006 | -0.034286 | -0.038701 | -0.025651 | -0.016739 | -0.031677 | -0.016612 | 0.005724 | -0.0026605 | 0.0043951 | 0.020878 | 0.01229 | -0.034378 | -0.02399 | -0.047806 | -0.030793 | -0.044402 | -0.028875 | -0.029484 | -0.08056 | -0.030592 | ten times!
14 | 0.0056695 | -0.058378 | -0.055996 | -0.084326 | -0.0888 | -0.084116 | -0.083799 | -0.079093 | -0.056054 | -0.081038 | -0.083972 | -0.085771 | -0.08255 | -0.06182 | -0.073287 | -0.067639 | -0.067974 | -0.074581 | -0.072459 | -0.022125 | -0.06821 | ten times!
15 | 0.066545 | -0.067268 | -0.047481 | -0.037946 | -0.020505 | -0.012496 | -0.0096189 | -0.03844 | -0.045747 | -0.0047926 | -0.006215 | -0.014528 | -0.066661 | -0.037567 | -0.04678 | -0.045735 | -0.046525 | -0.019243 | -0.042345 | -0.062109 | -0.031488 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -3.5019 | 0.12419 | -0.233 | -0.094078 | -0.84855 | 0.020536 | -0.50777 | 0.20597 | 0.16091 | -0.3686 | 0.0030509 | 0.020871 | -0.17731 | -0.41155 | -0.30814 | -0.39092 | 0.019093 | 0.022599 | -0.42688 | 0.82376 | -0.24645
3 | 37.9448 | -0.4859 | 0.42685 | 0.99566 | 1.3344 | 0.43388 | 1.255 | 1.2895 | 1.2353 | 0.95063 | -0.16122 | -0.17026 | 0.59807 | -0.38882 | 0.16833 | 1.1165 | 1.2604 | 1.1957 | 1.96 | 0.2113 | 1.9403
4 | -0.56302 | 0.11429 | -0.092814 | 0.096596 | -0.13264 | 0.28157 | 0.15043 | 0.22313 | 0.020704 | 0.11236 | 0.058712 | -0.0012087 | 0.046837 | 0.16971 | 0.10608 | -0.13053 | -0.045908 | -0.14684 | -0.12038 | -0.15573 | 0.017438
5 | -2.3914 | 0.076798 | -0.40796 | 0.29998 | 0.40567 | 0.38889 | 0.35377 | 0.2738 | 0.42092 | 0.42642 | 1.0023 | 0.65729 | 0.49944 | 0.4353 | 0.42825 | 0.4357 | 0.45128 | 0.62574 | 0.66818 | -1.8288 | 0.24664
6 | -0.23355 | 0.036705 | 0.022942 | 0.027219 | 0.033463 | -0.064224 | -0.018477 | -0.011333 | -0.065765 | 0.056633 | -0.018948 | 0.023987 | -0.11611 | 0.014662 | 0.0039093 | -0.046766 | -0.094635 | -0.039907 | -0.017252 | -0.0028575 | -0.016672
7 | 3.1039 | 0.92032 | 0.97819 | 1.0904 | 1.1048 | 1.1351 | 1.0735 | 1.066 | 1.0698 | 1.0444 | 1.0596 | 1.0771 | 1.1132 | 1.1413 | 1.1742 | 1.0948 | 1.1184 | 1.187 | 1.1163 | 0.89384 | 1.1477
8 | -0.057936 | 0.0071778 | 0.0013128 | -0.012275 | -0.011663 | 0.04583 | 0.021578 | 0.014686 | 0.032331 | 0.01921 | 0.012313 | -0.0051671 | -0.0074184 | 0.018182 | 0.0013452 | -0.001622 | 0.052493 | 0.023009 | 0.035085 | 0.053245 | 0.011136
9 | 0.56227 | 0.52171 | 0.49236 | 0.51809 | 0.52555 | 0.48626 | 0.49441 | 0.50561 | 0.511 | 0.49787 | 0.46838 | 0.45639 | 0.48311 | 0.49131 | 0.49659 | 0.48874 | 0.49454 | 0.5188 | 0.52751 | 0.56329 | 0.50415
10 | 0.0002019 | 3.0086e-05 | -1.674e-06 | 5.1612e-07 | 5.6158e-07 | -3.1665e-05 | -1.38e-05 | 0.00010949 | 7.0645e-06 | -0.00018514 | 3.9998e-06 | -3.3295e-06 | 8.963e-06 | -0.0015881 | 1.8864e-06 | 7.4835e-06 | 2.047e-05 | 3.5934e-06 | -2.9658e-06 | -1.5307e-05 | 0
11 | 0.66855 | 0.73228 | 0.72951 | 0.74304 | 0.74356 | 0.74751 | 0.7463 | 0.7453 | 0.75202 | 0.74806 | 0.76412 | 0.75161 | 0.74677 | 0.75133 | 0.75062 | 0.74758 | 0.7418 | 0.74521 | 0.74527 | 0.68638 | 0.74162
12 | 0.0085121 | -0.022093 | -0.019517 | -0.039216 | 0.0077791 | 0.013413 | 0.0037599 | -0.010584 | 0.062709 | 0.0072047 | 0.063961 | 0.02183 | 0.011581 | -0.014301 | -0.017646 | 0.0093297 | 0.025157 | 0.017742 | 0.016203 | 0.018251 | 0.006263 | ten times!
13 | 1.0397 | 0.7296 | 0.76445 | 0.72234 | 0.72341 | 0.73658 | 0.75908 | 0.76459 | 0.74388 | 0.73152 | 0.72206 | 0.73708 | 0.74215 | 0.73736 | 0.72483 | 0.74485 | 0.71259 | 0.73846 | 0.74325 | 0.59713 | 0.74503 | ten times!
14 | 0.063373 | -0.047373 | -0.04303 | -0.043773 | -0.033776 | -0.030095 | -0.028868 | -0.03786 | -0.026306 | -0.047222 | -0.036548 | -0.048165 | -0.039435 | -0.035614 | -0.037327 | -0.041247 | -0.036444 | -0.0297 | -0.036919 | -0.042825 | -0.033852 | ten times!
15 | 0.0045834 | 0.3113 | 0.25508 | 0.26261 | 0.26333 | 0.27901 | 0.27836 | 0.26746 | 0.29506 | 0.29106 | 0.28347 | 0.25583 | 0.25183 | 0.26715 | 0.27979 | 0.26823 | 0.24443 | 0.25045 | 0.28069 | 0.12742 | 0.25779 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 1.8454 | 2.0033 | 2.3822 | 3.2613 | 2.023 | 0.57284 | 1.4879 | 0.91878 | 2.0355 | 1.1127 | 1.0886 | 1.9457 | 0.88452 | 1.2463 | 1.842 | 2.4405 | 1.0415 | 2.0332 | 1.6128 | 2.194 | 1.645
3 | -1.4812 | 1.0727 | 1.0191 | 0.6668 | 0.83322 | 0.67151 | 0.63107 | 0.46866 | 0.60779 | 0.83666 | 0.77706 | 0.64732 | 0.44322 | 0.56019 | 0.25282 | 0.44089 | 0.29331 | 0.50073 | 0.46023 | 0.65472 | 0.55258
4 | 0.77952 | -0.39045 | -0.19705 | -0.028057 | 0.25438 | 0.34567 | -0.25156 | -0.25068 | -0.23281 | 0.47945 | 0.73927 | 0.57676 | 0.41796 | 0.38251 | -0.066416 | 0.056258 | -0.0074262 | 0.017551 | -0.26896 | -0.22251 | 0.10977
5 | 2.2191 | 0.015995 | 0.14302 | 0.19007 | 0.0049607 | -0.095086 | 0.1081 | 0.094648 | 0.13674 | -0.057611 | 0.084523 | 0.12845 | 0.051529 | -0.079027 | -0.19219 | -0.074494 | -0.10424 | 0.11948 | 0.015218 | -0.36427 | 0.087496
6 | -0.048551 | -0.10712 | -0.15412 | -0.037628 | -0.079734 | -0.013817 | -0.093865 | -0.23224 | -0.15832 | -0.10701 | -0.075703 | -0.037152 | -0.060988 | -0.036857 | 0.0028445 | -0.044433 | -0.069728 | 0.012059 | -0.057899 | -0.30346 | -0.087396
7 | 1.315 | -0.040483 | -0.037846 | -0.016376 | 0.02278 | 0.013027 | 0.058889 | 0.0024313 | 0.039768 | -0.029295 | -0.034764 | -0.0094568 | -0.011395 | -0.0071457 | -0.069631 | 0.0045074 | -0.028936 | -0.059245 | -0.034205 | 0.019092 | 0.032636
8 | 0.0061125 | -0.0067352 | -0.050587 | -0.011993 | -0.039681 | -0.045744 | -0.016453 | -0.084022 | -0.044175 | -0.057101 | -0.071481 | -0.076581 | -0.09035 | -0.063065 | -0.079676 | -0.066692 | -0.053826 | -0.015696 | -0.017895 | 0.028526 | -0.045489
9 | -0.23723 | 0.03248 | 0.0063461 | -0.0052925 | 0.0018037 | -0.041926 | -0.029222 | -0.013763 | -0.025164 | -0.028892 | -0.0067181 | -0.016574 | -0.022287 | -0.034084 | -0.046595 | -0.031375 | -0.050792 | -0.023503 | -0.019842 | -0.070643 | -0.028603
10 | -0.00039682 | 5.6446e-06 | -8.1733e-07 | 1.9653e-06 | -1.8327e-05 | -1.3727e-05 | -1.8106e-05 | 0.0026428 | -2.3689e-05 | -0.0032557 | 3.5538e-06 | -5.0595e-06 | 3.541e-05 | -0.003954 | -3.1948e-06 | 5.2645e-06 | 0.00011117 | -1.7609e-06 | -7.0168e-07 | 0.00010544 | 0
11 | 0.13475 | -0.056635 | -0.041093 | -0.043798 | -0.035788 | -0.048424 | -0.028896 | -0.030874 | -0.023017 | -0.042016 | -0.027975 | -0.038731 | -0.028973 | -0.037798 | -0.040916 | -0.047385 | -0.044561 | -0.03999 | -0.036848 | -0.01234 | -0.031794
12 | -0.10467 | -0.052426 | -0.048634 | -0.020853 | -0.034174 | -0.025983 | 0.010942 | -0.037954 | -0.0333 | -0.048784 | -0.03789 | -0.020119 | -0.046757 | -0.052003 | -0.036502 | -0.040031 | -0.09025 | -0.055987 | -0.051296 | -0.14123 | -0.048116 | ten times!
13 | -0.2812 | 0.0018061 | 0.0058398 | -0.020855 | -0.020107 | -0.044005 | -0.023099 | -0.006521 | -0.018985 | -0.043826 | -0.047269 | -0.05116 | -0.040183 | -0.052455 | -0.074698 | -0.068703 | -0.053498 | -0.034801 | -0.019827 | -0.036297 | -0.042467 | ten times!
14 | 0.0024115 | -0.1021 | -0.079187 | -0.072896 | -0.086587 | -0.090407 | -0.078334 | -0.08637 | -0.081797 | -0.075755 | -0.082814 | -0.089368 | -0.10617 | -0.089343 | -0.1027 | -0.11606 | -0.11524 | -0.10046 | -0.11418 | -0.037212 | -0.084361 | ten times!
15 | 0.04882 | -0.04219 | -0.034436 | -0.035125 | -0.01694 | -0.0068938 | -0.0011247 | -0.029121 | -0.047481 | -0.033207 | -0.033286 | -0.050282 | -0.081029 | -0.04852 | -0.04859 | -0.060989 | -0.078435 | -0.036256 | -0.020466 | -0.055846 | -0.036771 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -6.5742e-05 | -6.5742e-05 | 0.00013148 | -6.5742e-05 | 0.00013148 | 0.00013148 | -0.00026297 | -0.00026297 | -6.5742e-05 | -6.5742e-05 | -6.5742e-05 | -6.5742e-05 | -6.5742e-05 | -6.5742e-05 | 0.00013148 | 0.00013148 | 0.00013148 | 0.00032871 |  | 596.5098
|  | -0.00012054 | -0.0003178 | -0.00012054 | -0.0003178 | -0.00012054 | 7.671e-05 | -0.00012054 | -0.00012054 | -0.00012054 | 7.671e-05 | 7.671e-05 | 7.671e-05 | -0.00012054 | 7.671e-05 | 7.671e-05 | 0.00027396 | 0.00027396 | 0.00047122 |  | 596.4248
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3624.5747
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.5 | -3613.7271 | -0.49307 | -0.49307 | 1.4792 | -0.49307 | 1.4792 | 1.4792 | -2.4654 | -2.4654 | -0.49307 | -0.49307 | -0.49307 | -0.49307 | -0.49307 | -0.49307 | 1.4792 | 1.4792 | 1.4792 | 3.4515 | -3635.4223
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.073131 | 0.28116 | -0.50933 | -0.69801 | -0.57504 | -1.0817 | -0.42026 | -0.17324 | 0.69904 | 0.47921 | 0.82064 | 0.28768 | 1.3545 | 0.0096787 | 1.0651 | -0.12143 | -0.08512 | -0.55925 | 0.21409 | -0.70658 | 1.9406
db2 | 0.55884 | 1.617 | 0.34542 | 0.38457 | 0.92322 | 0.52053 | -0.17353 | 0.061891 | -0.33415 | -0.34728 | 0.043732 | -0.82576 | -0.9744 | 0.22017 | 0.47147 | 0.45962 | 0.33885 | -0.20083 | -0.22632 | -0.68721 | 0.68155
db3 | 1.2527 | 39.0575 | -1.2503 | -0.40513 | -1.4078 | 0.88169 | 0.60772 | 0.75326 | -0.3491 | 0.26596 | 0.12385 | -0.011092 | -0.088697 | 0.23149 | -0.73599 | -0.44027 | 0.34287 | 0.23633 | 0.26803 | 0.97712 | 2.0915
db4 | -0.033991 | 0.45086 | 0.017881 | -0.10942 | 0.04362 | -0.2284 | -0.045266 | 0.067271 | -0.23006 | -0.19953 | 0.073506 | 0.034262 | -0.0031947 | 0.17147 | -0.061722 | -0.080937 | 0.1863 | 0.076428 | 0.17254 | 0.11525 | -0.037381
db5 | 0.21419 | -2.5709 | -0.24658 | -0.6965 | 0.026319 | 0.015853 | 0.17762 | 0.15526 | 0.011863 | -0.026797 | 0.10252 | 0.36321 | 0.0064442 | -0.20166 | -0.18575 | 0.070481 | 0.032301 | -0.011585 | 0.20864 | 0.19837 | -1.9497
db6 | -0.01965 | 0.36437 | 0.058672 | 0.0048895 | -0.17234 | -0.033759 | -0.0066107 | 0.0092651 | -0.012825 | 0.04241 | 0.084127 | 0.071219 | -0.016077 | 0.053133 | -0.050553 | -0.00026017 | -0.053551 | -0.076062 | 0.065303 | 0.033017 | 0.051704
db7 | 0.93717 | 3.0103 | -0.025827 | -0.070597 | 0.12174 | 0.053911 | 0.05792 | 0.0098563 | -0.009068 | 0.02084 | 0.035452 | -0.00056591 | -0.066743 | -0.12063 | -0.030236 | -0.045614 | -0.059936 | -0.015638 | 0.076546 | 0.068592 | 0.96177
db8 | 0.015636 | 0.073097 | -0.031141 | -0.027181 | -0.025673 | 0.0097429 | 0.015528 | 0.014136 | 0.0072104 | 0.015329 | 0.037482 | 0.01425 | -0.00053106 | 0.0059411 | -0.0006273 | 0.0051655 | -0.018294 | -0.024673 | 0.035646 | -0.032308 | -0.07916
db9 | 0.51006 | 0.52584 | 0.018777 | -0.02695 | -0.0024314 | 0.016669 | -0.00036455 | -0.012932 | -0.00094675 | 0.024697 | 0.0079831 | -0.00073035 | -0.016261 | -0.0089977 | -0.0093685 | -0.0063577 | -0.010006 | -0.0099489 | 0.01239 | 0.024779 | 0.5613
db10 | -0.00095993 | 0.089889 | -0.041545 | -0.029331 | -0.051077 | 0.015314 | 0.032293 | 0.032386 | -0.018055 | -0.0073419 | 0.068634 | 0.056758 | 0.013863 | -0.022079 | 0.0072208 | -0.0021069 | -0.017892 | -0.05205 | 0.035475 | -0.020466 | -0.063053
db11 | 0.75427 | 0.68057 | -0.0093312 | -0.0098417 | -0.0021753 | -0.0012505 | 0.0053005 | 0.0038598 | 0.0027504 | -0.0018789 | -0.0021339 | 0.0097116 | 0.0066377 | 0.0025062 | -4.83e-05 | 0.0045243 | 0.0030275 | -0.0022893 | -0.0045308 | -0.0048381 | 0.6847
db12 | -3.9138e-05 | 0.017478 | -0.0068936 | -0.0024142 | -0.00020592 | 0.0028081 | 0.0040289 | 0.0041595 | -0.002086 | 0.0029386 | 0.0042463 | 0.0019007 | 0.001533 | -0.0011067 | 0.0020028 | 0.00024931 | -0.0025448 | -0.0047082 | 0.0023405 | -0.0062483 | -0.012245
db13 | 0.072528 | 0.10449 | 0.0032992 | 0.0046786 | -0.0022467 | 0.00045178 | 0.00070991 | 0.00029597 | -0.0008332 | -0.0025123 | -0.002198 | -0.0035832 | -0.00027946 | -0.0011316 | -0.00079239 | -0.0030033 | 0.0011515 | 0.0026802 | 0.00039238 | 0.0029205 | 0.065135
db14 | -0.003271 | 0.0062452 | -0.0016896 | -0.00034434 | 0.00068718 | 0.00088244 | 0.0011173 | 0.00058234 | -0.00036633 | -0.00023756 | 0.0032836 | 0.002663 | 0.00095166 | -0.0012201 | 5.0675e-05 | -0.0010609 | -0.0013277 | -0.0024579 | 0.00091478 | -0.0024285 | -0.0062915
db15 | 0.026787 | -0.0039762 | -0.00026759 | -0.0010397 | 9.5298e-05 | -0.0014035 | -0.0014229 | 0.0008289 | 0.0011761 | 0.0015612 | 0.00088646 | 7.1438e-05 | -0.0001029 | 0.0013526 | 0.0021789 | 0.0025197 | -0.00045972 | -0.0032999 | -0.0011493 | -0.0015251 | 0.010631
da2 | 0.14348 | -1.9131 | -0.35886 | 0.7312 | 1.5153 | -0.41242 | -0.21559 | 0.57923 | 0.17804 | -0.21043 | 0.58466 | 0.10114 | -0.3507 | 0.31085 | -0.14675 | -0.45259 | -0.44514 | -0.021772 | -0.63962 | -0.74654 | 4.699
da3 | 0.22275 | -2.2351 | 0.28482 | -0.042104 | 0.45248 | 0.045732 | -0.27584 | 0.10829 | -0.073347 | 0.037359 | -0.038759 | 0.52846 | 0.19132 | -0.74159 | -0.59513 | -0.42835 | -0.070765 | 0.15264 | 0.57949 | -0.11471 | 0.064941
da4 | 0.35555 | -0.0020932 | -0.018368 | -0.073605 | -0.54218 | -0.3251 | -0.093338 | -0.12036 | -0.29181 | 0.0097839 | -0.23138 | -0.13765 | -0.035422 | 0.0031876 | 0.67931 | 0.39272 | 0.41386 | -0.03285 | 0.18019 | 0.22302 | 0.54052
da5 | 0.072229 | 2.031 | -0.089373 | -0.21517 | 0.21484 | -0.044351 | -0.095587 | 0.079258 | 0.075368 | 0.13297 | 0.063969 | 0.19661 | 0.094527 | -0.092124 | -0.0033765 | -0.22689 | -0.050237 | -0.092421 | 0.088409 | -0.036412 | -0.29066
da6 | 0.11361 | -0.16389 | -0.096387 | -0.0058312 | -0.16787 | -0.094664 | -0.033734 | -0.00085741 | -0.098278 | -0.047394 | -0.036308 | 0.031075 | 0.0393 | 0.016251 | 0.046546 | 0.10985 | 0.14792 | 0.068834 | 0.057866 | 0.063682 | 0.30478
da7 | -0.053356 | 1.2976 | 0.002116 | -0.021535 | -0.019318 | -0.0089766 | 0.0056103 | 0.025314 | -0.024443 | -0.007763 | -0.027395 | -0.032876 | 0.029683 | -0.032152 | 0.076974 | 0.0073464 | 0.032547 | 0.0094713 | 0.024387 | -0.038991 | 0.022659
da8 | 0.0037137 | -0.11334 | -0.050002 | -0.048208 | -0.029845 | -0.042863 | -0.026394 | -0.024574 | -0.050337 | -0.019838 | -0.013729 | 0.040449 | 0.033971 | 0.015451 | 0.031479 | 0.03408 | 0.063064 | 0.058488 | 0.039561 | -0.010754 | 0.048099
da9 | -0.016119 | -0.22297 | 0.029348 | -0.001821 | 0.011807 | 0.02009 | 0.0021562 | 0.0015214 | 0.0059658 | 0.0034871 | -0.00075996 | 0.015611 | 0.0046254 | -0.020017 | -0.00061644 | -0.031918 | 0.0085571 | -0.019334 | -0.0070679 | -0.021636 | -0.048624
da10 | 0.0025545 | -0.040511 | -0.024197 | 0.016368 | -0.022512 | 0.020164 | 0.054691 | 0.028993 | -0.048793 | -0.065209 | -0.024373 | 0.025187 | 0.030123 | -0.054441 | -0.037664 | -0.021738 | 0.038 | 0.038297 | 0.041192 | 0.0059113 | -0.031684
da11 | -0.038277 | 0.13376 | -0.0049754 | -0.0073411 | -0.0082222 | -0.01033 | 0.0033279 | 0.0010381 | 0.0076258 | 0.0094133 | 0.0073285 | 0.00059711 | 0.0096577 | -0.0028841 | 0.011166 | -0.0043657 | 0.0048234 | -0.0080107 | 0.0016579 | -0.010506 | -0.019515
da12 | 0.0011567 | -0.012017 | -0.0056189 | 0.0025342 | -0.0064369 | -0.002112 | 0.0052571 | 0.0052525 | -0.0079885 | -0.0069737 | -0.0017198 | 0.0047422 | 0.0047886 | -0.0073815 | -0.0068234 | 0.00043367 | 0.0079684 | 0.0075617 | 0.0035979 | 0.0029186 | -0.0012824
da13 | -0.0020077 | -0.028359 | -0.00074528 | -0.0017237 | 0.00035309 | 6.3497e-05 | -0.001657 | -0.00015047 | 0.0031733 | 0.0022132 | 0.0013648 | 0.0034211 | 0.0029666 | -0.00064876 | -0.00030278 | -0.0026297 | -0.0012198 | -0.0023562 | -0.0013675 | -0.00075437 | -0.0074019
da14 | -0.0073029 | 0.0021054 | 0.00078652 | 0.0028065 | -0.0019278 | -0.00051468 | 0.0015631 | 0.00055385 | -0.0029114 | -0.0015684 | -0.001841 | 0.00034146 | 0.00032033 | -0.0034612 | -0.0008045 | -0.0010692 | 0.0026423 | 0.0026229 | 0.0018986 | 0.00056254 | -0.0021518
da15 | -0.0033883 | 0.0066545 | -0.0033385 | -0.0013598 | -0.00040625 | 0.0013378 | 0.0021387 | 0.0024264 | -0.00045567 | -0.0011864 | 0.002909 | 0.0027668 | 0.0019355 | -0.0032778 | -0.00036845 | -0.0012898 | -0.0011852 | -0.0012642 | 0.001464 | -0.00084625 | -0.0062109
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3614.7549
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.4118 | -3612.7823 | -0.98629 | -2.9589 | -0.98629 | -2.9589 | -0.98629 | 0.98629 | -0.98629 | -0.98629 | -0.98629 | 0.98629 | 0.98629 | 0.98629 | -0.98629 | 0.98629 | 0.98629 | 2.9589 | 2.9589 | 4.9315 | -3616.7275
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.13105 | 1.4208 | -1.0071 | -0.40941 | -1.1085 | -0.14329 | -0.96924 | 0.765 | -0.21001 | 1.3499 | -0.29902 | 0.46683 | -0.35235 | 1.4861 | 0.24405 | 0.88285 | -0.55234 | -0.14743 | -0.12019 | 0.1242 | 2.5293
db2 | -0.17805 | -3.2418 | 0.30831 | -0.060114 | 0.090241 | -0.66474 | 0.20271 | -0.31195 | 0.38594 | 0.34434 | -0.18361 | 0.17615 | 0.1947 | -0.0087664 | -0.23351 | -0.12806 | -0.22746 | 0.16458 | 0.19096 | -0.23972 | 0.82187
db3 | 0.7224 | 37.9901 | -1.2128 | -0.29406 | 0.27453 | 0.60483 | -0.28336 | 0.53811 | 0.56756 | 0.50779 | 0.22953 | -0.89216 | -0.90114 | -0.12265 | -1.1112 | -0.55437 | 0.39657 | 0.54085 | 0.47394 | 1.238 | 0.21266
db4 | 0.037412 | -0.44996 | 0.078857 | -0.13378 | 0.063028 | -0.16027 | 0.25932 | 0.12368 | 0.18654 | -0.014856 | 0.081184 | 0.011917 | -0.057246 | -0.009113 | 0.1323 | 0.068366 | -0.17861 | -0.10888 | -0.19183 | -0.15059 | -0.077885
db5 | 0.41484 | -2.4207 | -0.33993 | -0.82058 | -0.11365 | -0.0057914 | -0.026144 | -0.05888 | -0.14107 | -0.00046831 | 0.013026 | 0.58872 | 0.24302 | 0.092224 | 0.020461 | 0.013159 | 0.023737 | 0.047927 | 0.21048 | 0.25376 | -1.8287
db6 | -0.014899 | -0.021278 | 0.082533 | 0.045183 | 0.067848 | 0.087784 | 0.012564 | 0.048214 | 0.0083272 | -0.057036 | 0.094399 | -0.016235 | -0.00040306 | -0.1639 | 0.029561 | 0.026326 | -0.07275 | -0.16972 | -0.038996 | 0.016299 | -0.061887
db7 | 1.0877 | 3.0925 | -0.16716 | -0.10867 | 0.0028333 | 0.018256 | 0.053385 | -0.010976 | -0.021635 | -0.02159 | -0.042463 | -0.02675 | -0.0074527 | 0.027041 | 0.053585 | 0.084435 | 0.0085193 | 0.03005 | 0.09897 | 0.029625 | 0.89408
db8 | 0.014506 | -0.074395 | 0.018371 | -0.0088997 | -0.010481 | -0.00098732 | 0.067856 | 0.040005 | 0.0031906 | 0.01156 | 0.019252 | -0.0095659 | -0.041134 | -0.058862 | 0.0036757 | -0.0093614 | -0.040334 | -0.016989 | 0.00033433 | 0.032372 | 0.0016765
db9 | 0.50026 | 0.54721 | 0.021343 | -0.008385 | 0.018801 | 0.025431 | -0.0087753 | -0.0016377 | 0.0053852 | 0.00766 | -0.0015075 | -0.032546 | -0.04328 | -0.011431 | -0.0089533 | -0.0042862 | -0.0090054 | 0.0061642 | 0.017444 | 0.027579 | 0.5706
db10 | 0.00092105 | 0.050164 | 0.038795 | 0.0085819 | 0.028672 | 0.046399 | 0.068284 | 0.06047 | 0.0047352 | -0.011663 | 0.026174 | -0.015418 | -0.0448 | -0.071458 | -0.0025092 | 0.0072359 | -0.046993 | -0.10115 | -0.015424 | 0.020065 | -0.076087
db11 | 0.74638 | 0.66219 | -0.014534 | -0.016815 | -0.0033467 | -0.0025218 | 0.0021059 | 0.00055666 | -0.0010852 | 0.0048572 | 0.0018399 | 0.017647 | 0.0052701 | 0.0010082 | 0.0049525 | 0.0039992 | 0.0015428 | -0.0034606 | -0.0010981 | -0.00091714 | 0.68679
db12 | 0.00090029 | -0.00019328 | 0.0018987 | -0.001822 | -0.0013962 | 0.0053713 | 0.0086761 | 0.0069904 | -0.0012761 | 0.0041814 | 0.0030141 | 0.0039338 | -0.0038795 | -0.0082153 | -0.0023304 | -0.0018643 | -0.0055079 | -0.01009 | -0.0008938 | 0.0032101 | -0.0062753
db13 | 0.073935 | 0.10446 | -0.0021031 | 0.0027247 | -0.0017413 | -0.0011473 | 8.2417e-05 | 0.0022032 | 0.0024909 | -0.00046971 | -0.00093583 | -0.0012958 | 0.00073902 | 0.00092303 | -0.0001984 | -0.0019912 | 0.0011765 | -0.001498 | 0.0004973 | 0.0005436 | 0.060545
db14 | -0.0036652 | 0.0075407 | 0.0011531 | -8.8035e-05 | 0.00078374 | 0.0026714 | 0.0042254 | 0.0039529 | 0.0001578 | 7.3802e-05 | 0.00042862 | -0.00062558 | -0.0029819 | -0.0038062 | 0.00010373 | 0.00020902 | -0.0027246 | -0.0047844 | 0.0001555 | 0.0010957 | -0.0064187
db15 | 0.02714 | 0.00045834 | 0.00399 | -0.0016317 | -0.00087819 | -0.00080694 | 0.00076099 | 0.000696 | -0.00039371 | 0.0023664 | 0.0019665 | 0.001207 | -0.0015567 | -0.0019568 | -0.00042462 | 0.0008399 | -0.00031683 | -0.0026963 | -0.0020941 | 0.00092923 | 0.012742
da2 | 1.6658 | 0.52486 | 0.35537 | 0.72123 | 1.603 | 0.38378 | -1.0781 | -0.16222 | -0.74633 | 0.35129 | -0.54695 | -0.58159 | 0.27031 | -0.79177 | -0.41949 | 0.17574 | 0.77094 | -0.63748 | 0.3759 | -0.043643 | 2.1785
da3 | 0.61965 | -1.446 | 0.44423 | 0.39785 | 0.047216 | 0.21611 | 0.063784 | 0.0064956 | -0.15155 | -0.011435 | 0.22224 | 0.15336 | 0.016636 | -0.18817 | -0.05946 | -0.36812 | -0.1802 | -0.32565 | -0.12117 | -0.16217 | 0.66157
da4 | 0.089504 | 0.98242 | -0.48138 | -0.29079 | -0.1132 | 0.17515 | 0.25709 | -0.33451 | -0.3399 | -0.33441 | 0.38908 | 0.66371 | 0.49735 | 0.32228 | 0.29301 | -0.16237 | -0.030039 | -0.095604 | -0.061803 | -0.35366 | -0.19168
da5 | 0.027422 | 2.2141 | -0.017386 | 0.11528 | 0.16236 | -0.021356 | -0.12227 | 0.077922 | 0.066362 | 0.11303 | -0.086997 | 0.057062 | 0.10077 | 0.029362 | -0.10645 | -0.21992 | -0.10103 | -0.12427 | 0.090361 | -0.01283 | -0.34616
da6 | -0.070482 | -0.3358 | -0.068829 | -0.066207 | 0.03999 | 0.032622 | 0.076555 | -0.0063263 | -0.16175 | -0.1322 | -0.037201 | 0.018401 | 0.066192 | -0.0036197 | 0.033625 | 0.050295 | 0.030952 | -0.011325 | 0.11513 | 0.023687 | -0.31599
da7 | -0.012556 | 1.3218 | -0.030378 | -0.025479 | -0.00399 | 0.035355 | 0.024819 | 0.071743 | 0.014549 | 0.05168 | -0.018219 | -0.021262 | 0.004604 | 0.0081173 | 0.0054102 | -0.05801 | 0.019617 | -0.010993 | -0.046275 | -0.021288 | 0.017148
da8 | -0.04685 | -0.077414 | 0.018259 | 0.0081745 | 0.039484 | 0.033513 | 0.0093295 | 0.037389 | -0.037254 | -0.025084 | -0.011098 | -0.010863 | -0.010346 | -0.049816 | -0.016215 | -0.045993 | -0.015154 | -0.010305 | 0.050098 | 0.035879 | 0.024826
da9 | -0.01939 | -0.24901 | 0.047564 | 0.026213 | 0.014645 | 0.026223 | -0.019429 | -0.007591 | 0.0056436 | -0.0045438 | -0.0096591 | 0.011677 | -0.0013722 | -0.00059489 | -0.014694 | -0.027718 | -0.012803 | -0.028363 | -0.0053689 | 0.00017091 | -0.067862
da10 | 0.0043409 | -0.10902 | -0.04835 | 0.017161 | 0.0024887 | 0.039511 | 0.01229 | 0.0097058 | -0.0019281 | -0.055833 | -0.0091359 | 0.024084 | 0.035811 | -0.016599 | -0.0082948 | -0.028864 | 0.0044733 | -0.013975 | 0.02984 | 0.0076158 | -0.01334
da11 | -0.038414 | 0.13328 | -0.019041 | -0.0027355 | -0.0054426 | 0.0031725 | -0.0097825 | 0.009959 | 0.0075093 | 0.015138 | -0.0038055 | 0.010644 | -0.00067344 | 0.010238 | 0.00061648 | -0.0025581 | -0.0088163 | -0.0044188 | -0.0015392 | 0.0015353 | -0.010338
da12 | -0.0034222 | -0.033515 | -0.006652 | 0.0013334 | 0.0022162 | 0.0052175 | 0.002738 | 0.006223 | -0.00037977 | -0.0059605 | -0.001664 | 0.0029272 | 0.0062806 | -0.0027588 | -0.0017781 | -0.0031034 | 0.00057922 | -0.0068918 | 0.0019059 | -0.0002326 | -0.01527
da13 | -0.0033852 | -0.029787 | 0.0027589 | 0.0037207 | 0.00089736 | 0.00078014 | -0.0019418 | 0.00034389 | 0.0026493 | 0.0019816 | -0.0013645 | -0.0013671 | -0.0016332 | 0.00080985 | -0.0018602 | -0.0041312 | -0.0026799 | 0.00013623 | -6.0959e-05 | 0.00096078 | -0.0029878
da14 | -0.0089341 | 0.00040936 | -0.0040424 | 0.0020361 | 0.0018749 | 0.0024118 | 0.00090265 | 0.0019425 | 0.00026679 | -0.0020104 | 0.0011891 | 0.0021943 | 0.0022265 | -0.0013252 | -2.0211e-07 | -0.0026742 | -0.0018463 | -0.0018338 | 0.00056624 | -0.0018784 | -0.0032366
da15 | -0.0039132 | 0.004882 | -0.00030584 | 0.00046954 | 0.00040069 | 0.0022192 | 0.0032238 | 0.0038007 | 0.0010011 | -0.00083495 | 0.00059245 | 0.00058454 | -0.001115 | -0.0041897 | -0.0009388 | -0.0009458 | -0.0021857 | -0.0039303 | 0.00028755 | 0.0018666 | -0.0055846
Field Colinearity | 0.025219
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 0.20803 | -0.58246 | -0.77114 | -0.64817 | -1.1548 | -0.49339 | -0.24637 | 0.62591 | 0.40608 | 0.74751 | 0.21455 | 1.2814 | -0.063452 | 0.99196 | -0.19456 | -0.15825 | -0.63238 | 0.14096 | -0.77971 | 1.8674
Twist (m2 mrad) | 1.5602 | -8.3 | -10.4104 | -8.2642 | -13.858 | -5.5506 | -2.5869 | 6.1026 | 3.6548 | 6.1669 | 1.6092 | 8.6493 | -0.38071 | 5.2078 | -0.87551 | -0.59344 | -1.8971 | 0.31716 | -1.1696 | 0.70028
Twist Integral (I0) (m2 rad) | -0.014939
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 1.2897 | -1.1381 | -0.54045 | -1.2396 | -0.27434 | -1.1003 | 0.63396 | -0.34105 | 1.2188 | -0.43007 | 0.33579 | -0.48339 | 1.355 | 0.113 | 0.75181 | -0.68339 | -0.27848 | -0.25123 | -0.006847 | 2.3982
Twist (m2 mrad) | 9.6731 | -16.2179 | -7.2961 | -15.8046 | -3.2921 | -12.3782 | 6.6565 | -3.3253 | 10.9693 | -3.5481 | 2.5184 | -3.2629 | 8.1303 | 0.59326 | 3.3831 | -2.5627 | -0.83544 | -0.56527 | -0.010271 | 0.89933
Twist Integral (I0) (m2 rad) | -0.019707

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.55884 | 1.2527 | -0.033991 | 0.21419 | 0.93717 | -0.73131 | 0.14348 | 0.22275 | 0.35555 | 18.0061 | 0.59651 | -0.17805 | 0.7224 | 0.037412 | 0.41484 | 1.0877 | -1.3105 | 1.6658 | 0.61965 | 0.089504 | 17.8404 | 0.59642 | 0.59794 | 2.533 | -0.018028 | 0.050484 | 1.0067 | 0 | 0.22571 | 0.13612 | 0.34981 | 14422.323 | -0.24645 | 1.9403 | 0.017438 | 0.24664 | 1.1477 | 0 | 1.645 | 0.55258 | 0.10977 | 14423.743
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8