## Alarm sheet
File name | HCMB__A001-03000045_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000045
Date of test Ap 1 | 2003-06-30 |  |  |  | Date of test Ap 2 | 2003-06-30
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | yellow alarm | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | yellow alarm | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
120 | HCMB__A001 | 3000045 |  | 2 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2003-06-30 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-06-30 | Answer to manufacturer |  |  | 2003-07-01
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 11.07.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)

## Original data
File name |  | HCMB__A001-03000045_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000045 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 2.0 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | noell3045cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | noell3045cct2 |  |  |  | a even | 1
Test Operator |  | Ortega |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-06-30 |  |  | Time of test |  |  | 6:15 - 8:45 |  |  |  | Date of test |  | 2003-06-30 |  |  | Time of test |  |  | 9:15 - 12:30
Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03 |  |  |  | Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 25 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 25
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | -6.796e-05 |  |  | Dy average |  |  | 0.0006982 |  |  |  | Dx average |  | 5.4195e-05 |  |  | Dy average |  |  | 0.00064685
Magnetic length |  | 14.4504 |  |  | Tranfs. Function (mT/KA) |  |  | 596.719 |  |  |  | Magnetic length |  | 14.4512 |  |  | Tranfs. Function (mT/KA) |  |  | 596.8105
Phase |  | -3.8614 |  |  | Integral field (T.m) |  |  | 0.073294 |  |  |  | Phase |  | -4.2269 |  |  | Integral field (T.m) |  |  | 0.07331
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003196 | 0.005075 | 0.005072 | 0.005072 | 0.005071 | 0.005072 | 0.005072 | 0.005071 | 0.005072 | 0.005071 | 0.005071 | 0.005072 | 0.005072 | 0.005072 | 0.005071 | 0.005072 | 0.005072 | 0.005073 | 0.005075 | 0.003244 | 0.073294 | C1 (T) | 0.003221 | 0.005075 | 0.005073 | 0.005072 | 0.005072 | 0.005073 | 0.005073 | 0.005073 | 0.005072 | 0.005073 | 0.005072 | 0.005073 | 0.005073 | 0.005072 | 0.005073 | 0.005072 | 0.005073 | 0.005073 | 0.005075 | 0.003226 | 0.07331
Angle (mrad) | 2.0904 | 0.02292 | -0.8311 | 0.47376 | -0.59695 | 0.68968 | -0.98929 | 1.6076 | -0.19001 | 0.65768 | -0.89924 | 0.43989 | -0.9157 | 0.53588 | -0.55606 | 0.67781 | -1.0408 | 0.56931 | -1.0278 | -0.062689 | -3.8614 | Angle (mrad) | 5.4556 | -1.2617 | -1.5 | -0.65263 | -0.43474 | 0.50146 | 0.29176 | 1.3386 | 0.0644 | 0.40248 | -1.0112 | 0.18743 | -0.25584 | 0.60667 | -0.36217 | -0.40436 | -0.90387 | 0.027631 | -0.17845 | 0.2128 | -4.2269
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.95401 | -1.6885 | -1.0561 | -0.65712 | -0.80065 | -1.2324 | -1.2124 | -0.98462 | -1.3363 | -1.187 | -1.7762 | -2.0555 | -2.4337 | -1.2277 | -0.8854 | -0.43993 | -0.82169 | -1.4286 | -1.5601 | -0.36795 | -1.1634 | b2 | -0.58471 | 1.1094 | 1.1667 | 0.66643 | 1.6571 | 1.5681 | 1.5321 | 0.94116 | 0.59267 | -0.24962 | 0.95907 | 0.78169 | 0.5866 | 1.1883 | 0.9385 | 0.94394 | 0.99168 | 2.1243 | 0.92937 | 0.027404 | 0.93788
b3 | 37.2182 | 0.53611 | 1.0712 | 0.27313 | 0.63359 | 0.29595 | 1.346 | 0.8714 | 1.9598 | 2.1369 | 0.78112 | 0.53099 | 0.64045 | -0.72285 | -0.025242 | 1.3304 | 2.1578 | 1.7063 | 1.9902 | -1.3032 | 2.0825 | b3 | 35.3426 | 0.15692 | 0.59621 | 0.54839 | 1.3021 | 0.17915 | 1.5326 | 0.99936 | 0.93501 | 0.821 | 0.31704 | 0.92856 | 0.7805 | 0.70049 | 0.84541 | 1.5617 | 1.6406 | 0.36862 | 2.2212 | 0.054062 | 2.0191
b4 | -0.6856 | 0.057688 | 0.090458 | 0.12664 | 0.12492 | 0.03039 | -0.01394 | -0.14617 | -0.10157 | 0.0089864 | 0.076109 | -0.11305 | -0.069602 | -0.018434 | 0.030895 | 0.097731 | 0.02764 | -0.055288 | 0.074264 | -0.32323 | -0.021331 | b4 | 0.13846 | 0.3976 | 0.21994 | 0.013383 | 0.20731 | 0.31223 | 0.30724 | 0.27694 | 0.039208 | -0.14578 | 0.014298 | 0.21379 | -0.027256 | -0.00083824 | 0.11403 | -0.20291 | -0.15303 | 0.2194 | -0.075404 | 0.050758 | 0.096019
b5 | -3.9892 | -0.46687 | -0.66859 | -0.14402 | -0.22763 | -0.020241 | -0.051769 | -0.14211 | -0.21524 | -0.087456 | 0.18454 | 0.086737 | -0.25459 | -0.19868 | -0.20964 | -0.2961 | -0.22357 | -0.10559 | -0.078238 | -2.3262 | -0.36952 | b5 | -4.5368 | -0.4176 | -0.65105 | -0.42579 | -0.35187 | -0.22065 | -0.30084 | -0.28651 | -0.37474 | -0.10887 | -0.012987 | -0.23198 | -0.24939 | -0.41146 | -0.41802 | -0.44188 | -0.33228 | -0.10329 | -0.29597 | -2.9582 | -0.53951
b6 | -0.64866 | 0.096155 | 0.06791 | -0.020979 | 0.15375 | 0.22798 | 0.06234 | 0.092256 | 0.072683 | -0.15916 | -0.089115 | 0.049315 | 0.1925 | 0.10721 | 0.11259 | 0.14228 | 0.054866 | 0.052709 | 0.15268 | 0.047728 | 0.051371 | b6 | -0.085709 | 0.003707 | -0.037538 | -0.11856 | 0.04142 | -0.067043 | -0.13609 | -0.10035 | 0.0009054 | 0.087317 | -0.091293 | -0.044311 | 0.1711 | 0.071202 | 0.016293 | -0.021074 | -0.10295 | -0.29166 | 0.034884 | 0.027579 | -0.032219
b7 | 3.0072 | 0.9252 | 1.0556 | 1.1972 | 1.123 | 1.1797 | 1.1252 | 1.0436 | 1.1108 | 1.1381 | 1.1069 | 1.0147 | 1.0939 | 1.0062 | 1.0067 | 1.1362 | 1.1588 | 1.231 | 1.1894 | 0.85054 | 1.1563 | b7 | 2.7403 | 0.99346 | 1.049 | 1.1856 | 1.1599 | 1.188 | 1.1183 | 1.1712 | 1.1709 | 1.153 | 1.167 | 1.1 | 1.1463 | 1.1661 | 1.151 | 1.17 | 1.1606 | 1.1987 | 1.1997 | 0.94969 | 1.1931
b8 | -0.18757 | 0.015267 | -0.0053549 | -0.017842 | 0.012115 | 0.066942 | 0.02941 | 0.034215 | 0.0030858 | -0.056473 | -0.066028 | -0.082711 | -0.0037473 | -0.027439 | 0.051221 | 0.071893 | 0.036687 | -0.0057154 | -0.0065728 | 0.067258 | -0.0013602 | b8 | -0.0055027 | -0.0037293 | -0.033807 | -0.058535 | -0.024231 | -0.043753 | 0.00013109 | 0.011022 | -0.016776 | -0.080106 | -0.015227 | -0.041064 | 0.018881 | 0.015553 | -0.0031223 | 0.026568 | -0.010216 | -0.063124 | -0.0083294 | -0.023866 | -0.018086
b9 | 0.43672 | 0.4771 | 0.47761 | 0.4929 | 0.49636 | 0.48665 | 0.47892 | 0.48304 | 0.4963 | 0.50369 | 0.49042 | 0.48729 | 0.4839 | 0.49067 | 0.4847 | 0.50365 | 0.50389 | 0.50181 | 0.50555 | 0.46731 | 0.48878 | b9 | 0.34142 | 0.4835 | 0.47362 | 0.47116 | 0.46699 | 0.45199 | 0.45916 | 0.46125 | 0.47274 | 0.45069 | 0.45522 | 0.45949 | 0.46496 | 0.46229 | 0.47086 | 0.47445 | 0.47101 | 0.47279 | 0.49632 | 0.44215 | 0.46269
b10 | -0.076707 | -0.026976 | -0.032943 | -0.070278 | 0.041666 | 0.052943 | 0.044599 | 0.056142 | 0.054388 | -0.12672 | -0.10699 | -0.10058 | -0.039155 | -0.022238 | 0.054111 | 0.10852 | 0.071383 | -0.0078079 | 0.06111 | 0.058045 | 0 | b10 | -0.038188 | 0.016815 | -0.048945 | -0.055584 | 0.049113 | 0.00025426 | 0.0080903 | 0.021055 | 0.01703 | -0.057957 | -0.0302 | -0.00083092 | 0.074309 | 0.057201 | 0.056186 | 0.052346 | -0.0020494 | -0.099991 | -0.0019164 | -0.048241 | 0
b11 | 0.63456 | 0.74992 | 0.74396 | 0.75384 | 0.75697 | 0.76389 | 0.76049 | 0.76456 | 0.75628 | 0.74808 | 0.75906 | 0.76352 | 0.75965 | 0.77178 | 0.76692 | 0.75197 | 0.74601 | 0.75204 | 0.74753 | 0.6635 | 0.7494 | b11 | 0.59754 | 0.73676 | 0.73698 | 0.74557 | 0.74059 | 0.75299 | 0.74882 | 0.75145 | 0.74309 | 0.73912 | 0.74578 | 0.74638 | 0.74633 | 0.74729 | 0.74191 | 0.73878 | 0.74041 | 0.74255 | 0.73448 | 0.65576 | 0.7356
b12 | -0.025595 | -0.0070938 | -0.0068406 | -0.013886 | -0.00171 | 0.0032098 | 0.0037381 | 0.0073407 | 0.0040324 | -0.017927 | -0.016137 | -0.022838 | -0.010315 | -0.003025 | 0.0056743 | 0.013192 | 0.0058219 | -0.0029492 | 0.0013455 | 0.0043922 | -0.0037197 | b12 | -0.0078592 | -0.00051532 | -0.0064585 | -0.0048125 | 0.01126 | 0.0032445 | 0.0042 | 0.0025886 | 0.0016197 | -0.0094112 | -0.00298 | -0.0010039 | 0.012052 | 0.0088462 | 0.01322 | 0.01023 | 0.0017601 | -0.0085017 | -0.00035797 | -0.0045459 | 0.0014059
b13 | 0.090225 | 0.076752 | 0.079799 | 0.075914 | 0.075191 | 0.073048 | 0.075677 | 0.073097 | 0.075197 | 0.075001 | 0.071127 | 0.070204 | 0.077281 | 0.071536 | 0.072697 | 0.073994 | 0.074995 | 0.074817 | 0.074085 | 0.050534 | 0.074188 | b13 | 0.085264 | 0.075112 | 0.076622 | 0.076644 | 0.079059 | 0.075502 | 0.075049 | 0.077797 | 0.074969 | 0.0721 | 0.072908 | 0.072771 | 0.076006 | 0.077269 | 0.069904 | 0.073694 | 0.072798 | 0.072471 | 0.074867 | 0.057732 | 0.074537
b14 | -0.0010875 | -0.00064368 | -0.0016638 | -0.0041351 | 0.0028218 | 0.0042664 | 0.0043156 | 0.0056704 | 0.0044257 | -0.0066203 | -0.005691 | -0.0074508 | -0.0023235 | 0.0011625 | 0.005298 | 0.0097533 | 0.0072947 | 0.001909 | 0.0061577 | 0.0065257 | 0.0014561 | b14 | -0.00028777 | -0.0017193 | -0.0061845 | -0.0051562 | 0.0038135 | 0.00094759 | -0.00086672 | 0.00052049 | -0.00013451 | -0.0046717 | -0.0036222 | -0.0017117 | 0.0049542 | 0.0022823 | 0.0020094 | 0.0015217 | -0.0017447 | -0.0068937 | -0.0027226 | -0.0037355 | -0.001139
b15 | -0.00064037 | 0.043959 | 0.038468 | 0.035798 | 0.038212 | 0.035022 | 0.036833 | 0.037297 | 0.035569 | 0.032474 | 0.03451 | 0.036953 | 0.036619 | 0.040606 | 0.03982 | 0.037033 | 0.037492 | 0.039418 | 0.039099 | 0.034113 | 0.03615 | b15 | -0.0056473 | 0.037247 | 0.035231 | 0.030739 | 0.034474 | 0.032473 | 0.036734 | 0.034788 | 0.035688 | 0.035704 | 0.037411 | 0.038751 | 0.038029 | 0.035415 | 0.036037 | 0.03348 | 0.034374 | 0.032036 | 0.028301 | 0.018934 | 0.03297
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 20.9039 | 0.2292 | -8.311 | 4.7376 | -5.9695 | 6.8968 | -9.8929 | 16.0761 | -1.9001 | 6.5768 | -8.9924 | 4.3989 | -9.157 | 5.3588 | -5.5606 | 6.7781 | -10.4083 | 5.6931 | -10.2777 | -0.62689 | 0 | a1 | 54.5569 | -12.6174 | -15.0004 | -6.5263 | -4.3474 | 5.0146 | 2.9176 | 13.3864 | 0.644 | 4.0248 | -10.1123 | 1.8743 | -2.5584 | 6.0667 | -3.6217 | -4.0436 | -9.0387 | 0.27631 | -1.7845 | 2.128 | 0
a2 | 3.9697 | -0.52675 | -0.4032 | -1.2731 | -1.0522 | -0.084677 | 0.61471 | -0.38424 | 1.1719 | -0.48143 | -1.0499 | -0.38628 | -0.70564 | -1.4401 | -1.0432 | 0.40847 | -0.24424 | -0.35565 | -0.48441 | 13.8086 | 0.1875 | a2 | 4.0477 | -0.49171 | -0.7941 | -0.067756 | -0.48445 | -0.50115 | 0.3806 | -0.19886 | -0.31269 | -0.068797 | -0.92106 | -0.2413 | -0.31512 | -0.29597 | 0.15566 | -0.50558 | -1.1322 | 0.55973 | -1.5753 | 8.8103 | 0.070641
a3 | -0.65792 | -0.024798 | -0.12829 | 0.11719 | 0.6878 | 0.14464 | 0.26616 | -0.063505 | 0.058834 | -0.093429 | 0.16469 | -0.061023 | -0.31233 | 0.37636 | -0.14431 | 0.020451 | 0.34737 | 0.72792 | 0.14049 | 0.056541 | 0.095827 | a3 | 2.9669 | 0.88403 | 0.36987 | 0.30249 | 0.69672 | 0.11519 | 0.034489 | -0.26498 | -0.24904 | -0.1211 | 0.089953 | -0.06771 | 0.017497 | 0.22697 | 0.046279 | -0.18648 | 0.041618 | 0.52616 | 0.065699 | 0.44751 | 0.24368
a4 | 1.4258 | 0.18044 | -0.37873 | -0.025842 | -0.25584 | -0.42264 | -0.23975 | -0.12203 | 0.28894 | 0.35079 | 0.045963 | -0.030229 | 0.16653 | 0.70039 | 0.40445 | -0.018209 | 0.21336 | 0.17406 | -0.06706 | 1.6884 | 0.15272 | a4 | 0.53416 | -0.25441 | -0.55546 | -0.413 | -0.2116 | -0.50822 | -0.60527 | -0.20881 | -0.19654 | -0.40946 | -0.38788 | -0.56195 | -0.23517 | -0.099675 | -0.20117 | -0.37103 | -0.012279 | -0.87534 | -0.3052 | 0.58002 | -0.29602
a5 | 2.5314 | -0.037203 | -0.076254 | -0.041614 | -0.14298 | -0.054386 | -0.033659 | 0.023986 | 0.035269 | -0.16124 | 0.10835 | 0.14397 | 0.15676 | 0.096966 | -0.11679 | -0.082363 | -0.076137 | 0.1597 | 0.092864 | -0.16808 | 0.076937 | a5 | 3.5672 | 0.026794 | -0.077014 | 0.16501 | 0.046696 | -0.010895 | 0.056264 | 0.15226 | -0.029697 | 0.094819 | -0.15523 | 0.053449 | 0.055155 | 0.091655 | -0.00097531 | 0.010418 | 0.1398 | 0.5238 | 0.056368 | -0.15175 | 0.17474
a6 | 0.31231 | -0.12901 | -0.11202 | -0.057501 | -0.017567 | 0.012351 | 0.029543 | 0.057562 | 0.047528 | 0.026105 | -0.042778 | 0.0052962 | 0.026739 | -0.12627 | -0.0012065 | 0.039169 | -0.021852 | -0.044131 | -0.095937 | -0.16428 | -0.016211 | a6 | 0.35048 | -0.066655 | -0.060307 | 0.02972 | 0.0083842 | 0.017293 | -0.11133 | -0.044346 | 0.0419 | 0.031857 | -0.077095 | -0.055481 | -0.069919 | -0.045509 | -0.053684 | -0.054689 | -0.02242 | 0.013962 | 0.025273 | 0.099688 | -0.010747
a7 | 1.6519 | -0.072986 | -0.086441 | -0.027023 | -0.019404 | -0.075284 | -0.028438 | -0.023167 | -0.050071 | -0.064932 | -0.071348 | -0.11931 | -0.071664 | -0.065926 | -0.09391 | -0.086936 | -0.065138 | -0.04323 | -0.049848 | 0.036081 | -0.0026179 | a7 | 1.9967 | -0.025845 | -0.065265 | -0.015107 | 0.031127 | -0.042426 | -0.086034 | -0.089839 | -0.0071468 | -0.0017477 | -0.075173 | -0.031669 | -0.0093378 | -0.025267 | -0.041899 | -0.08062 | -0.010079 | -0.091631 | -0.047563 | 0.060495 | 0.03063
a8 | 0.10599 | -0.044853 | -0.029481 | -0.011628 | -0.016858 | 0.021088 | -0.012327 | -0.03516 | 0.013051 | 0.084061 | 0.0277 | 0.064241 | 0.012523 | -0.013861 | 0.023652 | 0.029253 | -0.0026149 | -0.031337 | -0.0070722 | 0.0042709 | 0.0072567 | a8 | 0.062046 | 0.008259 | -0.025965 | -0.060184 | -0.050207 | -0.036251 | -0.048282 | -0.029882 | -0.037626 | -0.0061873 | -0.056407 | -0.058369 | -0.050586 | -0.026944 | -0.011828 | -0.013973 | -0.032474 | 0.0070707 | -0.010715 | 0.01359 | -0.025555
a9 | -0.2145 | 0.0089127 | -0.025923 | -0.045346 | -0.041051 | -0.049031 | -0.039355 | -0.040861 | -0.021169 | -0.054037 | -0.033574 | -0.033958 | -0.040088 | -0.025012 | -0.057177 | -0.049496 | -0.054724 | -0.014019 | -0.0058878 | -0.10461 | -0.042729 | a9 | -0.13377 | 0.039744 | -0.022585 | -0.029239 | -0.028699 | -0.035865 | -0.029271 | -0.02619 | -0.025512 | -0.026084 | -0.042753 | -0.043721 | -0.021813 | -0.0037321 | -0.038512 | -0.032615 | -0.037422 | -0.0040588 | -0.034652 | -0.070656 | -0.029736
a10 | 0.028462 | -0.046265 | -0.020071 | 0.0069602 | -0.0014125 | 0.020099 | 0.0061632 | 0.022044 | 0.030151 | 0.083466 | 0.03707 | 0.043186 | 0.029162 | -0.019494 | -0.011988 | -0.026965 | -0.021126 | -0.036725 | -0.033513 | -0.12289 | 0 | a10 | 0.073376 | 0.00026769 | -0.0051571 | 0.011297 | -0.0026591 | 0.011978 | -0.021481 | 0.0082686 | -0.0021632 | 0.0052793 | -0.017764 | -0.038006 | -0.030635 | -0.021575 | 0.010494 | 0.011697 | -0.0053512 | -0.0027242 | 0.062103 | -0.032231 | 0
a11 | 0.16887 | -0.039522 | -0.045794 | -0.036654 | -0.044237 | -0.034303 | -0.041891 | -0.014441 | -0.035598 | -0.032192 | -0.048871 | -0.041407 | -0.038728 | -0.025773 | -0.046991 | -0.038573 | -0.050045 | -0.031924 | -0.045626 | -0.029318 | -0.031352 | a11 | 0.19193 | -0.056738 | -0.06093 | -0.041849 | -0.044225 | -0.039993 | -0.045887 | -0.025586 | -0.027841 | -0.024931 | -0.047154 | -0.036345 | -0.034969 | -0.021411 | -0.045629 | -0.039349 | -0.042424 | -0.034387 | -0.033764 | -0.028913 | -0.031155
a12 | 0.0072074 | -0.011666 | -0.0081624 | -0.004117 | -0.00075648 | -0.00068353 | -0.0020569 | 0.0034341 | 0.0017899 | 0.0055448 | -0.0037886 | -0.0014767 | -1.0117e-06 | -0.015963 | -0.010141 | -0.0077641 | -0.0083724 | -0.01074 | -0.0093596 | -0.015338 | -0.004648 | a12 | 0.017172 | -0.00081075 | -0.0018627 | -0.00061022 | -0.0026936 | -0.00035836 | -0.0072396 | -0.0039239 | -0.0063784 | -0.0052025 | -0.0081902 | -0.0077816 | -0.0070284 | -0.0092198 | -0.0016269 | -0.0031689 | -0.0049505 | -0.0026497 | 0.0028917 | 0.00051108 | -0.0030913
a13 | -0.017275 | -0.00068057 | -0.0053369 | -0.0025908 | -0.0049367 | -0.0056694 | -0.0064517 | -0.0032836 | -0.0048776 | -0.0027238 | -0.0028987 | -0.0023629 | -0.0035128 | -0.0013424 | -0.0068113 | -0.0073559 | -0.007228 | -0.00068023 | -0.0039702 | -0.0085297 | -0.0046165 | a13 | -0.012997 | -3.1064e-05 | -0.0054641 | -0.0058855 | -0.0063287 | -0.0066389 | -0.0058561 | -0.0038264 | -0.0062548 | -0.0042936 | -0.0054611 | -0.0036454 | -0.0063073 | -0.0047942 | -0.0045803 | -0.0073398 | -0.0055007 | -0.0040401 | -0.0053617 | -0.0044879 | -0.0053327
a14 | 0.0040982 | -0.020538 | -0.01882 | -0.014847 | -0.016947 | -0.015197 | -0.016934 | -0.015929 | -0.014639 | -0.010462 | -0.014803 | -0.012443 | -0.014984 | -0.016091 | -0.014516 | -0.014699 | -0.015381 | -0.019051 | -0.01682 | -0.010059 | -0.014892 | a14 | 0.0036666 | -0.014389 | -0.015001 | -0.013408 | -0.013916 | -0.013515 | -0.016915 | -0.012515 | -0.015726 | -0.015651 | -0.017732 | -0.017794 | -0.015931 | -0.014653 | -0.015727 | -0.013456 | -0.015069 | -0.014581 | -0.011825 | -0.0038083 | -0.013902
a15 | 0.0031201 | -0.00066866 | -0.0018064 | -0.0049504 | 0.0034805 | 0.0061546 | 0.0047333 | 0.0067539 | 0.0039755 | -0.005542 | -0.0063819 | -0.0086177 | -0.0031998 | -0.00010444 | 0.0057125 | 0.01084 | 0.0073331 | 0.0018699 | 0.0052866 | 0.011558 | 0.001779 | a15 | 0.0020479 | -0.0026039 | -0.0089856 | -0.0084706 | 0.0013193 | -0.0025736 | -0.0031822 | -0.0012838 | -0.0035855 | -0.0073619 | -0.0068822 | -0.0046008 | 0.0030356 | -4.5522e-05 | 0.00083582 | -0.00078703 | -0.0042995 | -0.012083 | -0.0048159 | -0.0066664 | -0.003598
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 0.00017291 | 5.5399e-05 | 7.2163e-05 | 0.00015899 | -9.3428e-05 | -0.00011559 | -9.865e-05 | -0.00012387 | -0.0001188 | 0.00029582 | 0.00024431 | 0.00022892 | 9.0741e-05 | 4.7471e-05 | -0.00012113 | -0.00024799 | -0.00016515 | 1.4084e-05 | -0.00014302 | -0.00016165 | -6.796e-05 | Dx (m) | 3.7494e-05 | -3.8522e-05 | 0.00011118 | 0.0001278 | -0.00011275 | 0 | -2.1267e-05 | -4.6946e-05 | -3.9091e-05 | 0.00013362 | 6.5999e-05 | 0 | -0.00017221 | -0.00013143 | -0.00012686 | -0.00011874 | 0 | 0.00022823 | 1.1016e-05 | 0.00012117 | 5.4195e-05
Dy (m) | -0.00012278 | 0.00010779 | 5.0331e-05 | 0 | 0 | -4.992e-05 | -1.9205e-05 | -5.1396e-05 | -7.3383e-05 | -0.00017758 | -6.7343e-05 | -8.3995e-05 | -6.0674e-05 | 4.453e-05 | 1.9218e-05 | 4.8498e-05 | 3.7162e-05 | 8.359e-05 | 6.7531e-05 | 0.00030748 | 0.0006982 | Dy (m) | -0.00022054 | 0 | 2.1105e-05 | -1.8583e-05 | 0 | -2.6997e-05 | 4.7462e-05 | -2.0302e-05 | 0 | 0 | 4.4678e-05 | 8.6432e-05 | 6.1859e-05 | 4.5402e-05 | -3.1874e-05 | -3.3249e-05 | 1.2515e-05 | 1.688e-05 | -0.00014323 | 8.8921e-05 | 0.00064685
Magnetic length (m) |  | 14.4504 | [19 empty] | Magnetic length (m) |  | 14.4512
Transfer funct. (mT/KA) |  | 596.719 | [19 empty] | Transfer funct. (mT/KA) |  | 596.8105
Coil positioning (mm) |  | 0.016807 | [19 empty] | Coil positioning (mm) |  | 0.016718
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMB__A001-3000045 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000045 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 376 | 597.0588 | 596.7059 | 596.7059 | 596.5882 | 596.7059 | 596.7059 | 596.5882 | 596.7059 | 596.5882 | 596.5882 | 596.7059 | 596.7059 | 596.7059 | 596.5882 | 596.7059 | 596.7059 | 596.8235 | 597.0588 | 381.6471 | 0.073294 | C1/i (mT/KA) | 378.9412 | 597.0588 | 596.8235 | 596.7059 | 596.7059 | 596.8235 | 596.8235 | 596.8235 | 596.7059 | 596.8235 | 596.7059 | 596.8235 | 596.8235 | 596.7059 | 596.8235 | 596.7059 | 596.8235 | 596.8235 | 597.0588 | 379.5294 | 0.07331
Angle (mrad) | 2.0904 | 0.02292 | -0.8311 | 0.47376 | -0.59695 | 0.68968 | -0.98929 | 1.6076 | -0.19001 | 0.65768 | -0.89924 | 0.43989 | -0.9157 | 0.53588 | -0.55606 | 0.67781 | -1.0408 | 0.56931 | -1.0278 | -0.062689 | -3.8614 | Angle (mrad) | 5.4556 | -1.2617 | -1.5 | -0.65263 | -0.43474 | 0.50146 | 0.29176 | 1.3386 | 0.0644 | 0.40248 | -1.0112 | 0.18743 | -0.25584 | 0.60667 | -0.36217 | -0.40436 | -0.90387 | 0.027631 | -0.17845 | 0.2128 | -4.2269
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.95401 | -1.6885 | -1.0561 | -0.65712 | -0.80065 | -1.2324 | -1.2124 | -0.98462 | -1.3363 | -1.187 | -1.7762 | -2.0555 | -2.4337 | -1.2277 | -0.8854 | -0.43993 | -0.82169 | -1.4286 | -1.5601 | -0.36795 | -1.1634 | b2 | -0.58471 | 1.1094 | 1.1667 | 0.66643 | 1.6571 | 1.5681 | 1.5321 | 0.94116 | 0.59267 | -0.24962 | 0.95907 | 0.78169 | 0.5866 | 1.1883 | 0.9385 | 0.94394 | 0.99168 | 2.1243 | 0.92937 | 0.027404 | 0.93788
b3 | 37.2182 | 0.53611 | 1.0712 | 0.27313 | 0.63359 | 0.29595 | 1.346 | 0.8714 | 1.9598 | 2.1369 | 0.78112 | 0.53099 | 0.64045 | -0.72285 | -0.025242 | 1.3304 | 2.1578 | 1.7063 | 1.9902 | -1.3032 | 2.0825 | b3 | 35.3426 | 0.15692 | 0.59621 | 0.54839 | 1.3021 | 0.17915 | 1.5326 | 0.99936 | 0.93501 | 0.821 | 0.31704 | 0.92856 | 0.7805 | 0.70049 | 0.84541 | 1.5617 | 1.6406 | 0.36862 | 2.2212 | 0.054062 | 2.0191
b4 | -0.6856 | 0.057688 | 0.090458 | 0.12664 | 0.12492 | 0.03039 | -0.01394 | -0.14617 | -0.10157 | 0.0089864 | 0.076109 | -0.11305 | -0.069602 | -0.018434 | 0.030895 | 0.097731 | 0.02764 | -0.055288 | 0.074264 | -0.32323 | -0.021331 | b4 | 0.13846 | 0.3976 | 0.21994 | 0.013383 | 0.20731 | 0.31223 | 0.30724 | 0.27694 | 0.039208 | -0.14578 | 0.014298 | 0.21379 | -0.027256 | -0.00083824 | 0.11403 | -0.20291 | -0.15303 | 0.2194 | -0.075404 | 0.050758 | 0.096019
b5 | -3.9892 | -0.46687 | -0.66859 | -0.14402 | -0.22763 | -0.020241 | -0.051769 | -0.14211 | -0.21524 | -0.087456 | 0.18454 | 0.086737 | -0.25459 | -0.19868 | -0.20964 | -0.2961 | -0.22357 | -0.10559 | -0.078238 | -2.3262 | -0.36952 | b5 | -4.5368 | -0.4176 | -0.65105 | -0.42579 | -0.35187 | -0.22065 | -0.30084 | -0.28651 | -0.37474 | -0.10887 | -0.012987 | -0.23198 | -0.24939 | -0.41146 | -0.41802 | -0.44188 | -0.33228 | -0.10329 | -0.29597 | -2.9582 | -0.53951
b6 | -0.64866 | 0.096155 | 0.06791 | -0.020979 | 0.15375 | 0.22798 | 0.06234 | 0.092256 | 0.072683 | -0.15916 | -0.089115 | 0.049315 | 0.1925 | 0.10721 | 0.11259 | 0.14228 | 0.054866 | 0.052709 | 0.15268 | 0.047728 | 0.051371 | b6 | -0.085709 | 0.003707 | -0.037538 | -0.11856 | 0.04142 | -0.067043 | -0.13609 | -0.10035 | 0.0009054 | 0.087317 | -0.091293 | -0.044311 | 0.1711 | 0.071202 | 0.016293 | -0.021074 | -0.10295 | -0.29166 | 0.034884 | 0.027579 | -0.032219
b7 | 3.0072 | 0.9252 | 1.0556 | 1.1972 | 1.123 | 1.1797 | 1.1252 | 1.0436 | 1.1108 | 1.1381 | 1.1069 | 1.0147 | 1.0939 | 1.0062 | 1.0067 | 1.1362 | 1.1588 | 1.231 | 1.1894 | 0.85054 | 1.1563 | b7 | 2.7403 | 0.99346 | 1.049 | 1.1856 | 1.1599 | 1.188 | 1.1183 | 1.1712 | 1.1709 | 1.153 | 1.167 | 1.1 | 1.1463 | 1.1661 | 1.151 | 1.17 | 1.1606 | 1.1987 | 1.1997 | 0.94969 | 1.1931
b8 | -0.18757 | 0.015267 | -0.0053549 | -0.017842 | 0.012115 | 0.066942 | 0.02941 | 0.034215 | 0.0030858 | -0.056473 | -0.066028 | -0.082711 | -0.0037473 | -0.027439 | 0.051221 | 0.071893 | 0.036687 | -0.0057154 | -0.0065728 | 0.067258 | -0.0013602 | b8 | -0.0055027 | -0.0037293 | -0.033807 | -0.058535 | -0.024231 | -0.043753 | 0.00013109 | 0.011022 | -0.016776 | -0.080106 | -0.015227 | -0.041064 | 0.018881 | 0.015553 | -0.0031223 | 0.026568 | -0.010216 | -0.063124 | -0.0083294 | -0.023866 | -0.018086
b9 | 0.43672 | 0.4771 | 0.47761 | 0.4929 | 0.49636 | 0.48665 | 0.47892 | 0.48304 | 0.4963 | 0.50369 | 0.49042 | 0.48729 | 0.4839 | 0.49067 | 0.4847 | 0.50365 | 0.50389 | 0.50181 | 0.50555 | 0.46731 | 0.48878 | b9 | 0.34142 | 0.4835 | 0.47362 | 0.47116 | 0.46699 | 0.45199 | 0.45916 | 0.46125 | 0.47274 | 0.45069 | 0.45522 | 0.45949 | 0.46496 | 0.46229 | 0.47086 | 0.47445 | 0.47101 | 0.47279 | 0.49632 | 0.44215 | 0.46269
b10 | -0.076707 | -0.026976 | -0.032943 | -0.070278 | 0.041666 | 0.052943 | 0.044599 | 0.056142 | 0.054388 | -0.12672 | -0.10699 | -0.10058 | -0.039155 | -0.022238 | 0.054111 | 0.10852 | 0.071383 | -0.0078079 | 0.06111 | 0.058045 | 0 | b10 | -0.038188 | 0.016815 | -0.048945 | -0.055584 | 0.049113 | 0.00025426 | 0.0080903 | 0.021055 | 0.01703 | -0.057957 | -0.0302 | -0.00083092 | 0.074309 | 0.057201 | 0.056186 | 0.052346 | -0.0020494 | -0.099991 | -0.0019164 | -0.048241 | 0
b11 | 0.63456 | 0.74992 | 0.74396 | 0.75384 | 0.75697 | 0.76389 | 0.76049 | 0.76456 | 0.75628 | 0.74808 | 0.75906 | 0.76352 | 0.75965 | 0.77178 | 0.76692 | 0.75197 | 0.74601 | 0.75204 | 0.74753 | 0.6635 | 0.7494 | b11 | 0.59754 | 0.73676 | 0.73698 | 0.74557 | 0.74059 | 0.75299 | 0.74882 | 0.75145 | 0.74309 | 0.73912 | 0.74578 | 0.74638 | 0.74633 | 0.74729 | 0.74191 | 0.73878 | 0.74041 | 0.74255 | 0.73448 | 0.65576 | 0.7356
b12 | -0.025595 | -0.0070938 | -0.0068406 | -0.013886 | -0.00171 | 0.0032098 | 0.0037381 | 0.0073407 | 0.0040324 | -0.017927 | -0.016137 | -0.022838 | -0.010315 | -0.003025 | 0.0056743 | 0.013192 | 0.0058219 | -0.0029492 | 0.0013455 | 0.0043922 | -0.0037197 | b12 | -0.0078592 | -0.00051532 | -0.0064585 | -0.0048125 | 0.01126 | 0.0032445 | 0.0042 | 0.0025886 | 0.0016197 | -0.0094112 | -0.00298 | -0.0010039 | 0.012052 | 0.0088462 | 0.01322 | 0.01023 | 0.0017601 | -0.0085017 | -0.00035797 | -0.0045459 | 0.0014059
b13 | 0.090225 | 0.076752 | 0.079799 | 0.075914 | 0.075191 | 0.073048 | 0.075677 | 0.073097 | 0.075197 | 0.075001 | 0.071127 | 0.070204 | 0.077281 | 0.071536 | 0.072697 | 0.073994 | 0.074995 | 0.074817 | 0.074085 | 0.050534 | 0.074188 | b13 | 0.085264 | 0.075112 | 0.076622 | 0.076644 | 0.079059 | 0.075502 | 0.075049 | 0.077797 | 0.074969 | 0.0721 | 0.072908 | 0.072771 | 0.076006 | 0.077269 | 0.069904 | 0.073694 | 0.072798 | 0.072471 | 0.074867 | 0.057732 | 0.074537
b14 | -0.0010875 | -0.00064368 | -0.0016638 | -0.0041351 | 0.0028218 | 0.0042664 | 0.0043156 | 0.0056704 | 0.0044257 | -0.0066203 | -0.005691 | -0.0074508 | -0.0023235 | 0.0011625 | 0.005298 | 0.0097533 | 0.0072947 | 0.001909 | 0.0061577 | 0.0065257 | 0.0014561 | b14 | -0.00028777 | -0.0017193 | -0.0061845 | -0.0051562 | 0.0038135 | 0.00094759 | -0.00086672 | 0.00052049 | -0.00013451 | -0.0046717 | -0.0036222 | -0.0017117 | 0.0049542 | 0.0022823 | 0.0020094 | 0.0015217 | -0.0017447 | -0.0068937 | -0.0027226 | -0.0037355 | -0.001139
b15 | -0.00064037 | 0.043959 | 0.038468 | 0.035798 | 0.038212 | 0.035022 | 0.036833 | 0.037297 | 0.035569 | 0.032474 | 0.03451 | 0.036953 | 0.036619 | 0.040606 | 0.03982 | 0.037033 | 0.037492 | 0.039418 | 0.039099 | 0.034113 | 0.03615 | b15 | -0.0056473 | 0.037247 | 0.035231 | 0.030739 | 0.034474 | 0.032473 | 0.036734 | 0.034788 | 0.035688 | 0.035704 | 0.037411 | 0.038751 | 0.038029 | 0.035415 | 0.036037 | 0.03348 | 0.034374 | 0.032036 | 0.028301 | 0.018934 | 0.03297
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 20.9039 | 0.2292 | -8.311 | 4.7376 | -5.9695 | 6.8968 | -9.8929 | 16.0761 | -1.9001 | 6.5768 | -8.9924 | 4.3989 | -9.157 | 5.3588 | -5.5606 | 6.7781 | -10.4083 | 5.6931 | -10.2777 | -0.62689 | 0 | a1 | 54.5569 | -12.6174 | -15.0004 | -6.5263 | -4.3474 | 5.0146 | 2.9176 | 13.3864 | 0.644 | 4.0248 | -10.1123 | 1.8743 | -2.5584 | 6.0667 | -3.6217 | -4.0436 | -9.0387 | 0.27631 | -1.7845 | 2.128 | 0
a2 | 3.9697 | -0.52675 | -0.4032 | -1.2731 | -1.0522 | -0.084677 | 0.61471 | -0.38424 | 1.1719 | -0.48143 | -1.0499 | -0.38628 | -0.70564 | -1.4401 | -1.0432 | 0.40847 | -0.24424 | -0.35565 | -0.48441 | 13.8086 | 0.1875 | a2 | 4.0477 | -0.49171 | -0.7941 | -0.067756 | -0.48445 | -0.50115 | 0.3806 | -0.19886 | -0.31269 | -0.068797 | -0.92106 | -0.2413 | -0.31512 | -0.29597 | 0.15566 | -0.50558 | -1.1322 | 0.55973 | -1.5753 | 8.8103 | 0.070641
a3 | -0.65792 | -0.024798 | -0.12829 | 0.11719 | 0.6878 | 0.14464 | 0.26616 | -0.063505 | 0.058834 | -0.093429 | 0.16469 | -0.061023 | -0.31233 | 0.37636 | -0.14431 | 0.020451 | 0.34737 | 0.72792 | 0.14049 | 0.056541 | 0.095827 | a3 | 2.9669 | 0.88403 | 0.36987 | 0.30249 | 0.69672 | 0.11519 | 0.034489 | -0.26498 | -0.24904 | -0.1211 | 0.089953 | -0.06771 | 0.017497 | 0.22697 | 0.046279 | -0.18648 | 0.041618 | 0.52616 | 0.065699 | 0.44751 | 0.24368
a4 | 1.4258 | 0.18044 | -0.37873 | -0.025842 | -0.25584 | -0.42264 | -0.23975 | -0.12203 | 0.28894 | 0.35079 | 0.045963 | -0.030229 | 0.16653 | 0.70039 | 0.40445 | -0.018209 | 0.21336 | 0.17406 | -0.06706 | 1.6884 | 0.15272 | a4 | 0.53416 | -0.25441 | -0.55546 | -0.413 | -0.2116 | -0.50822 | -0.60527 | -0.20881 | -0.19654 | -0.40946 | -0.38788 | -0.56195 | -0.23517 | -0.099675 | -0.20117 | -0.37103 | -0.012279 | -0.87534 | -0.3052 | 0.58002 | -0.29602
a5 | 2.5314 | -0.037203 | -0.076254 | -0.041614 | -0.14298 | -0.054386 | -0.033659 | 0.023986 | 0.035269 | -0.16124 | 0.10835 | 0.14397 | 0.15676 | 0.096966 | -0.11679 | -0.082363 | -0.076137 | 0.1597 | 0.092864 | -0.16808 | 0.076937 | a5 | 3.5672 | 0.026794 | -0.077014 | 0.16501 | 0.046696 | -0.010895 | 0.056264 | 0.15226 | -0.029697 | 0.094819 | -0.15523 | 0.053449 | 0.055155 | 0.091655 | -0.00097531 | 0.010418 | 0.1398 | 0.5238 | 0.056368 | -0.15175 | 0.17474
a6 | 0.31231 | -0.12901 | -0.11202 | -0.057501 | -0.017567 | 0.012351 | 0.029543 | 0.057562 | 0.047528 | 0.026105 | -0.042778 | 0.0052962 | 0.026739 | -0.12627 | -0.0012065 | 0.039169 | -0.021852 | -0.044131 | -0.095937 | -0.16428 | -0.016211 | a6 | 0.35048 | -0.066655 | -0.060307 | 0.02972 | 0.0083842 | 0.017293 | -0.11133 | -0.044346 | 0.0419 | 0.031857 | -0.077095 | -0.055481 | -0.069919 | -0.045509 | -0.053684 | -0.054689 | -0.02242 | 0.013962 | 0.025273 | 0.099688 | -0.010747
a7 | 1.6519 | -0.072986 | -0.086441 | -0.027023 | -0.019404 | -0.075284 | -0.028438 | -0.023167 | -0.050071 | -0.064932 | -0.071348 | -0.11931 | -0.071664 | -0.065926 | -0.09391 | -0.086936 | -0.065138 | -0.04323 | -0.049848 | 0.036081 | -0.0026179 | a7 | 1.9967 | -0.025845 | -0.065265 | -0.015107 | 0.031127 | -0.042426 | -0.086034 | -0.089839 | -0.0071468 | -0.0017477 | -0.075173 | -0.031669 | -0.0093378 | -0.025267 | -0.041899 | -0.08062 | -0.010079 | -0.091631 | -0.047563 | 0.060495 | 0.03063
a8 | 0.10599 | -0.044853 | -0.029481 | -0.011628 | -0.016858 | 0.021088 | -0.012327 | -0.03516 | 0.013051 | 0.084061 | 0.0277 | 0.064241 | 0.012523 | -0.013861 | 0.023652 | 0.029253 | -0.0026149 | -0.031337 | -0.0070722 | 0.0042709 | 0.0072567 | a8 | 0.062046 | 0.008259 | -0.025965 | -0.060184 | -0.050207 | -0.036251 | -0.048282 | -0.029882 | -0.037626 | -0.0061873 | -0.056407 | -0.058369 | -0.050586 | -0.026944 | -0.011828 | -0.013973 | -0.032474 | 0.0070707 | -0.010715 | 0.01359 | -0.025555
a9 | -0.2145 | 0.0089127 | -0.025923 | -0.045346 | -0.041051 | -0.049031 | -0.039355 | -0.040861 | -0.021169 | -0.054037 | -0.033574 | -0.033958 | -0.040088 | -0.025012 | -0.057177 | -0.049496 | -0.054724 | -0.014019 | -0.0058878 | -0.10461 | -0.042729 | a9 | -0.13377 | 0.039744 | -0.022585 | -0.029239 | -0.028699 | -0.035865 | -0.029271 | -0.02619 | -0.025512 | -0.026084 | -0.042753 | -0.043721 | -0.021813 | -0.0037321 | -0.038512 | -0.032615 | -0.037422 | -0.0040588 | -0.034652 | -0.070656 | -0.029736
a10 | 0.028462 | -0.046265 | -0.020071 | 0.0069602 | -0.0014125 | 0.020099 | 0.0061632 | 0.022044 | 0.030151 | 0.083466 | 0.03707 | 0.043186 | 0.029162 | -0.019494 | -0.011988 | -0.026965 | -0.021126 | -0.036725 | -0.033513 | -0.12289 | 0 | a10 | 0.073376 | 0.00026769 | -0.0051571 | 0.011297 | -0.0026591 | 0.011978 | -0.021481 | 0.0082686 | -0.0021632 | 0.0052793 | -0.017764 | -0.038006 | -0.030635 | -0.021575 | 0.010494 | 0.011697 | -0.0053512 | -0.0027242 | 0.062103 | -0.032231 | 0
a11 | 0.16887 | -0.039522 | -0.045794 | -0.036654 | -0.044237 | -0.034303 | -0.041891 | -0.014441 | -0.035598 | -0.032192 | -0.048871 | -0.041407 | -0.038728 | -0.025773 | -0.046991 | -0.038573 | -0.050045 | -0.031924 | -0.045626 | -0.029318 | -0.031352 | a11 | 0.19193 | -0.056738 | -0.06093 | -0.041849 | -0.044225 | -0.039993 | -0.045887 | -0.025586 | -0.027841 | -0.024931 | -0.047154 | -0.036345 | -0.034969 | -0.021411 | -0.045629 | -0.039349 | -0.042424 | -0.034387 | -0.033764 | -0.028913 | -0.031155
a12 | 0.0072074 | -0.011666 | -0.0081624 | -0.004117 | -0.00075648 | -0.00068353 | -0.0020569 | 0.0034341 | 0.0017899 | 0.0055448 | -0.0037886 | -0.0014767 | -1.0117e-06 | -0.015963 | -0.010141 | -0.0077641 | -0.0083724 | -0.01074 | -0.0093596 | -0.015338 | -0.004648 | a12 | 0.017172 | -0.00081075 | -0.0018627 | -0.00061022 | -0.0026936 | -0.00035836 | -0.0072396 | -0.0039239 | -0.0063784 | -0.0052025 | -0.0081902 | -0.0077816 | -0.0070284 | -0.0092198 | -0.0016269 | -0.0031689 | -0.0049505 | -0.0026497 | 0.0028917 | 0.00051108 | -0.0030913
a13 | -0.017275 | -0.00068057 | -0.0053369 | -0.0025908 | -0.0049367 | -0.0056694 | -0.0064517 | -0.0032836 | -0.0048776 | -0.0027238 | -0.0028987 | -0.0023629 | -0.0035128 | -0.0013424 | -0.0068113 | -0.0073559 | -0.007228 | -0.00068023 | -0.0039702 | -0.0085297 | -0.0046165 | a13 | -0.012997 | -3.1064e-05 | -0.0054641 | -0.0058855 | -0.0063287 | -0.0066389 | -0.0058561 | -0.0038264 | -0.0062548 | -0.0042936 | -0.0054611 | -0.0036454 | -0.0063073 | -0.0047942 | -0.0045803 | -0.0073398 | -0.0055007 | -0.0040401 | -0.0053617 | -0.0044879 | -0.0053327
a14 | 0.0040982 | -0.020538 | -0.01882 | -0.014847 | -0.016947 | -0.015197 | -0.016934 | -0.015929 | -0.014639 | -0.010462 | -0.014803 | -0.012443 | -0.014984 | -0.016091 | -0.014516 | -0.014699 | -0.015381 | -0.019051 | -0.01682 | -0.010059 | -0.014892 | a14 | 0.0036666 | -0.014389 | -0.015001 | -0.013408 | -0.013916 | -0.013515 | -0.016915 | -0.012515 | -0.015726 | -0.015651 | -0.017732 | -0.017794 | -0.015931 | -0.014653 | -0.015727 | -0.013456 | -0.015069 | -0.014581 | -0.011825 | -0.0038083 | -0.013902
a15 | 0.0031201 | -0.00066866 | -0.0018064 | -0.0049504 | 0.0034805 | 0.0061546 | 0.0047333 | 0.0067539 | 0.0039755 | -0.005542 | -0.0063819 | -0.0086177 | -0.0031998 | -0.00010444 | 0.0057125 | 0.01084 | 0.0073331 | 0.0018699 | 0.0052866 | 0.011558 | 0.001779 | a15 | 0.0020479 | -0.0026039 | -0.0089856 | -0.0084706 | 0.0013193 | -0.0025736 | -0.0031822 | -0.0012838 | -0.0035855 | -0.0073619 | -0.0068822 | -0.0046008 | 0.0030356 | -4.5522e-05 | 0.00083582 | -0.00078703 | -0.0042995 | -0.012083 | -0.0048159 | -0.0066664 | -0.003598
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.17291 | -0.055399 | -0.072163 | -0.15899 | 0.093428 | 0.11559 | 0.09865 | 0.12387 | 0.1188 | -0.29582 | -0.24431 | -0.22892 | -0.090741 | -0.047471 | 0.12113 | 0.24799 | 0.16515 | -0.014084 | 0.14302 | 0.16165 |  | Dx (mm) | -0.037494 | 0.038522 | -0.11118 | -0.1278 | 0.11275 | 0 | 0.021267 | 0.046946 | 0.039091 | -0.13362 | -0.065999 | 0 | 0.17221 | 0.13143 | 0.12686 | 0.11874 | 0 | -0.22823 | -0.011016 | -0.12117
Dy (mm) | 0.12278 | -0.10779 | -0.050331 | 0 | 0 | 0.04992 | 0.019205 | 0.051396 | 0.073383 | 0.17758 | 0.067343 | 0.083995 | 0.060674 | -0.04453 | -0.019217 | -0.048498 | -0.037162 | -0.08359 | -0.067531 | -0.30748 |  | Dy (mm) | 0.22054 | 0 | -0.021105 | 0.018583 | 0 | 0.026997 | -0.047462 | 0.020302 | 0 | 0 | -0.044678 | -0.086432 | -0.061859 | -0.045402 | 0.031874 | 0.033249 | -0.012515 | -0.01688 | 0.14323 | -0.088921
Mag. len. (m) | 14.4184 | [20 empty] | Mag. len. (m) | 14.4192
TF (mT/KA) | 596.719 | [20 empty] | TF (mT/KA) | 596.8105
Coil pos. (mm) | 0.016807 | [20 empty] | Coil pos. (mm) | 0.016718
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5
Twist (m2 rad) | 0.0079501 | [20 empty] | Twist (m2 rad) | 0.0031151

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -1.1634 | 0.70852 | -1.269 | 0.49534 | 0.1875 | 3.3418 | -0.43157 | 0.66221 | 0.93788 | 0.63585 | 1.0209 | 0.50981 | 0.070641 | 2.2858 | -0.37875 | 0.52287 |  | 0
3 | 2.0825 | 8.183 | 0.97291 | 0.80861 | 0.095827 | 0.31437 | 0.12535 | 0.27642 | 2.0191 | 7.7292 | 0.91367 | 0.55559 | 0.24368 | 0.6988 | 0.13873 | 0.31192 |  | 0
4 | -0.021331 | 0.18689 | 0.014224 | 0.083098 | 0.15272 | 0.53886 | 0.05321 | 0.2883 | 0.096019 | 0.17116 | 0.097702 | 0.17905 | -0.29602 | 0.34416 | -0.35493 | 0.20584 |  | 0
5 | -0.36952 | 0.97425 | -0.17538 | 0.18892 | 0.076937 | 0.57795 | -0.00022618 | 0.10436 | -0.53951 | 1.0968 | -0.3146 | 0.14789 | 0.17474 | 0.79775 | 0.06707 | 0.13927 |  | 0
6 | 0.051371 | 0.1846 | 0.074695 | 0.067712 | -0.016211 | 0.10079 | -0.025606 | 0.050582 | -0.032219 | 0.099374 | -0.035236 | 0.082058 | -0.010747 | 0.097962 | -0.026537 | 0.039165 |  | 0
7 | 1.1563 | 0.43914 | 1.1007 | 0.079979 | -0.0026179 | 0.38354 | -0.060684 | 0.0263 | 1.1931 | 0.36472 | 1.1464 | 0.052403 | 0.03063 | 0.45597 | -0.03954 | 0.035955 |  | 0
8 | -0.0013602 | 0.061163 | 0.0019524 | 0.021207 | 0.0072567 | 0.039529 | 0.0024154 | 0.024194 | -0.018086 | 0.028509 | -0.019928 | 0.017292 | -0.025555 | 0.030377 | -0.029325 | 0.019731 |  | 0
9 | 0.48878 | 0.015949 | 0.48705 | 0.0088467 | -0.042729 | 0.045671 | -0.032007 | 0.016902 | 0.46269 | 0.030426 | 0.46569 | 0.012269 | -0.029736 | 0.031832 | -0.024403 | 0.019789 |  | 0
10 | 0 | 0.069542 | 3.476e-05 | 8.0411e-05 | 0 | 0.043231 | 0.000239 | 0.00086071 | 0 | 0.046839 | -5.2072e-05 | 0.00049639 | 0 | 0.027704 | 0.0002123 | 0.00094053 |  | 0
11 | 0.7494 | 0.034222 | 0.75575 | 0.0079217 | -0.031352 | 0.047048 | -0.038166 | 0.0086852 | 0.7356 | 0.037443 | 0.74294 | 0.0052746 | -0.031155 | 0.052539 | -0.039009 | 0.010676 |  | 0
12 | -0.0037197 | 0.010796 | -0.0033377 | 0.0029811 | -0.004648 | 0.0066429 | -0.0048664 | 0.0032876 | 0.0014059 | 0.0071202 | 0.0015295 | 0.0025489 | -0.0030913 | 0.0056943 | -0.0037068 | 0.0017663 |  | 0
13 | 0.074188 | 0.0069204 | 0.073908 | 0.0029858 | -0.0046165 | 0.0036904 | -0.0038273 | 0.0015753 | 0.074537 | 0.0050769 | 0.074593 | 0.0024028 | -0.0053327 | 0.0023526 | -0.0050057 | 0.0018792 |  | 0
14 | 0.0014561 | 0.0049846 | 0.0011244 | 0.00092059 | -0.014892 | 0.0050553 | -0.016486 | 0.0011558 | -0.001139 | 0.0032678 | -0.001351 | 0.0010522 | -0.013902 | 0.0049675 | -0.015 | 0.0010958 |  | 0
15 | 0.03615 | 0.0088671 | 0.03751 | 0.0025802 | 0.001779 | 0.0057054 | 0.0013816 | 0.0055253 | 0.03297 | 0.0098732 | 0.034828 | 0.0026577 | -0.003598 | 0.0040385 | -0.0036873 | 0.0039741 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.55733 | -0.66106
3 | [12 empty] | 3 | 9 | -0.74913 | -0.87637
4 | [12 empty] | 4 | 16 | -1.8657 | -1.6504
5 | [12 empty] | 5 | 25 | -1.9632 | -1.9413
6 | [12 empty] | 6 | 36 | -2.8383 | -2.8702
7 | [12 empty] | 7 | 49 | -3.0821 | -3.1371
8 | [12 empty] | 8 | 64 | -3.7875 | -3.9915
9 | [12 empty] | 9 | 81 | -4.404 | -4.1616
10 | [12 empty] | sigma (mm) |  | 0.016807 | 0.016718
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | 1.7016 | -1.6847 | -1.0463 | -0.65201 | -0.80761 | -1.2355 | -1.2274 | -0.99771 | -1.3632 | -1.1146 | -1.7524 | -2.0418 | -2.4291 | -1.2337 | -0.88472 | -0.47886 | -0.86513 | -1.4329 | -1.5947 | -0.34521 | -1.1634
3 | 37.2281 | 0.53324 | 1.0757 | 0.27669 | 0.63153 | 0.29161 | 1.3454 | 0.87349 | 1.9656 | 2.1484 | 0.78495 | 0.52598 | 0.64112 | -0.72851 | -0.027274 | 1.3263 | 2.1555 | 1.7035 | 1.9891 | -1.3856 | 2.0825
4 | -0.77477 | 0.052546 | 0.080008 | 0.12125 | 0.12993 | 0.030301 | -0.012891 | -0.14174 | -0.094944 | -0.0038382 | 0.088434 | -0.10554 | -0.0728 | -0.021669 | 0.037398 | 0.11595 | 0.036994 | -0.058779 | 0.075422 | -0.22259 | -0.021331
5 | -4.0109 | -0.46122 | -0.66549 | -0.145 | -0.23185 | -0.02781 | -0.053411 | -0.1446 | -0.21676 | -0.099941 | 0.17729 | 0.090188 | -0.24898 | -0.19553 | -0.21364 | -0.30703 | -0.226 | -0.10429 | -0.082755 | -2.3136 | -0.36952
6 | -0.39355 | 0.11702 | 0.096332 | 0.046205 | 0.11672 | 0.17853 | 0.022972 | 0.04621 | 0.024808 | -0.044403 | 0.0046336 | 0.12777 | 0.226 | 0.12511 | 0.070187 | 0.044316 | -0.011824 | 0.060104 | 0.093834 | -0.004715 | 0.051371
7 | 2.9992 | 0.92753 | 1.0561 | 1.1961 | 1.1226 | 1.1769 | 1.1239 | 1.0411 | 1.1111 | 1.1374 | 1.101 | 1.0092 | 1.0941 | 1.0059 | 1.004 | 1.1283 | 1.1563 | 1.232 | 1.19 | 0.84552 | 1.1563
8 | -0.16443 | 0.027252 | 0.011478 | 0.019038 | -0.0097085 | 0.039318 | 0.0068216 | 0.0050681 | -0.025392 | 0.0091297 | -0.010709 | -0.031558 | 0.015771 | -0.015954 | 0.024109 | 0.014246 | -0.0015181 | -0.0018379 | -0.040411 | 0.046846 | -0.0013602
9 | 0.43155 | 0.47895 | 0.47689 | 0.48699 | 0.4943 | 0.48394 | 0.47665 | 0.47996 | 0.49405 | 0.49169 | 0.4779 | 0.47702 | 0.48295 | 0.49057 | 0.48135 | 0.4901 | 0.49806 | 0.50338 | 0.50212 | 0.48235 | 0.48878
10 | 3.1726e-05 | -3.1993e-05 | -7.345e-06 | 0.00022611 | 6.4975e-05 | -5.4478e-06 | -5.1113e-06 | -6.086e-06 | -1.0099e-06 | 9.2307e-05 | 0.00015891 | 0.0001869 | 1.0033e-05 | -1.2136e-05 | -1.3858e-06 | -7.1056e-05 | 2.7564e-06 | -7.6248e-06 | 3.288e-05 | 0.00025481 | 0
11 | 0.63227 | 0.75048 | 0.7439 | 0.75241 | 0.75708 | 0.76363 | 0.76023 | 0.76409 | 0.75605 | 0.74529 | 0.75634 | 0.76005 | 0.75904 | 0.77215 | 0.7666 | 0.75009 | 0.74559 | 0.7526 | 0.74781 | 0.66609 | 0.7494
12 | -0.1608 | -0.040407 | -0.025862 | -0.053663 | -0.066688 | -0.029504 | -0.016191 | 0.0083 | -0.025263 | -0.026072 | -0.040086 | -0.11634 | -0.055157 | -0.0058568 | -0.00449 | 0.0049121 | -0.027314 | -0.021652 | -0.059445 | 0.0047713 | -0.037197 | ten times!
13 | 0.90466 | 0.78417 | 0.80432 | 0.75411 | 0.7499 | 0.72091 | 0.75103 | 0.71933 | 0.73973 | 0.72082 | 0.69302 | 0.681 | 0.76424 | 0.72126 | 0.72419 | 0.7269 | 0.74511 | 0.76056 | 0.7428 | 0.52093 | 0.74188 | ten times!
14 | -0.0086317 | 0.014212 | 0.0069712 | 0.0055228 | -0.0011831 | 0.011855 | 0.013981 | 0.021515 | 0.011859 | 0.0048046 | 0.0089841 | -0.010805 | 0.0025311 | 0.027537 | 0.012355 | 0.017572 | 0.01971 | 0.022375 | 0.012586 | -0.0094235 | 0.014561 | ten times!
15 | -0.0064037 | 0.43959 | 0.38468 | 0.35798 | 0.38212 | 0.35022 | 0.36833 | 0.37297 | 0.35569 | 0.32474 | 0.3451 | 0.36953 | 0.36619 | 0.40606 | 0.3982 | 0.37033 | 0.37492 | 0.39418 | 0.39099 | 0.34113 | 0.3615 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 3.4187 | -0.52011 | -0.39795 | -1.271 | -1.0598 | -0.088382 | 0.60858 | -0.38858 | 1.1541 | -0.52933 | -1.0514 | -0.39317 | -0.71354 | -1.4417 | -1.0412 | 0.41546 | -0.24156 | -0.33767 | -0.47096 | 13.7604 | 0.1875
3 | -0.59956 | -0.021936 | -0.13231 | 0.11647 | 0.69202 | 0.15299 | 0.27038 | -0.059512 | 0.054092 | -0.075398 | 0.16576 | -0.060568 | -0.30892 | 0.38208 | -0.15285 | 0.022084 | 0.34133 | 0.72754 | 0.14307 | -0.0091637 | 0.095827
4 | 1.6441 | 0.16811 | -0.38794 | -0.027398 | -0.25269 | -0.42092 | -0.23874 | -0.12101 | 0.29167 | 0.34322 | 0.049268 | -0.024189 | 0.17351 | 0.6994 | 0.40683 | -0.016782 | 0.21436 | 0.17251 | -0.071428 | 1.5265 | 0.15272
5 | 2.5708 | -0.036257 | -0.077626 | -0.044303 | -0.1425 | -0.058154 | -0.034868 | 0.020494 | 0.03204 | -0.15066 | 0.10704 | 0.1431 | 0.15404 | 0.096607 | -0.11611 | -0.08319 | -0.074476 | 0.16082 | 0.099932 | -0.15595 | 0.076937
6 | 0.2828 | -0.095234 | -0.095474 | -0.059017 | -0.016927 | -0.0053624 | 0.022907 | 0.039644 | 0.020856 | -0.052008 | -0.075239 | -0.034426 | 0.0010192 | -0.11156 | 0.0096363 | 0.066226 | -0.0028562 | -0.0080284 | -0.065072 | -0.074031 | -0.016211
7 | 1.6689 | -0.073331 | -0.087428 | -0.027784 | -0.018755 | -0.077664 | -0.02817 | -0.022098 | -0.050803 | -0.050564 | -0.06673 | -0.11039 | -0.071103 | -0.0667 | -0.094684 | -0.088487 | -0.064399 | -0.043609 | -0.049614 | 0.044312 | -0.0026179
8 | 0.063306 | -0.02042 | -0.019049 | -0.01502 | -0.015054 | 0.012323 | -0.014828 | -0.044461 | -0.0029048 | 0.034446 | 0.0082981 | 0.041321 | -0.0030054 | -0.0041379 | 0.031294 | 0.046523 | 0.01045 | -0.011691 | 0.0093902 | 0.079848 | 0.0072567
9 | -0.20691 | 0.0060164 | -0.027568 | -0.04476 | -0.040981 | -0.051661 | -0.040131 | -0.043834 | -0.025178 | -0.029052 | -0.024965 | -0.024251 | -0.037429 | -0.026026 | -0.055857 | -0.043169 | -0.051473 | -0.014638 | -0.0011656 | -0.084648 | -0.042729
10 | -0.00019217 | -2.0257e-06 | 1.1213e-05 | 0.003532 | 0.0010186 | 3.3049e-07 | 2.6029e-06 | -1.8931e-05 | -7.1802e-06 | -0.00028041 | -2.2373e-05 | -0.00011405 | -1.7252e-05 | 2.5586e-06 | 2.9409e-05 | 0.00011425 | 4.3614e-05 | -1.0776e-05 | 2.0365e-05 | -8.985e-05 | 0
11 | 0.17171 | -0.040435 | -0.046397 | -0.037078 | -0.044192 | -0.034356 | -0.041806 | -0.014961 | -0.035927 | -0.029071 | -0.048767 | -0.040385 | -0.038323 | -0.026351 | -0.046125 | -0.036913 | -0.04901 | -0.032181 | -0.044701 | -0.02684 | -0.031352 | ten times!
12 | -0.027208 | -0.058529 | -0.055991 | -0.044077 | -0.0043092 | -0.02795 | -0.026336 | 0.010693 | -0.016962 | -0.044254 | -0.076697 | -0.06021 | -0.035359 | -0.1376 | -0.085729 | -0.039433 | -0.055625 | -0.063323 | -0.054272 | -0.033964 | -0.04648 | ten times!
13 | -0.16631 | -0.016037 | -0.064394 | -0.04396 | -0.037259 | -0.044889 | -0.052376 | -0.019975 | -0.03796 | -0.041914 | -0.053713 | -0.040625 | -0.044448 | -0.018869 | -0.05389 | -0.042067 | -0.050782 | -0.0076339 | -0.018126 | -0.057518 | -0.046165 | ten times!
14 | 0.046072 | -0.16666 | -0.17333 | -0.15495 | -0.17215 | -0.17223 | -0.17901 | -0.18196 | -0.17178 | -0.16561 | -0.18001 | -0.16623 | -0.17053 | -0.14605 | -0.14455 | -0.15434 | -0.15231 | -0.16316 | -0.15268 | -0.029599 | -0.14892 | ten times!
15 | 0.031201 | -0.0066866 | -0.018064 | -0.049504 | 0.034805 | 0.061546 | 0.047333 | 0.067539 | 0.039755 | -0.05542 | -0.063819 | -0.086177 | -0.031998 | -0.0010444 | 0.057125 | 0.1084 | 0.073331 | 0.018699 | 0.052866 | 0.11558 | 0.01779 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -0.35183 | 1.1087 | 1.1736 | 0.67534 | 1.6398 | 1.5685 | 1.5281 | 0.93501 | 0.58837 | -0.23671 | 0.96106 | 0.78238 | 0.57065 | 1.1762 | 0.92606 | 0.92139 | 0.99162 | 2.1331 | 0.93336 | 0.023493 | 0.93788
3 | 35.3644 | 0.15422 | 0.6026 | 0.54734 | 1.298 | 0.17673 | 1.5365 | 0.99632 | 0.93474 | 0.81757 | 0.32026 | 0.93713 | 0.7839 | 0.70131 | 0.84172 | 1.5637 | 1.6406 | 0.38006 | 2.2133 | 0.046046 | 2.0191
4 | 0.28356 | 0.40139 | 0.2033 | 0.0013007 | 0.21664 | 0.31216 | 0.30812 | 0.28083 | 0.042655 | -0.1492 | 0.015728 | 0.2127 | -0.017953 | 0.010907 | 0.1265 | -0.19049 | -0.15345 | 0.21177 | -0.074271 | -0.030403 | 0.096019
5 | -4.5151 | -0.41765 | -0.65191 | -0.43009 | -0.35324 | -0.22051 | -0.29843 | -0.28539 | -0.37475 | -0.10543 | -0.013746 | -0.23057 | -0.25679 | -0.41361 | -0.41913 | -0.44168 | -0.3322 | -0.12294 | -0.29479 | -2.9598 | -0.53951
6 | 0.10598 | -0.0098002 | 0.0041088 | -0.06518 | -0.0047381 | -0.067448 | -0.14305 | -0.12039 | -0.015249 | 0.14169 | -0.062924 | -0.043345 | 0.10163 | 0.017517 | -0.035714 | -0.071053 | -0.10291 | -0.19455 | 0.037144 | 0.066293 | -0.032219
7 | 2.7458 | 0.99352 | 1.0476 | 1.1821 | 1.1611 | 1.1876 | 1.1193 | 1.1707 | 1.1712 | 1.1486 | 1.1676 | 1.1021 | 1.1463 | 1.1657 | 1.151 | 1.1685 | 1.1607 | 1.1928 | 1.199 | 0.948 | 1.1931
8 | -0.013362 | -0.012494 | -0.0088032 | -0.030454 | -0.049009 | -0.044208 | -0.0038104 | 0.00058132 | -0.025473 | -0.051767 | -0.00018937 | -0.039285 | -0.018165 | -0.012959 | -0.03181 | -0.0004543 | -0.0099951 | -0.012313 | -0.0080921 | 0.004301 | -0.018086
9 | 0.34923 | 0.48315 | 0.4708 | 0.46751 | 0.46405 | 0.45216 | 0.45961 | 0.46082 | 0.47239 | 0.44659 | 0.45459 | 0.46123 | 0.45919 | 0.45883 | 0.46727 | 0.47136 | 0.47104 | 0.46073 | 0.50102 | 0.44057 | 0.46269
10 | -0.00010969 | 0.00011989 | 8.5775e-06 | 8.7524e-06 | -6.2149e-06 | -0.00038086 | 3.7811e-06 | -2.1293e-06 | -5.7074e-05 | 0.00013777 | -7.6217e-06 | 0.0010169 | -2.1515e-05 | -9.7694e-08 | -3.4966e-05 | -2.5413e-05 | -0.0017371 | 4.1525e-05 | -1.554e-06 | 1.0149e-05 | 0
11 | 0.5998 | 0.73678 | 0.73654 | 0.74516 | 0.73977 | 0.75299 | 0.74899 | 0.75132 | 0.74305 | 0.7383 | 0.74589 | 0.74682 | 0.74527 | 0.74681 | 0.74079 | 0.73792 | 0.74045 | 0.74133 | 0.73475 | 0.65538 | 0.7356
12 | -0.07626 | -0.025578 | -0.0036387 | 0.020246 | 0.049679 | 0.03118 | 0.032696 | -0.00044301 | -0.0044899 | -0.026107 | 0.0058881 | -0.0078154 | 0.030878 | 0.018314 | 0.068568 | 0.038808 | 0.018087 | 0.032219 | -0.0031788 | 0.0067349 | 0.014059 | ten times!
13 | 0.85874 | 0.75163 | 0.76339 | 0.7595 | 0.7873 | 0.75223 | 0.75677 | 0.77584 | 0.74973 | 0.71623 | 0.73331 | 0.73947 | 0.76107 | 0.77549 | 0.69326 | 0.73214 | 0.72943 | 0.71456 | 0.73549 | 0.57645 | 0.74537 | ten times!
14 | -0.00090195 | -0.02901 | -0.028027 | -0.020506 | 0.0061256 | 0.0089037 | -0.013857 | -0.0084593 | -0.012834 | -0.0074283 | -0.013356 | -0.013842 | -0.005936 | -0.015491 | -0.017336 | -0.017738 | -0.017004 | -0.0070434 | -0.030339 | -0.01358 | -0.01139 | ten times!
15 | -0.056473 | 0.37247 | 0.35231 | 0.30739 | 0.34474 | 0.32473 | 0.36734 | 0.34788 | 0.35688 | 0.35704 | 0.37411 | 0.38751 | 0.38029 | 0.35415 | 0.36037 | 0.3348 | 0.34374 | 0.32036 | 0.28301 | 0.18934 | 0.3297 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 3.1438 | -0.49571 | -0.78778 | -0.064407 | -0.49369 | -0.50172 | 0.38907 | -0.19979 | -0.31155 | -0.070701 | -0.91869 | -0.23185 | -0.3098 | -0.29573 | 0.1518 | -0.50908 | -1.1298 | 0.57459 | -1.6126 | 8.8172 | 0.070641
3 | 2.965 | 0.88576 | 0.35979 | 0.29313 | 0.70093 | 0.1137 | 0.039334 | -0.26424 | -0.24768 | -0.13075 | 0.085548 | -0.064449 | 0.024347 | 0.22927 | 0.050141 | -0.17751 | 0.04128 | 0.49156 | 0.067012 | 0.46071 | 0.24368
4 | 0.80106 | -0.25465 | -0.56071 | -0.40618 | -0.21284 | -0.50682 | -0.60891 | -0.20913 | -0.19627 | -0.40648 | -0.39043 | -0.56666 | -0.24104 | -0.1069 | -0.198 | -0.36786 | -0.013257 | -0.84762 | -0.29507 | 0.5138 | -0.29602
5 | 3.5767 | 0.027549 | -0.079219 | 0.16678 | 0.046418 | -0.010362 | 0.055061 | 0.15347 | -0.030179 | 0.096071 | -0.15792 | 0.052322 | 0.06181 | 0.094365 | 0.00087501 | 0.012534 | 0.13942 | 0.52329 | 0.05498 | -0.14747 | 0.17474
6 | 0.1636 | -0.066304 | -0.055055 | 0.021262 | 0.0071456 | 0.0059737 | -0.091946 | -0.05125 | 0.041999 | 0.031775 | -0.060444 | -0.021926 | -0.044324 | -0.025651 | -0.064756 | -0.06504 | -0.017294 | 0.013722 | -0.03556 | 0.13208 | -0.010747
7 | 1.9982 | -0.025976 | -0.066747 | -0.017826 | 0.033457 | -0.04194 | -0.085609 | -0.089354 | -0.0065411 | -0.0020881 | -0.076986 | -0.033131 | -0.0052699 | -0.023518 | -0.04124 | -0.0803 | -0.010132 | -0.091405 | -0.047121 | 0.060299 | 0.03063
8 | 0.024251 | 0.0075385 | -0.022443 | -0.066062 | -0.048684 | -0.041993 | -0.037734 | -0.03371 | -0.037157 | -0.0078275 | -0.048164 | -0.03968 | -0.035283 | -0.016836 | -0.016592 | -0.019574 | -0.0297 | 0.01039 | -0.044348 | 0.028063 | -0.025555
9 | -0.12785 | 0.039739 | -0.023436 | -0.027928 | -0.02854 | -0.035869 | -0.028825 | -0.026622 | -0.025467 | -0.025711 | -0.044088 | -0.043759 | -0.016586 | -0.00085601 | -0.040165 | -0.034272 | -0.037436 | -0.0052815 | -0.034144 | -0.074994 | -0.029736
10 | 8.9055e-05 | 0.0015534 | 7.6469e-06 | 1.1263e-06 | 0.00027407 | 2.017e-05 | -6.877e-07 | 8.3056e-07 | -0.001523 | 0.0033198 | 5.3834e-06 | -5.8586e-05 | 6.4622e-05 | 3.8653e-05 | -1.1294e-05 | -3.4351e-06 | 9.9369e-05 | 3.2278e-05 | 1.164e-06 | 9.4567e-06 | 0
11 | 0.19347 | -0.056718 | -0.061152 | -0.041841 | -0.044028 | -0.040049 | -0.045659 | -0.025501 | -0.02768 | -0.02538 | -0.04759 | -0.036401 | -0.033703 | -0.020367 | -0.045768 | -0.039325 | -0.04241 | -0.034871 | -0.03371 | -0.029135 | -0.031155
12 | 0.035538 | -0.008099 | -0.0115 | -0.021466 | -0.021899 | -0.017972 | -0.046373 | -0.04912 | -0.062058 | -0.056075 | -0.061453 | -0.033418 | -0.029429 | -0.062986 | -0.027896 | -0.042833 | -0.043074 | -0.024371 | -0.047194 | 0.03751 | -0.030913 | ten times!
13 | -0.12844 | 0.0039282 | -0.068393 | -0.071226 | -0.051288 | -0.066585 | -0.056124 | -0.033852 | -0.057847 | -0.058928 | -0.064798 | -0.037585 | -0.03975 | -0.032422 | -0.031036 | -0.061567 | -0.055174 | -0.066739 | -0.051631 | -0.050948 | -0.053327 | ten times!
14 | 0.047555 | -0.14307 | -0.15211 | -0.1477 | -0.14038 | -0.14237 | -0.15424 | -0.13047 | -0.15611 | -0.16461 | -0.1673 | -0.15036 | -0.14424 | -0.13324 | -0.1676 | -0.14296 | -0.14715 | -0.16407 | -0.15207 | -0.03087 | -0.13902 | ten times!
15 | 0.020479 | -0.026039 | -0.089856 | -0.084706 | 0.013193 | -0.025736 | -0.031822 | -0.012838 | -0.035855 | -0.073619 | -0.068822 | -0.046008 | 0.030356 | -0.00045522 | 0.0083582 | -0.0078703 | -0.042995 | -0.12083 | -0.048159 | -0.066664 | -0.03598 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00056956 | -2.1906e-05 | -2.1906e-05 | -0.00021906 | -2.1906e-05 | -2.1906e-05 | -0.00021906 | -2.1906e-05 | -0.00021906 | -0.00021906 | -2.1906e-05 | -2.1906e-05 | -2.1906e-05 | -0.00021906 | -2.1906e-05 | -2.1906e-05 | 0.00017525 | 0.00056956 |  | 596.719
|  | 0.00041616 | 2.1903e-05 | -0.00017522 | -0.00017522 | 2.1903e-05 | 2.1903e-05 | 2.1903e-05 | -0.00017522 | 2.1903e-05 | -0.00017522 | 2.1903e-05 | 2.1903e-05 | -0.00017522 | 2.1903e-05 | -0.00017522 | 2.1903e-05 | 2.1903e-05 | 0.00041616 |  | 596.8105
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3651.1066
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.6765 | -3698.4276 | 6.408 | 0.49293 | 0.49293 | -1.4788 | 0.49293 | 0.49293 | -1.4788 | 0.49293 | -1.4788 | -1.4788 | 0.49293 | 0.49293 | 0.49293 | -1.4788 | 0.49293 | 0.49293 | 2.4646 | 6.408 | -3603.7857
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.076244 | 2.1666 | 0.099164 | -0.75486 | 0.55001 | -0.5207 | 0.76593 | -0.91304 | 1.6839 | -0.11376 | 0.73392 | -0.82299 | 0.51614 | -0.83946 | 0.61212 | -0.47982 | 0.75405 | -0.96459 | 0.64556 | -0.95153 | 0.013555
db2 | -1.2658 | 0.95401 | -0.42277 | 0.20964 | 0.60865 | 0.46512 | 0.033401 | 0.053361 | 0.28115 | -0.070513 | 0.07877 | -0.51038 | -0.78971 | -1.168 | 0.038086 | 0.38037 | 0.82584 | 0.44409 | -0.16282 | -0.29436 | -0.36795
db3 | 0.97295 | 37.2182 | -0.43684 | 0.098276 | -0.69981 | -0.33936 | -0.677 | 0.37303 | -0.10155 | 0.98681 | 1.164 | -0.19183 | -0.44196 | -0.3325 | -1.6958 | -0.99819 | 0.35744 | 1.1848 | 0.7333 | 1.0172 | -1.3032
db4 | 0.012648 | -0.6856 | 0.04504 | 0.077809 | 0.11399 | 0.11228 | 0.017742 | -0.026588 | -0.15882 | -0.11422 | -0.0036617 | 0.063461 | -0.1257 | -0.08225 | -0.031082 | 0.018247 | 0.085082 | 0.014992 | -0.067936 | 0.061616 | -0.32323
db5 | -0.17328 | -3.9892 | -0.29359 | -0.49531 | 0.029262 | -0.054345 | 0.15304 | 0.12151 | 0.031174 | -0.041961 | 0.085825 | 0.35782 | 0.26002 | -0.081311 | -0.025397 | -0.036357 | -0.12282 | -0.050289 | 0.067689 | 0.095043 | -2.3262
db6 | 0.075998 | -0.64866 | 0.020157 | -0.0080883 | -0.096977 | 0.077749 | 0.15198 | -0.013658 | 0.016258 | -0.0033151 | -0.23516 | -0.16511 | -0.026683 | 0.1165 | 0.031215 | 0.036589 | 0.066277 | -0.021132 | -0.023289 | 0.076686 | 0.047728
db7 | 1.1024 | 3.0072 | -0.17716 | -0.04671 | 0.094878 | 0.02067 | 0.07735 | 0.022813 | -0.058759 | 0.0084902 | 0.035787 | 0.0045432 | -0.087608 | -0.0084648 | -0.096159 | -0.09564 | 0.033864 | 0.056416 | 0.12864 | 0.087041 | 0.85054
db8 | 0.0027195 | -0.18757 | 0.012547 | -0.0080744 | -0.020561 | 0.0093951 | 0.064222 | 0.026691 | 0.031495 | 0.00036632 | -0.059192 | -0.068748 | -0.08543 | -0.0064668 | -0.030159 | 0.048501 | 0.069174 | 0.033967 | -0.0084348 | -0.0092922 | 0.067258
db9 | 0.49136 | 0.43672 | -0.014262 | -0.013746 | 0.0015424 | 0.0050045 | -0.0047115 | -0.012443 | -0.0083162 | 0.0049413 | 0.012332 | -0.0009394 | -0.0040667 | -0.0074604 | -0.0006851 | -0.0066604 | 0.012295 | 0.012526 | 0.010456 | 0.014192 | 0.46731
db10 | 0.00062072 | -0.076707 | -0.027597 | -0.033564 | -0.070898 | 0.041046 | 0.052322 | 0.043978 | 0.055521 | 0.053767 | -0.12734 | -0.10761 | -0.1012 | -0.039776 | -0.022859 | 0.053491 | 0.1079 | 0.070762 | -0.0084286 | 0.060489 | 0.058045
db11 | 0.75647 | 0.63456 | -0.0065464 | -0.012513 | -0.0026324 | 0.00050393 | 0.0074189 | 0.0040221 | 0.0080928 | -0.00019247 | -0.0083887 | 0.0025868 | 0.0070454 | 0.0031752 | 0.015311 | 0.010452 | -0.0045028 | -0.010462 | -0.0044277 | -0.0089431 | 0.6635
db12 | -0.0032426 | -0.025595 | -0.0038512 | -0.003598 | -0.010644 | 0.0015327 | 0.0064524 | 0.0069808 | 0.010583 | 0.0072751 | -0.014684 | -0.012894 | -0.019596 | -0.0070726 | 0.00021769 | 0.0089169 | 0.016435 | 0.0090645 | 0.00029344 | 0.0045882 | 0.0043922
db13 | 0.074467 | 0.090225 | 0.0022842 | 0.0053319 | 0.0014463 | 0.00072383 | -0.0014194 | 0.00121 | -0.0013707 | 0.00072963 | 0.00053324 | -0.0033398 | -0.0042635 | 0.0028133 | -0.0029311 | -0.0017705 | -0.00047316 | 0.00052785 | 0.00035008 | -0.00038212 | 0.050534
db14 | 0.0013637 | -0.0010875 | -0.0020074 | -0.0030275 | -0.0054988 | 0.001458 | 0.0029027 | 0.0029518 | 0.0043067 | 0.003062 | -0.007984 | -0.0070547 | -0.0088145 | -0.0036872 | -0.00020126 | 0.0039343 | 0.0083896 | 0.005931 | 0.00054531 | 0.004794 | 0.0065257
db15 | 0.03751 | -0.00064037 | 0.0064491 | 0.00095756 | -0.0017118 | 0.000702 | -0.0024883 | -0.00067767 | -0.00021329 | -0.001941 | -0.0050359 | -0.0030001 | -0.00055761 | -0.00089085 | 0.003096 | 0.0023097 | -0.00047683 | -1.7997e-05 | 0.0019082 | 0.0015889 | 0.034113
da2 | -0.42889 | 3.9697 | -0.097854 | 0.025689 | -0.84426 | -0.62335 | 0.34422 | 1.0436 | 0.044655 | 1.6008 | -0.052541 | -0.62105 | 0.042611 | -0.27675 | -1.0112 | -0.61432 | 0.83736 | 0.18465 | 0.073237 | -0.05552 | 13.8086
da3 | 0.12357 | -0.65792 | -0.14836 | -0.25186 | -0.0063761 | 0.56423 | 0.02107 | 0.14259 | -0.18707 | -0.064733 | -0.217 | 0.041119 | -0.18459 | -0.4359 | 0.25279 | -0.26787 | -0.10312 | 0.2238 | 0.60435 | 0.016928 | 0.056541
da4 | 0.053589 | 1.4258 | 0.12685 | -0.43231 | -0.07943 | -0.30942 | -0.47622 | -0.29334 | -0.17562 | 0.23535 | 0.2972 | -0.0076255 | -0.083818 | 0.11294 | 0.64681 | 0.35086 | -0.071798 | 0.15977 | 0.12047 | -0.12065 | 1.6884
da5 | -0.00026495 | 2.5314 | -0.036938 | -0.075989 | -0.041349 | -0.14272 | -0.054121 | -0.033394 | 0.024251 | 0.035534 | -0.16098 | 0.10862 | 0.14423 | 0.15702 | 0.097231 | -0.11653 | -0.082098 | -0.075872 | 0.15997 | 0.093129 | -0.16808
da6 | -0.022443 | 0.31231 | -0.10656 | -0.089581 | -0.035058 | 0.0048764 | 0.034794 | 0.051986 | 0.080005 | 0.069971 | 0.048548 | -0.020335 | 0.027739 | 0.049182 | -0.10382 | 0.021237 | 0.061612 | 0.00059142 | -0.021688 | -0.073494 | -0.16428
da7 | -0.061947 | 1.6519 | -0.011038 | -0.024494 | 0.034924 | 0.042544 | -0.013337 | 0.033509 | 0.03878 | 0.011876 | -0.0029851 | -0.0094004 | -0.057358 | -0.0097169 | -0.0039788 | -0.031962 | -0.024988 | -0.0031908 | 0.018717 | 0.0121 | 0.036081
da8 | 0.0039097 | 0.10599 | -0.048763 | -0.033391 | -0.015537 | -0.020768 | 0.017178 | -0.016236 | -0.03907 | 0.0091409 | 0.080151 | 0.02379 | 0.060331 | 0.0086132 | -0.017771 | 0.019742 | 0.025343 | -0.0065246 | -0.035247 | -0.010982 | 0.0042709
da9 | -0.034544 | -0.2145 | 0.043457 | 0.0086212 | -0.010802 | -0.0065071 | -0.014487 | -0.004811 | -0.0063169 | 0.013376 | -0.019493 | 0.00096987 | 0.00058615 | -0.0055433 | 0.0095323 | -0.022632 | -0.014952 | -0.02018 | 0.020526 | 0.028656 | -0.10461
da10 | 0.0033746 | 0.028462 | -0.04964 | -0.023446 | 0.0035856 | -0.0047871 | 0.016725 | 0.0027886 | 0.018669 | 0.026776 | 0.080091 | 0.033695 | 0.039812 | 0.025788 | -0.022869 | -0.015363 | -0.03034 | -0.024501 | -0.040099 | -0.036888 | -0.12289
da11 | -0.038476 | 0.16887 | -0.0010462 | -0.0073173 | 0.0018218 | -0.0057612 | 0.004173 | -0.0034151 | 0.024035 | 0.002878 | 0.0062841 | -0.010395 | -0.0029311 | -0.00025171 | 0.012703 | -0.0085143 | -9.7051e-05 | -0.011568 | 0.0065521 | -0.0071495 | -0.029318
da12 | -0.0046823 | 0.0072074 | -0.0069839 | -0.0034801 | 0.00056531 | 0.0039258 | 0.0039987 | 0.0026253 | 0.0081164 | 0.0064722 | 0.010227 | 0.00089365 | 0.0032056 | 0.0046813 | -0.011281 | -0.0054592 | -0.0030818 | -0.0036901 | -0.0060579 | -0.0046773 | -0.015338
da13 | -0.0040396 | -0.017275 | 0.0033591 | -0.0012972 | 0.0014489 | -0.00089704 | -0.0016298 | -0.0024121 | 0.00075607 | -0.00083799 | 0.0013158 | 0.0011409 | 0.0016768 | 0.00052682 | 0.0026972 | -0.0027717 | -0.0033162 | -0.0031883 | 0.0033594 | 6.9433e-05 | -0.0085297
da14 | -0.015728 | 0.0040982 | -0.0048098 | -0.0030919 | 0.0008806 | -0.0012193 | 0.00053059 | -0.0012064 | -0.00020098 | 0.0010884 | 0.0052661 | 0.00092433 | 0.0032852 | 0.00074358 | -0.00036276 | 0.001212 | 0.0010289 | 0.00034702 | -0.0033233 | -0.0010922 | -0.010059
da15 | 0.0013816 | 0.0031201 | -0.0020503 | -0.003188 | -0.006332 | 0.0020989 | 0.004773 | 0.0033518 | 0.0053723 | 0.002594 | -0.0069236 | -0.0077635 | -0.0099993 | -0.0045814 | -0.001486 | 0.0043309 | 0.0094584 | 0.0059515 | 0.0004883 | 0.003905 | 0.011558
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3645.302
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.7794 | -3650.2304 | 4.682 | 0.73926 | -1.2321 | -1.2321 | 0.73926 | 0.73926 | 0.73926 | -1.2321 | 0.73926 | -1.2321 | 0.73926 | 0.73926 | -1.2321 | 0.73926 | -1.2321 | 0.73926 | 0.73926 | 4.682 | -3640.3736
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.19692 | 5.6526 | -1.0648 | -1.3031 | -0.4557 | -0.23782 | 0.69839 | 0.48868 | 1.5356 | 0.26132 | 0.5994 | -0.81431 | 0.38435 | -0.058915 | 0.80359 | -0.16525 | -0.20744 | -0.70695 | 0.22455 | 0.018467 | 0.40972
db2 | 1.0238 | -0.58471 | 0.085641 | 0.143 | -0.35732 | 0.63332 | 0.54436 | 0.50839 | -0.082589 | -0.43108 | -1.2734 | -0.064682 | -0.24206 | -0.43716 | 0.16452 | -0.085251 | -0.079813 | -0.032073 | 1.1005 | -0.094376 | 0.027404
db3 | 0.91304 | 35.3426 | -0.75612 | -0.31683 | -0.36465 | 0.38906 | -0.73389 | 0.61951 | 0.086325 | 0.021975 | -0.092036 | -0.596 | 0.015518 | -0.13254 | -0.21255 | -0.067632 | 0.64862 | 0.72752 | -0.54442 | 1.3081 | 0.054062
db4 | 0.096119 | 0.13846 | 0.30148 | 0.12383 | -0.082736 | 0.11119 | 0.21611 | 0.21112 | 0.18082 | -0.056911 | -0.2419 | -0.081822 | 0.11767 | -0.12337 | -0.096957 | 0.017911 | -0.29903 | -0.24915 | 0.12328 | -0.17152 | 0.050758
db5 | -0.31307 | -4.5368 | -0.10454 | -0.33799 | -0.11273 | -0.038801 | 0.092417 | 0.012229 | 0.026555 | -0.061674 | 0.2042 | 0.30008 | 0.081085 | 0.063674 | -0.098397 | -0.10496 | -0.12881 | -0.019216 | 0.20977 | 0.017094 | -2.9582
db6 | -0.032447 | -0.085709 | 0.036154 | -0.0050909 | -0.086115 | 0.073867 | -0.034597 | -0.10365 | -0.067898 | 0.033352 | 0.11976 | -0.058846 | -0.011864 | 0.20355 | 0.10365 | 0.048739 | 0.011373 | -0.070507 | -0.25921 | 0.067331 | 0.027579
db7 | 1.1472 | 2.7403 | -0.15369 | -0.098196 | 0.038491 | 0.012776 | 0.040821 | -0.028833 | 0.024021 | 0.02375 | 0.00588 | 0.019833 | -0.047169 | -0.00081903 | 0.018926 | 0.00383 | 0.022825 | 0.013406 | 0.051593 | 0.052559 | 0.94969
db8 | -0.018326 | -0.0055027 | 0.014596 | -0.015481 | -0.040209 | -0.0059055 | -0.025427 | 0.018457 | 0.029347 | 0.0015495 | -0.06178 | 0.0030992 | -0.022738 | 0.037206 | 0.033879 | 0.015204 | 0.044893 | 0.0081103 | -0.044798 | 0.0099964 | -0.023866
db9 | 0.46769 | 0.34142 | 0.015803 | 0.0059295 | 0.0034657 | -0.00070719 | -0.015701 | -0.0085344 | -0.0064433 | 0.0050452 | -0.017007 | -0.012473 | -0.008207 | -0.0027311 | -0.0053989 | 0.0031707 | 0.0067562 | 0.0033149 | 0.0050929 | 0.028625 | 0.44215
db10 | 0.0030513 | -0.038188 | 0.013764 | -0.051997 | -0.058636 | 0.046061 | -0.0027971 | 0.005039 | 0.018004 | 0.013978 | -0.061008 | -0.033252 | -0.0038823 | 0.071258 | 0.05415 | 0.053134 | 0.049295 | -0.0051007 | -0.10304 | -0.0049677 | -0.048241
db11 | 0.74329 | 0.59754 | -0.0065304 | -0.0063124 | 0.0022754 | -0.0027004 | 0.0096974 | 0.0055283 | 0.0081565 | -0.00020705 | -0.0041762 | 0.0024849 | 0.0030884 | 0.0030391 | 0.0039945 | -0.0013825 | -0.0045191 | -0.0028851 | -0.00073935 | -0.0088128 | 0.65576
db12 | 0.0019433 | -0.0078592 | -0.0024586 | -0.0084018 | -0.0067558 | 0.0093168 | 0.0013012 | 0.0022567 | 0.00064531 | -0.00032363 | -0.011354 | -0.0049233 | -0.0029472 | 0.010108 | 0.0069029 | 0.011276 | 0.0082867 | -0.0001832 | -0.010445 | -0.0023013 | -0.0045459
db13 | 0.074752 | 0.085264 | 0.00035948 | 0.0018699 | 0.001892 | 0.0043062 | 0.00074918 | 0.00029654 | 0.0030447 | 0.00021704 | -0.0026524 | -0.0018442 | -0.0019811 | 0.0012537 | 0.0025169 | -0.0048482 | -0.0010584 | -0.0019539 | -0.0022818 | 0.00011423 | 0.057732
db14 | -0.0010766 | -0.00028777 | -0.00064276 | -0.0051079 | -0.0040796 | 0.0048901 | 0.0020242 | 0.00020987 | 0.0015971 | 0.00094209 | -0.0035952 | -0.0025456 | -0.0006351 | 0.0060308 | 0.0033589 | 0.003086 | 0.0025983 | -0.00066809 | -0.0058171 | -0.001646 | -0.0037355
db15 | 0.034828 | -0.0056473 | 0.0024184 | 0.00040226 | -0.0040896 | -0.00035485 | -0.0023553 | 0.0019054 | -4.0493e-05 | 0.0008594 | 0.000876 | 0.0025829 | 0.0039224 | 0.0032002 | 0.00058619 | 0.001209 | -0.0013481 | -0.00045403 | -0.0027921 | -0.0065276 | 0.018934
da2 | -0.37834 | 4.0477 | -0.11337 | -0.41576 | 0.31058 | -0.10611 | -0.12282 | 0.75894 | 0.17947 | 0.065644 | 0.30954 | -0.54272 | 0.13704 | 0.063212 | 0.08237 | 0.534 | -0.12724 | -0.75389 | 0.93807 | -1.197 | 8.8103
da3 | 0.14043 | 2.9669 | 0.7436 | 0.22945 | 0.16206 | 0.55629 | -0.025236 | -0.10594 | -0.4054 | -0.38946 | -0.26152 | -0.050472 | -0.20814 | -0.12293 | 0.086543 | -0.094147 | -0.3269 | -0.098808 | 0.38574 | -0.074726 | 0.44751
da4 | -0.35625 | 0.53416 | 0.10184 | -0.19922 | -0.056753 | 0.14464 | -0.15198 | -0.24902 | 0.14743 | 0.15971 | -0.053209 | -0.031633 | -0.2057 | 0.12107 | 0.25657 | 0.15508 | -0.014779 | 0.34397 | -0.51909 | 0.051052 | 0.58002
da5 | 0.066593 | 3.5672 | -0.039799 | -0.14361 | 0.098421 | -0.019897 | -0.077488 | -0.010329 | 0.085665 | -0.09629 | 0.028225 | -0.22182 | -0.013144 | -0.011438 | 0.025061 | -0.067569 | -0.056176 | 0.073204 | 0.45721 | -0.010226 | -0.15175
da6 | -0.027391 | 0.35048 | -0.039264 | -0.032916 | 0.057111 | 0.035775 | 0.044685 | -0.083934 | -0.016955 | 0.069292 | 0.059248 | -0.049703 | -0.02809 | -0.042528 | -0.018118 | -0.026292 | -0.027298 | 0.0049713 | 0.041353 | 0.052664 | 0.099688
da7 | -0.039751 | 1.9967 | 0.013906 | -0.025513 | 0.024644 | 0.070878 | -0.0026748 | -0.046283 | -0.050088 | 0.032605 | 0.038004 | -0.035421 | 0.0080819 | 0.030413 | 0.014484 | -0.0021475 | -0.040868 | 0.029672 | -0.051879 | -0.0078119 | 0.060495
da8 | -0.03003 | 0.062046 | 0.038289 | 0.0040654 | -0.030153 | -0.020176 | -0.0062205 | -0.018252 | 0.00014886 | -0.0075956 | 0.023843 | -0.026377 | -0.028338 | -0.020555 | 0.0030868 | 0.018202 | 0.016058 | -0.0024435 | 0.037101 | 0.019316 | 0.01359
da9 | -0.02461 | -0.13377 | 0.064354 | 0.0020247 | -0.0046291 | -0.0040887 | -0.011255 | -0.0046605 | -0.0015802 | -0.0009018 | -0.0014745 | -0.018143 | -0.019111 | 0.0027974 | 0.020878 | -0.013902 | -0.0080055 | -0.012812 | 0.020551 | -0.010042 | -0.070656
da10 | -0.0014517 | 0.073376 | 0.0017194 | -0.0037054 | 0.012749 | -0.0012074 | 0.01343 | -0.020029 | 0.0097202 | -0.00071152 | 0.006731 | -0.016312 | -0.036555 | -0.029184 | -0.020123 | 0.011946 | 0.013149 | -0.0038995 | -0.0012725 | 0.063555 | -0.032231
da11 | -0.039078 | 0.19193 | -0.01766 | -0.021852 | -0.0027705 | -0.0051463 | -0.00091446 | -0.006809 | 0.013492 | 0.011237 | 0.014148 | -0.0080761 | 0.0027332 | 0.0041098 | 0.017667 | -0.0065503 | -0.00027041 | -0.0033454 | 0.0046916 | 0.0053145 | -0.028913
da12 | -0.0039336 | 0.017172 | 0.0031228 | 0.0020708 | 0.0033234 | 0.00124 | 0.0035752 | -0.003306 | 9.6531e-06 | -0.0024448 | -0.0012689 | -0.0042566 | -0.003848 | -0.0030948 | -0.0052862 | 0.0023066 | 0.0007647 | -0.0010169 | 0.0012839 | 0.0068253 | 0.00051108
da13 | -0.0050894 | -0.012997 | 0.0050584 | -0.00037472 | -0.00079603 | -0.0012392 | -0.0015495 | -0.00076663 | 0.0012631 | -0.0011654 | 0.0007958 | -0.00037165 | 0.001444 | -0.0012178 | 0.00029525 | 0.00050908 | -0.0022504 | -0.00041123 | 0.0010493 | -0.0002723 | -0.0044879
da14 | -0.014878 | 0.0036666 | 0.00048873 | -0.00012299 | 0.0014698 | 0.00096233 | 0.0013631 | -0.0020371 | 0.0023628 | -0.00084824 | -0.00077257 | -0.0028544 | -0.0029163 | -0.001053 | 0.00022466 | -0.00084854 | 0.0014223 | -0.00019083 | 0.00029727 | 0.003053 | -0.0038083
da15 | -0.0036873 | 0.0020479 | 0.0010833 | -0.0052984 | -0.0047833 | 0.0050066 | 0.0011137 | 0.00050507 | 0.0024034 | 0.00010175 | -0.0036747 | -0.0031949 | -0.00091349 | 0.0067229 | 0.0036417 | 0.0045231 | 0.0029002 | -0.00061226 | -0.0083962 | -0.0011286 | -0.0066664
Field Colinearity | 0.3655
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.0904 | 0.02292 | -0.8311 | 0.47376 | -0.59695 | 0.68968 | -0.98929 | 1.6076 | -0.19001 | 0.65768 | -0.89924 | 0.43989 | -0.9157 | 0.53588 | -0.55606 | 0.67781 | -1.0408 | 0.56931 | -1.0278 | -0.062689
Twist (m2 mrad) | 15.6779 | 0.32661 | -11.2198 | 6.0405 | -7.1634 | 7.7589 | -10.3875 | 15.6742 | -1.7101 | 5.4258 | -6.7443 | 2.9693 | -5.4942 | 2.8134 | -2.5023 | 2.5418 | -3.1225 | 1.281 | -1.5417 | -0.023508
Twist Integral (I0) (m2 rad) | 0.0079501
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.4556 | -1.2617 | -1.5 | -0.65263 | -0.43474 | 0.50146 | 0.29176 | 1.3386 | 0.0644 | 0.40248 | -1.0112 | 0.18743 | -0.25584 | 0.60667 | -0.36217 | -0.40436 | -0.90387 | 0.027631 | -0.17845 | 0.2128
Twist (m2 mrad) | 40.9173 | -17.9798 | -20.2505 | -8.321 | -5.2169 | 5.6415 | 3.0634 | 13.0517 | 0.5796 | 3.3205 | -7.5842 | 1.2652 | -1.535 | 3.185 | -1.6298 | -1.5163 | -2.7116 | 0.06217 | -0.26768 | 0.079799
Twist Integral (I0) (m2 rad) | 0.0031151

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-1.2658 | 0.97295 | 0.012648 | -0.17328 | 1.1024 | -0.76244 | -0.42889 | 0.12357 | 0.053589 | 16.8067 | 0.59672 | 1.0238 | 0.91304 | 0.096119 | -0.31307 | 1.1472 | -1.9692 | -0.37834 | 0.14043 | -0.35625 | 16.7182 | 0.59681 | -1.1634 | 2.0825 | -0.021331 | -0.36952 | 1.1563 | 0 | 0.1875 | 0.095827 | 0.15272 | 14418.357 | 0.93788 | 2.0191 | 0.096019 | -0.53951 | 1.1931 | 0 | 0.070641 | 0.24368 | -0.29602 | 14419.207
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8