## Alarm sheet
File name | HCMB__A001-03000057_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000057
Date of test Ap 1 | 2003-08-01 |  |  |  | Date of test Ap 2 | 2003-08-01
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | yellow alarm | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Original data
File name |  | HCMB__A001-03000057_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000057 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 2.0 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | noell3057cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | noell3057cct2 |  |  |  | a even | 1
Test Operator |  | Ortega |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-08-01 |  |  | Time of test |  |  | 9:15 - 11:47 |  |  |  | Date of test |  | 2003-08-01 |  |  | Time of test |  |  | 12:23 - 14:34
Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03 |  |  |  | Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 26 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 26
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00014929 |  |  | Dy average |  |  | 0.0006996 |  |  |  | Dx average |  | 0 |  |  | Dy average |  |  | 0.0007583
Magnetic length |  | 14.4461 |  |  | Tranfs. Function (mT/KA) |  |  | 596.6144 |  |  |  | Magnetic length |  | 14.447 |  |  | Tranfs. Function (mT/KA) |  |  | 596.6732
Phase |  | -3.4621 |  |  | Integral field (T.m) |  |  | 0.073259 |  |  |  | Phase |  | -4.0023 |  |  | Integral field (T.m) |  |  | 0.073271
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003164 | 0.005074 | 0.00507 | 0.00507 | 0.00507 | 0.005071 | 0.005071 | 0.00507 | 0.005071 | 0.00507 | 0.005071 | 0.005071 | 0.005071 | 0.005071 | 0.005071 | 0.005072 | 0.005072 | 0.005072 | 0.005074 | 0.003246 | 0.073259 | C1 (T) | 0.003156 | 0.005074 | 0.005071 | 0.005071 | 0.005071 | 0.005072 | 0.005072 | 0.005071 | 0.005072 | 0.005071 | 0.005071 | 0.005071 | 0.005071 | 0.005071 | 0.005072 | 0.005072 | 0.005072 | 0.005072 | 0.005074 | 0.003261 | 0.073271
Angle (mrad) | 3.7522 | 0.21098 | -0.53454 | 0.78522 | 0.22378 | 0.23129 | 0.033251 | 0.39613 | -0.030066 | -0.78445 | -1.3528 | -0.2473 | 0.21721 | 0.038931 | -0.55105 | 0.019647 | -0.33474 | -0.10966 | -0.57254 | -0.090046 | -3.4621 | Angle (mrad) | 3.6574 | 0.040379 | 0.23941 | 0.38213 | -0.23445 | 0.16532 | 0.090492 | 0.35895 | 0.36554 | 0.17346 | -0.57186 | -0.14507 | -0.44319 | -0.37952 | -0.63189 | -0.41419 | -0.54584 | -0.13022 | -0.013492 | -0.7626 | -4.0023
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -4.6954 | -1.0974 | -0.70051 | -0.36437 | -0.32852 | -1.2184 | -1.4724 | -1.1514 | -1.4493 | -1.0057 | -0.87271 | -0.99096 | -0.77625 | -0.57321 | -0.71659 | -0.63281 | -1.4918 | -1.6763 | -1.3594 | -0.74174 | -1.1048 | b2 | -0.24879 | 0.93881 | 1.1914 | 0.53585 | 0.4419 | 1.5247 | 1.11 | 0.92736 | 1.5588 | 1.0014 | 1.4892 | 0.74683 | 0.90124 | 0.87752 | 1.3316 | 1.01 | 1.3898 | 1.4539 | 1.3377 | 1.115 | 1.0553
b3 | 33.9626 | 1.8268 | 2.2019 | 1.665 | 3.1322 | 2.4361 | 2.9573 | 3.1115 | 2.9898 | 2.8617 | 1.3278 | 1.5764 | 2.8649 | 1.3358 | 1.1909 | 2.5736 | 3.168 | 2.4436 | 2.7677 | 0.92475 | 3.3333 | b3 | 38.9361 | 1.5003 | 1.5468 | 0.93078 | 1.0979 | 0.72123 | 1.582 | 1.7075 | 1.6626 | 1.727 | -0.35359 | 0.36056 | 1.6663 | 0.18851 | 0.82808 | 1.6208 | 1.8598 | 2.0927 | 2.3281 | 1.3225 | 2.4994
b4 | -1.5218 | 0.14003 | -0.011481 | 0.096796 | 0.11714 | 0.33643 | 0.25589 | -0.012939 | 0.18249 | 0.30706 | 0.19367 | 0.093906 | -0.042597 | 0.09971 | 0.17155 | -0.027835 | 0.020991 | 0.048786 | 0.055196 | -0.095908 | 0.05263 | b4 | -0.72725 | 0.039597 | -0.15888 | -0.26052 | -0.12896 | 0.12254 | 0.018641 | -0.074257 | -0.072915 | -0.14179 | -0.098764 | -0.21506 | -0.055586 | -0.21907 | 0.029117 | -0.046627 | 0.027954 | 0.063852 | 0.040958 | -0.28175 | -0.091527
b5 | -5.843 | -0.61632 | -0.6031 | -0.40304 | -0.36317 | -0.21588 | -0.26385 | -0.4286 | -0.53816 | -0.30572 | -0.095168 | -0.18161 | -0.44684 | -0.49297 | -0.20136 | -0.5037 | -0.5171 | -0.5986 | -0.50662 | -3.2664 | -0.6758 | b5 | -4.9105 | -0.60892 | -0.77757 | -0.36462 | -0.13981 | -0.18116 | -0.25194 | -0.26522 | -0.39148 | -0.33799 | 0.12141 | -0.16168 | -0.38958 | -0.47401 | -0.3502 | -0.54374 | -0.55132 | -0.42417 | -0.51404 | -3.38 | -0.61433
b6 | -0.32634 | -0.051115 | -0.005821 | -0.038387 | 0.034317 | -0.0042943 | 0.0066373 | -0.019376 | -0.019295 | -0.0058654 | -0.0084928 | 0.078778 | 0.10324 | -0.043788 | -0.079614 | -0.072655 | 0.17053 | 0.127 | 0.0058353 | 0.0090864 | -0.0010476 | b6 | -0.42518 | 0.0089523 | -0.098124 | -0.019716 | -0.16281 | -0.023512 | 0.039209 | -0.070966 | -0.095424 | -0.084468 | -0.090394 | -0.034577 | 0.044652 | 0.024426 | -0.073314 | -0.037276 | 0.028786 | 0.062601 | -0.031839 | -0.12415 | -0.049739
b7 | 2.5374 | 1.0697 | 1.1836 | 1.2448 | 1.2678 | 1.2695 | 1.2344 | 1.2167 | 1.2171 | 1.2753 | 1.1995 | 1.2 | 1.1896 | 1.2042 | 1.2069 | 1.2669 | 1.2543 | 1.3509 | 1.334 | 1.0147 | 1.2675 | b7 | 2.6804 | 1.0632 | 1.1082 | 1.0886 | 1.1715 | 1.2079 | 1.1487 | 1.1683 | 1.1491 | 1.208 | 1.2499 | 1.2308 | 1.1783 | 1.111 | 1.1363 | 1.1087 | 1.1599 | 1.2768 | 1.1373 | 0.8484 | 1.1999
b8 | -0.074199 | 0.0055852 | 0.0093356 | 0.058033 | 0.0067908 | -0.018788 | -0.002234 | 0.0024829 | 0.00114 | 0.030083 | 0.013724 | -0.023551 | 0.0004377 | 0.016143 | -0.0066983 | -0.023262 | 0.042164 | 0.041667 | -0.019801 | -0.07443 | 0.0020395 | b8 | -0.18301 | -0.012667 | -0.029033 | -0.029724 | -0.047822 | -0.0054291 | 0.044888 | -0.0052039 | -0.026211 | -0.014805 | 0.028342 | 0.043817 | 0.036897 | 0.044059 | -0.0098811 | 0.010767 | 0.014107 | 0.0047176 | -0.06756 | -0.067834 | -0.0092557
b9 | 0.373 | 0.53667 | 0.52138 | 0.55922 | 0.57171 | 0.53915 | 0.54578 | 0.55061 | 0.56529 | 0.55563 | 0.54384 | 0.52575 | 0.55432 | 0.54847 | 0.52258 | 0.54599 | 0.57219 | 0.60201 | 0.58603 | 0.51116 | 0.5454 | b9 | 0.36728 | 0.54674 | 0.52723 | 0.55195 | 0.54353 | 0.51786 | 0.53589 | 0.54473 | 0.54952 | 0.54696 | 0.52281 | 0.52081 | 0.53555 | 0.54235 | 0.52137 | 0.53053 | 0.54391 | 0.57203 | 0.5505 | 0.45483 | 0.53076
b10 | -0.027621 | -0.02248 | -0.034021 | 0.036656 | 0.042661 | -0.018883 | -0.011445 | -0.0051572 | 0.012289 | 0.067866 | -0.01682 | -0.010088 | -0.0068505 | -0.038934 | -0.031825 | -0.025561 | 0.078049 | 0.065462 | -0.026232 | -0.058446 | 0 | b10 | -0.042904 | -0.0044217 | -0.010428 | 0.058751 | -0.00678 | -0.037273 | 0.033323 | -0.02458 | -0.022018 | -0.021636 | -0.0048378 | 0.036276 | 0.037843 | 0.028886 | -0.0046178 | -0.00039596 | 0.059216 | 0.0020387 | -0.071942 | -0.052596 | -0.00068323
b11 | 0.6289 | 0.7517 | 0.74279 | 0.7518 | 0.74776 | 0.75388 | 0.751 | 0.75114 | 0.75423 | 0.74852 | 0.76164 | 0.75872 | 0.75236 | 0.76348 | 0.76514 | 0.75007 | 0.7562 | 0.76272 | 0.75777 | 0.66884 | 0.74758 | b11 | 0.60969 | 0.74989 | 0.74978 | 0.77142 | 0.76507 | 0.76791 | 0.76687 | 0.76025 | 0.76511 | 0.75993 | 0.7703 | 0.76327 | 0.76421 | 0.7722 | 0.76899 | 0.76364 | 0.76154 | 0.76204 | 0.75963 | 0.66087 | 0.75505
b12 | 0.0016175 | -0.0030668 | -0.002226 | 0.00042534 | 0.0011497 | -0.0062801 | -0.00061549 | -0.0016238 | -0.0019551 | 0.0053561 | -0.0032773 | -0.0017269 | -0.0021507 | -0.0077063 | -0.0098302 | -0.0071459 | 0.0068777 | 0.0038951 | -0.0059715 | -0.013168 | -0.0022473 | b12 | -0.01375 | -0.002464 | -0.0012976 | 0.0070952 | -0.0064243 | -0.005084 | 0.0037206 | -0.0042592 | -0.004494 | -0.0051407 | 0.00064403 | 0.0062331 | 0.0050733 | 0.0075455 | 0.0021436 | 0.0019818 | 0.010292 | 0.0012618 | -0.0096771 | -0.0039061 | -0.00020409
b13 | 0.089317 | 0.080484 | 0.080749 | 0.07328 | 0.073054 | 0.073993 | 0.073998 | 0.074079 | 0.071071 | 0.071168 | 0.07236 | 0.076449 | 0.07501 | 0.072378 | 0.07307 | 0.073328 | 0.072552 | 0.075939 | 0.074141 | 0.055992 | 0.074163 | b13 | 0.08874 | 0.074471 | 0.076079 | 0.071363 | 0.070772 | 0.070476 | 0.071702 | 0.069777 | 0.069531 | 0.071129 | 0.071428 | 0.071917 | 0.07582 | 0.070556 | 0.073073 | 0.073769 | 0.075385 | 0.075641 | 0.075632 | 0.057968 | 0.072722
b14 | 0.0041559 | -0.0044545 | -0.0059119 | -0.0016043 | -0.0027234 | -0.0074728 | -0.0063461 | -0.004471 | -0.0049497 | -0.0013423 | -0.0071242 | -0.0039144 | -0.0042702 | -0.0080544 | -0.0080119 | -0.0067808 | 0.0001185 | -0.00047044 | -0.0056819 | -0.0082707 | -0.0044722 | b14 | 0.0020275 | -0.0033275 | -0.0019387 | 0.0025871 | -0.00234 | -0.0032524 | 0.0011071 | -0.0037941 | -0.0031612 | -0.0020376 | 0.00020308 | 0.00076311 | 0.0019345 | 0.0022269 | -0.00063789 | -0.00036141 | 0.0042129 | 0.0016442 | -0.0035287 | -0.0043842 | -0.00058545
b15 | -0.00011352 | 0.032369 | 0.030582 | 0.031515 | 0.03338 | 0.03282 | 0.031996 | 0.033284 | 0.033684 | 0.034225 | 0.035622 | 0.032787 | 0.03072 | 0.031235 | 0.029306 | 0.030387 | 0.028797 | 0.032224 | 0.028878 | 0.015972 | 0.030314 | b15 | 0.0063014 | 0.036907 | 0.033427 | 0.036808 | 0.0357 | 0.036804 | 0.036568 | 0.037054 | 0.034579 | 0.033828 | 0.038152 | 0.038516 | 0.03662 | 0.037369 | 0.037676 | 0.038345 | 0.033392 | 0.034673 | 0.0342 | 0.028386 | 0.034922
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 37.5222 | 2.1098 | -5.3454 | 7.8523 | 2.2378 | 2.3129 | 0.33251 | 3.9613 | -0.30066 | -7.8445 | -13.5285 | -2.473 | 2.1721 | 0.38931 | -5.5105 | 0.19647 | -3.3474 | -1.0966 | -5.7254 | -0.90046 | 0 | a1 | 36.5741 | 0.40379 | 2.3941 | 3.8213 | -2.3445 | 1.6532 | 0.90492 | 3.5895 | 3.6554 | 1.7346 | -5.7186 | -1.4507 | -4.4319 | -3.7952 | -6.3189 | -4.1419 | -5.4584 | -1.3022 | -0.13492 | -7.626 | 0
a2 | 14.177 | 1.0023 | 0.94373 | 1.6174 | 1.4738 | 2.4252 | 1.9005 | 0.87109 | 1.7523 | 1.5666 | 2.5384 | 2.4521 | 2.1523 | 2.1792 | 1.5035 | 1.8983 | 1.7013 | 1.4783 | 1.3358 | -3.4283 | 1.9436 | a2 | 3.3318 | 2.1062 | 0.76069 | -0.63538 | -0.59074 | 0.93604 | 1.0593 | 0.54741 | 1.5314 | 1.5212 | 0.068061 | -0.72106 | 0.1516 | 0.54715 | 0.58407 | 0.19651 | 0.67309 | 1.0875 | 1.7421 | -0.11975 | 0.70402
a3 | 0.84616 | 1.0025 | 0.86369 | 0.72655 | 0.83609 | 0.0028603 | -0.021639 | 0.15921 | -0.10869 | -0.51003 | 0.43902 | 0.30423 | 0.36856 | 0.39539 | 0.5423 | 0.24541 | -0.033386 | 0.51464 | 0.2687 | 0.022273 | 0.33942 | a3 | 2.345 | 0.95372 | 0.90709 | 0.54371 | 0.55766 | 0.67716 | 0.45972 | 0.29221 | 0.76465 | 0.29823 | 0.67641 | 0.24847 | 0.20067 | 0.55236 | 0.76026 | 0.6783 | 0.70461 | 0.97138 | 0.79896 | 1.0807 | 0.68514
a4 | 2.6223 | 0.042174 | 0.064075 | 0.033458 | 0.12591 | -0.37652 | -0.40615 | -0.14852 | -0.24785 | -0.087915 | -0.29469 | -0.24206 | -0.0054413 | -0.32291 | -0.25729 | -0.34434 | -0.4275 | -0.32614 | -0.013551 | -0.023215 | -0.083785 | a4 | 1.4991 | 0.26715 | -0.14434 | 0.256 | 0.25096 | -0.406 | -0.32377 | -0.21956 | -0.46685 | -0.30793 | 0.12722 | 0.12617 | -0.17924 | 0.10731 | -0.082225 | 0.32221 | 0.095071 | -0.21477 | 0.21704 | 0.29701 | 0.028476
a5 | 2.3487 | 0.24717 | 0.17819 | 0.32751 | 0.22101 | 0.0089305 | -0.060267 | -0.1625 | -0.055993 | 0.087271 | 0.10688 | -0.10121 | -0.24029 | 0.012204 | 0.045352 | 0.12438 | -0.084144 | 0.026553 | 0.2224 | -0.098067 | 0.1197 | a5 | 3.3738 | 0.13004 | 0.042182 | 0.073551 | 0.21089 | -0.020032 | -0.16706 | -0.091418 | 0.10428 | -0.011466 | 0.15958 | 0.067171 | 0.020008 | 0.21712 | 0.076563 | 0.18571 | 0.15236 | 0.24719 | 0.25132 | 0.0040727 | 0.19469
a6 | 0.44045 | 0.0063112 | -0.011661 | -0.014444 | -0.06019 | -0.06785 | -0.082446 | -0.12184 | -0.067988 | -0.075836 | -0.0489 | 0.007427 | -0.036342 | -0.013889 | 0.020121 | -0.048861 | 0.014922 | 0.082811 | 0.23011 | 0.18969 | 0.0056015 | a6 | 0.10166 | 0.026637 | 0.023724 | -0.29536 | -0.27498 | -0.23326 | -0.16326 | -0.15842 | -0.11003 | -0.071533 | -0.17588 | -0.106 | -0.12232 | -0.15019 | -0.1618 | -0.1051 | -0.088004 | -0.069396 | -0.12753 | -0.082906 | -0.12212
a7 | 1.5207 | -0.09995 | -0.027137 | -0.041924 | -0.018353 | -0.0090982 | -0.019494 | -0.027595 | -0.047536 | -0.049128 | -0.073423 | -0.0068967 | -0.085264 | -0.039328 | -0.086373 | -0.067774 | -0.087671 | -0.086647 | -0.10873 | 0.055044 | 0.00012006 | a7 | 1.878 | -0.038272 | -0.028078 | -0.08612 | -0.031085 | -0.086978 | -0.067529 | -0.088512 | -0.072448 | -0.044186 | 0.00018927 | -0.058978 | -0.096118 | -0.08575 | -0.10757 | -0.053138 | -0.063177 | -0.046155 | -0.075988 | 0.024535 | 0.0027962
a8 | 0.016172 | 0.014804 | -0.017093 | -0.014712 | -0.016135 | -0.031667 | -0.0094288 | -0.022579 | -0.0051982 | -0.022332 | -0.013063 | 0.0093509 | -0.018096 | 0.023587 | 0.055531 | 0.0083938 | 0.033118 | 0.031908 | 0.063668 | -0.010569 | 0.0038133 | a8 | -0.021208 | 0.011277 | 0.041144 | -0.013078 | 0.0063972 | -0.045031 | -0.039114 | -0.018662 | 0.016638 | -0.0038784 | -0.0093191 | -0.041148 | -0.008541 | -0.071707 | -0.058297 | -0.069539 | -0.031254 | -0.057886 | -0.033107 | -0.064135 | -0.024891
a9 | -0.2672 | -0.02111 | -0.016828 | -0.014357 | -0.030107 | -0.02594 | -0.02831 | -0.015479 | -0.031643 | -0.023077 | -0.0034133 | -0.009919 | 0.013485 | -0.016443 | -0.033902 | -0.025861 | -0.029753 | -0.022435 | -0.025945 | -0.085986 | -0.030234 | a9 | -0.16102 | 0.022377 | 0.0067891 | 0.019213 | -0.0012889 | -0.041657 | -0.037821 | -0.031469 | -0.031111 | -0.042969 | -0.026273 | -0.033897 | -0.044525 | -0.05544 | -0.045069 | -0.02358 | -0.033845 | -0.016954 | -0.016545 | -0.066088 | -0.02996
a10 | 0.038717 | 0.040202 | 0.023735 | 0.0048726 | -0.014849 | -0.031087 | -0.01417 | -0.034373 | -0.020622 | -0.018197 | -0.011145 | 0.005786 | 0.015473 | -0.013803 | 0.015795 | -0.012199 | 0.016526 | 0.0098741 | 0.024378 | -0.016246 | 0 | a10 | -0.043063 | 0.036477 | 0.051136 | 0.010075 | 0.024773 | -0.011884 | 0.0081148 | -0.013942 | 0.022166 | 0.017812 | -0.010711 | -0.026856 | -0.017427 | -0.021777 | -0.019034 | -0.018425 | 0.026308 | 0.015803 | 0.019441 | -0.1007 | 2.7665e-05
a11 | 0.14797 | -0.041301 | -0.04046 | -0.019655 | -0.028205 | -0.033459 | -0.038045 | -0.027558 | -0.029469 | -0.041107 | -0.045912 | -0.047008 | -0.039304 | -0.0422 | -0.047891 | -0.032566 | -0.031857 | -0.028907 | -0.025437 | -0.020877 | -0.029108 | a11 | 0.1819 | -0.033777 | -0.031733 | -0.041149 | -0.034241 | -0.036059 | -0.039825 | -0.031571 | -0.030045 | -0.033848 | -0.030441 | -0.033179 | -0.035649 | -0.040122 | -0.0461 | -0.035134 | -0.031837 | -0.028466 | -0.03045 | -0.030378 | -0.027548
a12 | 0.0058655 | -0.0022318 | -0.0027102 | -0.0031884 | -0.0087189 | -0.007562 | -0.0078431 | -0.011397 | -0.0071211 | -0.0079296 | -0.006828 | -0.0026797 | -0.0021535 | -0.0059189 | -0.0013581 | -0.0054903 | -0.0022991 | -0.0090207 | -0.0043058 | -0.003446 | -0.0050509 | a12 | -0.015468 | -0.0020509 | -0.0007335 | -0.011679 | -0.0072927 | -0.0074305 | -0.0030976 | -0.0072232 | -0.0019934 | -0.00097474 | -0.0096679 | -0.0084702 | -0.0069118 | -0.010734 | -0.0084254 | -0.0062759 | 0.0020893 | 0.0056208 | -0.00082001 | -0.010638 | -0.005322
a13 | -0.021638 | 0.00014732 | -0.0016177 | -0.0041681 | -0.0028521 | -0.00069515 | -0.0023836 | -0.001957 | -0.0021877 | -0.0050245 | -0.00020593 | -0.0019196 | -0.0021336 | -0.001829 | -0.0034965 | -0.0026078 | -0.0018195 | -0.00044975 | 0.0004769 | 0.00083815 | -0.0024714 | a13 | -0.014953 | 0.0019307 | 0.0013847 | -0.0023908 | 0.0014965 | -0.0013984 | -0.0032167 | -0.0008736 | -0.00071858 | -0.0014609 | -0.0017309 | -0.0039533 | -0.0031879 | -0.0034987 | -0.0034588 | -0.0026468 | -0.0037801 | -0.00038583 | -0.0003067 | -0.0039971 | -0.0020844
a14 | 7.3157e-05 | -0.016 | -0.015272 | -0.016083 | -0.016434 | -0.017107 | -0.015081 | -0.016949 | -0.018077 | -0.017339 | -0.018124 | -0.015358 | -0.014994 | -0.014382 | -0.011179 | -0.013975 | -0.013251 | -0.017037 | -0.013711 | -0.0087832 | -0.014843 | a14 | 0.0015203 | -0.016486 | -0.015316 | -0.015326 | -0.013603 | -0.017997 | -0.016668 | -0.01605 | -0.014711 | -0.013003 | -0.014381 | -0.016372 | -0.014488 | -0.016416 | -0.01566 | -0.016296 | -0.013173 | -0.015337 | -0.012616 | -0.0075144 | -0.014419
a15 | 0.00040643 | -0.0077462 | -0.0086136 | -0.0021824 | -0.0037843 | -0.0085588 | -0.0065014 | -0.0062188 | -0.0061519 | -0.0019128 | -0.008163 | -0.0056693 | -0.0051553 | -0.0092366 | -0.010029 | -0.0098851 | -0.00027789 | -0.00084763 | -0.0081811 | -0.01042 | -0.0059956 | a15 | 0.0026301 | -0.002779 | -0.0035727 | 0.0016732 | -0.0036417 | -0.0041757 | 0.0026628 | -0.0038813 | -0.0041521 | -0.003558 | -0.0012401 | 0.00026749 | 0.0014202 | 0.0012427 | -0.0024219 | -0.0016233 | 0.003631 | 0.00034404 | -0.0063859 | -0.0060412 | -0.0014788
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 4.7415e-05 | 5.5679e-05 | 8.0596e-05 | -8.2544e-05 | -9.81e-05 | 3.9368e-05 | 2.4216e-05 | 0 | -2.9463e-05 | -0.00015593 | 3.5908e-05 | 2.3319e-05 | 1.7259e-05 | 8.4734e-05 | 7.2643e-05 | 5.6623e-05 | -0.00017366 | -0.00014491 | 6.0634e-05 | 0.00014721 | 0.00014929 | Dx (m) | 0.00014273 | 1.3719e-05 | 2.8504e-05 | -0.000128 | 1.7494e-05 | 8.1095e-05 | -7.2751e-05 | 5.3568e-05 | 5.0784e-05 | 5.0081e-05 | 0 | -8.3219e-05 | -8.5798e-05 | -6.5888e-05 | 0 | 0 | -0.00012954 | 0 | 0.00016256 | 0.00012286 | 0
Dy (m) | -0.00011578 | -8.7875e-05 | -4.9928e-05 | -1.3166e-05 | 3.0087e-05 | 7.1852e-05 | 3.3302e-05 | 7.8091e-05 | 4.5325e-05 | 3.2848e-05 | 2.7045e-05 | -1.1519e-05 | -3.4062e-05 | 3.5448e-05 | -3.056e-05 | 3.0121e-05 | -4.4457e-05 | -2.7423e-05 | -5.2667e-05 | 4.5987e-05 | 0.0006996 | Dy (m) | 7.7722e-05 | -8.2086e-05 | -0.00011475 | -2.8967e-05 | -5.4277e-05 | 3.0141e-05 | -2.1763e-05 | 3.3411e-05 | -4.7263e-05 | -3.7621e-05 | 2.402e-05 | 5.6229e-05 | 3.4789e-05 | 4.4543e-05 | 4.2532e-05 | 4.0966e-05 | -6.4205e-05 | -3.5373e-05 | -3.7011e-05 | 0.00026455 | 0.0007583
Magnetic length (m) |  | 14.4461 | [19 empty] | Magnetic length (m) |  | 14.447
Transfer funct. (mT/KA) |  | 596.6144 | [19 empty] | Transfer funct. (mT/KA) |  | 596.6732
Coil positioning (mm) |  | 0.016862 | [19 empty] | Coil positioning (mm) |  | 0.017624
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
144 | HCMB__A001 | 3000057 |  | 2 |  |  | 3
| COMMENTS: | Non-nominal shims on the outer layers of both apertures.
| HOLDING POINT | OK | Data at CERN |  | 2003-08-04 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-08-04 | Answer to manufacturer |  |  | 2003-08-06
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.85 | 0.8 | 0.05 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.85 | 0.2 | 0.2 | 0.85 | Up | 0.85 | 0.2 | 0.2 | 0.85
[8 empty] | 0.85 | 0.2 | 0.2 | 0.85 | Down | 0.85 | 0.2 | 0.2 | 0.85
Seen from connection side
| Collared coil template - version 11.07.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)

## Summary Data
File | HCMB__A001-3000057 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000057 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 372.2353 | 596.9412 | 596.4706 | 596.4706 | 596.4706 | 596.5882 | 596.5882 | 596.4706 | 596.5882 | 596.4706 | 596.5882 | 596.5882 | 596.5882 | 596.5882 | 596.5882 | 596.7059 | 596.7059 | 596.7059 | 596.9412 | 381.8824 | 0.073259 | C1/i (mT/KA) | 371.2941 | 596.9412 | 596.5882 | 596.5882 | 596.5882 | 596.7059 | 596.7059 | 596.5882 | 596.7059 | 596.5882 | 596.5882 | 596.5882 | 596.5882 | 596.5882 | 596.7059 | 596.7059 | 596.7059 | 596.7059 | 596.9412 | 383.6471 | 0.073271
Angle (mrad) | 3.7522 | 0.21098 | -0.53454 | 0.78522 | 0.22378 | 0.23129 | 0.033251 | 0.39613 | -0.030066 | -0.78445 | -1.3528 | -0.2473 | 0.21721 | 0.038931 | -0.55105 | 0.019647 | -0.33474 | -0.10966 | -0.57254 | -0.090046 | -3.4621 | Angle (mrad) | 3.6574 | 0.040379 | 0.23941 | 0.38213 | -0.23445 | 0.16532 | 0.090492 | 0.35895 | 0.36554 | 0.17346 | -0.57186 | -0.14507 | -0.44319 | -0.37952 | -0.63189 | -0.41419 | -0.54584 | -0.13022 | -0.013492 | -0.7626 | -4.0023
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -4.6954 | -1.0974 | -0.70051 | -0.36437 | -0.32852 | -1.2184 | -1.4724 | -1.1514 | -1.4493 | -1.0057 | -0.87271 | -0.99096 | -0.77625 | -0.57321 | -0.71659 | -0.63281 | -1.4918 | -1.6763 | -1.3594 | -0.74174 | -1.1048 | b2 | -0.24879 | 0.93881 | 1.1914 | 0.53585 | 0.4419 | 1.5247 | 1.11 | 0.92736 | 1.5588 | 1.0014 | 1.4892 | 0.74683 | 0.90124 | 0.87752 | 1.3316 | 1.01 | 1.3898 | 1.4539 | 1.3377 | 1.115 | 1.0553
b3 | 33.9626 | 1.8268 | 2.2019 | 1.665 | 3.1322 | 2.4361 | 2.9573 | 3.1115 | 2.9898 | 2.8617 | 1.3278 | 1.5764 | 2.8649 | 1.3358 | 1.1909 | 2.5736 | 3.168 | 2.4436 | 2.7677 | 0.92475 | 3.3333 | b3 | 38.9361 | 1.5003 | 1.5468 | 0.93078 | 1.0979 | 0.72123 | 1.582 | 1.7075 | 1.6626 | 1.727 | -0.35359 | 0.36056 | 1.6663 | 0.18851 | 0.82808 | 1.6208 | 1.8598 | 2.0927 | 2.3281 | 1.3225 | 2.4994
b4 | -1.5218 | 0.14003 | -0.011481 | 0.096796 | 0.11714 | 0.33643 | 0.25589 | -0.012939 | 0.18249 | 0.30706 | 0.19367 | 0.093906 | -0.042597 | 0.09971 | 0.17155 | -0.027835 | 0.020991 | 0.048786 | 0.055196 | -0.095908 | 0.05263 | b4 | -0.72725 | 0.039597 | -0.15888 | -0.26052 | -0.12896 | 0.12254 | 0.018641 | -0.074257 | -0.072915 | -0.14179 | -0.098764 | -0.21506 | -0.055586 | -0.21907 | 0.029117 | -0.046627 | 0.027954 | 0.063852 | 0.040958 | -0.28175 | -0.091527
b5 | -5.843 | -0.61632 | -0.6031 | -0.40304 | -0.36317 | -0.21588 | -0.26385 | -0.4286 | -0.53816 | -0.30572 | -0.095168 | -0.18161 | -0.44684 | -0.49297 | -0.20136 | -0.5037 | -0.5171 | -0.5986 | -0.50662 | -3.2664 | -0.6758 | b5 | -4.9105 | -0.60892 | -0.77757 | -0.36462 | -0.13981 | -0.18116 | -0.25194 | -0.26522 | -0.39148 | -0.33799 | 0.12141 | -0.16168 | -0.38958 | -0.47401 | -0.3502 | -0.54374 | -0.55132 | -0.42417 | -0.51404 | -3.38 | -0.61433
b6 | -0.32634 | -0.051115 | -0.005821 | -0.038387 | 0.034317 | -0.0042943 | 0.0066373 | -0.019376 | -0.019295 | -0.0058654 | -0.0084928 | 0.078778 | 0.10324 | -0.043788 | -0.079614 | -0.072655 | 0.17053 | 0.127 | 0.0058353 | 0.0090864 | -0.0010476 | b6 | -0.42518 | 0.0089523 | -0.098124 | -0.019716 | -0.16281 | -0.023512 | 0.039209 | -0.070966 | -0.095424 | -0.084468 | -0.090394 | -0.034577 | 0.044652 | 0.024426 | -0.073314 | -0.037276 | 0.028786 | 0.062601 | -0.031839 | -0.12415 | -0.049739
b7 | 2.5374 | 1.0697 | 1.1836 | 1.2448 | 1.2678 | 1.2695 | 1.2344 | 1.2167 | 1.2171 | 1.2753 | 1.1995 | 1.2 | 1.1896 | 1.2042 | 1.2069 | 1.2669 | 1.2543 | 1.3509 | 1.334 | 1.0147 | 1.2675 | b7 | 2.6804 | 1.0632 | 1.1082 | 1.0886 | 1.1715 | 1.2079 | 1.1487 | 1.1683 | 1.1491 | 1.208 | 1.2499 | 1.2308 | 1.1783 | 1.111 | 1.1363 | 1.1087 | 1.1599 | 1.2768 | 1.1373 | 0.8484 | 1.1999
b8 | -0.074199 | 0.0055852 | 0.0093356 | 0.058033 | 0.0067908 | -0.018788 | -0.002234 | 0.0024829 | 0.00114 | 0.030083 | 0.013724 | -0.023551 | 0.0004377 | 0.016143 | -0.0066983 | -0.023262 | 0.042164 | 0.041667 | -0.019801 | -0.07443 | 0.0020395 | b8 | -0.18301 | -0.012667 | -0.029033 | -0.029724 | -0.047822 | -0.0054291 | 0.044888 | -0.0052039 | -0.026211 | -0.014805 | 0.028342 | 0.043817 | 0.036897 | 0.044059 | -0.0098811 | 0.010767 | 0.014107 | 0.0047176 | -0.06756 | -0.067834 | -0.0092557
b9 | 0.373 | 0.53667 | 0.52138 | 0.55922 | 0.57171 | 0.53915 | 0.54578 | 0.55061 | 0.56529 | 0.55563 | 0.54384 | 0.52575 | 0.55432 | 0.54847 | 0.52258 | 0.54599 | 0.57219 | 0.60201 | 0.58603 | 0.51116 | 0.5454 | b9 | 0.36728 | 0.54674 | 0.52723 | 0.55195 | 0.54353 | 0.51786 | 0.53589 | 0.54473 | 0.54952 | 0.54696 | 0.52281 | 0.52081 | 0.53555 | 0.54235 | 0.52137 | 0.53053 | 0.54391 | 0.57203 | 0.5505 | 0.45483 | 0.53076
b10 | -0.027621 | -0.02248 | -0.034021 | 0.036656 | 0.042661 | -0.018883 | -0.011445 | -0.0051572 | 0.012289 | 0.067866 | -0.01682 | -0.010088 | -0.0068505 | -0.038934 | -0.031825 | -0.025561 | 0.078049 | 0.065462 | -0.026232 | -0.058446 | 0 | b10 | -0.042904 | -0.0044217 | -0.010428 | 0.058751 | -0.00678 | -0.037273 | 0.033323 | -0.02458 | -0.022018 | -0.021636 | -0.0048378 | 0.036276 | 0.037843 | 0.028886 | -0.0046178 | -0.00039596 | 0.059216 | 0.0020387 | -0.071942 | -0.052596 | -0.00068323
b11 | 0.6289 | 0.7517 | 0.74279 | 0.7518 | 0.74776 | 0.75388 | 0.751 | 0.75114 | 0.75423 | 0.74852 | 0.76164 | 0.75872 | 0.75236 | 0.76348 | 0.76514 | 0.75007 | 0.7562 | 0.76272 | 0.75777 | 0.66884 | 0.74758 | b11 | 0.60969 | 0.74989 | 0.74978 | 0.77142 | 0.76507 | 0.76791 | 0.76687 | 0.76025 | 0.76511 | 0.75993 | 0.7703 | 0.76327 | 0.76421 | 0.7722 | 0.76899 | 0.76364 | 0.76154 | 0.76204 | 0.75963 | 0.66087 | 0.75505
b12 | 0.0016175 | -0.0030668 | -0.002226 | 0.00042534 | 0.0011497 | -0.0062801 | -0.00061549 | -0.0016238 | -0.0019551 | 0.0053561 | -0.0032773 | -0.0017269 | -0.0021507 | -0.0077063 | -0.0098302 | -0.0071459 | 0.0068777 | 0.0038951 | -0.0059715 | -0.013168 | -0.0022473 | b12 | -0.01375 | -0.002464 | -0.0012976 | 0.0070952 | -0.0064243 | -0.005084 | 0.0037206 | -0.0042592 | -0.004494 | -0.0051407 | 0.00064403 | 0.0062331 | 0.0050733 | 0.0075455 | 0.0021436 | 0.0019818 | 0.010292 | 0.0012618 | -0.0096771 | -0.0039061 | -0.00020409
b13 | 0.089317 | 0.080484 | 0.080749 | 0.07328 | 0.073054 | 0.073993 | 0.073998 | 0.074079 | 0.071071 | 0.071168 | 0.07236 | 0.076449 | 0.07501 | 0.072378 | 0.07307 | 0.073328 | 0.072552 | 0.075939 | 0.074141 | 0.055992 | 0.074163 | b13 | 0.08874 | 0.074471 | 0.076079 | 0.071363 | 0.070772 | 0.070476 | 0.071702 | 0.069777 | 0.069531 | 0.071129 | 0.071428 | 0.071917 | 0.07582 | 0.070556 | 0.073073 | 0.073769 | 0.075385 | 0.075641 | 0.075632 | 0.057968 | 0.072722
b14 | 0.0041559 | -0.0044545 | -0.0059119 | -0.0016043 | -0.0027234 | -0.0074728 | -0.0063461 | -0.004471 | -0.0049497 | -0.0013423 | -0.0071242 | -0.0039144 | -0.0042702 | -0.0080544 | -0.0080119 | -0.0067808 | 0.0001185 | -0.00047044 | -0.0056819 | -0.0082707 | -0.0044722 | b14 | 0.0020275 | -0.0033275 | -0.0019387 | 0.0025871 | -0.00234 | -0.0032524 | 0.0011071 | -0.0037941 | -0.0031612 | -0.0020376 | 0.00020308 | 0.00076311 | 0.0019345 | 0.0022269 | -0.00063789 | -0.00036141 | 0.0042129 | 0.0016442 | -0.0035287 | -0.0043842 | -0.00058545
b15 | -0.00011352 | 0.032369 | 0.030582 | 0.031515 | 0.03338 | 0.03282 | 0.031996 | 0.033284 | 0.033684 | 0.034225 | 0.035622 | 0.032787 | 0.03072 | 0.031235 | 0.029306 | 0.030387 | 0.028797 | 0.032224 | 0.028878 | 0.015972 | 0.030314 | b15 | 0.0063014 | 0.036907 | 0.033427 | 0.036808 | 0.0357 | 0.036804 | 0.036568 | 0.037054 | 0.034579 | 0.033828 | 0.038152 | 0.038516 | 0.03662 | 0.037369 | 0.037676 | 0.038345 | 0.033392 | 0.034673 | 0.0342 | 0.028386 | 0.034922
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 37.5222 | 2.1098 | -5.3454 | 7.8523 | 2.2378 | 2.3129 | 0.33251 | 3.9613 | -0.30066 | -7.8445 | -13.5285 | -2.473 | 2.1721 | 0.38931 | -5.5105 | 0.19647 | -3.3474 | -1.0966 | -5.7254 | -0.90046 | 0 | a1 | 36.5741 | 0.40379 | 2.3941 | 3.8213 | -2.3445 | 1.6532 | 0.90492 | 3.5895 | 3.6554 | 1.7346 | -5.7186 | -1.4507 | -4.4319 | -3.7952 | -6.3189 | -4.1419 | -5.4584 | -1.3022 | -0.13492 | -7.626 | 0
a2 | 14.177 | 1.0023 | 0.94373 | 1.6174 | 1.4738 | 2.4252 | 1.9005 | 0.87109 | 1.7523 | 1.5666 | 2.5384 | 2.4521 | 2.1523 | 2.1792 | 1.5035 | 1.8983 | 1.7013 | 1.4783 | 1.3358 | -3.4283 | 1.9436 | a2 | 3.3318 | 2.1062 | 0.76069 | -0.63538 | -0.59074 | 0.93604 | 1.0593 | 0.54741 | 1.5314 | 1.5212 | 0.068061 | -0.72106 | 0.1516 | 0.54715 | 0.58407 | 0.19651 | 0.67309 | 1.0875 | 1.7421 | -0.11975 | 0.70402
a3 | 0.84616 | 1.0025 | 0.86369 | 0.72655 | 0.83609 | 0.0028603 | -0.021639 | 0.15921 | -0.10869 | -0.51003 | 0.43902 | 0.30423 | 0.36856 | 0.39539 | 0.5423 | 0.24541 | -0.033386 | 0.51464 | 0.2687 | 0.022273 | 0.33942 | a3 | 2.345 | 0.95372 | 0.90709 | 0.54371 | 0.55766 | 0.67716 | 0.45972 | 0.29221 | 0.76465 | 0.29823 | 0.67641 | 0.24847 | 0.20067 | 0.55236 | 0.76026 | 0.6783 | 0.70461 | 0.97138 | 0.79896 | 1.0807 | 0.68514
a4 | 2.6223 | 0.042174 | 0.064075 | 0.033458 | 0.12591 | -0.37652 | -0.40615 | -0.14852 | -0.24785 | -0.087915 | -0.29469 | -0.24206 | -0.0054413 | -0.32291 | -0.25729 | -0.34434 | -0.4275 | -0.32614 | -0.013551 | -0.023215 | -0.083785 | a4 | 1.4991 | 0.26715 | -0.14434 | 0.256 | 0.25096 | -0.406 | -0.32377 | -0.21956 | -0.46685 | -0.30793 | 0.12722 | 0.12617 | -0.17924 | 0.10731 | -0.082225 | 0.32221 | 0.095071 | -0.21477 | 0.21704 | 0.29701 | 0.028476
a5 | 2.3487 | 0.24717 | 0.17819 | 0.32751 | 0.22101 | 0.0089305 | -0.060267 | -0.1625 | -0.055993 | 0.087271 | 0.10688 | -0.10121 | -0.24029 | 0.012204 | 0.045352 | 0.12438 | -0.084144 | 0.026553 | 0.2224 | -0.098067 | 0.1197 | a5 | 3.3738 | 0.13004 | 0.042182 | 0.073551 | 0.21089 | -0.020032 | -0.16706 | -0.091418 | 0.10428 | -0.011466 | 0.15958 | 0.067171 | 0.020008 | 0.21712 | 0.076563 | 0.18571 | 0.15236 | 0.24719 | 0.25132 | 0.0040727 | 0.19469
a6 | 0.44045 | 0.0063112 | -0.011661 | -0.014444 | -0.06019 | -0.06785 | -0.082446 | -0.12184 | -0.067988 | -0.075836 | -0.0489 | 0.007427 | -0.036342 | -0.013889 | 0.020121 | -0.048861 | 0.014922 | 0.082811 | 0.23011 | 0.18969 | 0.0056015 | a6 | 0.10166 | 0.026637 | 0.023724 | -0.29536 | -0.27498 | -0.23326 | -0.16326 | -0.15842 | -0.11003 | -0.071533 | -0.17588 | -0.106 | -0.12232 | -0.15019 | -0.1618 | -0.1051 | -0.088004 | -0.069396 | -0.12753 | -0.082906 | -0.12212
a7 | 1.5207 | -0.09995 | -0.027137 | -0.041924 | -0.018353 | -0.0090982 | -0.019494 | -0.027595 | -0.047536 | -0.049128 | -0.073423 | -0.0068967 | -0.085264 | -0.039328 | -0.086373 | -0.067774 | -0.087671 | -0.086647 | -0.10873 | 0.055044 | 0.00012006 | a7 | 1.878 | -0.038272 | -0.028078 | -0.08612 | -0.031085 | -0.086978 | -0.067529 | -0.088512 | -0.072448 | -0.044186 | 0.00018927 | -0.058978 | -0.096118 | -0.08575 | -0.10757 | -0.053138 | -0.063177 | -0.046155 | -0.075988 | 0.024535 | 0.0027962
a8 | 0.016172 | 0.014804 | -0.017093 | -0.014712 | -0.016135 | -0.031667 | -0.0094288 | -0.022579 | -0.0051982 | -0.022332 | -0.013063 | 0.0093509 | -0.018096 | 0.023587 | 0.055531 | 0.0083938 | 0.033118 | 0.031908 | 0.063668 | -0.010569 | 0.0038133 | a8 | -0.021208 | 0.011277 | 0.041144 | -0.013078 | 0.0063972 | -0.045031 | -0.039114 | -0.018662 | 0.016638 | -0.0038784 | -0.0093191 | -0.041148 | -0.008541 | -0.071707 | -0.058297 | -0.069539 | -0.031254 | -0.057886 | -0.033107 | -0.064135 | -0.024891
a9 | -0.2672 | -0.02111 | -0.016828 | -0.014357 | -0.030107 | -0.02594 | -0.02831 | -0.015479 | -0.031643 | -0.023077 | -0.0034133 | -0.009919 | 0.013485 | -0.016443 | -0.033902 | -0.025861 | -0.029753 | -0.022435 | -0.025945 | -0.085986 | -0.030234 | a9 | -0.16102 | 0.022377 | 0.0067891 | 0.019213 | -0.0012889 | -0.041657 | -0.037821 | -0.031469 | -0.031111 | -0.042969 | -0.026273 | -0.033897 | -0.044525 | -0.05544 | -0.045069 | -0.02358 | -0.033845 | -0.016954 | -0.016545 | -0.066088 | -0.02996
a10 | 0.038717 | 0.040202 | 0.023735 | 0.0048726 | -0.014849 | -0.031087 | -0.01417 | -0.034373 | -0.020622 | -0.018197 | -0.011145 | 0.005786 | 0.015473 | -0.013803 | 0.015795 | -0.012199 | 0.016526 | 0.0098741 | 0.024378 | -0.016246 | 0 | a10 | -0.043063 | 0.036477 | 0.051136 | 0.010075 | 0.024773 | -0.011884 | 0.0081148 | -0.013942 | 0.022166 | 0.017812 | -0.010711 | -0.026856 | -0.017427 | -0.021777 | -0.019034 | -0.018425 | 0.026308 | 0.015803 | 0.019441 | -0.1007 | 2.7665e-05
a11 | 0.14797 | -0.041301 | -0.04046 | -0.019655 | -0.028205 | -0.033459 | -0.038045 | -0.027558 | -0.029469 | -0.041107 | -0.045912 | -0.047008 | -0.039304 | -0.0422 | -0.047891 | -0.032566 | -0.031857 | -0.028907 | -0.025437 | -0.020877 | -0.029108 | a11 | 0.1819 | -0.033777 | -0.031733 | -0.041149 | -0.034241 | -0.036059 | -0.039825 | -0.031571 | -0.030045 | -0.033848 | -0.030441 | -0.033179 | -0.035649 | -0.040122 | -0.0461 | -0.035134 | -0.031837 | -0.028466 | -0.03045 | -0.030378 | -0.027548
a12 | 0.0058655 | -0.0022318 | -0.0027102 | -0.0031884 | -0.0087189 | -0.007562 | -0.0078431 | -0.011397 | -0.0071211 | -0.0079296 | -0.006828 | -0.0026797 | -0.0021535 | -0.0059189 | -0.0013581 | -0.0054903 | -0.0022991 | -0.0090207 | -0.0043058 | -0.003446 | -0.0050509 | a12 | -0.015468 | -0.0020509 | -0.0007335 | -0.011679 | -0.0072927 | -0.0074305 | -0.0030976 | -0.0072232 | -0.0019934 | -0.00097474 | -0.0096679 | -0.0084702 | -0.0069118 | -0.010734 | -0.0084254 | -0.0062759 | 0.0020893 | 0.0056208 | -0.00082001 | -0.010638 | -0.005322
a13 | -0.021638 | 0.00014732 | -0.0016177 | -0.0041681 | -0.0028521 | -0.00069515 | -0.0023836 | -0.001957 | -0.0021877 | -0.0050245 | -0.00020593 | -0.0019196 | -0.0021336 | -0.001829 | -0.0034965 | -0.0026078 | -0.0018195 | -0.00044975 | 0.0004769 | 0.00083815 | -0.0024714 | a13 | -0.014953 | 0.0019307 | 0.0013847 | -0.0023908 | 0.0014965 | -0.0013984 | -0.0032167 | -0.0008736 | -0.00071858 | -0.0014609 | -0.0017309 | -0.0039533 | -0.0031879 | -0.0034987 | -0.0034588 | -0.0026468 | -0.0037801 | -0.00038583 | -0.0003067 | -0.0039971 | -0.0020844
a14 | 7.3157e-05 | -0.016 | -0.015272 | -0.016083 | -0.016434 | -0.017107 | -0.015081 | -0.016949 | -0.018077 | -0.017339 | -0.018124 | -0.015358 | -0.014994 | -0.014382 | -0.011179 | -0.013975 | -0.013251 | -0.017037 | -0.013711 | -0.0087832 | -0.014843 | a14 | 0.0015203 | -0.016486 | -0.015316 | -0.015326 | -0.013603 | -0.017997 | -0.016668 | -0.01605 | -0.014711 | -0.013003 | -0.014381 | -0.016372 | -0.014488 | -0.016416 | -0.01566 | -0.016296 | -0.013173 | -0.015337 | -0.012616 | -0.0075144 | -0.014419
a15 | 0.00040643 | -0.0077462 | -0.0086136 | -0.0021824 | -0.0037843 | -0.0085588 | -0.0065014 | -0.0062188 | -0.0061519 | -0.0019128 | -0.008163 | -0.0056693 | -0.0051553 | -0.0092366 | -0.010029 | -0.0098851 | -0.00027789 | -0.00084763 | -0.0081811 | -0.01042 | -0.0059956 | a15 | 0.0026301 | -0.002779 | -0.0035727 | 0.0016732 | -0.0036417 | -0.0041757 | 0.0026628 | -0.0038813 | -0.0041521 | -0.003558 | -0.0012401 | 0.00026749 | 0.0014202 | 0.0012427 | -0.0024219 | -0.0016233 | 0.003631 | 0.00034404 | -0.0063859 | -0.0060412 | -0.0014788
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.047415 | -0.055679 | -0.080596 | 0.082544 | 0.0981 | -0.039368 | -0.024216 | 0 | 0.029463 | 0.15593 | -0.035908 | -0.023319 | -0.017259 | -0.084734 | -0.072643 | -0.056623 | 0.17366 | 0.14491 | -0.060634 | -0.14721 |  | Dx (mm) | -0.14273 | -0.013719 | -0.028504 | 0.128 | -0.017494 | -0.081095 | 0.072751 | -0.053568 | -0.050784 | -0.050081 | 0 | 0.083219 | 0.085798 | 0.065888 | 0 | 0 | 0.12954 | 0 | -0.16256 | -0.12286
Dy (mm) | 0.11578 | 0.087875 | 0.049928 | 0.013166 | -0.030087 | -0.071852 | -0.033302 | -0.078091 | -0.045325 | -0.032848 | -0.027045 | 0.011519 | 0.034062 | -0.035448 | 0.03056 | -0.030121 | 0.044457 | 0.027423 | 0.052667 | -0.045987 |  | Dy (mm) | -0.077722 | 0.082086 | 0.11475 | 0.028967 | 0.054277 | -0.030141 | 0.021763 | -0.033411 | 0.047263 | 0.037621 | -0.02402 | -0.056229 | -0.034789 | -0.044543 | -0.042532 | -0.040966 | 0.064205 | 0.035373 | 0.037011 | -0.26455
Mag. len. (m) | 14.4141 | [20 empty] | Mag. len. (m) | 14.415
TF (mT/KA) | 596.6144 | [20 empty] | TF (mT/KA) | 596.6732
Coil pos. (mm) | 0.016862 | [20 empty] | Coil pos. (mm) | 0.017624
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5
Twist (m2 rad) | 0.016382 | [20 empty] | Twist (m2 rad) | 0.020831

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -1.1048 | 0.91523 | -0.99739 | 0.40776 | 1.9436 | 3.1076 | 1.7128 | 0.5083 | 1.0553 | 0.4374 | 1.1006 | 0.33886 | 0.70402 | 0.99521 | 0.63889 | 0.81037 |  | 0
3 | 3.3333 | 7.1216 | 2.3576 | 0.69293 | 0.33942 | 0.39112 | 0.33372 | 0.38554 | 2.4994 | 8.4458 | 1.282 | 0.70535 | 0.68514 | 0.45566 | 0.61304 | 0.23955 |  | 0
4 | 0.05263 | 0.38115 | 0.114 | 0.11345 | -0.083785 | 0.64893 | -0.17921 | 0.18328 | -0.091527 | 0.1855 | -0.061916 | 0.10807 | 0.028476 | 0.42309 | -0.029645 | 0.25684 |  | 0
5 | -0.6758 | 1.3526 | -0.40484 | 0.161 | 0.1197 | 0.5367 | 0.050138 | 0.15413 | -0.61433 | 1.205 | -0.3675 | 0.20615 | 0.19469 | 0.74343 | 0.092272 | 0.11623 |  | 0
6 | -0.0010476 | 0.09933 | 0.0055199 | 0.055251 | 0.0056015 | 0.1321 | -0.014829 | 0.06511 | -0.049739 | 0.10588 | -0.036486 | 0.054626 | -0.12212 | 0.096176 | -0.1357 | 0.078977 |  | 0
7 | 1.2675 | 0.30405 | 1.2322 | 0.061622 | 0.00012006 | 0.35318 | -0.054554 | 0.032972 | 1.1999 | 0.35455 | 1.161 | 0.056773 | 0.0027962 | 0.43417 | -0.062185 | 0.028107 |  | 0
8 | 0.0020395 | 0.033507 | 0.0050011 | 0.017496 | 0.0038133 | 0.02682 | 0.0044972 | 0.02323 | -0.0092557 | 0.052607 | -0.0024481 | 0.025658 | -0.024891 | 0.031134 | -0.02671 | 0.0245 |  | 0
9 | 0.5454 | 0.045463 | 0.55144 | 0.020328 | -0.030234 | 0.057499 | -0.019955 | 0.011689 | 0.53076 | 0.04392 | 0.53844 | 0.013613 | -0.02996 | 0.038054 | -0.023872 | 0.022509 |  | 0
10 | 0 | 0.039005 | -0.00021735 | 0.00091695 | 0 | 0.022051 | 1.2661e-05 | 3.2119e-05 | -0.00068323 | 0.0359 | -0.00033171 | 0.0013666 | 2.7665e-05 | 0.033535 | 1.8035e-05 | 6.5e-05 |  | 0
11 | 0.74758 | 0.033643 | 0.75436 | 0.0059601 | -0.029108 | 0.041716 | -0.035519 | 0.0082109 | 0.75505 | 0.040748 | 0.76326 | 0.0063045 | -0.027548 | 0.048575 | -0.03453 | 0.0044655 |  | 0
12 | -0.0022473 | 0.0050155 | -0.002313 | 0.0020801 | -0.0050509 | 0.0037946 | -0.0054028 | 0.0018808 | -0.00020409 | 0.0062165 | 0.0001139 | 0.0023609 | -0.005322 | 0.0052047 | -0.0053632 | 0.0036004 |  | 0
13 | 0.074163 | 0.005999 | 0.074215 | 0.0023366 | -0.0024714 | 0.0046867 | -0.0018363 | 0.0014115 | 0.072722 | 0.0054463 | 0.072459 | 0.002133 | -0.0020844 | 0.0034965 | -0.0014646 | 0.0015681 |  | 0
14 | -0.0044722 | 0.0032543 | -0.0048317 | 0.00095218 | -0.014843 | 0.0041177 | -0.01565 | 0.0015642 | -0.00058545 | 0.002577 | -0.00077495 | 0.0012149 | -0.014419 | 0.0042663 | -0.015672 | 0.0015965 |  | 0
15 | 0.030314 | 0.008019 | 0.031878 | 0.0018787 | -0.0059956 | 0.0034399 | -0.0060619 | 0.003108 | 0.034922 | 0.0069968 | 0.036146 | 0.0017143 | -0.0014788 | 0.0030209 | -0.001455 | 0.0028251 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.78689 | -0.64622
3 | [12 empty] | 3 | 9 | -0.65996 | -0.88902
4 | [12 empty] | 4 | 16 | -1.9366 | -1.7921
5 | [12 empty] | 5 | 25 | -1.8482 | -1.8657
6 | [12 empty] | 6 | 36 | -2.8138 | -2.7229
7 | [12 empty] | 7 | 49 | -3.0994 | -3.2202
8 | [12 empty] | 8 | 64 | -3.904 | -3.686
9 | [12 empty] | 9 | 81 | -4.1724 | -4.0453
10 | [12 empty] | sigma (mm) |  | 0.016862 | 0.017624
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -4.4945 | -1.0751 | -0.67456 | -0.37941 | -0.36763 | -1.2071 | -1.4638 | -1.1528 | -1.459 | -1.0563 | -0.8685 | -0.98622 | -0.76895 | -0.56154 | -0.70447 | -0.61653 | -1.5567 | -1.7163 | -1.338 | -0.72585 | -1.1048
3 | 34.0035 | 1.8289 | 2.2023 | 1.6637 | 3.1295 | 2.4432 | 2.9608 | 3.1135 | 2.9909 | 2.8538 | 1.3305 | 1.5763 | 2.8647 | 1.3393 | 1.1917 | 2.5752 | 3.164 | 2.4408 | 2.7681 | 0.92245 | 3.3333
4 | -1.523 | 0.13707 | -0.020824 | 0.10564 | 0.12395 | 0.33428 | 0.25486 | -0.0099529 | 0.18682 | 0.3176 | 0.19218 | 0.092635 | -0.046338 | 0.089779 | 0.16843 | -0.035428 | 0.04124 | 0.069368 | 0.050725 | -0.20799 | 0.05263
5 | -5.8326 | -0.61699 | -0.60341 | -0.40216 | -0.36363 | -0.2145 | -0.263 | -0.42581 | -0.53709 | -0.30472 | -0.094869 | -0.18104 | -0.44668 | -0.49391 | -0.20288 | -0.50447 | -0.52562 | -0.60335 | -0.50295 | -3.2686 | -0.6758
6 | -0.22173 | -0.033195 | 0.02737 | -0.074847 | -0.0093847 | 0.013576 | 0.017417 | -0.018616 | -0.03119 | -0.075481 | 0.0074095 | 0.088626 | 0.10946 | -0.0072821 | -0.049602 | -0.046617 | 0.092283 | 0.057069 | 0.032362 | 0.060911 | -0.0010476
7 | 2.5367 | 1.0703 | 1.1836 | 1.2427 | 1.2677 | 1.2702 | 1.2345 | 1.2174 | 1.2171 | 1.2736 | 1.1998 | 1.1998 | 1.1893 | 1.2045 | 1.2074 | 1.2662 | 1.2519 | 1.3488 | 1.3349 | 1.0104 | 1.2675
8 | -0.080435 | 0.018774 | 0.028715 | 0.036222 | -0.019176 | -0.0079225 | 0.0044292 | 0.0030518 | -0.0060228 | -0.010333 | 0.022957 | -0.017835 | 0.0051561 | 0.038287 | 0.010678 | -0.0083465 | -0.0052179 | 0.00032557 | -0.003722 | -0.037159 | 0.0020395
9 | 0.37468 | 0.53787 | 0.52055 | 0.55765 | 0.56974 | 0.53994 | 0.54589 | 0.55203 | 0.5656 | 0.55035 | 0.54368 | 0.52566 | 0.55454 | 0.54698 | 0.52161 | 0.54542 | 0.5654 | 0.59713 | 0.58587 | 0.507 | 0.5454
10 | -2.3727e-06 | 5.4793e-06 | 6.2001e-06 | -1.1162e-07 | 1.0005e-05 | -1.0639e-05 | -2.2555e-06 | -0.0038913 | 2.6801e-06 | 1.7486e-06 | -2.0758e-06 | 4.0002e-07 | 3.098e-07 | 9.2373e-07 | 8.7524e-06 | -5.5867e-07 | -2.9665e-05 | -1.9399e-05 | 7.1836e-06 | 3.4558e-05 | 0
11 | 0.62939 | 0.75146 | 0.74258 | 0.75175 | 0.74786 | 0.75407 | 0.75116 | 0.75171 | 0.75448 | 0.74815 | 0.76168 | 0.75867 | 0.75228 | 0.7632 | 0.76465 | 0.74991 | 0.75536 | 0.7622 | 0.75739 | 0.66769 | 0.74758
12 | 0.028384 | 0.0010561 | 0.023109 | -0.038832 | -0.038485 | -0.041886 | 0.0070543 | -0.015159 | -0.033632 | -0.02361 | -0.014393 | -0.0048414 | -0.012881 | -0.033314 | -0.061588 | -0.041596 | -0.020728 | -0.038813 | -0.027805 | -0.073768 | -0.022473 | ten times!
13 | 0.89474 | 0.79219 | 0.79802 | 0.7322 | 0.73636 | 0.74708 | 0.74265 | 0.75091 | 0.71809 | 0.71764 | 0.72539 | 0.76244 | 0.74563 | 0.72246 | 0.72364 | 0.73357 | 0.72086 | 0.75634 | 0.73325 | 0.55369 | 0.74163 | ten times!
14 | 0.041903 | -0.035308 | -0.042363 | -0.037702 | -0.053264 | -0.059024 | -0.055297 | -0.04071 | -0.055374 | -0.056856 | -0.05889 | -0.033385 | -0.039782 | -0.056051 | -0.065111 | -0.051186 | -0.040099 | -0.04335 | -0.045947 | -0.059398 | -0.044722 | ten times!
15 | -0.0011352 | 0.32369 | 0.30582 | 0.31515 | 0.3338 | 0.3282 | 0.31996 | 0.33284 | 0.33684 | 0.34225 | 0.35622 | 0.32787 | 0.3072 | 0.31235 | 0.29306 | 0.30387 | 0.28797 | 0.32224 | 0.28878 | 0.15972 | 0.30314 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 13.7191 | 0.98999 | 0.93899 | 1.6078 | 1.4752 | 2.4458 | 1.912 | 0.89968 | 1.7686 | 1.587 | 2.5445 | 2.4508 | 2.1416 | 2.1887 | 1.5039 | 1.909 | 1.6854 | 1.4616 | 1.3206 | -3.4229 | 1.9436
3 | 0.8992 | 1.0007 | 0.86471 | 0.72583 | 0.83453 | 0.0045104 | -0.02187 | 0.15904 | -0.10594 | -0.50584 | 0.43808 | 0.30304 | 0.3688 | 0.39119 | 0.53808 | 0.24182 | -0.02045 | 0.52275 | 0.26804 | 0.020892 | 0.33942
4 | 2.8077 | 0.058156 | 0.07454 | 0.028345 | 0.11824 | -0.38008 | -0.40856 | -0.1564 | -0.2532 | -0.09348 | -0.29439 | -0.24213 | -0.0028359 | -0.32678 | -0.25507 | -0.34625 | -0.41866 | -0.32318 | -0.0040996 | -0.061956 | -0.083785
5 | 2.366 | 0.2486 | 0.178 | 0.32801 | 0.22305 | 0.0080541 | -0.06079 | -0.16295 | -0.055661 | 0.090692 | 0.1063 | -0.10142 | -0.24151 | 0.011402 | 0.046498 | 0.12293 | -0.087136 | 0.021999 | 0.22641 | -0.089732 | 0.1197
6 | 0.36221 | -0.028828 | -0.03329 | -0.019006 | -0.046092 | -0.035782 | -0.068104 | -0.088303 | -0.048025 | -0.058348 | -0.038382 | 0.0024913 | -0.051162 | 6.2948e-07 | 0.0048894 | -0.036748 | 0.00061452 | 0.074167 | 0.20298 | 0.20902 | 0.0056015
7 | 1.5246 | -0.099813 | -0.027896 | -0.041739 | -0.017617 | -0.010167 | -0.019619 | -0.027516 | -0.047452 | -0.047287 | -0.073463 | -0.0066952 | -0.085398 | -0.038269 | -0.084628 | -0.067867 | -0.090811 | -0.089021 | -0.10671 | 0.052994 | 0.00012006
8 | -0.010114 | -0.0079415 | -0.029982 | -0.017619 | -0.0066506 | -0.013918 | -0.0011982 | -0.0023445 | 0.0072978 | -0.01205 | -0.006199 | 0.0063921 | -0.026871 | 0.032081 | 0.046857 | 0.015444 | 0.023579 | 0.025669 | 0.048403 | -0.005464 | 0.0038133
9 | -0.26454 | -0.018879 | -0.014916 | -0.014825 | -0.028656 | -0.027306 | -0.028694 | -0.015693 | -0.031027 | -0.020395 | -0.003866 | -0.009786 | 0.01375 | -0.017793 | -0.03278 | -0.026634 | -0.033109 | -0.024143 | -0.024431 | -0.088675 | -0.030234
10 | 1.0985e-05 | -7.1662e-06 | 1.7127e-06 | 4.6348e-06 | 1.2914e-05 | 1.6688e-06 | -6.6681e-07 | 0.00013149 | -2.669e-06 | 3.7266e-05 | 1.6617e-06 | 9.3804e-08 | -8.7269e-07 | 1.3176e-05 | -5.6654e-06 | 6.184e-06 | 4.5756e-06 | 3.4632e-05 | -5.0713e-06 | 3.9429e-05 | 0
11 | 0.14803 | -0.041207 | -0.040529 | -0.019488 | -0.027629 | -0.033944 | -0.038181 | -0.02764 | -0.029391 | -0.040193 | -0.046128 | -0.047036 | -0.039281 | -0.042701 | -0.04776 | -0.032907 | -0.031796 | -0.02813 | -0.025403 | -0.021597 | -0.029108 | ten times!
12 | -0.021586 | -0.072184 | -0.056481 | -0.036266 | -0.0697 | -0.038285 | -0.061443 | -0.073134 | -0.048018 | -0.057263 | -0.054518 | -0.03333 | -0.039831 | -0.042172 | -0.031136 | -0.040354 | -0.043528 | -0.10445 | -0.070416 | -0.015414 | -0.050509 | ten times!
13 | -0.22004 | -0.0023458 | -0.023332 | -0.031368 | -0.016819 | -0.016208 | -0.028244 | -0.022239 | -0.01952 | -0.029906 | -0.0085095 | -0.02159 | -0.022202 | -0.029792 | -0.039303 | -0.033691 | -0.00063903 | 0.014481 | 0.00070013 | -0.0044142 | -0.024714 | ten times!
14 | 0.00099851 | -0.18697 | -0.17101 | -0.16276 | -0.15301 | -0.15443 | -0.14333 | -0.14809 | -0.16671 | -0.16168 | -0.17572 | -0.15778 | -0.15929 | -0.14115 | -0.12516 | -0.13682 | -0.14265 | -0.17664 | -0.15372 | -0.094415 | -0.14843 | ten times!
15 | 0.0040643 | -0.077462 | -0.086136 | -0.021824 | -0.037843 | -0.085588 | -0.065014 | -0.062188 | -0.061519 | -0.019128 | -0.08163 | -0.056693 | -0.051553 | -0.092366 | -0.10029 | -0.098851 | -0.0027789 | -0.0084763 | -0.081811 | -0.1042 | -0.059956 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 0.38357 | 0.95045 | 1.2088 | 0.52369 | 0.44772 | 1.5291 | 1.0976 | 0.93697 | 1.5729 | 1.0129 | 1.4873 | 0.74165 | 0.8836 | 0.87317 | 1.3278 | 1.0067 | 1.3668 | 1.4579 | 1.3857 | 1.1004 | 1.0553
3 | 38.8972 | 1.5043 | 1.5431 | 0.93797 | 1.0999 | 0.72514 | 1.5805 | 1.7081 | 1.6581 | 1.7237 | -0.35413 | 0.36246 | 1.6682 | 0.19021 | 0.82869 | 1.6185 | 1.8602 | 2.0914 | 2.3306 | 1.3026 | 2.4994
4 | -0.95386 | 0.040143 | -0.16296 | -0.24903 | -0.12684 | 0.11923 | 0.022098 | -0.076882 | -0.076433 | -0.14588 | -0.099666 | -0.21278 | -0.047885 | -0.214 | 0.028351 | -0.048417 | 0.04706 | 0.065909 | 0.023484 | -0.37971 | -0.091527
5 | -4.9306 | -0.60824 | -0.77759 | -0.36639 | -0.14504 | -0.17965 | -0.25383 | -0.26478 | -0.39444 | -0.34002 | 0.12265 | -0.15908 | -0.38945 | -0.47252 | -0.34817 | -0.54247 | -0.55408 | -0.42489 | -0.51695 | -3.378 | -0.61433
6 | -0.34167 | 0.012991 | -0.088113 | -0.069772 | -0.15617 | 0.011986 | 0.0091947 | -0.047834 | -0.076036 | -0.063702 | -0.090396 | -0.069559 | 0.010151 | -6.0651e-05 | -0.0717 | -0.036507 | -0.025677 | 0.062025 | 0.032423 | -0.089656 | -0.049739
7 | 2.6704 | 1.0635 | 1.1098 | 1.09 | 1.1713 | 1.2083 | 1.147 | 1.1685 | 1.1489 | 1.2077 | 1.2499 | 1.2303 | 1.1771 | 1.1111 | 1.1373 | 1.1099 | 1.1583 | 1.276 | 1.1323 | 0.85195 | 1.1999
8 | -0.15245 | -0.0082733 | -0.021595 | -0.062708 | -0.04338 | 0.014925 | 0.026154 | 0.0090226 | -0.013771 | -0.0026755 | 0.028639 | 0.024318 | 0.016003 | 0.028405 | -0.0089791 | 0.011222 | -0.020073 | 0.0044354 | -0.025734 | -0.033311 | -0.0092557
9 | 0.36581 | 0.54829 | 0.53018 | 0.54812 | 0.54418 | 0.51645 | 0.5347 | 0.54427 | 0.54948 | 0.54674 | 0.52294 | 0.52001 | 0.53415 | 0.54185 | 0.52179 | 0.53093 | 0.54074 | 0.57233 | 0.54469 | 0.46551 | 0.53076
10 | -3.1931e-05 | -1.233e-06 | 1.6844e-06 | -3.1643e-05 | -1.1231e-07 | -2.4501e-06 | -4.3255e-06 | -3.4322e-06 | 2.7788e-06 | 2.0602e-06 | -0.0044077 | 9.7615e-06 | 2.8968e-06 | 8.1935e-06 | -0.0034644 | 0.0004507 | -1.5762e-05 | 0.0014464 | 3.5819e-05 | -0.00010894 | -0.00068323
11 | 0.6092 | 0.74976 | 0.7497 | 0.77062 | 0.76474 | 0.76779 | 0.76665 | 0.76026 | 0.7649 | 0.75974 | 0.77045 | 0.76325 | 0.76409 | 0.77219 | 0.76922 | 0.76381 | 0.76076 | 0.76217 | 0.75859 | 0.66238 | 0.75505
12 | -0.039893 | -0.01631 | 0.0034536 | 0.0059874 | -0.05493 | -0.010199 | -0.00010983 | -0.016001 | -0.020255 | -0.02665 | 0.0067337 | 0.021654 | 0.0055971 | 0.04374 | 0.022475 | 0.020584 | 0.032274 | 0.012522 | -0.010063 | 0.018674 | -0.0020409 | ten times!
13 | 0.88871 | 0.73401 | 0.74693 | 0.7077 | 0.70176 | 0.70689 | 0.71363 | 0.70032 | 0.68876 | 0.70677 | 0.71693 | 0.72573 | 0.76078 | 0.71003 | 0.73582 | 0.74279 | 0.74321 | 0.75226 | 0.74836 | 0.59076 | 0.72722 | ten times!
14 | 0.025998 | -0.030984 | -0.014916 | -0.012528 | -0.019884 | -0.0069084 | -0.01036 | -0.020527 | -0.018766 | -0.0075264 | 0.0022761 | -0.018889 | -0.0069369 | 0.0015363 | -0.0055306 | -0.0030664 | 0.0084257 | 0.016543 | 0.0085518 | -0.0019603 | -0.0058545 | ten times!
15 | 0.063014 | 0.36907 | 0.33427 | 0.36808 | 0.357 | 0.36804 | 0.36568 | 0.37054 | 0.34579 | 0.33828 | 0.38152 | 0.38516 | 0.3662 | 0.37369 | 0.37676 | 0.38345 | 0.33392 | 0.34673 | 0.342 | 0.28386 | 0.34922 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 3.7272 | 2.0932 | 0.74285 | -0.64674 | -0.5966 | 0.94506 | 1.0513 | 0.55596 | 1.5267 | 1.5153 | 0.067062 | -0.72111 | 0.1564 | 0.54385 | 0.58821 | 0.20432 | 0.6483 | 1.0788 | 1.7472 | -0.062967 | 0.70402
3 | 2.3728 | 0.9538 | 0.90958 | 0.53925 | 0.55967 | 0.672 | 0.46381 | 0.28969 | 0.76108 | 0.29645 | 0.67599 | 0.24449 | 0.20304 | 0.54939 | 0.76048 | 0.67797 | 0.70212 | 0.97098 | 0.80492 | 1.074 | 0.68514
4 | 1.5226 | 0.27933 | -0.12306 | 0.25627 | 0.25361 | -0.40766 | -0.31962 | -0.2228 | -0.46125 | -0.30507 | 0.12791 | 0.12271 | -0.18283 | 0.098973 | -0.08573 | 0.31697 | 0.098756 | -0.21124 | 0.23113 | 0.08673 | 0.028476
5 | 3.3683 | 0.12993 | 0.045693 | 0.084838 | 0.21208 | -0.025804 | -0.16381 | -0.094611 | 0.10396 | -0.011585 | 0.15894 | 0.069194 | 0.023552 | 0.22035 | 0.075646 | 0.18526 | 0.15517 | 0.24654 | 0.24557 | -0.0085833 | 0.19469
6 | 0.26979 | -0.0043503 | -0.02144 | -0.30259 | -0.29762 | -0.2229 | -0.17035 | -0.14631 | -0.1305 | -0.088354 | -0.16529 | -0.079845 | -0.10494 | -0.13073 | -0.14475 | -0.08907 | -0.1114 | -0.085337 | -0.14675 | -0.0026271 | -0.12212
7 | 1.8709 | -0.03778 | -0.026224 | -0.085076 | -0.02997 | -0.088549 | -0.06676 | -0.088996 | -0.07159 | -0.044037 | 0.00046959 | -0.056553 | -0.095288 | -0.082997 | -0.10774 | -0.052956 | -0.061883 | -0.046223 | -0.077174 | 0.013901 | 0.0027962
8 | -0.01859 | -0.009699 | 0.012764 | -0.021759 | -0.0074963 | -0.039275 | -0.043307 | -0.010891 | 0.0036726 | -0.014574 | -0.0034094 | -0.02604 | 0.0020245 | -0.058619 | -0.047861 | -0.059311 | -0.045625 | -0.067408 | -0.043961 | -0.011333 | -0.024891
9 | -0.16604 | 0.022834 | 0.0081943 | 0.01763 | -0.00086461 | -0.042762 | -0.038518 | -0.0323 | -0.029964 | -0.042066 | -0.026334 | -0.031634 | -0.043036 | -0.054 | -0.045173 | -0.023588 | -0.037662 | -0.016992 | -0.013462 | -0.080004 | -0.02996
10 | 8.3454e-05 | -5.2758e-06 | -7.4796e-06 | 2.8826e-05 | -6.3245e-06 | 1.0977e-05 | 1.6359e-06 | 4.5235e-06 | -3.1763e-06 | -2.7517e-06 | 0.00017323 | 1.4013e-05 | 1.1342e-05 | 1.1711e-05 | 0.00020483 | -2.2446e-05 | -2.7699e-05 | -5.2885e-05 | -8.4258e-06 | -5.0225e-05 | 2.7665e-05
11 | 0.17978 | -0.033664 | -0.03165 | -0.040314 | -0.034098 | -0.036548 | -0.039732 | -0.031914 | -0.029973 | -0.033754 | -0.030431 | -0.032496 | -0.035151 | -0.039447 | -0.046041 | -0.035082 | -0.03244 | -0.028495 | -0.030304 | -0.031893 | -0.027548
12 | -0.12106 | -0.063473 | -0.068682 | -0.12922 | -0.099857 | -0.060112 | -0.04034 | -0.056106 | -0.043388 | -0.029153 | -0.084568 | -0.053835 | -0.048568 | -0.083531 | -0.062316 | -0.041427 | -0.0098159 | 0.037321 | -0.028311 | -0.0015924 | -0.05322 | ten times!
13 | -0.14667 | 0.019667 | 0.01221 | -0.0094802 | 0.014117 | -0.025894 | -0.023078 | -0.01628 | -0.011756 | -0.019003 | -0.017272 | -0.028786 | -0.021859 | -0.025958 | -0.034795 | -0.026582 | -0.02682 | -0.0043031 | -0.017752 | -0.0559 | -0.020844 | ten times!
14 | 0.022328 | -0.19012 | -0.18559 | -0.1638 | -0.15251 | -0.17362 | -0.17483 | -0.15202 | -0.16231 | -0.14197 | -0.13627 | -0.14607 | -0.13539 | -0.15113 | -0.14341 | -0.15002 | -0.15326 | -0.16347 | -0.14513 | -0.019412 | -0.14419 | ten times!
15 | 0.026301 | -0.02779 | -0.035727 | 0.016732 | -0.036417 | -0.041757 | 0.026628 | -0.038813 | -0.041521 | -0.03558 | -0.012401 | 0.0026749 | 0.014202 | 0.012427 | -0.024219 | -0.016233 | 0.03631 | 0.0034404 | -0.063859 | -0.060412 | -0.014788 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00054775 | -0.00024101 | -0.00024101 | -0.00024101 | -4.382e-05 | -4.382e-05 | -0.00024101 | -4.382e-05 | -0.00024101 | -4.382e-05 | -4.382e-05 | -4.382e-05 | -4.382e-05 | -4.382e-05 | 0.00015337 | 0.00015337 | 0.00015337 | 0.00054775 |  | 596.6144
|  | 0.00044911 | -0.0001424 | -0.0001424 | -0.0001424 | 5.477e-05 | 5.477e-05 | -0.0001424 | 5.477e-05 | -0.0001424 | -0.0001424 | -0.0001424 | -0.0001424 | -0.0001424 | 5.477e-05 | 5.477e-05 | 5.477e-05 | 5.477e-05 | 0.00044911 |  | 596.6732
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3679.5918
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.4536 | -3760.4457 | 6.1626 | -1.7255 | -1.7255 | -1.7255 | 0.24651 | 0.24651 | -1.7255 | 0.24651 | -1.7255 | 0.24651 | 0.24651 | 0.24651 | 0.24651 | 0.24651 | 2.2186 | 2.2186 | 2.2186 | 6.1626 | -3598.7379
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.13115 | 3.8834 | 0.34214 | -0.40339 | 0.91638 | 0.35494 | 0.36245 | 0.1644 | 0.52728 | 0.10109 | -0.6533 | -1.2217 | -0.11615 | 0.34836 | 0.17008 | -0.4199 | 0.1508 | -0.20359 | 0.021497 | -0.44139 | 0.041106
db2 | -0.99322 | -4.6954 | -0.1042 | 0.29271 | 0.62885 | 0.6647 | -0.22517 | -0.47913 | -0.15815 | -0.45604 | -0.012512 | 0.12051 | 0.0022648 | 0.21697 | 0.42002 | 0.27663 | 0.36042 | -0.49861 | -0.68312 | -0.36614 | -0.74174
db3 | 1.5473 | 33.9626 | -0.53045 | -0.15538 | -0.69228 | 0.77491 | 0.078806 | 0.60003 | 0.75418 | 0.63256 | 0.50442 | -1.0295 | -0.7809 | 0.50758 | -1.0214 | -1.1664 | 0.21633 | 0.81072 | 0.086366 | 0.41039 | 0.92475
db4 | 0.11249 | -1.5218 | 0.027541 | -0.12397 | -0.015692 | 0.0046484 | 0.22395 | 0.1434 | -0.12543 | 0.070002 | 0.19457 | 0.081179 | -0.018583 | -0.15509 | -0.012778 | 0.059061 | -0.14032 | -0.091497 | -0.063702 | -0.057292 | -0.095908
db5 | -0.36604 | -5.843 | -0.21177 | -0.19856 | 0.0015067 | 0.041371 | 0.18866 | 0.14069 | -0.024059 | -0.13362 | 0.098826 | 0.30938 | 0.22294 | -0.042298 | -0.088423 | 0.20319 | -0.099152 | -0.11256 | -0.19406 | -0.10207 | -3.2664
db6 | 0.0098686 | -0.32634 | -0.060984 | -0.01569 | -0.048256 | 0.024448 | -0.014163 | -0.0032313 | -0.029245 | -0.029163 | -0.015734 | -0.018361 | 0.068909 | 0.09337 | -0.053657 | -0.089482 | -0.082523 | 0.16066 | 0.11713 | -0.0040333 | 0.0090864
db7 | 1.2435 | 2.5374 | -0.16285 | -0.048891 | 0.012275 | 0.035319 | 0.037017 | 0.0019093 | -0.015812 | -0.015453 | 0.042759 | -0.033027 | -0.032479 | -0.042934 | -0.028267 | -0.025606 | 0.03435 | 0.021768 | 0.11842 | 0.10149 | 1.0147
db8 | 0.0074029 | -0.074199 | -0.0018177 | 0.0019326 | 0.05063 | -0.0006121 | -0.026191 | -0.0096369 | -0.00492 | -0.0062629 | 0.02268 | 0.006321 | -0.030954 | -0.0069652 | 0.0087397 | -0.014101 | -0.030665 | 0.034761 | 0.034265 | -0.027204 | -0.07443
db9 | 0.55104 | 0.373 | -0.015923 | -0.031214 | 0.0066296 | 0.019124 | -0.013442 | -0.0068074 | -0.0019847 | 0.012702 | 0.0030444 | -0.0087488 | -0.026841 | 0.001729 | -0.0041234 | -0.030012 | -0.006597 | 0.019601 | 0.049418 | 0.033445 | 0.51116
db10 | 0.0030382 | -0.027621 | -0.025518 | -0.037059 | 0.033618 | 0.039623 | -0.021921 | -0.014483 | -0.0081954 | 0.0092506 | 0.064828 | -0.019858 | -0.013127 | -0.0098887 | -0.041972 | -0.034864 | -0.028599 | 0.075011 | 0.062424 | -0.02927 | -0.058446
db11 | 0.75449 | 0.6289 | -0.0027955 | -0.011708 | -0.0026997 | -0.006731 | -0.00061335 | -0.0034935 | -0.0033581 | -0.00026095 | -0.005974 | 0.0071406 | 0.0042222 | -0.0021393 | 0.0089873 | 0.010644 | -0.0044272 | 0.0017063 | 0.0082288 | 0.0032714 | 0.66884
db12 | -0.0019929 | 0.0016175 | -0.0010739 | -0.00023315 | 0.0024182 | 0.0031426 | -0.0042872 | 0.0013774 | 0.00036915 | 3.7822e-05 | 0.007349 | -0.0012844 | 0.00026597 | -0.0001578 | -0.0057134 | -0.0078373 | -0.005153 | 0.0088706 | 0.005888 | -0.0039786 | -0.013168
db13 | 0.074284 | 0.089317 | 0.0062004 | 0.0064658 | -0.0010032 | -0.0012299 | -0.00029029 | -0.00028559 | -0.00020503 | -0.0032125 | -0.0031152 | -0.0019239 | 0.0021659 | 0.0007264 | -0.0019054 | -0.0012135 | -0.00095519 | -0.0017318 | 0.0016558 | -0.00014284 | 0.055992
db14 | -0.004637 | 0.0041559 | 0.00018248 | -0.001275 | 0.0030327 | 0.0019136 | -0.0028358 | -0.0017091 | 0.00016604 | -0.00031273 | 0.0032947 | -0.0024872 | 0.00072263 | 0.00036676 | -0.0034174 | -0.0033749 | -0.0021438 | 0.0047555 | 0.0041665 | -0.0010449 | -0.0082707
db15 | 0.031878 | -0.00011352 | 0.00049118 | -0.0012967 | -0.00036372 | 0.001502 | 0.00094122 | 0.00011743 | 0.0014058 | 0.0018054 | 0.0023467 | 0.0037432 | 0.00090898 | -0.0011582 | -0.00064315 | -0.0025727 | -0.001491 | -0.0030818 | 0.00034531 | -0.003 | 0.015972
da2 | 1.7107 | 14.177 | -0.70836 | -0.76694 | -0.093243 | -0.23689 | 0.7145 | 0.18982 | -0.83958 | 0.04161 | -0.14408 | 0.82772 | 0.74139 | 0.44163 | 0.46852 | -0.20713 | 0.18763 | -0.0093326 | -0.2324 | -0.37487 | -3.4283
da3 | 0.33308 | 0.84616 | 0.6694 | 0.53062 | 0.39347 | 0.50301 | -0.33022 | -0.35472 | -0.17386 | -0.44177 | -0.84311 | 0.10595 | -0.028849 | 0.035484 | 0.062313 | 0.20922 | -0.087671 | -0.36646 | 0.18157 | -0.064376 | 0.022273
da4 | -0.17974 | 2.6223 | 0.22191 | 0.24381 | 0.21319 | 0.30565 | -0.19678 | -0.22642 | 0.031217 | -0.068113 | 0.091821 | -0.11495 | -0.062326 | 0.1743 | -0.14318 | -0.077552 | -0.1646 | -0.24777 | -0.1464 | 0.16619 | -0.023215
da5 | 0.050192 | 2.3487 | 0.19698 | 0.128 | 0.27732 | 0.17081 | -0.041262 | -0.11046 | -0.21269 | -0.10619 | 0.037078 | 0.05669 | -0.1514 | -0.29048 | -0.037988 | -0.0048402 | 0.074193 | -0.13434 | -0.02364 | 0.17221 | -0.098067
da6 | -0.01603 | 0.44045 | 0.022341 | 0.004369 | 0.0015867 | -0.04416 | -0.05182 | -0.066416 | -0.10581 | -0.051958 | -0.059806 | -0.03287 | 0.023457 | -0.020312 | 0.0021408 | 0.036152 | -0.032831 | 0.030952 | 0.098841 | 0.24614 | 0.18969
da7 | -0.054574 | 1.5207 | -0.045376 | 0.027437 | 0.012649 | 0.036221 | 0.045475 | 0.035079 | 0.026978 | 0.0070375 | 0.0054458 | -0.018849 | 0.047677 | -0.03069 | 0.015246 | -0.031799 | -0.013201 | -0.033097 | -0.032073 | -0.05416 | 0.055044
da8 | 0.0038921 | 0.016172 | 0.010912 | -0.020986 | -0.018604 | -0.020027 | -0.035559 | -0.013321 | -0.026471 | -0.0090903 | -0.026224 | -0.016955 | 0.0054588 | -0.021988 | 0.019695 | 0.051639 | 0.0045018 | 0.029226 | 0.028016 | 0.059776 | -0.010569
da9 | -0.020058 | -0.2672 | -0.0010524 | 0.0032299 | 0.0057011 | -0.010049 | -0.0058824 | -0.0082526 | 0.0045782 | -0.011585 | -0.0030195 | 0.016644 | 0.010139 | 0.033543 | 0.0036144 | -0.013844 | -0.0058033 | -0.0096952 | -0.0023777 | -0.0058873 | -0.085986
da10 | -0.00076687 | 0.038717 | 0.040969 | 0.024502 | 0.0056395 | -0.014082 | -0.03032 | -0.013403 | -0.033606 | -0.019856 | -0.01743 | -0.010379 | 0.0065529 | 0.016239 | -0.013036 | 0.016562 | -0.011432 | 0.017293 | 0.010641 | 0.025145 | -0.016246
da11 | -0.035575 | 0.14797 | -0.0057269 | -0.0048851 | 0.015919 | 0.0073696 | 0.0021153 | -0.0024703 | 0.0080163 | 0.0061053 | -0.0055324 | -0.010337 | -0.011434 | -0.0037297 | -0.0066254 | -0.012316 | 0.0030084 | 0.003718 | 0.0066676 | 0.010137 | -0.020877
da12 | -0.0054864 | 0.0058655 | 0.0032546 | 0.0027763 | 0.0022981 | -0.0032325 | -0.0020756 | -0.0023566 | -0.0059104 | -0.0016347 | -0.0024431 | -0.0013416 | 0.0028067 | 0.0033329 | -0.00043243 | 0.0041284 | -3.8572e-06 | 0.0031874 | -0.0035343 | 0.0011807 | -0.003446
da13 | -0.0019291 | -0.021638 | 0.0020764 | 0.0003114 | -0.002239 | -0.00092302 | 0.0012339 | -0.00045449 | -2.7888e-05 | -0.00025866 | -0.0030954 | 0.0017231 | 9.4458e-06 | -0.00020451 | 0.00010007 | -0.0015675 | -0.00067869 | 0.00010955 | 0.0014793 | 0.002406 | 0.00083815
da14 | -0.015575 | 7.3157e-05 | -0.00042436 | 0.00030292 | -0.00050758 | -0.00085833 | -0.0015321 | 0.00049419 | -0.001374 | -0.0025022 | -0.0017641 | -0.0025492 | 0.00021675 | 0.00058163 | 0.0011931 | 0.0043965 | 0.0016001 | 0.0023244 | -0.0014622 | 0.0018646 | -0.0087832
da15 | -0.0060619 | 0.00040643 | -0.0016842 | -0.0025517 | 0.0038796 | 0.0022776 | -0.0024968 | -0.0004395 | -0.00015686 | -8.9962e-05 | 0.0041491 | -0.002101 | 0.00039264 | 0.00090663 | -0.0031747 | -0.003967 | -0.0038231 | 0.0057841 | 0.0052143 | -0.0021191 | -0.01042
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3673.3914
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.5198 | -3776.9124 | 5.0528 | -0.86267 | -0.86267 | -0.86267 | 1.1092 | 1.1092 | -0.86267 | 1.1092 | -0.86267 | -0.86267 | -0.86267 | -0.86267 | -0.86267 | 1.1092 | 1.1092 | 1.1092 | 1.1092 | 5.0528 | -3569.8705
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.094113 | 3.7515 | 0.13449 | 0.33352 | 0.47625 | -0.14033 | 0.25943 | 0.18461 | 0.45306 | 0.45965 | 0.26757 | -0.47774 | -0.050958 | -0.34908 | -0.2854 | -0.53778 | -0.32008 | -0.45172 | -0.036107 | 0.080621 | -0.66849
db2 | 1.0982 | -0.24879 | -0.15941 | 0.093167 | -0.56237 | -0.65632 | 0.42644 | 0.011781 | -0.17086 | 0.46054 | -0.096774 | 0.39102 | -0.35139 | -0.19698 | -0.2207 | 0.23339 | -0.08825 | 0.29162 | 0.35564 | 0.23946 | 1.115
db3 | 0.47151 | 38.9361 | 0.21881 | 0.26527 | -0.35074 | -0.18364 | -0.56029 | 0.30045 | 0.42599 | 0.38109 | 0.44545 | -1.6351 | -0.92095 | 0.38476 | -1.093 | -0.45343 | 0.33932 | 0.57828 | 0.8112 | 1.0465 | 1.3225
db4 | -0.062764 | -0.72725 | 0.10236 | -0.096115 | -0.19775 | -0.066192 | 0.18531 | 0.081406 | -0.011493 | -0.010151 | -0.079026 | -0.036 | -0.15229 | 0.0071783 | -0.15631 | 0.091881 | 0.016138 | 0.090718 | 0.12662 | 0.10372 | -0.28175
db5 | -0.3285 | -4.9105 | -0.24192 | -0.41057 | 0.0023847 | 0.22719 | 0.18585 | 0.11506 | 0.10178 | -0.024481 | 0.029013 | 0.48841 | 0.20532 | -0.022578 | -0.10701 | 0.016803 | -0.17674 | -0.18432 | -0.057167 | -0.14703 | -3.38
db6 | -0.0341 | -0.42518 | 0.043052 | -0.064024 | 0.014384 | -0.12871 | 0.010588 | 0.073309 | -0.036867 | -0.061325 | -0.050369 | -0.056294 | -0.00047779 | 0.078752 | 0.058526 | -0.039215 | -0.0031761 | 0.062886 | 0.096701 | 0.0022608 | -0.12415
db7 | 1.1723 | 2.6804 | -0.098092 | -0.053101 | -0.072698 | 0.010247 | 0.046651 | -0.012507 | 0.0070808 | -0.012125 | 0.046771 | 0.088599 | 0.069596 | 0.017058 | -0.050281 | -0.024939 | -0.052561 | -0.0013462 | 0.11558 | -0.023935 | 0.8484
db8 | -0.0011524 | -0.18301 | -0.011515 | -0.027881 | -0.028572 | -0.04667 | -0.0042767 | 0.046041 | -0.0040515 | -0.025059 | -0.013653 | 0.029494 | 0.044969 | 0.03805 | 0.045211 | -0.0087287 | 0.011919 | 0.015259 | 0.00587 | -0.066407 | -0.067834
db9 | 0.53758 | 0.36728 | 0.0076162 | -0.011898 | 0.012823 | 0.0044022 | -0.021264 | -0.0032312 | 0.0055994 | 0.010394 | 0.0078316 | -0.01632 | -0.018313 | -0.0035731 | 0.0032219 | -0.01776 | -0.0085968 | 0.0047842 | 0.032904 | 0.011379 | 0.45483
db10 | 0.0026336 | -0.042904 | -0.0070553 | -0.013061 | 0.056118 | -0.0094135 | -0.039907 | 0.03069 | -0.027214 | -0.024652 | -0.02427 | -0.0074714 | 0.033643 | 0.035209 | 0.026252 | -0.0072513 | -0.0030295 | 0.056583 | -0.0005949 | -0.074575 | -0.052596
db11 | 0.76345 | 0.60969 | -0.013558 | -0.013666 | 0.0079777 | 0.0016196 | 0.0044605 | 0.0034188 | -0.0031934 | 0.001666 | -0.0035199 | 0.0068524 | -0.00017217 | 0.00076423 | 0.0087562 | 0.0055397 | 0.00019353 | -0.0019109 | -0.0014092 | -0.0038189 | 0.66087
db12 | 0.00039723 | -0.01375 | -0.0028612 | -0.0016948 | 0.006698 | -0.0068216 | -0.0054812 | 0.0033234 | -0.0046564 | -0.0048912 | -0.0055379 | 0.0002468 | 0.0058359 | 0.0046761 | 0.0071482 | 0.0017464 | 0.0015846 | 0.0098948 | 0.00086461 | -0.010074 | -0.0039061
db13 | 0.072696 | 0.08874 | 0.0017754 | 0.0033835 | -0.0013329 | -0.0019233 | -0.0022198 | -0.0009935 | -0.0029185 | -0.0031646 | -0.0015667 | -0.0012671 | -0.0007781 | 0.003124 | -0.0021394 | 0.00037719 | 0.0010732 | 0.0026891 | 0.0029456 | 0.002936 | 0.057968
db14 | -0.00053891 | 0.0020275 | -0.0027886 | -0.0013998 | 0.0031261 | -0.0018011 | -0.0027135 | 0.001646 | -0.0032552 | -0.0026222 | -0.0014987 | 0.000742 | 0.001302 | 0.0024734 | 0.0027658 | -9.8981e-05 | 0.00017751 | 0.0047518 | 0.0021831 | -0.0029898 | -0.0043842
db15 | 0.036146 | 0.0063014 | 0.00076158 | -0.0027189 | 0.00066258 | -0.0004459 | 0.00065883 | 0.00042217 | 0.00090867 | -0.0015663 | -0.0023179 | 0.0020068 | 0.0023708 | 0.00047471 | 0.0012239 | 0.0015307 | 0.0021999 | -0.0027531 | -0.0014726 | -0.0019459 | 0.028386
da2 | 0.6425 | 3.3318 | 1.4637 | 0.11819 | -1.2779 | -1.2332 | 0.29354 | 0.41678 | -0.09509 | 0.88889 | 0.87867 | -0.57444 | -1.3636 | -0.4909 | -0.095352 | -0.058433 | -0.44599 | 0.030586 | 0.445 | 1.0996 | -0.11975
da3 | 0.61364 | 2.345 | 0.34008 | 0.29345 | -0.069939 | -0.055982 | 0.063516 | -0.15392 | -0.32144 | 0.15101 | -0.31542 | 0.062769 | -0.36517 | -0.41297 | -0.061286 | 0.14662 | 0.064659 | 0.090971 | 0.35774 | 0.18532 | 1.0807
da4 | -0.031976 | 1.4991 | 0.29912 | -0.11237 | 0.28797 | 0.28293 | -0.37402 | -0.2918 | -0.18758 | -0.43488 | -0.27595 | 0.1592 | 0.15814 | -0.14727 | 0.13928 | -0.050249 | 0.35418 | 0.12705 | -0.1828 | 0.24902 | 0.29701
da5 | 0.091555 | 3.3738 | 0.038484 | -0.049373 | -0.018004 | 0.11934 | -0.11159 | -0.25861 | -0.18297 | 0.012725 | -0.10302 | 0.068022 | -0.024384 | -0.071547 | 0.12556 | -0.014992 | 0.094155 | 0.060803 | 0.15563 | 0.15977 | 0.0040727
da6 | -0.13126 | 0.10166 | 0.1579 | 0.15499 | -0.16409 | -0.14372 | -0.102 | -0.031999 | -0.027154 | 0.021231 | 0.059728 | -0.044621 | 0.025258 | 0.0089421 | -0.018925 | -0.030542 | 0.026161 | 0.043258 | 0.061865 | 0.0037276 | -0.082906
da7 | -0.062772 | 1.878 | 0.0245 | 0.034693 | -0.023348 | 0.031687 | -0.024206 | -0.0047576 | -0.025741 | -0.0096758 | 0.018586 | 0.062961 | 0.0037939 | -0.033347 | -0.022979 | -0.044799 | 0.009634 | -0.00040492 | 0.016617 | -0.013216 | 0.024535
da8 | -0.023617 | -0.021208 | 0.034894 | 0.064761 | 0.010539 | 0.030014 | -0.021414 | -0.015497 | 0.0049548 | 0.040255 | 0.019739 | 0.014298 | -0.017531 | 0.015076 | -0.04809 | -0.03468 | -0.045922 | -0.0076375 | -0.034269 | -0.0094902 | -0.064135
da9 | -0.024115 | -0.16102 | 0.046492 | 0.030904 | 0.043328 | 0.022826 | -0.017542 | -0.013706 | -0.0073547 | -0.0069965 | -0.018855 | -0.0021581 | -0.0097826 | -0.02041 | -0.031326 | -0.020954 | 0.00053496 | -0.0097305 | 0.0071608 | 0.0075699 | -0.066088
da10 | 0.0051139 | -0.043063 | 0.031363 | 0.046022 | 0.0049613 | 0.019659 | -0.016998 | 0.0030009 | -0.019056 | 0.017052 | 0.012698 | -0.015825 | -0.03197 | -0.022541 | -0.026891 | -0.024148 | -0.023538 | 0.021194 | 0.010689 | 0.014327 | -0.1007
da11 | -0.034646 | 0.1819 | 0.00086889 | 0.0029128 | -0.0065026 | 0.00040468 | -0.0014133 | -0.0051791 | 0.0030745 | 0.0046006 | 0.00079825 | 0.0042054 | 0.0014668 | -0.0010033 | -0.0054763 | -0.011454 | -0.00048831 | 0.0028088 | 0.0061803 | 0.0041963 | -0.030378
da12 | -0.0047817 | -0.015468 | 0.0027308 | 0.0040482 | -0.0068971 | -0.002511 | -0.0026488 | 0.0016841 | -0.0024415 | 0.0027883 | 0.003807 | -0.0048862 | -0.0036884 | -0.0021301 | -0.0059526 | -0.0036437 | -0.0014942 | 0.006871 | 0.010403 | 0.0039617 | -0.010638
da13 | -0.0015664 | -0.014953 | 0.0034972 | 0.0029512 | -0.00082432 | 0.003063 | 0.00016809 | -0.0016503 | 0.00069285 | 0.00084786 | 0.00010551 | -0.00016445 | -0.0023869 | -0.0016215 | -0.0019323 | -0.0018923 | -0.0010804 | -0.0022136 | 0.0011806 | 0.0012597 | -0.0039971
da14 | -0.015217 | 0.0015203 | -0.0012689 | -9.9746e-05 | -0.00010895 | 0.0016135 | -0.00278 | -0.0014513 | -0.00083365 | 0.00050549 | 0.0022139 | 0.00083529 | -0.0011558 | 0.00072832 | -0.0011997 | -0.00044365 | -0.0010791 | 0.0020434 | -0.0001199 | 0.0026008 | -0.0075144
da15 | -0.001455 | 0.0026301 | -0.001324 | -0.0021176 | 0.0031282 | -0.0021867 | -0.0027206 | 0.0041178 | -0.0024263 | -0.0026971 | -0.002103 | 0.00021493 | 0.0017225 | 0.0028752 | 0.0026977 | -0.00096693 | -0.00016826 | 0.005086 | 0.0017991 | -0.0049309 | -0.0060412
Field Colinearity | 0.54021
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.7522 | 0.21098 | -0.53454 | 0.78522 | 0.22378 | 0.23129 | 0.033251 | 0.39613 | -0.030066 | -0.78445 | -1.3528 | -0.2473 | 0.21721 | 0.038931 | -0.55105 | 0.019647 | -0.33474 | -0.10966 | -0.57254 | -0.090046
Twist (m2 mrad) | 28.1415 | 3.0065 | -7.2163 | 10.0116 | 2.6854 | 2.6021 | 0.34914 | 3.8623 | -0.27059 | -6.4717 | -10.1464 | -1.6693 | 1.3032 | 0.20439 | -2.4797 | 0.073676 | -1.0042 | -0.24672 | -0.85881 | -0.033767
Twist Integral (I0) (m2 rad) | 0.016382
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.6574 | 0.040379 | 0.23941 | 0.38213 | -0.23445 | 0.16532 | 0.090492 | 0.35895 | 0.36554 | 0.17346 | -0.57186 | -0.14507 | -0.44319 | -0.37952 | -0.63189 | -0.41419 | -0.54584 | -0.13022 | -0.013492 | -0.7626
Twist (m2 mrad) | 27.4305 | 0.5754 | 3.232 | 4.8722 | -2.8134 | 1.8598 | 0.95017 | 3.4998 | 3.2899 | 1.4311 | -4.2889 | -0.97923 | -2.6591 | -1.9925 | -2.8435 | -1.5532 | -1.6375 | -0.293 | -0.020238 | -0.28597
Twist Integral (I0) (m2 rad) | 0.020831

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.99322 | 2.3573 | 0.11249 | -0.40454 | 1.2325 | -1.3115 | 1.7107 | 0.33308 | -0.17974 | 16.8617 | 0.59661 | 1.0982 | 1.2815 | -0.062764 | -0.367 | 1.1613 | -0.94113 | 0.6425 | 0.61364 | -0.031976 | 17.6241 | 0.59667 | -1.1048 | 3.3333 | 0.05263 | -0.6758 | 1.2675 | 0 | 1.9436 | 0.33942 | -0.083785 | 14414.12 | 1.0553 | 2.4994 | -0.091527 | -0.61433 | 1.1999 | 0 | 0.70402 | 0.68514 | 0.028476 | 14415.03
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85