## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3064
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-08-25
Beginn der Messung / start of measuring: |  | TT.08.JJ 14:05 | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000064
Ende der Messung / end of measuring: |  | TT.08.JJ 22:17
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Weinberg |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-08-25 |  | 2003-08-26 | 2003-08-26
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000064_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000064
Date of test Ap 1 | 2003-08-25 |  |  |  | Date of test Ap 2 | 2003-08-25
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | yellow alarm | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | yellow alarm | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Original data
File name |  | HCMB__A001-03000064_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000064 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 2.0 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | noell3064cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | noell3064cct2 |  |  |  | a even | 1
Test Operator |  | Weinberg |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-08-25 |  |  | Time of test |  |  | 14:05 - !6:46 |  |  |  | Date of test |  | 2003-08-25 |  |  | Time of test |  |  | 17:30 - 22:17
Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03 |  |  |  | Mole name |  | Dipole 15 |  |  | Calibration date |  |  | 2003-03-03
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 26 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 26
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00013055 |  |  | Dy average |  |  | 0.00076885 |  |  |  | Dx average |  | 7.7419e-05 |  |  | Dy average |  |  | 0.00080716
Magnetic length |  | 14.4546 |  |  | Tranfs. Function (mT/KA) |  |  | 596.4444 |  |  |  | Magnetic length |  | 14.4553 |  |  | Tranfs. Function (mT/KA) |  |  | 596.4118
Phase |  | -3.7057 |  |  | Integral field (T.m) |  |  | 0.073282 |  |  |  | Phase |  | -4.015 |  |  | Integral field (T.m) |  |  | 0.073283
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003232 | 0.005072 | 0.005069 | 0.00507 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.00507 | 0.005069 | 0.00507 | 0.005069 | 0.00507 | 0.00507 | 0.00507 | 0.005074 | 0.003235 | 0.073282 | C1 (T) | 0.003231 | 0.005072 | 0.005069 | 0.005069 | 0.005068 | 0.005069 | 0.00507 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.005069 | 0.00507 | 0.00507 | 0.005073 | 0.003239 | 0.073283
Angle (mrad) | 2.7929 | -1.0206 | -0.90894 | -0.62681 | -0.67924 | -0.21012 | -0.28785 | 0.37513 | 0.16458 | -0.041501 | -0.51338 | -0.081788 | 0.31266 | 0.57715 | -0.046823 | -0.041395 | 0.3164 | 0.82818 | 0.55782 | -0.75882 | -3.7057 | Angle (mrad) | 2.6736 | -0.43744 | -0.24251 | -0.042776 | -0.32752 | 0.073181 | 0.16304 | 0.54684 | 0.24503 | -0.070187 | -0.54354 | -0.37744 | -0.21532 | -0.10903 | -0.18939 | -0.17646 | 0.058923 | 0.40207 | -0.10296 | -0.49769 | -4.015
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -1.9332 | -0.36022 | -0.72574 | -0.10108 | -0.086225 | 0.10892 | 0.04892 | 0.073 | -0.60918 | -0.32768 | -0.68721 | 0.10025 | -0.071561 | -0.10924 | -0.38749 | -0.58245 | -0.70487 | -0.91352 | -1.6873 | -1.0481 | -0.46303 | b2 | -0.92842 | 0.075593 | 0.42475 | 0.73749 | 0.82592 | 1.1213 | 0.6939 | 0.24554 | 0.88719 | -0.094447 | 0.73359 | 0.22232 | 0.54289 | 0.74521 | 0.99691 | 0.28677 | 0.64585 | 0.44535 | 0.57599 | 2.829 | 0.58764
b3 | 32.5389 | -0.66577 | -0.14794 | -0.63183 | 0.60405 | -1.2084 | -0.088205 | -1.0961 | -0.35196 | 0.1203 | -1.3607 | -1.3242 | -0.43744 | -2.5152 | -2.252 | -0.47254 | -0.076472 | -0.12003 | 0.48596 | -0.63819 | 0.45636 | b3 | 35.2322 | -0.54249 | -0.29373 | -0.77743 | 0.23889 | -0.74759 | 0.37451 | -0.68327 | -0.16543 | 0.22342 | -1.154 | -0.43044 | -0.51875 | -2.4964 | -1.1282 | 0.29103 | -0.076333 | 0.079646 | 0.76522 | -0.78087 | 0.77389
b4 | 0.48348 | 0.10909 | 0.077857 | -0.033867 | 0.075962 | 0.25122 | 0.084665 | 0.11143 | -0.041596 | -0.090746 | -0.05879 | 0.06107 | -0.15262 | 0.075815 | 0.018742 | 0.1074 | 0.1365 | 0.064055 | 0.19402 | -0.13539 | 0.062886 | b4 | -0.021038 | 0.30391 | 0.12917 | -0.10804 | -0.017883 | 0.1954 | 0.17603 | 0.08677 | 0.0031098 | -0.075531 | -0.0098859 | -0.15129 | -0.036275 | 0.071151 | 0.07156 | 0.028449 | 0.14512 | 0.076437 | -0.00067856 | 0.097278 | 0.048579
b5 | -5.0151 | -0.4459 | -0.29238 | -0.085486 | -0.34908 | -0.072476 | -0.085434 | 0.03938 | -0.055008 | 0.020385 | 0.27464 | -0.031843 | -0.29161 | -0.26351 | -0.029389 | -0.23465 | -0.20656 | -0.22675 | -0.33969 | -3.3796 | -0.41653 | b5 | -4.4657 | -0.59901 | -0.45953 | -0.13202 | -0.34156 | -0.089084 | -0.14239 | -0.097567 | -0.17672 | -0.10407 | 0.27965 | -0.24148 | -0.19499 | -0.12691 | -0.053946 | -0.29132 | -0.14126 | -0.17774 | -0.353 | -3.5476 | -0.44388
b6 | 0.055837 | 0.05431 | -0.0040801 | -0.080276 | 0.060316 | -0.013224 | 0.031065 | 0.10917 | -0.0017672 | -0.14183 | 0.10342 | 0.051798 | -0.0081148 | 0.051013 | 0.030572 | 0.1176 | -0.03431 | 0.19626 | 0.20531 | -0.047539 | 0.03801 | b6 | -0.074615 | 0.163 | 0.076967 | -0.014218 | 0.082706 | -0.11364 | -0.040339 | 0.096362 | 0.019023 | 0.037261 | 0.0073831 | 0.059513 | 0.072477 | -0.046289 | 0.041919 | 0.097379 | 0.0090744 | 0.0071095 | -0.030975 | -0.086986 | 0.021869
b7 | 2.5683 | 1.0444 | 1.1643 | 1.2782 | 1.213 | 1.3321 | 1.2911 | 1.3247 | 1.3301 | 1.2368 | 1.2855 | 1.1523 | 1.2351 | 1.1878 | 1.1668 | 1.2088 | 1.1981 | 1.2905 | 1.2254 | 0.98116 | 1.2673 | b7 | 2.6187 | 0.86269 | 1.0882 | 1.2539 | 1.2186 | 1.2813 | 1.2423 | 1.2099 | 1.2426 | 1.3103 | 1.2954 | 1.2369 | 1.196 | 1.2069 | 1.1557 | 1.1878 | 1.1615 | 1.1936 | 1.2125 | 0.98919 | 1.2376
b8 | -0.081903 | -0.025297 | -0.026808 | -0.070994 | 0.011199 | -0.036528 | 0.020203 | 0.02521 | 0.0044173 | -0.04819 | 0.0063898 | 0.056085 | 0.017043 | 0.02268 | 0.0020412 | -0.0029904 | -0.04321 | 0.013637 | 0.036004 | -0.024867 | -0.0055589 | b8 | -0.073188 | 0.06052 | 0.026498 | 0.031761 | 0.048257 | -0.0031009 | 0.013628 | 0.029999 | -0.058646 | -0.081643 | 0.020303 | -0.020031 | -0.015874 | -0.044416 | -0.013254 | -0.010114 | 0.014723 | -0.00049146 | -0.0077224 | -0.028096 | -0.0038489
b9 | 0.38175 | 0.54148 | 0.52114 | 0.53636 | 0.55939 | 0.51811 | 0.54414 | 0.54903 | 0.55639 | 0.54012 | 0.52814 | 0.51981 | 0.52322 | 0.53101 | 0.52367 | 0.55365 | 0.53341 | 0.56655 | 0.54572 | 0.50166 | 0.53201 | b9 | 0.39345 | 0.5242 | 0.52285 | 0.52045 | 0.55054 | 0.52755 | 0.54848 | 0.52153 | 0.53916 | 0.53671 | 0.50834 | 0.5135 | 0.49753 | 0.50881 | 0.50869 | 0.53962 | 0.51869 | 0.5409 | 0.52989 | 0.49648 | 0.52009
b10 | 0.035127 | 0.0032998 | -0.01806 | -0.14386 | 0.047019 | -0.033959 | 0.039188 | 0.090859 | 0.003581 | -0.10708 | 0.06835 | 0.019917 | -0.0031598 | -0.0045063 | -0.014464 | 0.024449 | -0.080997 | 0.050669 | 0.051291 | -0.023456 | 0 | b10 | 0.039918 | 0.074821 | 0.046879 | 0.010354 | 0.058092 | -0.06041 | 0.012488 | 0.019724 | -0.069097 | -0.095594 | -0.0017308 | -0.008189 | 0.0011202 | -0.031363 | -0.02099 | 0.014868 | 0.028376 | -0.0094501 | 0.035985 | -0.049115 | 0
b11 | 0.63282 | 0.75497 | 0.7538 | 0.76664 | 0.76186 | 0.76385 | 0.76155 | 0.76282 | 0.76471 | 0.76128 | 0.77115 | 0.77393 | 0.76488 | 0.77344 | 0.77571 | 0.76662 | 0.76107 | 0.76566 | 0.7639 | 0.66807 | 0.75731 | b11 | 0.64512 | 0.76299 | 0.76005 | 0.76997 | 0.76262 | 0.77171 | 0.76843 | 0.77152 | 0.76739 | 0.76372 | 0.774 | 0.76017 | 0.76533 | 0.77782 | 0.77864 | 0.76738 | 0.76909 | 0.76816 | 0.75623 | 0.65719 | 0.75981
b12 | -0.0025266 | -0.001217 | -0.010016 | -0.019424 | 0.0081897 | -0.0054829 | 0.0031411 | 0.010851 | 0.00014145 | -0.013156 | 0.0091387 | 0.0043672 | 0.00022294 | -0.006202 | -0.005171 | -0.0030832 | -0.016675 | 0.0022507 | -0.0006101 | -0.0079025 | -0.0025623 | b12 | 0.0062476 | 0.0021393 | 0.0041858 | 0.0029587 | 0.010189 | -0.007878 | 0.0024569 | 0.002033 | -0.0095216 | -0.012921 | 0.003374 | -0.0023829 | 0.00045475 | -0.0038667 | -0.0039305 | -0.00045567 | 0.00025616 | -0.003514 | 0.0057288 | -0.0029093 | -0.0004445
b13 | 0.084584 | 0.078814 | 0.077007 | 0.078209 | 0.076237 | 0.075353 | 0.075547 | 0.072116 | 0.073858 | 0.07554 | 0.074888 | 0.074604 | 0.079018 | 0.073495 | 0.072179 | 0.073262 | 0.076129 | 0.075766 | 0.078953 | 0.058497 | 0.07534 | b13 | 0.088381 | 0.078972 | 0.076932 | 0.07572 | 0.073831 | 0.075475 | 0.076313 | 0.073898 | 0.075822 | 0.077433 | 0.076095 | 0.075378 | 0.079458 | 0.074227 | 0.075862 | 0.075843 | 0.074445 | 0.07537 | 0.078157 | 0.059759 | 0.075935
b14 | 0.0049715 | -0.0062401 | -0.0093283 | -0.010357 | -0.00046398 | -0.007557 | -0.0018093 | 0.0012903 | -0.0045237 | -0.0097373 | 6.9138e-06 | -0.0022718 | -0.0040566 | -0.0068461 | -0.007163 | -0.0039792 | -0.0095978 | -0.0014491 | -0.0011029 | -0.0052012 | -0.0044268 | b14 | 0.0060566 | -0.00090254 | -0.0005599 | -0.002073 | 0.0024981 | -0.0081838 | -0.0026244 | -0.00089506 | -0.006989 | -0.0090936 | -0.0029179 | -0.0046916 | -0.0025075 | -0.0062459 | -0.0061335 | -0.0027054 | -0.0023835 | -0.0044538 | -0.00047795 | -0.0031584 | -0.0030869
b15 | 0.0031714 | 0.039941 | 0.038518 | 0.024919 | 0.037409 | 0.038351 | 0.03572 | 0.037912 | 0.036664 | 0.028377 | 0.037439 | 0.036944 | 0.035057 | 0.041015 | 0.040486 | 0.038497 | 0.027802 | 0.039219 | 0.033763 | 0.027758 | 0.034643 | b15 | 0.012655 | 0.042579 | 0.041466 | 0.037358 | 0.03833 | 0.036974 | 0.03464 | 0.037326 | 0.034013 | 0.031642 | 0.040573 | 0.036821 | 0.038441 | 0.040926 | 0.04053 | 0.037997 | 0.042219 | 0.039656 | 0.038808 | 0.030672 | 0.037246
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 27.9288 | -10.2061 | -9.0894 | -6.2681 | -6.7924 | -2.1012 | -2.8785 | 3.7513 | 1.6458 | -0.41501 | -5.1339 | -0.81788 | 3.1266 | 5.7715 | -0.46823 | -0.41395 | 3.164 | 8.2818 | 5.5782 | -7.5881 | 0 | a1 | 26.7356 | -4.3744 | -2.4251 | -0.42776 | -3.2752 | 0.73181 | 1.6304 | 5.4684 | 2.4503 | -0.70187 | -5.4354 | -3.7744 | -2.1532 | -1.0903 | -1.8939 | -1.7646 | 0.58923 | 4.0207 | -1.0296 | -4.9769 | 0
a2 | 4.4755 | -0.19467 | 0.41637 | 1.4357 | 1.3875 | 0.7926 | 1.1463 | 1.9574 | 1.4165 | 1.2 | 1.5416 | 3.1439 | 1.3995 | 2.1932 | 1.2032 | 0.66067 | 0.99695 | 1.4914 | 1.1873 | 2.4501 | 1.4418 | a2 | -2.4489 | -1.3101 | -1.5489 | -0.36246 | -1.2099 | -0.73264 | -1.3844 | -1.8586 | -1.9658 | -1.5832 | -1.4491 | -0.90651 | -0.95979 | -2.1859 | -1.5864 | -1.7918 | -1.5184 | -1.4594 | -0.9661 | 6.2491 | -1.1593
a3 | -0.02745 | 0.67984 | 0.44267 | 0.62416 | 0.39536 | 0.27605 | 0.47047 | 0.44977 | 0.22095 | 0.379 | 0.66557 | 0.34985 | 0.59761 | 0.51453 | 0.33704 | 0.023918 | 0.35813 | 0.60682 | 0.51476 | 1.3778 | 0.45489 | a3 | 1.1041 | 0.40699 | 0.51563 | 0.10747 | 0.62707 | 0.28164 | 0.11257 | 0.40935 | 0.12651 | 0.39532 | 0.24481 | 0.15283 | 0.040037 | 0.30304 | 0.60771 | 0.31906 | 0.55594 | 0.67468 | 0.12726 | 0.35158 | 0.35982
a4 | 1.5929 | 0.075771 | 0.011303 | -0.32351 | 0.30188 | -0.19879 | 0.046134 | -0.19709 | 0.080937 | 0.43671 | 0.037951 | 0.15029 | -0.018779 | -0.062823 | 0.075325 | 0.33767 | -0.020671 | 0.25288 | 0.26929 | 0.11629 | 0.12162 | a4 | 0.26782 | 0.51298 | 0.27423 | -0.36701 | -0.11421 | -0.58375 | -0.53716 | 0.17063 | -0.0025844 | 0.18997 | 0.17449 | -0.27634 | -0.21955 | 0.52669 | 0.66025 | 0.32012 | 0.40413 | 0.39584 | -0.11362 | 0.57467 | 0.10132
a5 | 2.5179 | 0.0039138 | -0.025107 | 0.1121 | 0.15172 | -0.05571 | 0.066755 | -0.048622 | 0.055937 | 0.14042 | 0.1148 | 0.062049 | -0.089571 | 0.14648 | 0.094673 | 0.045681 | -0.04597 | 0.015107 | 0.10776 | 0.27271 | 0.13651 | a5 | 2.8357 | -0.084502 | 0.0076394 | 0.17651 | 0.2712 | 0.18116 | 0.028031 | 0.045178 | 0.041234 | 0.22729 | 0.078663 | -0.098819 | -0.27437 | 0.22729 | 0.041788 | 0.095234 | 0.033926 | 0.16578 | 0.14661 | -0.053651 | 0.15994
a6 | 0.38583 | 0.024159 | -0.044271 | 0.059028 | -0.068502 | -0.026598 | 0.023227 | 0.0062051 | -0.011032 | 0.12146 | 0.00094596 | 0.050013 | -0.0044564 | -0.061588 | 0.018384 | 0.022781 | 0.15515 | -0.0074992 | 0.0040984 | 0.0031091 | 0.026431 | a6 | -0.086315 | 0.14667 | 0.075519 | 0.048638 | 0.0041118 | 0.0018149 | -0.047487 | 0.0017669 | -0.0086948 | 0.063343 | -0.084697 | -0.12938 | -0.10057 | -0.0059471 | 0.039418 | 0.08383 | 3.6152e-05 | -0.045268 | 0.036004 | -0.025205 | 0.00041706
a7 | 1.5449 | -0.097522 | -0.062731 | -0.042633 | -0.028571 | -0.024616 | -0.04692 | 0.05708 | 0.0038768 | 0.019558 | 0.053564 | -0.025491 | -0.043104 | -0.00213 | 0.013869 | 0.040916 | 0.033099 | 0.036319 | 0.062988 | 0.1358 | 0.052876 | a7 | 1.6312 | -0.04549 | -0.026169 | -0.065198 | 0.028469 | 0.0053781 | 0.0022514 | 0.039412 | 0.058876 | 0.019805 | 0.095325 | -0.028926 | -0.040064 | -0.027443 | -0.058077 | -0.006238 | 0.00044802 | 0.027427 | 0.033422 | 0.10544 | 0.058101
a8 | -0.0043672 | 0.0023834 | -0.019045 | 0.070675 | -0.042683 | -0.061119 | -0.041305 | 0.0023807 | -0.030865 | 0.0074473 | -0.032583 | -0.068953 | -0.0014965 | -0.013375 | -0.050337 | -0.029819 | 0.026862 | -0.050312 | -0.033081 | -0.080005 | -0.02174 | a8 | -0.015445 | -0.0030515 | 0.00056564 | 0.0077679 | -0.018426 | -0.013374 | 0.011505 | -0.053532 | -0.063357 | -0.040298 | -0.055525 | -0.0068386 | -0.011274 | -0.027094 | -0.0030914 | 0.013965 | -0.011614 | 0.017963 | 0.020113 | -0.040699 | -0.01408
a9 | -0.24985 | 0.0044856 | -0.016289 | -0.017177 | -0.01329 | -0.021349 | -0.0091219 | -0.0193 | 0.012785 | -0.019292 | -0.0068962 | -0.0059406 | 0.0033718 | -0.023289 | -0.048062 | -0.036345 | -0.033877 | -0.0076102 | 0.0060118 | -0.049218 | -0.022914 | a9 | -0.1698 | 0.061159 | 0.02687 | 0.018525 | -0.021627 | -0.019982 | -0.017829 | 0.012006 | 0.028423 | 0.017556 | -0.0039466 | -0.013915 | -0.013183 | -0.0029425 | -0.020938 | 0.023197 | 0.024711 | 0.00829 | 0.0050041 | -0.057218 | -0.0017416
a10 | -0.011518 | -0.0047849 | -0.033348 | 0.075969 | -0.018945 | -0.059915 | 0.0002508 | -0.0068917 | -0.0098639 | 0.082262 | -0.00053689 | 0.0081789 | 0.014944 | -0.047713 | -0.036514 | -0.0029816 | 0.11349 | -0.029482 | 0.022694 | -0.093255 | 0 | a10 | -0.0272 | -0.008014 | 0.0029828 | 0.019025 | 0.016759 | -0.01435 | 0.021444 | 0.014383 | 0.0049515 | 0.022768 | -0.021341 | 0.0086194 | -0.0020967 | -0.0045695 | -0.0033844 | 0.022398 | -0.010822 | 0.0095134 | 0.0048739 | -0.10298 | 0
a11 | 0.14872 | -0.055883 | -0.048724 | -0.037827 | -0.038409 | -0.032947 | -0.034554 | -0.021196 | -0.025945 | -0.029077 | -0.028198 | -0.031109 | -0.029704 | -0.02005 | -0.025787 | -0.015342 | -0.021322 | -0.01479 | -0.023469 | -0.024002 | -0.023585 | a11 | 0.1604 | -0.040629 | -0.034686 | -0.028245 | -0.03932 | -0.031969 | -0.028564 | -0.024313 | -0.022342 | -0.022329 | -0.019455 | -0.035299 | -0.038096 | -0.029136 | -0.028068 | -0.016152 | -0.022971 | -0.012943 | -0.022564 | -0.02553 | -0.021347
a12 | 0.0017251 | 0.00064014 | -0.0087245 | 0.008212 | -0.0058833 | -0.008066 | 0.0024801 | 0.0051544 | -0.00094115 | 0.0069156 | -0.0014165 | 0.0060048 | 0.0005019 | -0.0098463 | -0.0022132 | -0.0013708 | 0.016053 | -0.007034 | -0.00094323 | -0.01462 | -0.00045162 | a12 | -0.0092969 | 0.0006197 | 0.0014319 | 0.0049452 | 0.0017722 | -0.0018128 | -0.0019494 | 0.00051606 | -0.0017322 | 0.0021104 | -0.009027 | -0.00671 | -0.004658 | -0.005641 | -0.0038413 | 0.00012345 | -0.006855 | -0.0051868 | -0.00054573 | -0.015972 | -0.0027273
a13 | -0.027823 | -0.0019914 | -0.0027213 | -0.0018084 | -0.0036439 | -0.00035458 | -0.0025454 | -0.0037679 | 5.6252e-05 | 1.7155e-05 | -0.0019991 | -0.0020001 | -0.0021177 | -0.0012807 | -0.0029599 | -0.0030621 | -0.00061624 | -0.00051587 | -0.0013075 | -0.0010341 | -0.002645 | a13 | -0.024201 | 0.00032062 | -0.002701 | -0.0011291 | -0.002416 | -0.00062599 | -0.0037757 | -0.0025272 | -0.00036571 | 0.00077506 | -0.0019743 | -0.002364 | -0.0048964 | -0.0033863 | -0.0022652 | 0.00083856 | 0.0010081 | 0.0011537 | -0.0014343 | -0.0040564 | -0.0022733
a14 | 0.0032214 | -0.018078 | -0.019833 | -0.0095636 | -0.018914 | -0.020207 | -0.017595 | -0.016849 | -0.018495 | -0.012186 | -0.019939 | -0.018031 | -0.01686 | -0.019216 | -0.020031 | -0.018249 | -0.010944 | -0.019445 | -0.014465 | -0.013256 | -0.016358 | a14 | 0.0067196 | -0.019777 | -0.01877 | -0.017534 | -0.017905 | -0.017043 | -0.014644 | -0.017128 | -0.016811 | -0.014944 | -0.020108 | -0.015515 | -0.017706 | -0.018698 | -0.017569 | -0.015432 | -0.020078 | -0.016599 | -0.017174 | -0.011225 | -0.01641
a15 | -0.0043667 | -0.010926 | -0.012197 | -0.014221 | -0.0033211 | -0.012283 | -0.003991 | 0.001164 | -0.0073882 | -0.012385 | -0.0020751 | -0.0031688 | -0.0057606 | -0.0092538 | -0.0093221 | -0.0067214 | -0.010428 | -0.0030122 | -0.0012765 | -0.006418 | -0.0069221 | a15 | 0.0021693 | -0.0010971 | -0.001388 | -0.0039546 | 0.0011635 | -0.011241 | -0.0044295 | -0.00058666 | -0.011428 | -0.013808 | -0.0056158 | -0.0054061 | -0.0051882 | -0.010682 | -0.0093206 | -0.0045052 | -0.0033085 | -0.0055633 | -0.0021296 | -0.0086031 | -0.0053237
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -8.2545e-05 | 0 | 3.5683e-05 | 0.0003268 | -0.00010679 | 6.9628e-05 | -8.7278e-05 | -0.00020289 | 0 | 0.00024576 | -0.00015058 | -4.2956e-05 | 0 | 0 | 2.8992e-05 | -5.4323e-05 | 0.00018728 | -0.00011369 | -0.00011249 | 5.0782e-05 | 0.00013055 | Dx (m) | -8.2221e-05 | -0.00016717 | -0.00010434 | -2.1285e-05 | -0.00012727 | 0.00013157 | -2.583e-05 | -4.242e-05 | 0.00015333 | 0.00021425 | 0 | 1.9169e-05 | 0 | 6.8078e-05 | 4.5502e-05 | -3.1881e-05 | -6.3372e-05 | 2.1265e-05 | -8.0508e-05 | 0.00011613 | 7.7419e-05
Dy (m) | 5.0318e-05 | 1.0169e-05 | 7.7521e-05 | -0.00015309 | 3.6933e-05 | 0.00013635 | 0 | 0 | 2.1633e-05 | -0.00017468 | 0 | -1.9698e-05 | -3.2891e-05 | 0.00010502 | 8.0993e-05 | 0 | -0.00024865 | 6.3287e-05 | -5.3938e-05 | 0.00023891 | 0.00076885 | Dy (m) | 9.2154e-05 | 0 | -1.1442e-05 | -4.2786e-05 | -4.3964e-05 | 3.7092e-05 | -4.8405e-05 | -3.3031e-05 | 0 | -4.4513e-05 | 4.6931e-05 | -1.8386e-05 | 0 | 1.2549e-05 | 0 | -5.0288e-05 | 2.2037e-05 | -2.0697e-05 | -1.336e-05 | 0.00027056 | 0.00080716
Magnetic length (m) |  | 14.4546 | [19 empty] | Magnetic length (m) |  | 14.4553
Transfer funct. (mT/KA) |  | 596.4444 | [19 empty] | Transfer funct. (mT/KA) |  | 596.4118
Coil positioning (mm) |  | 0.015915 | [19 empty] | Coil positioning (mm) |  | 0.018139
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
156 | HCMB__A001 | 3000064 |  | 2 |  |  | 3
| COMMENTS: | Non-nominal shims on outer layers of both apertures.
| HOLDING POINT | OK | Data at CERN |  | 2003-08-26 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-08-26 | Answer to manufacturer |  |  | 2003-08-26
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.85 | 0.8 | 0.05 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.85 | 0.2 | 0.2 | 0.85 | Up | 0.85 | 0.2 | 0.2 | 0.85
[8 empty] | 0.85 | 0.2 | 0.2 | 0.85 | Down | 0.85 | 0.2 | 0.2 | 0.85
Seen from connection side
| Collared coil template - version 11.07.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)

## Summary Data
File | HCMB__A001-3000064 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000064 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 380.2353 | 596.7059 | 596.3529 | 596.4706 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.4706 | 596.3529 | 596.4706 | 596.3529 | 596.4706 | 596.4706 | 596.4706 | 596.9412 | 380.5882 | 0.073282 | C1/i (mT/KA) | 380.1176 | 596.7059 | 596.3529 | 596.3529 | 596.2353 | 596.3529 | 596.4706 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.3529 | 596.4706 | 596.4706 | 596.8235 | 381.0588 | 0.073283
Angle (mrad) | 2.7929 | -1.0206 | -0.90894 | -0.62681 | -0.67924 | -0.21012 | -0.28785 | 0.37513 | 0.16458 | -0.041501 | -0.51338 | -0.081788 | 0.31266 | 0.57715 | -0.046823 | -0.041395 | 0.3164 | 0.82818 | 0.55782 | -0.75882 | -3.7057 | Angle (mrad) | 2.6736 | -0.43744 | -0.24251 | -0.042776 | -0.32752 | 0.073181 | 0.16304 | 0.54684 | 0.24503 | -0.070187 | -0.54354 | -0.37744 | -0.21532 | -0.10903 | -0.18939 | -0.17646 | 0.058923 | 0.40207 | -0.10296 | -0.49769 | -4.015
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -1.9332 | -0.36022 | -0.72574 | -0.10108 | -0.086225 | 0.10892 | 0.04892 | 0.073 | -0.60918 | -0.32768 | -0.68721 | 0.10025 | -0.071561 | -0.10924 | -0.38749 | -0.58245 | -0.70487 | -0.91352 | -1.6873 | -1.0481 | -0.46303 | b2 | -0.92842 | 0.075593 | 0.42475 | 0.73749 | 0.82592 | 1.1213 | 0.6939 | 0.24554 | 0.88719 | -0.094447 | 0.73359 | 0.22232 | 0.54289 | 0.74521 | 0.99691 | 0.28677 | 0.64585 | 0.44535 | 0.57599 | 2.829 | 0.58764
b3 | 32.5389 | -0.66577 | -0.14794 | -0.63183 | 0.60405 | -1.2084 | -0.088205 | -1.0961 | -0.35196 | 0.1203 | -1.3607 | -1.3242 | -0.43744 | -2.5152 | -2.252 | -0.47254 | -0.076472 | -0.12003 | 0.48596 | -0.63819 | 0.45636 | b3 | 35.2322 | -0.54249 | -0.29373 | -0.77743 | 0.23889 | -0.74759 | 0.37451 | -0.68327 | -0.16543 | 0.22342 | -1.154 | -0.43044 | -0.51875 | -2.4964 | -1.1282 | 0.29103 | -0.076333 | 0.079646 | 0.76522 | -0.78087 | 0.77389
b4 | 0.48348 | 0.10909 | 0.077857 | -0.033867 | 0.075962 | 0.25122 | 0.084665 | 0.11143 | -0.041596 | -0.090746 | -0.05879 | 0.06107 | -0.15262 | 0.075815 | 0.018742 | 0.1074 | 0.1365 | 0.064055 | 0.19402 | -0.13539 | 0.062886 | b4 | -0.021038 | 0.30391 | 0.12917 | -0.10804 | -0.017883 | 0.1954 | 0.17603 | 0.08677 | 0.0031098 | -0.075531 | -0.0098859 | -0.15129 | -0.036275 | 0.071151 | 0.07156 | 0.028449 | 0.14512 | 0.076437 | -0.00067856 | 0.097278 | 0.048579
b5 | -5.0151 | -0.4459 | -0.29238 | -0.085486 | -0.34908 | -0.072476 | -0.085434 | 0.03938 | -0.055008 | 0.020385 | 0.27464 | -0.031843 | -0.29161 | -0.26351 | -0.029389 | -0.23465 | -0.20656 | -0.22675 | -0.33969 | -3.3796 | -0.41653 | b5 | -4.4657 | -0.59901 | -0.45953 | -0.13202 | -0.34156 | -0.089084 | -0.14239 | -0.097567 | -0.17672 | -0.10407 | 0.27965 | -0.24148 | -0.19499 | -0.12691 | -0.053946 | -0.29132 | -0.14126 | -0.17774 | -0.353 | -3.5476 | -0.44388
b6 | 0.055837 | 0.05431 | -0.0040801 | -0.080276 | 0.060316 | -0.013224 | 0.031065 | 0.10917 | -0.0017672 | -0.14183 | 0.10342 | 0.051798 | -0.0081148 | 0.051013 | 0.030572 | 0.1176 | -0.03431 | 0.19626 | 0.20531 | -0.047539 | 0.03801 | b6 | -0.074615 | 0.163 | 0.076967 | -0.014218 | 0.082706 | -0.11364 | -0.040339 | 0.096362 | 0.019023 | 0.037261 | 0.0073831 | 0.059513 | 0.072477 | -0.046289 | 0.041919 | 0.097379 | 0.0090744 | 0.0071095 | -0.030975 | -0.086986 | 0.021869
b7 | 2.5683 | 1.0444 | 1.1643 | 1.2782 | 1.213 | 1.3321 | 1.2911 | 1.3247 | 1.3301 | 1.2368 | 1.2855 | 1.1523 | 1.2351 | 1.1878 | 1.1668 | 1.2088 | 1.1981 | 1.2905 | 1.2254 | 0.98116 | 1.2673 | b7 | 2.6187 | 0.86269 | 1.0882 | 1.2539 | 1.2186 | 1.2813 | 1.2423 | 1.2099 | 1.2426 | 1.3103 | 1.2954 | 1.2369 | 1.196 | 1.2069 | 1.1557 | 1.1878 | 1.1615 | 1.1936 | 1.2125 | 0.98919 | 1.2376
b8 | -0.081903 | -0.025297 | -0.026808 | -0.070994 | 0.011199 | -0.036528 | 0.020203 | 0.02521 | 0.0044173 | -0.04819 | 0.0063898 | 0.056085 | 0.017043 | 0.02268 | 0.0020412 | -0.0029904 | -0.04321 | 0.013637 | 0.036004 | -0.024867 | -0.0055589 | b8 | -0.073188 | 0.06052 | 0.026498 | 0.031761 | 0.048257 | -0.0031009 | 0.013628 | 0.029999 | -0.058646 | -0.081643 | 0.020303 | -0.020031 | -0.015874 | -0.044416 | -0.013254 | -0.010114 | 0.014723 | -0.00049146 | -0.0077224 | -0.028096 | -0.0038489
b9 | 0.38175 | 0.54148 | 0.52114 | 0.53636 | 0.55939 | 0.51811 | 0.54414 | 0.54903 | 0.55639 | 0.54012 | 0.52814 | 0.51981 | 0.52322 | 0.53101 | 0.52367 | 0.55365 | 0.53341 | 0.56655 | 0.54572 | 0.50166 | 0.53201 | b9 | 0.39345 | 0.5242 | 0.52285 | 0.52045 | 0.55054 | 0.52755 | 0.54848 | 0.52153 | 0.53916 | 0.53671 | 0.50834 | 0.5135 | 0.49753 | 0.50881 | 0.50869 | 0.53962 | 0.51869 | 0.5409 | 0.52989 | 0.49648 | 0.52009
b10 | 0.035127 | 0.0032998 | -0.01806 | -0.14386 | 0.047019 | -0.033959 | 0.039188 | 0.090859 | 0.003581 | -0.10708 | 0.06835 | 0.019917 | -0.0031598 | -0.0045063 | -0.014464 | 0.024449 | -0.080997 | 0.050669 | 0.051291 | -0.023456 | 0 | b10 | 0.039918 | 0.074821 | 0.046879 | 0.010354 | 0.058092 | -0.06041 | 0.012488 | 0.019724 | -0.069097 | -0.095594 | -0.0017308 | -0.008189 | 0.0011202 | -0.031363 | -0.02099 | 0.014868 | 0.028376 | -0.0094501 | 0.035985 | -0.049115 | 0
b11 | 0.63282 | 0.75497 | 0.7538 | 0.76664 | 0.76186 | 0.76385 | 0.76155 | 0.76282 | 0.76471 | 0.76128 | 0.77115 | 0.77393 | 0.76488 | 0.77344 | 0.77571 | 0.76662 | 0.76107 | 0.76566 | 0.7639 | 0.66807 | 0.75731 | b11 | 0.64512 | 0.76299 | 0.76005 | 0.76997 | 0.76262 | 0.77171 | 0.76843 | 0.77152 | 0.76739 | 0.76372 | 0.774 | 0.76017 | 0.76533 | 0.77782 | 0.77864 | 0.76738 | 0.76909 | 0.76816 | 0.75623 | 0.65719 | 0.75981
b12 | -0.0025266 | -0.001217 | -0.010016 | -0.019424 | 0.0081897 | -0.0054829 | 0.0031411 | 0.010851 | 0.00014145 | -0.013156 | 0.0091387 | 0.0043672 | 0.00022294 | -0.006202 | -0.005171 | -0.0030832 | -0.016675 | 0.0022507 | -0.0006101 | -0.0079025 | -0.0025623 | b12 | 0.0062476 | 0.0021393 | 0.0041858 | 0.0029587 | 0.010189 | -0.007878 | 0.0024569 | 0.002033 | -0.0095216 | -0.012921 | 0.003374 | -0.0023829 | 0.00045475 | -0.0038667 | -0.0039305 | -0.00045567 | 0.00025616 | -0.003514 | 0.0057288 | -0.0029093 | -0.0004445
b13 | 0.084584 | 0.078814 | 0.077007 | 0.078209 | 0.076237 | 0.075353 | 0.075547 | 0.072116 | 0.073858 | 0.07554 | 0.074888 | 0.074604 | 0.079018 | 0.073495 | 0.072179 | 0.073262 | 0.076129 | 0.075766 | 0.078953 | 0.058497 | 0.07534 | b13 | 0.088381 | 0.078972 | 0.076932 | 0.07572 | 0.073831 | 0.075475 | 0.076313 | 0.073898 | 0.075822 | 0.077433 | 0.076095 | 0.075378 | 0.079458 | 0.074227 | 0.075862 | 0.075843 | 0.074445 | 0.07537 | 0.078157 | 0.059759 | 0.075935
b14 | 0.0049715 | -0.0062401 | -0.0093283 | -0.010357 | -0.00046398 | -0.007557 | -0.0018093 | 0.0012903 | -0.0045237 | -0.0097373 | 6.9138e-06 | -0.0022718 | -0.0040566 | -0.0068461 | -0.007163 | -0.0039792 | -0.0095978 | -0.0014491 | -0.0011029 | -0.0052012 | -0.0044268 | b14 | 0.0060566 | -0.00090254 | -0.0005599 | -0.002073 | 0.0024981 | -0.0081838 | -0.0026244 | -0.00089506 | -0.006989 | -0.0090936 | -0.0029179 | -0.0046916 | -0.0025075 | -0.0062459 | -0.0061335 | -0.0027054 | -0.0023835 | -0.0044538 | -0.00047795 | -0.0031584 | -0.0030869
b15 | 0.0031714 | 0.039941 | 0.038518 | 0.024919 | 0.037409 | 0.038351 | 0.03572 | 0.037912 | 0.036664 | 0.028377 | 0.037439 | 0.036944 | 0.035057 | 0.041015 | 0.040486 | 0.038497 | 0.027802 | 0.039219 | 0.033763 | 0.027758 | 0.034643 | b15 | 0.012655 | 0.042579 | 0.041466 | 0.037358 | 0.03833 | 0.036974 | 0.03464 | 0.037326 | 0.034013 | 0.031642 | 0.040573 | 0.036821 | 0.038441 | 0.040926 | 0.04053 | 0.037997 | 0.042219 | 0.039656 | 0.038808 | 0.030672 | 0.037246
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 27.9288 | -10.2061 | -9.0894 | -6.2681 | -6.7924 | -2.1012 | -2.8785 | 3.7513 | 1.6458 | -0.41501 | -5.1339 | -0.81788 | 3.1266 | 5.7715 | -0.46823 | -0.41395 | 3.164 | 8.2818 | 5.5782 | -7.5881 | 0 | a1 | 26.7356 | -4.3744 | -2.4251 | -0.42776 | -3.2752 | 0.73181 | 1.6304 | 5.4684 | 2.4503 | -0.70187 | -5.4354 | -3.7744 | -2.1532 | -1.0903 | -1.8939 | -1.7646 | 0.58923 | 4.0207 | -1.0296 | -4.9769 | 0
a2 | 4.4755 | -0.19467 | 0.41637 | 1.4357 | 1.3875 | 0.7926 | 1.1463 | 1.9574 | 1.4165 | 1.2 | 1.5416 | 3.1439 | 1.3995 | 2.1932 | 1.2032 | 0.66067 | 0.99695 | 1.4914 | 1.1873 | 2.4501 | 1.4418 | a2 | -2.4489 | -1.3101 | -1.5489 | -0.36246 | -1.2099 | -0.73264 | -1.3844 | -1.8586 | -1.9658 | -1.5832 | -1.4491 | -0.90651 | -0.95979 | -2.1859 | -1.5864 | -1.7918 | -1.5184 | -1.4594 | -0.9661 | 6.2491 | -1.1593
a3 | -0.02745 | 0.67984 | 0.44267 | 0.62416 | 0.39536 | 0.27605 | 0.47047 | 0.44977 | 0.22095 | 0.379 | 0.66557 | 0.34985 | 0.59761 | 0.51453 | 0.33704 | 0.023918 | 0.35813 | 0.60682 | 0.51476 | 1.3778 | 0.45489 | a3 | 1.1041 | 0.40699 | 0.51563 | 0.10747 | 0.62707 | 0.28164 | 0.11257 | 0.40935 | 0.12651 | 0.39532 | 0.24481 | 0.15283 | 0.040037 | 0.30304 | 0.60771 | 0.31906 | 0.55594 | 0.67468 | 0.12726 | 0.35158 | 0.35982
a4 | 1.5929 | 0.075771 | 0.011303 | -0.32351 | 0.30188 | -0.19879 | 0.046134 | -0.19709 | 0.080937 | 0.43671 | 0.037951 | 0.15029 | -0.018779 | -0.062823 | 0.075325 | 0.33767 | -0.020671 | 0.25288 | 0.26929 | 0.11629 | 0.12162 | a4 | 0.26782 | 0.51298 | 0.27423 | -0.36701 | -0.11421 | -0.58375 | -0.53716 | 0.17063 | -0.0025844 | 0.18997 | 0.17449 | -0.27634 | -0.21955 | 0.52669 | 0.66025 | 0.32012 | 0.40413 | 0.39584 | -0.11362 | 0.57467 | 0.10132
a5 | 2.5179 | 0.0039138 | -0.025107 | 0.1121 | 0.15172 | -0.05571 | 0.066755 | -0.048622 | 0.055937 | 0.14042 | 0.1148 | 0.062049 | -0.089571 | 0.14648 | 0.094673 | 0.045681 | -0.04597 | 0.015107 | 0.10776 | 0.27271 | 0.13651 | a5 | 2.8357 | -0.084502 | 0.0076394 | 0.17651 | 0.2712 | 0.18116 | 0.028031 | 0.045178 | 0.041234 | 0.22729 | 0.078663 | -0.098819 | -0.27437 | 0.22729 | 0.041788 | 0.095234 | 0.033926 | 0.16578 | 0.14661 | -0.053651 | 0.15994
a6 | 0.38583 | 0.024159 | -0.044271 | 0.059028 | -0.068502 | -0.026598 | 0.023227 | 0.0062051 | -0.011032 | 0.12146 | 0.00094596 | 0.050013 | -0.0044564 | -0.061588 | 0.018384 | 0.022781 | 0.15515 | -0.0074992 | 0.0040984 | 0.0031091 | 0.026431 | a6 | -0.086315 | 0.14667 | 0.075519 | 0.048638 | 0.0041118 | 0.0018149 | -0.047487 | 0.0017669 | -0.0086948 | 0.063343 | -0.084697 | -0.12938 | -0.10057 | -0.0059471 | 0.039418 | 0.08383 | 3.6152e-05 | -0.045268 | 0.036004 | -0.025205 | 0.00041706
a7 | 1.5449 | -0.097522 | -0.062731 | -0.042633 | -0.028571 | -0.024616 | -0.04692 | 0.05708 | 0.0038768 | 0.019558 | 0.053564 | -0.025491 | -0.043104 | -0.00213 | 0.013869 | 0.040916 | 0.033099 | 0.036319 | 0.062988 | 0.1358 | 0.052876 | a7 | 1.6312 | -0.04549 | -0.026169 | -0.065198 | 0.028469 | 0.0053781 | 0.0022514 | 0.039412 | 0.058876 | 0.019805 | 0.095325 | -0.028926 | -0.040064 | -0.027443 | -0.058077 | -0.006238 | 0.00044802 | 0.027427 | 0.033422 | 0.10544 | 0.058101
a8 | -0.0043672 | 0.0023834 | -0.019045 | 0.070675 | -0.042683 | -0.061119 | -0.041305 | 0.0023807 | -0.030865 | 0.0074473 | -0.032583 | -0.068953 | -0.0014965 | -0.013375 | -0.050337 | -0.029819 | 0.026862 | -0.050312 | -0.033081 | -0.080005 | -0.02174 | a8 | -0.015445 | -0.0030515 | 0.00056564 | 0.0077679 | -0.018426 | -0.013374 | 0.011505 | -0.053532 | -0.063357 | -0.040298 | -0.055525 | -0.0068386 | -0.011274 | -0.027094 | -0.0030914 | 0.013965 | -0.011614 | 0.017963 | 0.020113 | -0.040699 | -0.01408
a9 | -0.24985 | 0.0044856 | -0.016289 | -0.017177 | -0.01329 | -0.021349 | -0.0091219 | -0.0193 | 0.012785 | -0.019292 | -0.0068962 | -0.0059406 | 0.0033718 | -0.023289 | -0.048062 | -0.036345 | -0.033877 | -0.0076102 | 0.0060118 | -0.049218 | -0.022914 | a9 | -0.1698 | 0.061159 | 0.02687 | 0.018525 | -0.021627 | -0.019982 | -0.017829 | 0.012006 | 0.028423 | 0.017556 | -0.0039466 | -0.013915 | -0.013183 | -0.0029425 | -0.020938 | 0.023197 | 0.024711 | 0.00829 | 0.0050041 | -0.057218 | -0.0017416
a10 | -0.011518 | -0.0047849 | -0.033348 | 0.075969 | -0.018945 | -0.059915 | 0.0002508 | -0.0068917 | -0.0098639 | 0.082262 | -0.00053689 | 0.0081789 | 0.014944 | -0.047713 | -0.036514 | -0.0029816 | 0.11349 | -0.029482 | 0.022694 | -0.093255 | 0 | a10 | -0.0272 | -0.008014 | 0.0029828 | 0.019025 | 0.016759 | -0.01435 | 0.021444 | 0.014383 | 0.0049515 | 0.022768 | -0.021341 | 0.0086194 | -0.0020967 | -0.0045695 | -0.0033844 | 0.022398 | -0.010822 | 0.0095134 | 0.0048739 | -0.10298 | 0
a11 | 0.14872 | -0.055883 | -0.048724 | -0.037827 | -0.038409 | -0.032947 | -0.034554 | -0.021196 | -0.025945 | -0.029077 | -0.028198 | -0.031109 | -0.029704 | -0.02005 | -0.025787 | -0.015342 | -0.021322 | -0.01479 | -0.023469 | -0.024002 | -0.023585 | a11 | 0.1604 | -0.040629 | -0.034686 | -0.028245 | -0.03932 | -0.031969 | -0.028564 | -0.024313 | -0.022342 | -0.022329 | -0.019455 | -0.035299 | -0.038096 | -0.029136 | -0.028068 | -0.016152 | -0.022971 | -0.012943 | -0.022564 | -0.02553 | -0.021347
a12 | 0.0017251 | 0.00064014 | -0.0087245 | 0.008212 | -0.0058833 | -0.008066 | 0.0024801 | 0.0051544 | -0.00094115 | 0.0069156 | -0.0014165 | 0.0060048 | 0.0005019 | -0.0098463 | -0.0022132 | -0.0013708 | 0.016053 | -0.007034 | -0.00094323 | -0.01462 | -0.00045162 | a12 | -0.0092969 | 0.0006197 | 0.0014319 | 0.0049452 | 0.0017722 | -0.0018128 | -0.0019494 | 0.00051606 | -0.0017322 | 0.0021104 | -0.009027 | -0.00671 | -0.004658 | -0.005641 | -0.0038413 | 0.00012345 | -0.006855 | -0.0051868 | -0.00054573 | -0.015972 | -0.0027273
a13 | -0.027823 | -0.0019914 | -0.0027213 | -0.0018084 | -0.0036439 | -0.00035458 | -0.0025454 | -0.0037679 | 5.6252e-05 | 1.7155e-05 | -0.0019991 | -0.0020001 | -0.0021177 | -0.0012807 | -0.0029599 | -0.0030621 | -0.00061624 | -0.00051587 | -0.0013075 | -0.0010341 | -0.002645 | a13 | -0.024201 | 0.00032062 | -0.002701 | -0.0011291 | -0.002416 | -0.00062599 | -0.0037757 | -0.0025272 | -0.00036571 | 0.00077506 | -0.0019743 | -0.002364 | -0.0048964 | -0.0033863 | -0.0022652 | 0.00083856 | 0.0010081 | 0.0011537 | -0.0014343 | -0.0040564 | -0.0022733
a14 | 0.0032214 | -0.018078 | -0.019833 | -0.0095636 | -0.018914 | -0.020207 | -0.017595 | -0.016849 | -0.018495 | -0.012186 | -0.019939 | -0.018031 | -0.01686 | -0.019216 | -0.020031 | -0.018249 | -0.010944 | -0.019445 | -0.014465 | -0.013256 | -0.016358 | a14 | 0.0067196 | -0.019777 | -0.01877 | -0.017534 | -0.017905 | -0.017043 | -0.014644 | -0.017128 | -0.016811 | -0.014944 | -0.020108 | -0.015515 | -0.017706 | -0.018698 | -0.017569 | -0.015432 | -0.020078 | -0.016599 | -0.017174 | -0.011225 | -0.01641
a15 | -0.0043667 | -0.010926 | -0.012197 | -0.014221 | -0.0033211 | -0.012283 | -0.003991 | 0.001164 | -0.0073882 | -0.012385 | -0.0020751 | -0.0031688 | -0.0057606 | -0.0092538 | -0.0093221 | -0.0067214 | -0.010428 | -0.0030122 | -0.0012765 | -0.006418 | -0.0069221 | a15 | 0.0021693 | -0.0010971 | -0.001388 | -0.0039546 | 0.0011635 | -0.011241 | -0.0044295 | -0.00058666 | -0.011428 | -0.013808 | -0.0056158 | -0.0054061 | -0.0051882 | -0.010682 | -0.0093206 | -0.0045052 | -0.0033085 | -0.0055633 | -0.0021296 | -0.0086031 | -0.0053237
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.082545 | 0 | -0.035683 | -0.3268 | 0.10679 | -0.069628 | 0.087278 | 0.20289 | 0 | -0.24576 | 0.15058 | 0.042956 | 0 | 0 | -0.028992 | 0.054323 | -0.18728 | 0.11369 | 0.11249 | -0.050782 |  | Dx (mm) | 0.082221 | 0.16717 | 0.10434 | 0.021285 | 0.12727 | -0.13157 | 0.02583 | 0.04242 | -0.15333 | -0.21425 | 0 | -0.019169 | 0 | -0.068078 | -0.045502 | 0.031881 | 0.063372 | -0.021265 | 0.080508 | -0.11613
Dy (mm) | -0.050318 | -0.010169 | -0.077521 | 0.15309 | -0.036933 | -0.13635 | 0 | 0 | -0.021633 | 0.17468 | 0 | 0.019698 | 0.032891 | -0.10502 | -0.080993 | 0 | 0.24865 | -0.063287 | 0.053938 | -0.23891 |  | Dy (mm) | -0.092154 | 0 | 0.011442 | 0.042786 | 0.043964 | -0.037092 | 0.048405 | 0.033031 | 0 | 0.044513 | -0.046931 | 0.018386 | 0 | -0.012549 | 0 | 0.050288 | -0.022037 | 0.020697 | 0.01336 | -0.27056
Mag. len. (m) | 14.4226 | [20 empty] | Mag. len. (m) | 14.4233
TF (mT/KA) | 596.4444 | [20 empty] | TF (mT/KA) | 596.4118
Coil pos. (mm) | 0.015915 | [20 empty] | Coil pos. (mm) | 0.018139
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5
Twist (m2 rad) | -0.014324 | [20 empty] | Twist (m2 rad) | 0.0047246

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -0.46303 | 0.56991 | -0.38915 | 0.46013 | 1.4418 | 1.0022 | 1.2952 | 0.70961 | 0.58764 | 0.69275 | 0.56066 | 0.31935 | -1.1593 | 1.7881 | -1.3778 | 0.46005 |  | 0
3 | 0.45636 | 7.4628 | -0.6412 | 0.85083 | 0.45489 | 0.2878 | 0.43864 | 0.16786 | 0.77389 | 8.002 | -0.39192 | 0.74745 | 0.35982 | 0.25688 | 0.33306 | 0.20159 |  | 0
4 | 0.062886 | 0.14301 | 0.055123 | 0.1014 | 0.12162 | 0.38906 | 0.069771 | 0.19854 | 0.048579 | 0.10981 | 0.051857 | 0.11747 | 0.10132 | 0.37215 | 0.08027 | 0.37403 |  | 0
5 | -0.41653 | 1.285 | -0.15 | 0.17559 | 0.13651 | 0.55683 | 0.049494 | 0.077953 | -0.44388 | 1.1966 | -0.19201 | 0.18612 | 0.15994 | 0.63295 | 0.072038 | 0.13556 |  | 0
6 | 0.03801 | 0.084981 | 0.040862 | 0.050267 | 0.026431 | 0.09892 | 0.010345 | 0.030019 | 0.021869 | 0.070703 | 0.030874 | 0.060275 | 0.00041706 | 0.068813 | -0.00029355 | 0.067832 |  | 0
7 | 1.2673 | 0.31502 | 1.2308 | 0.074455 | 0.052876 | 0.34858 | -0.0012392 | 0.046294 | 1.2376 | 0.33683 | 1.1964 | 0.09919 | 0.058101 | 0.36633 | 0.00030446 | 0.042356 |  | 0
8 | -0.0055589 | 0.036172 | -0.0022125 | 0.02048 | -0.02174 | 0.035334 | -0.022729 | 0.025008 | -0.0038489 | 0.038407 | -0.00064147 | 0.019672 | -0.01408 | 0.025049 | -0.016302 | 0.023654 |  | 0
9 | 0.53201 | 0.038193 | 0.53649 | 0.014554 | -0.022914 | 0.054975 | -0.01045 | 0.016062 | 0.52009 | 0.033057 | 0.52346 | 0.014277 | -0.0017416 | 0.046317 | 0.0061009 | 0.023173 |  | 0
10 | 0 | 0.05854 | 1.3025e-05 | 0.00177 | 0 | 0.048502 | -0.0002053 | 0.0013872 | 0 | 0.044305 | -1.3978e-05 | 0.00040302 | 0 | 0.027781 | -0.00040755 | 0.0015957 |  | 0
11 | 0.75731 | 0.036118 | 0.76458 | 0.0059739 | -0.023585 | 0.04112 | -0.029184 | 0.010966 | 0.75981 | 0.036318 | 0.76723 | 0.0061783 | -0.021347 | 0.042688 | -0.027662 | 0.0080561 |  | 0
12 | -0.0025623 | 0.0081606 | -0.0022425 | 0.0033092 | -0.00045162 | 0.0071926 | -0.00047286 | 0.0031177 | -0.0004445 | 0.0056308 | -0.00064705 | 0.0024731 | -0.0027273 | 0.0049469 | -0.0026263 | 0.0030189 |  | 0
13 | 0.07534 | 0.0048728 | 0.07537 | 0.0022771 | -0.002645 | 0.0059364 | -0.0015019 | 0.0013596 | 0.075935 | 0.0049262 | 0.075739 | 0.0017323 | -0.0022733 | 0.0053688 | -0.0013583 | 0.0021113 |  | 0
14 | -0.0044268 | 0.00416 | -0.0051515 | 0.0011428 | -0.016358 | 0.0055223 | -0.017525 | 0.00099157 | -0.0030869 | 0.003556 | -0.0035969 | 0.0011642 | -0.01641 | 0.0057175 | -0.01805 | 0.00107 |  | 0
15 | 0.034643 | 0.0086262 | 0.036002 | 0.0045529 | -0.0069221 | 0.0043481 | -0.0070315 | 0.0045521 | 0.037246 | 0.0065224 | 0.03835 | 0.0029183 | -0.0053237 | 0.0044211 | -0.0054716 | 0.0042302 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.55964 | -0.95894
3 | [12 empty] | 3 | 9 | -0.97307 | -0.9463
4 | [12 empty] | 4 | 16 | -1.9527 | -1.5625
5 | [12 empty] | 5 | 25 | -2.1456 | -1.8399
6 | [12 empty] | 6 | 36 | -3.2482 | -2.7498
7 | [12 empty] | 7 | 49 | -2.8351 | -2.7362
8 | [12 empty] | 8 | 64 | -3.7884 | -3.8364
9 | [12 empty] | 9 | 81 | -4.1806 | -4.0069
10 | [12 empty] | sigma (mm) |  | 0.015915 | 0.018139
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -2.2491 | -0.36103 | -0.7304 | -0.11414 | -0.095532 | 0.094594 | 0.049825 | 0.099164 | -0.60974 | -0.31641 | -0.6631 | 0.10776 | -0.069249 | -0.1156 | -0.39838 | -0.57943 | -0.69608 | -0.91643 | -1.6905 | -1.0906 | -0.46303
3 | 32.5177 | -0.6659 | -0.1476 | -0.64253 | 0.60065 | -1.2005 | -0.089509 | -1.1001 | -0.35227 | 0.12982 | -1.3591 | -1.3241 | -0.43755 | -2.514 | -2.2529 | -0.47357 | -0.072868 | -0.12414 | 0.48467 | -0.6443 | 0.45636
4 | 0.55107 | 0.10908 | 0.07586 | -0.036402 | 0.083415 | 0.25182 | 0.086419 | 0.10955 | -0.041881 | -0.083795 | -0.06852 | 0.061679 | -0.15331 | 0.072195 | 0.016738 | 0.1104 | 0.12471 | 0.069895 | 0.20438 | -0.1911 | 0.062886
5 | -5.0222 | -0.44597 | -0.29141 | -0.090544 | -0.35023 | -0.07168 | -0.086232 | 0.032865 | -0.054938 | 0.016373 | 0.27006 | -0.032208 | -0.29165 | -0.26161 | -0.029566 | -0.23653 | -0.1971 | -0.23317 | -0.34642 | -3.3806 | -0.41653
6 | -0.046422 | 0.05466 | 0.012299 | 0.06485 | 0.014969 | 0.020696 | -0.0087062 | 0.01431 | -0.0017968 | -0.03335 | 0.035095 | 0.034151 | -0.0086151 | 0.051092 | 0.042115 | 0.094427 | 0.047788 | 0.14367 | 0.15786 | -0.041405 | 0.03801
7 | 2.5712 | 1.0444 | 1.1645 | 1.2731 | 1.2132 | 1.3345 | 1.2904 | 1.3226 | 1.3303 | 1.2325 | 1.2851 | 1.1507 | 1.2351 | 1.1884 | 1.1685 | 1.2088 | 1.1975 | 1.2912 | 1.223 | 0.98851 | 1.2673
8 | -0.090815 | -0.025318 | -0.017463 | 0.010255 | -0.016682 | -0.018181 | -0.0021455 | -0.027211 | 0.0042872 | 0.012689 | -0.031035 | 0.045522 | 0.017095 | 0.023831 | 0.011018 | -0.017144 | -0.00016353 | -0.016448 | 0.0072692 | -0.0073452 | -0.0055589
9 | 0.38052 | 0.54151 | 0.52217 | 0.51762 | 0.55711 | 0.52119 | 0.54233 | 0.53927 | 0.5565 | 0.5338 | 0.52269 | 0.51944 | 0.52348 | 0.53366 | 0.52501 | 0.55295 | 0.54032 | 0.56449 | 0.54331 | 0.51283 | 0.53201
10 | -2.0346e-06 | 0.0036341 | -1.6134e-05 | 0.00011403 | -4.322e-06 | -3.1488e-05 | 9.0122e-05 | -0.00018223 | 0.0039111 | -1.2277e-05 | 4.3214e-05 | 8.3255e-07 | -0.0037345 | -0.0032676 | -5.7967e-06 | -4.749e-05 | -0.00027344 | 1.5759e-05 | 6.054e-07 | -0.00012637 | 0
11 | 0.6329 | 0.75496 | 0.754 | 0.76334 | 0.76144 | 0.76432 | 0.76137 | 0.76139 | 0.76473 | 0.75997 | 0.77026 | 0.77388 | 0.76489 | 0.77411 | 0.77573 | 0.76672 | 0.76163 | 0.76579 | 0.76391 | 0.67007 | 0.75731
12 | -0.064668 | -0.012027 | -0.079277 | -0.015779 | 0.025379 | -0.017452 | -0.015132 | 0.0052264 | 0.0014059 | -0.00049748 | 0.011786 | 0.020772 | 0.0017377 | -0.061071 | -0.035247 | -0.058925 | -0.067195 | -0.038065 | -0.069289 | -0.056312 | -0.025623 | ten times!
13 | 0.84146 | 0.78955 | 0.77928 | 0.74501 | 0.76809 | 0.77058 | 0.75668 | 0.71916 | 0.74164 | 0.72083 | 0.74887 | 0.74407 | 0.78594 | 0.75038 | 0.73261 | 0.73428 | 0.72673 | 0.76833 | 0.78451 | 0.60717 | 0.7534 | ten times!
14 | 0.049368 | -0.061486 | -0.074178 | -0.054438 | -0.036528 | -0.039787 | -0.043768 | -0.050444 | -0.04392 | -0.057758 | -0.046359 | -0.036301 | -0.042126 | -0.060457 | -0.055746 | -0.057014 | -0.074453 | -0.049641 | -0.042873 | -0.027776 | -0.044268 | ten times!
15 | 0.031714 | 0.39941 | 0.38518 | 0.24919 | 0.37409 | 0.38351 | 0.3572 | 0.37912 | 0.36664 | 0.28377 | 0.37439 | 0.36944 | 0.35057 | 0.41015 | 0.40486 | 0.38497 | 0.27802 | 0.39219 | 0.33763 | 0.27758 | 0.34643 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 4.6684 | -0.19546 | 0.41688 | 1.4711 | 1.3852 | 0.77547 | 1.1415 | 1.9466 | 1.4156 | 1.2085 | 1.5298 | 3.1452 | 1.4012 | 2.1621 | 1.1829 | 0.66052 | 1.0071 | 1.4824 | 1.1774 | 2.4404 | 1.4418
3 | -0.046359 | 0.68004 | 0.44381 | 0.60642 | 0.39017 | 0.27965 | 0.46976 | 0.45682 | 0.22079 | 0.40074 | 0.66456 | 0.3485 | 0.59849 | 0.51594 | 0.3377 | 0.020681 | 0.35145 | 0.60246 | 0.50757 | 1.3731 | 0.45489
4 | 1.4846 | 0.074704 | 0.0057593 | -0.31182 | 0.29504 | -0.20203 | 0.044763 | -0.19477 | 0.080657 | 0.444 | 0.033883 | 0.14981 | -0.016522 | -0.069335 | 0.075411 | 0.33709 | -0.010611 | 0.2491 | 0.27075 | -0.070438 | 0.12162
5 | 2.5093 | 0.0040762 | -0.025664 | 0.12139 | 0.15453 | -0.056785 | 0.066158 | -0.048992 | 0.055926 | 0.15649 | 0.11476 | 0.061117 | -0.089492 | 0.14805 | 0.095558 | 0.045317 | -0.034915 | 0.019011 | 0.10437 | 0.26941 | 0.13651
6 | 0.38643 | 0.027907 | -0.013207 | -0.014952 | -0.051613 | 0.036902 | 0.024672 | 0.0021176 | -0.00087709 | 0.046908 | -0.0019008 | 0.042389 | -0.018794 | -0.017559 | 0.051881 | 0.021997 | 0.052195 | 0.019868 | -0.021731 | 0.088276 | 0.026431
7 | 1.5433 | -0.097628 | -0.063867 | -0.028647 | -0.026524 | -0.028419 | -0.045435 | 0.056882 | 0.0039162 | 0.023778 | 0.055584 | -0.024726 | -0.043335 | -0.0011492 | 0.013337 | 0.041583 | 0.039595 | 0.039029 | 0.063721 | 0.13168 | 0.052876
8 | 0.014378 | 0.0049746 | -0.00030675 | 0.029393 | -0.032293 | -0.028574 | -0.04093 | 0.0042234 | -0.025201 | -0.039183 | -0.032094 | -0.073652 | -0.0095949 | 0.012869 | -0.031033 | -0.02889 | -0.038539 | -0.033032 | -0.047251 | -0.02478 | -0.02174
9 | -0.24841 | 0.0045034 | -0.01766 | 0.0076259 | -0.0113 | -0.026009 | -0.0091335 | -0.018559 | 0.012826 | 0.0013127 | -0.0068534 | -0.0063343 | 0.0034268 | -0.023539 | -0.049243 | -0.036259 | -0.011963 | -0.004138 | 0.0031957 | -0.054691 | -0.022914
10 | -8.7346e-06 | -0.00026904 | 3.1938e-06 | -0.00034053 | 1.9271e-05 | 5.1831e-07 | 0.0020248 | -0.004362 | -0.00013273 | -0.0001656 | 0.0019608 | -2.7327e-06 | 0.00014504 | 6.9244e-05 | 3.2007e-06 | -0.0024914 | -0.00017793 | 1.0629e-05 | 9.8147e-06 | -8.4599e-05 | 0
11 | 0.14854 | -0.055891 | -0.049428 | -0.034166 | -0.037807 | -0.033794 | -0.034694 | -0.021873 | -0.025943 | -0.02649 | -0.02806 | -0.031332 | -0.029708 | -0.020472 | -0.0261 | -0.015293 | -0.016694 | -0.01418 | -0.023379 | -0.025704 | -0.023585 | ten times!
12 | 0.063506 | 0.012059 | -0.045792 | -0.0065664 | -0.036211 | -0.0083094 | 0.026369 | 0.05694 | 0.0018668 | -0.023959 | -0.01204 | 0.050281 | -0.013327 | -0.043978 | 0.018528 | -0.012534 | 0.02609 | -0.036079 | -0.038455 | -0.047918 | -0.0045162 | ten times!
13 | -0.27835 | -0.020399 | -0.038154 | -0.029859 | -0.021125 | -0.022184 | -0.013711 | -0.011537 | -0.00018582 | -0.009723 | 0.002969 | -0.013736 | -0.020157 | -0.018305 | -0.038477 | -0.02304 | -0.0035853 | 0.011045 | -0.00017743 | -0.024991 | -0.02645 | ten times!
14 | 0.036497 | -0.17743 | -0.17733 | -0.16533 | -0.17484 | -0.16605 | -0.17308 | -0.17044 | -0.17842 | -0.18775 | -0.19682 | -0.18519 | -0.1781 | -0.15669 | -0.17553 | -0.17948 | -0.18245 | -0.17119 | -0.15846 | -0.080628 | -0.16358 | ten times!
15 | -0.043667 | -0.10926 | -0.12197 | -0.14221 | -0.033211 | -0.12283 | -0.03991 | 0.01164 | -0.073882 | -0.12385 | -0.020751 | -0.031688 | -0.057606 | -0.092538 | -0.093221 | -0.067214 | -0.10428 | -0.030122 | -0.012765 | -0.06418 | -0.069221 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.2812 | 0.086262 | 0.42905 | 0.73998 | 0.82558 | 1.1085 | 0.69341 | 0.25054 | 0.88421 | -0.086745 | 0.73224 | 0.22168 | 0.54289 | 0.72477 | 0.99087 | 0.28757 | 0.64498 | 0.44719 | 0.56894 | 2.8071 | 0.58764
3 | 35.2282 | -0.55146 | -0.29555 | -0.77979 | 0.23841 | -0.73923 | 0.36912 | -0.68292 | -0.16534 | 0.22206 | -1.1555 | -0.43184 | -0.51875 | -2.4967 | -1.1277 | 0.29371 | -0.079527 | 0.081379 | 0.76496 | -0.80631 | 0.77389
4 | 0.0038694 | 0.32747 | 0.14047 | -0.1056 | -0.0048491 | 0.19106 | 0.17722 | 0.088095 | -0.0032658 | -0.078397 | -0.010755 | -0.1528 | -0.036275 | 0.068447 | 0.070983 | 0.031761 | 0.14705 | 0.076355 | 0.0064692 | 0.0037573 | 0.048579
5 | -4.4616 | -0.60703 | -0.46164 | -0.13132 | -0.3446 | -0.093501 | -0.14276 | -0.098752 | -0.17586 | -0.10089 | 0.28082 | -0.24184 | -0.19499 | -0.12782 | -0.053385 | -0.291 | -0.14143 | -0.17797 | -0.35213 | -3.5485 | -0.44388
6 | -0.20366 | 0.1121 | 0.036786 | -0.024622 | 0.028408 | -0.054215 | -0.051626 | 0.078707 | 0.086265 | 0.13665 | 0.0058042 | 0.067694 | 0.072477 | -0.017169 | 0.060479 | 0.083904 | -0.016907 | 0.016268 | -0.065271 | -0.056511 | 0.021869
7 | 2.6217 | 0.85853 | 1.0871 | 1.2537 | 1.2158 | 1.2813 | 1.2424 | 1.2087 | 1.2389 | 1.3024 | 1.2964 | 1.2367 | 1.196 | 1.2057 | 1.1555 | 1.1882 | 1.1612 | 1.1937 | 1.2129 | 0.99238 | 1.2376
8 | -0.081048 | 0.019281 | 0.00096997 | 0.026921 | 0.014837 | 0.029912 | 0.006555 | 0.019775 | -0.019743 | -0.027163 | 0.02039 | -0.015519 | -0.015874 | -0.028098 | -0.0023619 | -0.017661 | -0.0010021 | 0.005002 | -0.027766 | 0.0063208 | -0.0038489
9 | 0.39304 | 0.51758 | 0.52028 | 0.52076 | 0.54702 | 0.52362 | 0.54886 | 0.52134 | 0.53355 | 0.5264 | 0.50887 | 0.5135 | 0.49753 | 0.50771 | 0.50819 | 0.53997 | 0.51787 | 0.5409 | 0.52839 | 0.50821 | 0.52009
10 | 2.1608e-05 | -0.00020906 | -4.0956e-06 | 2.1316e-06 | -1.9294e-05 | 1.4611e-05 | -7.1431e-07 | -1.9206e-07 | 0.00011635 | 7.0544e-05 | -0.0011937 | 9.5627e-07 | 0.0011202 | 4.9768e-07 | -0.00014856 | -7.697e-07 | 4.0218e-06 | 9.4308e-07 | -5.4036e-06 | -0.00015818 | 0
11 | 0.64534 | 0.76276 | 0.75978 | 0.77007 | 0.76183 | 0.77108 | 0.76833 | 0.77147 | 0.76644 | 0.76199 | 0.77428 | 0.76006 | 0.76533 | 0.7777 | 0.77852 | 0.7674 | 0.76918 | 0.76804 | 0.75593 | 0.65977 | 0.75981
12 | 0.026923 | -0.071797 | -0.015022 | 0.017869 | 0.034811 | -0.0085192 | 0.0093652 | -0.0023867 | -0.013152 | -0.011861 | 0.034394 | -0.013937 | 0.0045475 | -0.0026973 | -0.014939 | -0.021327 | -0.030897 | -0.023659 | 0.012737 | 0.027641 | -0.004445 | ten times!
13 | 0.87527 | 0.79087 | 0.76812 | 0.7518 | 0.72986 | 0.75135 | 0.75823 | 0.73494 | 0.75003 | 0.75434 | 0.76816 | 0.75091 | 0.79458 | 0.74081 | 0.75649 | 0.75315 | 0.74899 | 0.75035 | 0.7801 | 0.61801 | 0.75935 | ten times!
14 | 0.050351 | -0.067645 | -0.041361 | -0.028672 | -0.014773 | -0.038342 | -0.035378 | -0.02215 | -0.026942 | -0.040169 | -0.027008 | -0.041922 | -0.025075 | -0.038411 | -0.046148 | -0.038896 | -0.045269 | -0.038542 | -0.030744 | 0.016918 | -0.030869 | ten times!
15 | 0.12655 | 0.42579 | 0.41466 | 0.37358 | 0.3833 | 0.36974 | 0.3464 | 0.37326 | 0.34013 | 0.31642 | 0.40573 | 0.36821 | 0.38441 | 0.40926 | 0.4053 | 0.37997 | 0.42219 | 0.39656 | 0.38808 | 0.30672 | 0.37246 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -2.0776 | -1.3181 | -1.5548 | -0.35881 | -1.2206 | -0.73155 | -1.3868 | -1.858 | -1.9636 | -1.5744 | -1.4555 | -0.90523 | -0.95979 | -2.1871 | -1.5831 | -1.7947 | -1.5227 | -1.4579 | -0.96851 | 6.229 | -1.1593
3 | 1.0998 | 0.39186 | 0.51032 | 0.10966 | 0.62977 | 0.26937 | 0.11351 | 0.40757 | 0.12644 | 0.40309 | 0.24473 | 0.15238 | 0.040037 | 0.30952 | 0.61301 | 0.31701 | 0.55198 | 0.67589 | 0.12887 | 0.368 | 0.35982
4 | 0.11613 | 0.5163 | 0.27528 | -0.36657 | -0.1188 | -0.57892 | -0.53571 | 0.17093 | -0.0010968 | 0.20252 | 0.17757 | -0.27574 | -0.21955 | 0.52996 | 0.66069 | 0.32286 | 0.40289 | 0.39753 | -0.11529 | 0.34736 | 0.10132
5 | 2.8358 | -0.091713 | 0.0050628 | 0.17639 | 0.26997 | 0.17999 | 0.028966 | 0.04422 | 0.040842 | 0.23079 | 0.078765 | -0.09987 | -0.27437 | 0.227 | 0.042316 | 0.093008 | 0.033984 | 0.16545 | 0.14588 | -0.061434 | 0.15994
6 | -0.04848 | 0.14935 | 0.072088 | 0.030193 | -0.016076 | 0.018838 | -0.068732 | -0.012928 | -0.0055087 | 0.044254 | -0.063241 | -0.13761 | -0.10057 | -0.0012612 | 0.038486 | 0.062818 | 0.0090598 | -0.053781 | 0.029337 | 0.073576 | 0.00041706
7 | 1.629 | -0.04528 | -0.026318 | -0.065826 | 0.028562 | 0.0046062 | 0.0018574 | 0.039939 | 0.054876 | 0.017746 | 0.095717 | -0.028829 | -0.040064 | -0.028432 | -0.058135 | -0.0062119 | 0.00088467 | 0.027589 | 0.032798 | 0.10037 | 0.058101
8 | 0.0081869 | -0.0078629 | -0.0035689 | -0.0028966 | -0.028521 | -0.0054026 | -0.00077188 | -0.061878 | -0.061306 | -0.049771 | -0.044298 | -0.011407 | -0.011274 | -0.024183 | -0.0035397 | 0.0008468 | -0.0069715 | 0.012778 | 0.016592 | 0.019386 | -0.01408
9 | -0.16666 | 0.061869 | 0.026422 | 0.018076 | -0.024109 | -0.022167 | -0.018442 | 0.011338 | 0.028825 | 0.022391 | -0.0039896 | -0.013748 | -0.013183 | -0.0033155 | -0.02102 | 0.022423 | 0.025405 | 0.0085006 | 0.0045419 | -0.070584 | -0.0017416
10 | 1.2328e-05 | -0.0040187 | -3.6794e-06 | 2.1211e-07 | -1.9866e-05 | 1.385e-05 | -1.8201e-06 | -6.7267e-07 | 0.0029363 | -4.3864e-05 | 2.6452e-05 | -2.3742e-07 | -0.0020967 | 5.5362e-06 | -0.0041357 | 3.9975e-07 | 3.9057e-06 | -4.9893e-07 | -8.1746e-07 | -0.00012741 | 0
11 | 0.16127 | -0.040696 | -0.034814 | -0.028395 | -0.039756 | -0.032312 | -0.028608 | -0.024371 | -0.022514 | -0.021665 | -0.019353 | -0.035353 | -0.038096 | -0.029416 | -0.028181 | -0.01614 | -0.022687 | -0.012968 | -0.022585 | -0.02724 | -0.021347
12 | -0.021431 | 0.0058186 | 0.010095 | 0.026753 | -0.0030197 | 0.0010525 | -0.04488 | -0.011313 | -0.017718 | -0.0020538 | -0.065062 | -0.077203 | -0.04658 | -0.051462 | -0.039141 | -0.025877 | -0.057421 | -0.062706 | -0.012013 | -0.048919 | -0.027273 | ten times!
13 | -0.24196 | 0.028488 | -0.011985 | -0.0077591 | -0.0075736 | -0.025729 | -0.033893 | -0.019489 | -0.023369 | -0.013638 | -0.02079 | -0.025255 | -0.048964 | -0.044196 | -0.028766 | 0.013188 | 0.019409 | 0.0095424 | -0.0037206 | -0.057067 | -0.022733 | ten times!
14 | 0.075331 | -0.19626 | -0.19041 | -0.18781 | -0.19415 | -0.17131 | -0.15931 | -0.18123 | -0.18254 | -0.18541 | -0.1854 | -0.16158 | -0.17706 | -0.18874 | -0.17919 | -0.16888 | -0.19139 | -0.17372 | -0.1746 | -0.052138 | -0.1641 | ten times!
15 | 0.021693 | -0.010971 | -0.01388 | -0.039546 | 0.011635 | -0.11241 | -0.044295 | -0.0058666 | -0.11428 | -0.13808 | -0.056158 | -0.054061 | -0.051882 | -0.10682 | -0.093206 | -0.045052 | -0.033085 | -0.055633 | -0.021296 | -0.086031 | -0.053237 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00043833 | -0.00015341 | 4.3833e-05 | -0.00015341 | -0.00015341 | -0.00015341 | -0.00015341 | -0.00015341 | -0.00015341 | -0.00015341 | 4.3833e-05 | -0.00015341 | 4.3833e-05 | -0.00015341 | 4.3833e-05 | 4.3833e-05 | 4.3833e-05 | 0.00083282 |  | 596.4444
|  | 0.00049315 | -9.8629e-05 | -9.8629e-05 | -0.00029589 | -9.8629e-05 | 9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | -9.8629e-05 | 9.8629e-05 | 9.8629e-05 | 0.0006904 |  | 596.4118
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3621.5017
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.2772 | -3624.4606 | 5.1782 | -0.73974 | 1.2329 | -0.73974 | -0.73974 | -0.73974 | -0.73974 | -0.73974 | -0.73974 | -0.73974 | 1.2329 | -0.73974 | 1.2329 | -0.73974 | 1.2329 | 1.2329 | 1.2329 | 9.1234 | -3618.5427
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.073697 | 2.8666 | -0.94691 | -0.83525 | -0.55312 | -0.60555 | -0.13642 | -0.21416 | 0.44883 | 0.23828 | 0.032196 | -0.43969 | -0.0080908 | 0.38636 | 0.65085 | 0.026874 | 0.032302 | 0.3901 | 0.90188 | 0.63152 | -0.68512
db2 | -0.39015 | -1.9332 | 0.029933 | -0.33559 | 0.28906 | 0.30392 | 0.49907 | 0.43907 | 0.46315 | -0.21903 | 0.062473 | -0.29706 | 0.4904 | 0.31859 | 0.28091 | 0.0026612 | -0.1923 | -0.31472 | -0.52337 | -1.2972 | -1.0481
db3 | -1.451 | 32.5389 | -0.02475 | 0.49308 | 0.0091847 | 1.2451 | -0.56736 | 0.55281 | -0.45508 | 0.28906 | 0.76131 | -0.71965 | -0.68315 | 0.20358 | -1.8742 | -1.6109 | 0.16848 | 0.56455 | 0.52099 | 1.127 | -0.63819
db4 | 0.055011 | 0.48348 | 0.05408 | 0.022846 | -0.088878 | 0.020951 | 0.19621 | 0.029653 | 0.056416 | -0.096608 | -0.14576 | -0.1138 | 0.0060585 | -0.20763 | 0.020804 | -0.036269 | 0.052391 | 0.081487 | 0.0090434 | 0.13901 | -0.13539
db5 | -0.11013 | -5.0151 | -0.29727 | -0.14374 | 0.063145 | -0.20045 | 0.076155 | 0.063197 | 0.18801 | 0.093623 | 0.16902 | 0.42327 | 0.11679 | -0.14298 | -0.11488 | 0.11924 | -0.086016 | -0.057929 | -0.078115 | -0.19106 | -3.3796
db6 | 0.040402 | 0.055837 | 0.013909 | -0.044482 | -0.12068 | 0.019914 | -0.053626 | -0.0093362 | 0.068771 | -0.042169 | -0.18224 | 0.063016 | 0.011396 | -0.048516 | 0.010611 | -0.0098297 | 0.077201 | -0.074711 | 0.15586 | 0.16491 | -0.047539
db7 | 1.2424 | 2.5683 | -0.18698 | -0.067141 | 0.046795 | -0.018361 | 0.10069 | 0.059729 | 0.093347 | 0.098671 | 0.0054317 | 0.05415 | -0.079132 | 0.0037227 | -0.04359 | -0.064553 | -0.022622 | -0.033291 | 0.059099 | -0.0059683 | 0.98116
db8 | -0.0021727 | -0.081903 | -0.023124 | -0.024635 | -0.068821 | 0.013372 | -0.034355 | 0.022376 | 0.027382 | 0.00659 | -0.046018 | 0.0085625 | 0.058257 | 0.019215 | 0.024853 | 0.0042138 | -0.00081772 | -0.041037 | 0.015809 | 0.038177 | -0.024867
db9 | 0.53686 | 0.38175 | 0.0030743 | -0.017269 | -0.0020508 | 0.020987 | -0.020295 | 0.0057301 | 0.010624 | 0.017978 | 0.001715 | -0.010269 | -0.018597 | -0.015187 | -0.0073994 | -0.014739 | 0.015242 | -0.0049967 | 0.028143 | 0.0073098 | 0.50166
db10 | -0.00041412 | 0.035127 | 0.0037139 | -0.017646 | -0.14344 | 0.047433 | -0.033545 | 0.039602 | 0.091273 | 0.0039951 | -0.10666 | 0.068764 | 0.020331 | -0.0027457 | -0.0040922 | -0.014049 | 0.024863 | -0.080583 | 0.051084 | 0.051706 | -0.023456
db11 | 0.76488 | 0.63282 | -0.0099124 | -0.011081 | 0.0017563 | -0.0030161 | -0.001028 | -0.0033312 | -0.002061 | -0.00016489 | -0.0035986 | 0.0062721 | 0.0090463 | -1.8889e-07 | 0.0085585 | 0.010831 | 0.0017361 | -0.0038141 | 0.00078461 | -0.00097699 | 0.66807
db12 | -0.0023742 | -0.0025266 | 0.0011572 | -0.0076421 | -0.01705 | 0.010564 | -0.0031087 | 0.0055153 | 0.013225 | 0.0025156 | -0.010782 | 0.011513 | 0.0067414 | 0.0025971 | -0.0038278 | -0.0027969 | -0.00070904 | -0.014301 | 0.0046249 | 0.0017641 | -0.0079025
db13 | 0.07561 | 0.084584 | 0.0032047 | 0.0013969 | 0.0025989 | 0.0006275 | -0.00025639 | -6.2762e-05 | -0.0034938 | -0.0017515 | -6.9482e-05 | -0.0007221 | -0.0010058 | 0.0034085 | -0.0021152 | -0.0034306 | -0.0023474 | 0.00051901 | 0.00015663 | 0.0033429 | 0.058497
db14 | -0.0047326 | 0.0049715 | -0.0015076 | -0.0045957 | -0.0056248 | 0.0042686 | -0.0028244 | 0.0029233 | 0.0060229 | 0.00020891 | -0.0050047 | 0.0047395 | 0.0024608 | 0.00067602 | -0.0021135 | -0.0024304 | 0.0007534 | -0.0048653 | 0.0032834 | 0.0036296 | -0.0052012
db15 | 0.036002 | 0.0031714 | 0.0039393 | 0.0025159 | -0.011083 | 0.0014067 | 0.0023488 | -0.00028139 | 0.0019105 | 0.0006625 | -0.007625 | 0.0014371 | 0.00094237 | -0.00094461 | 0.0050129 | 0.0044838 | 0.002495 | -0.0081997 | 0.0032174 | -0.0022386 | 0.027758
da2 | 1.2986 | 4.4755 | -1.4933 | -0.88227 | 0.13705 | 0.088885 | -0.50604 | -0.15234 | 0.65873 | 0.11789 | -0.098672 | 0.243 | 1.8453 | 0.10086 | 0.89453 | -0.0954 | -0.63797 | -0.30169 | 0.19279 | -0.1113 | 2.4501
da3 | 0.43925 | -0.02745 | 0.24059 | 0.0034232 | 0.18491 | -0.043889 | -0.1632 | 0.031218 | 0.010516 | -0.2183 | -0.060251 | 0.22632 | -0.089401 | 0.15836 | 0.075285 | -0.1022 | -0.41533 | -0.081123 | 0.16757 | 0.075511 | 1.3778
da4 | 0.069693 | 1.5929 | 0.0060777 | -0.05839 | -0.39321 | 0.23219 | -0.26848 | -0.023559 | -0.26679 | 0.011243 | 0.36702 | -0.031742 | 0.080599 | -0.088472 | -0.13252 | 0.0056319 | 0.26798 | -0.090364 | 0.18319 | 0.19959 | 0.11629
da5 | 0.047356 | 2.5179 | -0.043442 | -0.072463 | 0.064744 | 0.10436 | -0.10307 | 0.019398 | -0.095978 | 0.0085807 | 0.093066 | 0.067444 | 0.014693 | -0.13693 | 0.099121 | 0.047316 | -0.0016748 | -0.093327 | -0.03225 | 0.060403 | 0.27271
da6 | 0.014528 | 0.38583 | 0.0096306 | -0.0588 | 0.044499 | -0.08303 | -0.041127 | 0.0086984 | -0.0083233 | -0.025561 | 0.10694 | -0.013582 | 0.035485 | -0.018985 | -0.076116 | 0.0038558 | 0.0082528 | 0.14062 | -0.022028 | -0.01043 | 0.0031091
da7 | -0.0029137 | 1.5449 | -0.094609 | -0.059817 | -0.039719 | -0.025657 | -0.021702 | -0.044006 | 0.059994 | 0.0067905 | 0.022472 | 0.056478 | -0.022577 | -0.040191 | 0.00078368 | 0.016783 | 0.04383 | 0.036013 | 0.039232 | 0.065902 | 0.1358
da8 | -0.02029 | -0.0043672 | 0.022674 | 0.0012455 | 0.090965 | -0.022393 | -0.040829 | -0.021015 | 0.022671 | -0.010575 | 0.027738 | -0.012293 | -0.048663 | 0.018794 | 0.0069155 | -0.030046 | -0.0095284 | 0.047153 | -0.030022 | -0.012791 | -0.080005
da9 | -0.013955 | -0.24985 | 0.01844 | -0.0023342 | -0.0032221 | 0.00066464 | -0.0073947 | 0.0048328 | -0.0053449 | 0.02674 | -0.0053376 | 0.0070585 | 0.0080141 | 0.017326 | -0.0093339 | -0.034108 | -0.02239 | -0.019923 | 0.0063445 | 0.019967 | -0.049218
da10 | 0.0037117 | -0.011518 | -0.0084966 | -0.03706 | 0.072257 | -0.022657 | -0.063627 | -0.0034609 | -0.010603 | -0.013576 | 0.07855 | -0.0042486 | 0.0044672 | 0.011232 | -0.051424 | -0.040226 | -0.0066933 | 0.10978 | -0.033194 | 0.018983 | -0.093255
da11 | -0.029685 | 0.14872 | -0.026198 | -0.019039 | -0.0081416 | -0.008724 | -0.0032614 | -0.0048692 | 0.0084887 | 0.0037403 | 0.00060845 | 0.0014872 | -0.0014241 | -1.8345e-05 | 0.0096348 | 0.0038979 | 0.014344 | 0.0083633 | 0.014895 | 0.0062161 | -0.024002
da12 | -2.6537e-05 | 0.0017251 | 0.00066668 | -0.008698 | 0.0082385 | -0.0058568 | -0.0080395 | 0.0025067 | 0.0051809 | -0.00091462 | 0.0069421 | -0.00139 | 0.0060313 | 0.00052844 | -0.0098198 | -0.0021866 | -0.0013443 | 0.016079 | -0.0070074 | -0.00091669 | -0.01462
da13 | -0.0018121 | -0.027823 | -0.00017925 | -0.00090911 | 3.7626e-06 | -0.0018317 | 0.0014576 | -0.00073324 | -0.0019558 | 0.0018684 | 0.0018293 | -0.00018699 | -0.00018796 | -0.00030558 | 0.00053148 | -0.0011478 | -0.00125 | 0.0011959 | 0.0012963 | 0.00050467 | -0.0010341
da14 | -0.017161 | 0.0032214 | -0.00091679 | -0.0026723 | 0.0075975 | -0.0017525 | -0.0030455 | -0.00043356 | 0.00031197 | -0.0013339 | 0.0049749 | -0.0027782 | -0.00087043 | 0.00030074 | -0.002055 | -0.0028695 | -0.0010877 | 0.0062175 | -0.0022835 | 0.0026963 | -0.013256
da15 | -0.0070315 | -0.0043667 | -0.0038949 | -0.0051652 | -0.0071896 | 0.0037104 | -0.0052516 | 0.0030405 | 0.0081955 | -0.00035671 | -0.0053531 | 0.0049564 | 0.0038627 | 0.0012709 | -0.0022223 | -0.0022906 | 0.00031008 | -0.0033966 | 0.0040193 | 0.005755 | -0.006418
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3618.228
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.2478 | -3626.1189 | 5.6716 | -0.24659 | -0.24659 | -2.2193 | -0.24659 | 1.7261 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | -0.24659 | 1.7261 | 1.7261 | 7.6443 | -3610.3371
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.07475 | 2.7483 | -0.36269 | -0.16776 | 0.031974 | -0.25277 | 0.14793 | 0.23779 | 0.62159 | 0.31978 | 0.004563 | -0.4688 | -0.30269 | -0.14057 | -0.034283 | -0.11464 | -0.10171 | 0.13367 | 0.47682 | -0.028209 | -0.42294
db2 | 0.56178 | -0.92842 | -0.48619 | -0.13703 | 0.17571 | 0.26413 | 0.55951 | 0.13212 | -0.31625 | 0.32541 | -0.65623 | 0.17181 | -0.33946 | -0.018891 | 0.18343 | 0.43512 | -0.27502 | 0.084068 | -0.11644 | 0.014203 | 2.829
db3 | -1.2012 | 35.2322 | -0.1513 | 0.097461 | -0.38624 | 0.63008 | -0.3564 | 0.7657 | -0.29208 | 0.22576 | 0.61461 | -0.76283 | -0.039248 | -0.12756 | -2.1052 | -0.73705 | 0.68222 | 0.31486 | 0.47084 | 1.1564 | -0.78087
db4 | 0.049307 | -0.021038 | 0.2546 | 0.07986 | -0.15735 | -0.067189 | 0.14609 | 0.12673 | 0.037464 | -0.046197 | -0.12484 | -0.059192 | -0.20059 | -0.085581 | 0.021845 | 0.022254 | -0.020858 | 0.095809 | 0.027131 | -0.049985 | 0.097278
db5 | -0.15278 | -4.4657 | -0.40774 | -0.26825 | 0.059257 | -0.15029 | 0.10219 | 0.048883 | 0.093708 | 0.014555 | 0.087204 | 0.47093 | -0.050202 | -0.0037183 | 0.064361 | 0.13733 | -0.10005 | 0.050018 | 0.013534 | -0.16173 | -3.5476
db6 | 0.029151 | -0.074615 | 0.13385 | 0.047817 | -0.043369 | 0.053556 | -0.14279 | -0.06949 | 0.067211 | -0.010128 | 0.0081106 | -0.021767 | 0.030363 | 0.043326 | -0.075439 | 0.012768 | 0.068229 | -0.020076 | -0.022041 | -0.060126 | -0.086986
db7 | 1.2086 | 2.6187 | -0.33487 | -0.10932 | 0.056303 | 0.021083 | 0.083717 | 0.044788 | 0.012351 | 0.045012 | 0.11276 | 0.097798 | 0.039345 | -0.0015482 | 0.0092948 | -0.041833 | -0.0097752 | -0.036089 | -0.0039642 | 0.014948 | 0.98919
db8 | -0.00053355 | -0.073188 | 0.061053 | 0.027032 | 0.032295 | 0.048791 | -0.0025673 | 0.014161 | 0.030533 | -0.058112 | -0.081109 | 0.020836 | -0.019498 | -0.01534 | -0.043883 | -0.012721 | -0.0095808 | 0.015256 | 4.2086e-05 | -0.0071889 | -0.028096
db9 | 0.52386 | 0.39345 | -0.001209 | -0.0025632 | -0.0049668 | 0.025127 | 0.0021347 | 0.023071 | -0.0038816 | 0.013742 | 0.011295 | -0.017072 | -0.011913 | -0.027888 | -0.016603 | -0.016719 | 0.014208 | -0.0067211 | 0.015486 | 0.0044724 | 0.49648
db10 | 0.00032686 | 0.039918 | 0.074494 | 0.046552 | 0.010027 | 0.057765 | -0.060737 | 0.012161 | 0.019397 | -0.069424 | -0.095921 | -0.0020577 | -0.0085159 | 0.0007933 | -0.03169 | -0.021316 | 0.014541 | 0.02805 | -0.009777 | 0.035658 | -0.049115
db11 | 0.76751 | 0.64512 | -0.0045259 | -0.0074594 | 0.0024613 | -0.0048899 | 0.004197 | 0.00091968 | 0.0040028 | -0.00012412 | -0.0037968 | 0.0064905 | -0.0073455 | -0.0021792 | 0.010307 | 0.011127 | -0.00013082 | 0.0015814 | 0.00064818 | -0.011283 | 0.65719
db12 | -0.0005941 | 0.0062476 | 0.0027334 | 0.0047799 | 0.0035528 | 0.010783 | -0.0072839 | 0.003051 | 0.0026271 | -0.0089275 | -0.012327 | 0.0039681 | -0.0017888 | 0.0010489 | -0.0032726 | -0.0033364 | 0.00013844 | 0.00085026 | -0.0029199 | 0.0063229 | -0.0029093
db13 | 0.076068 | 0.088381 | 0.0029037 | 0.00086363 | -0.00034863 | -0.0022375 | -0.00059325 | 0.00024457 | -0.0021703 | -0.00024596 | 0.0013644 | 2.6244e-05 | -0.00069008 | 0.00339 | -0.0018417 | -0.000206 | -0.00022554 | -0.0016235 | -0.00069828 | 0.0020882 | 0.059759
db14 | -0.0034078 | 0.0060566 | 0.0025053 | 0.0028479 | 0.0013348 | 0.0059059 | -0.004776 | 0.00078338 | 0.0025127 | -0.0035812 | -0.0056858 | 0.00048992 | -0.0012838 | 0.0009003 | -0.0028382 | -0.0027257 | 0.00070241 | 0.0010243 | -0.001046 | 0.0029298 | -0.0031584
db15 | 0.03835 | 0.012655 | 0.0042292 | 0.0031161 | -0.00099203 | -1.9562e-05 | -0.0013762 | -0.0037103 | -0.0010236 | -0.0043372 | -0.0067081 | 0.002223 | -0.0015288 | 9.0668e-05 | 0.0025758 | 0.0021799 | -0.00035259 | 0.0038692 | 0.0013064 | 0.00045805 | 0.030672
da2 | -1.3766 | -2.4489 | 0.066566 | -0.17223 | 1.0142 | 0.1667 | 0.64398 | -0.0077421 | -0.48201 | -0.58922 | -0.20654 | -0.072487 | 0.47012 | 0.41683 | -0.80923 | -0.20975 | -0.41519 | -0.14175 | -0.082745 | 0.41052 | 6.2491
da3 | 0.33377 | 1.1041 | 0.073217 | 0.18185 | -0.22631 | 0.2933 | -0.052128 | -0.2212 | 0.075579 | -0.20726 | 0.061546 | -0.088963 | -0.18094 | -0.29374 | -0.030735 | 0.27394 | -0.014712 | 0.22216 | 0.34091 | -0.20651 | 0.35158
da4 | 0.078616 | 0.26782 | 0.43436 | 0.19562 | -0.44563 | -0.19283 | -0.66237 | -0.61577 | 0.09201 | -0.0812 | 0.11135 | 0.095871 | -0.35496 | -0.29816 | 0.44807 | 0.58163 | 0.24151 | 0.32551 | 0.31722 | -0.19223 | 0.57467
da5 | 0.072769 | 2.8357 | -0.15727 | -0.06513 | 0.10374 | 0.19843 | 0.1084 | -0.044738 | -0.027591 | -0.031535 | 0.15452 | 0.0058942 | -0.17159 | -0.34714 | 0.15452 | -0.030981 | 0.022465 | -0.038843 | 0.093007 | 0.073843 | -0.053651
da6 | 0.0043942 | -0.086315 | 0.14227 | 0.071125 | 0.044244 | -0.00028241 | -0.0025793 | -0.051881 | -0.0026272 | -0.013089 | 0.058948 | -0.089092 | -0.13378 | -0.10497 | -0.010341 | 0.035024 | 0.079436 | -0.004358 | -0.049662 | 0.03161 | -0.025205
da7 | 0.00073387 | 1.6312 | -0.046224 | -0.026902 | -0.065932 | 0.027736 | 0.0046443 | 0.0015175 | 0.038678 | 0.058142 | 0.019071 | 0.094591 | -0.02966 | -0.040798 | -0.028177 | -0.058811 | -0.0069718 | -0.00028584 | 0.026693 | 0.032689 | 0.10544
da8 | -0.013089 | -0.015445 | 0.010037 | 0.013654 | 0.020856 | -0.0053374 | -0.0002853 | 0.024594 | -0.040443 | -0.050268 | -0.02721 | -0.042436 | 0.00625 | 0.0018142 | -0.014005 | 0.0099972 | 0.027054 | 0.001475 | 0.031052 | 0.033202 | -0.040699
da9 | 0.0061877 | -0.1698 | 0.054972 | 0.020683 | 0.012337 | -0.027815 | -0.026169 | -0.024016 | 0.005818 | 0.022235 | 0.011368 | -0.010134 | -0.020103 | -0.019371 | -0.0091302 | -0.027126 | 0.017009 | 0.018524 | 0.0021023 | -0.0011836 | -0.057218
da10 | 0.0046189 | -0.0272 | -0.012633 | -0.0016361 | 0.014407 | 0.01214 | -0.018969 | 0.016825 | 0.0097642 | 0.00033259 | 0.018149 | -0.02596 | 0.0040005 | -0.0067156 | -0.0091883 | -0.0080033 | 0.017779 | -0.015441 | 0.0048945 | 0.00025501 | -0.10298
da11 | -0.027616 | 0.1604 | -0.013014 | -0.0070702 | -0.00062958 | -0.011704 | -0.0043528 | -0.00094832 | 0.0033024 | 0.0052734 | 0.0052863 | 0.0081605 | -0.0076828 | -0.010481 | -0.0015205 | -0.00045258 | 0.011464 | 0.0046445 | 0.014673 | 0.0050518 | -0.02553
da12 | -0.0020245 | -0.0092969 | 0.0026442 | 0.0034563 | 0.0069697 | 0.0037966 | 0.0002117 | 7.5116e-05 | 0.0025405 | 0.00029225 | 0.0041349 | -0.0070025 | -0.0046856 | -0.0026335 | -0.0036166 | -0.0018169 | 0.0021479 | -0.0048305 | -0.0031623 | 0.0014787 | -0.015972
da13 | -0.0014314 | -0.024201 | 0.001752 | -0.0012696 | 0.00030226 | -0.00098459 | 0.00080542 | -0.0023443 | -0.0010958 | 0.0010657 | 0.0022065 | -0.00054293 | -0.00093265 | -0.003465 | -0.0019549 | -0.00083384 | 0.00227 | 0.0024395 | 0.0025851 | -2.8647e-06 | -0.0040564
da14 | -0.017413 | 0.0067196 | -0.0023636 | -0.0013565 | -0.00012118 | -0.00049183 | 0.00037033 | 0.002769 | 0.00028475 | 0.00060183 | 0.002469 | -0.0026951 | 0.0018977 | -0.00029307 | -0.0012849 | -0.00015619 | 0.0019808 | -0.0026646 | 0.00081434 | 0.00023913 | -0.011225
da15 | -0.0054716 | 0.0021693 | 0.0043744 | 0.0040836 | 0.001517 | 0.0066351 | -0.0057689 | 0.0010421 | 0.0048849 | -0.0059562 | -0.0083369 | -0.00014425 | 6.5513e-05 | 0.00028342 | -0.0052107 | -0.003849 | 0.00096641 | 0.0021631 | -9.1678e-05 | 0.003342 | -0.0086031
Field Colinearity | 0.30936
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.7929 | -1.0206 | -0.90894 | -0.62681 | -0.67924 | -0.21012 | -0.28785 | 0.37513 | 0.16458 | -0.041501 | -0.51338 | -0.081788 | 0.31266 | 0.57715 | -0.046823 | -0.041395 | 0.3164 | 0.82818 | 0.55782 | -0.75882
Twist (m2 mrad) | 20.9466 | -14.5436 | -12.2707 | -7.9919 | -8.1509 | -2.3638 | -3.0225 | 3.6575 | 1.4813 | -0.34238 | -3.8504 | -0.55207 | 1.8759 | 3.0301 | -0.2107 | -0.15523 | 0.94919 | 1.8634 | 0.83673 | -0.28456
Twist Integral (I0) (m2 rad) | -0.014324
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.6736 | -0.43744 | -0.24251 | -0.042776 | -0.32752 | 0.073181 | 0.16304 | 0.54684 | 0.24503 | -0.070187 | -0.54354 | -0.37744 | -0.21532 | -0.10903 | -0.18939 | -0.17646 | 0.058923 | 0.40207 | -0.10296 | -0.49769
Twist (m2 mrad) | 20.0516 | -6.2335 | -3.2738 | -0.54539 | -3.9302 | 0.82329 | 1.7119 | 5.3317 | 2.2052 | -0.57904 | -4.0766 | -2.5477 | -1.2919 | -0.57242 | -0.85227 | -0.66172 | 0.17677 | 0.90465 | -0.15444 | -0.18663
Twist Integral (I0) (m2 rad) | 0.0047246

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.39015 | -0.64102 | 0.055011 | -0.14863 | 1.2314 | -0.73697 | 1.2986 | 0.43925 | 0.069693 | 15.9152 | 0.59644 | 0.56178 | -0.39119 | 0.049307 | -0.19128 | 1.1976 | -0.7475 | -1.3766 | 0.33377 | 0.078616 | 18.1394 | 0.59641 | -0.46303 | 0.45636 | 0.062886 | -0.41653 | 1.2673 | 0 | 1.4418 | 0.45489 | 0.12162 | 14422.645 | 0.58764 | 0.77389 | 0.048579 | -0.44388 | 1.2376 | 0 | -1.1593 | 0.35982 | 0.10132 | 14423.346
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85 | 0.2 | 0.2 | 0.85 | 0.85