## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0_rN-3075C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-11-11
Beginn der Messung / start of measuring: |  | TT.11.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000075_cc
Ende der Messung / end of measuring: |  | TT.11.JJ
ITP: |  | 11 | Status / status: | after repair
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Kenklies |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-11-11 |  | 2003-11-11 | 2003-11-11
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000075_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000075
Date of test Ap 1 | 2003-11-11 |  |  |  | Date of test Ap 2 | 2003-11-11
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | yellow alarm | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | yellow alarm | status ok | status ok | b7 | status ok | yellow alarm | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | yellow alarm | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | yellow alarm | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
174 | HCMB__A001 | 3000075 |  | 3 |  |  | 3
| COMMENTS: | Meas. after re-coll. Warning due to yellow al. on ap.1, b9 pos. 1 (5.8 sig), b7 pos. 2 (4.4 sig), b3 pos. 14 (4.45 sig) and ap. 2, b9 pos. 1 (5.9 sig), b7 pos. 2 (5.4 sig).
| HOLDING POINT | OK-w | Data at CERN |  | 2003-11-11 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-11-12 | Answer to manufacturer |  |  | 2003-11-14
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000075_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000075 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3075_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3075_cct2 |  |  |  | a even | 1
Test Operator |  | Kenklies |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-11-11 |  |  | Time of test |  |  | 7:59 |  |  |  | Date of test |  | 2003-11-11 |  |  | Time of test |  |  | 7:59
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-02-05
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00010281 |  |  | Dy average |  |  | 0.00072798 |  |  |  | Dx average |  | 0.00011399 |  |  | Dy average |  |  | 0.00064939
Magnetic length |  | 14.4223 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1729 |  |  |  | Magnetic length |  | 14.4204 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1794
Phase |  | -14.6953 |  |  | Integral field (T.m) |  |  | 85.8631 |  |  |  | Phase |  | -33.7026 |  |  | Integral field (T.m) |  |  | 85.7867
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.5091 | 5.9646 | 5.9622 | 5.9615 | 5.9609 | 5.9612 | 5.9614 | 5.9615 | 5.961 | 5.9616 | 5.9609 | 5.961 | 5.9606 | 5.961 | 5.9611 | 5.9617 | 5.9622 | 5.9615 | 5.9653 | 3.6796 | 85.8631 | C1 (T) | 3.4795 | 5.9647 | 5.9614 | 5.9615 | 5.961 | 5.9615 | 5.9621 | 5.9617 | 5.9611 | 5.9615 | 5.9609 | 5.9609 | 5.9615 | 5.9613 | 5.9615 | 5.9609 | 5.9618 | 5.9615 | 5.9654 | 3.6068 | 85.7867
Angle (mrad) | 3.8904 | -0.95757 | -0.43382 | -0.6591 | -0.38632 | -0.15506 | 0.43637 | 1.0616 | 0.35503 | -0.15707 | -0.085016 | 0.19735 | 0.30831 | 0.27972 | -0.13885 | -0.51634 | -0.55318 | -0.133 | -0.3101 | -0.71602 | -14.6953 | Angle (mrad) | 3.7923 | -1.0322 | -0.38556 | -0.66582 | -0.62438 | -0.25467 | 0.46898 | 1.0484 | 0.31292 | -0.12207 | -0.10795 | -0.061159 | 0.22732 | 0.49587 | -0.15404 | -0.47525 | -0.34102 | -0.004294 | -0.1329 | -0.66945 | -33.7026
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.96052 | 0.70179 | 0.86818 | 0.21402 | 0.59831 | 0.94735 | 0.84229 | 1.2612 | 0.67934 | -0.56039 | 0.55989 | 0.81401 | 1.1192 | 0.79002 | 0.1077 | 0.73825 | -0.042429 | 0.90791 | 0.49576 | 0.19374 | 0.6106 | b2 | 1.599 | 0.60025 | 0.1817 | 0.37762 | 0.22388 | 0.1231 | 0.40133 | 0.71793 | 0.10315 | -0.23543 | 0.41395 | 0.1789 | 0.4937 | 0.05998 | -0.46133 | -0.21328 | -0.017123 | 0.096406 | -0.065497 | 1.5995 | 0.25433
b3 | 32.7919 | -4.0613 | -3.4475 | -3.6019 | -3.5326 | -5.3452 | -4.1483 | -2.9929 | -2.6782 | -4.2814 | -5.2352 | -4.0019 | -3.9485 | -6.7933 | -5.5711 | -4.6372 | -3.7443 | -3.4887 | -2.7747 | -2.2636 | -2.9356 | b3 | 33.6779 | -3.6734 | -3.4156 | -4.2133 | -3.8157 | -4.748 | -3.6429 | -4.551 | -3.6111 | -3.3819 | -4.2859 | -3.8304 | -3.1176 | -5.0002 | -4.3987 | -3.3944 | -3.9448 | -4.3018 | -2.8332 | -3.8594 | -2.7537
b4 | -0.97516 | -0.11458 | 0.17238 | 0.085877 | 0.0020324 | 0.14666 | 0.069828 | 0.29641 | 0.095792 | -0.016535 | 0.040729 | -0.0075303 | -0.013346 | -0.073833 | 0.096835 | 0.084314 | -0.06699 | 0.027797 | -0.0069976 | 0.12348 | 0.016713 | b4 | -1.3872 | -0.18517 | -0.099553 | -0.040006 | -0.018825 | -0.24224 | -0.23827 | -0.29777 | -0.26356 | -0.18592 | -0.15834 | -0.23464 | -0.2885 | -0.42734 | -0.24926 | -0.36497 | -0.37778 | -0.36031 | -0.10628 | 0.036873 | -0.25671
b5 | -6.0724 | -0.44039 | -0.58019 | -0.26301 | -0.27026 | -0.1539 | -0.16582 | -0.15621 | -0.28179 | -0.017634 | -0.0050925 | -0.072987 | -0.2182 | -0.049364 | 0.058478 | -0.42767 | -0.37629 | -0.38753 | -0.27113 | -3.4357 | -0.5089 | b5 | -6.1453 | -0.5857 | -0.42248 | -0.13394 | 0.076638 | 0.21049 | 0.28393 | 0.15541 | 0.1322 | -0.029075 | 0.14453 | -0.063111 | -0.25054 | -0.079107 | -0.00086768 | -0.40925 | -0.24249 | -0.26683 | -0.49013 | -3.2283 | -0.39139
b6 | 0.13351 | 0.063316 | 0.080556 | 0.044375 | 0.0061567 | -0.036952 | -0.0064989 | -0.079278 | 0.070191 | 0.1981 | 0.055691 | 0.1542 | 0.095475 | -0.033219 | 0.18362 | 0.175 | 0.13813 | 0.060457 | 0.1173 | 0.022074 | 0.071794 | b6 | -0.20486 | -0.0093148 | -0.077161 | -0.079665 | 0.047809 | 0.030929 | -0.04616 | 0.016862 | -0.017079 | -0.0057918 | -0.098936 | -0.034573 | -0.16825 | -0.12246 | 0.089617 | 0.057834 | 0.051182 | -0.066443 | 0.048188 | -0.32053 | -0.036314
b7 | 2.6222 | 0.92039 | 1.0531 | 1.1824 | 1.2062 | 1.2304 | 1.2174 | 1.1842 | 1.1617 | 1.2197 | 1.2378 | 1.1644 | 1.1625 | 1.1759 | 1.0748 | 1.2216 | 1.1793 | 1.1591 | 1.0949 | 0.86186 | 1.1934 | b7 | 2.6037 | 0.82195 | 0.99421 | 1.1513 | 1.1618 | 1.2218 | 1.1791 | 1.2183 | 1.154 | 1.133 | 1.1492 | 1.1346 | 1.1128 | 1.124 | 1.1268 | 1.0915 | 1.0852 | 1.1507 | 1.0433 | 0.77612 | 1.1487
b8 | 0.027688 | 0.063021 | 0.034195 | 0.036441 | 0.037423 | 0.041028 | -0.0051053 | -0.02977 | 0.043838 | 0.011876 | 0.015189 | 0.02588 | -0.0024734 | -0.056412 | -0.076593 | -0.0093033 | 0.0092754 | -0.015677 | -0.018344 | -0.047534 | 0.0047626 | b8 | -0.086617 | 0.0073846 | -0.020646 | 0.01223 | -0.045456 | 0.0050984 | -0.042429 | -0.044125 | -0.041545 | -0.030734 | -0.046624 | -0.01142 | 0.0011444 | -0.016997 | 0.0097314 | 0.023758 | 0.0061949 | -0.021672 | 0.039395 | 0.019642 | -0.013306
b9 | 0.29228 | 0.54151 | 0.51316 | 0.50755 | 0.52718 | 0.49559 | 0.50231 | 0.50722 | 0.50724 | 0.49197 | 0.51223 | 0.50968 | 0.51613 | 0.50581 | 0.49504 | 0.50643 | 0.5051 | 0.53082 | 0.51835 | 0.45852 | 0.50237 | b9 | 0.28809 | 0.52478 | 0.50926 | 0.50839 | 0.53374 | 0.50826 | 0.50811 | 0.50158 | 0.51803 | 0.52842 | 0.51212 | 0.50664 | 0.51603 | 0.51079 | 0.51093 | 0.52741 | 0.4999 | 0.52034 | 0.52672 | 0.47152 | 0.5068
b10 | 0.10365 | 0.095723 | 0.048902 | 0.030023 | 0.013821 | -0.00025639 | -0.056336 | -0.059277 | 0.0039779 | 0.009329 | -0.041908 | 0.036013 | -0.0049741 | -0.11681 | -0.0091854 | -0.0044566 | 0.014285 | 0.018588 | 0.004146 | -0.069142 | 0 | b10 | 0.040517 | 0.010604 | -0.031852 | 0.01179 | -0.0042179 | 0.012152 | -0.067105 | -0.046924 | -0.017981 | 0.010401 | -0.052041 | -0.011147 | -0.024258 | -0.037577 | 0.079021 | 0.082547 | 0.05635 | -0.012277 | 0.059817 | -0.067756 | 0
b11 | 0.59019 | 0.75135 | 0.74101 | 0.75218 | 0.74884 | 0.75316 | 0.74583 | 0.75014 | 0.75031 | 0.74937 | 0.75592 | 0.75127 | 0.75021 | 0.76682 | 0.7695 | 0.75188 | 0.75294 | 0.75111 | 0.74703 | 0.65251 | 0.74399 | b11 | 0.58557 | 0.75934 | 0.75669 | 0.76509 | 0.76531 | 0.76819 | 0.76475 | 0.76184 | 0.76647 | 0.75831 | 0.7668 | 0.75871 | 0.76568 | 0.77081 | 0.77064 | 0.75918 | 0.75847 | 0.76394 | 0.76099 | 0.66995 | 0.75504
b12 | 0.01914 | 0.013916 | 0.002849 | 0.0002815 | -0.0012934 | -0.0014109 | -0.0058092 | -0.0084929 | -0.0028206 | -0.0032589 | -0.0042591 | 0.0068525 | 0.0001346 | -0.0099577 | -0.0022968 | 0.001071 | 0.0063571 | 0.00026812 | -0.001627 | -0.0065148 | -0.00011679 | b12 | 0.0033677 | 0.0014092 | -0.0053136 | 0.00019488 | -0.0025159 | -0.001634 | -0.0072785 | -0.0039561 | 0.001755 | 0.00047163 | -0.0053128 | 0.00073265 | 0.0007393 | -0.00053915 | 0.0084068 | 0.0096878 | 0.0093756 | 0.0019871 | 0.0062757 | 0.0018915 | 0.00091716
b13 | 0.068796 | 0.063528 | 0.067723 | 0.06709 | 0.067078 | 0.064199 | 0.065466 | 0.065564 | 0.065044 | 0.063947 | 0.06272 | 0.06276 | 0.06571 | 0.061453 | 0.060358 | 0.065807 | 0.065333 | 0.06676 | 0.06584 | 0.049023 | 0.064415 | b13 | 0.079928 | 0.064603 | 0.066341 | 0.065881 | 0.064925 | 0.063905 | 0.065236 | 0.065763 | 0.062332 | 0.063128 | 0.065082 | 0.068118 | 0.071304 | 0.067304 | 0.0654 | 0.06835 | 0.065791 | 0.068987 | 0.069004 | 0.04607 | 0.065975
b14 | 0.0018038 | 0.0031107 | 0.00033835 | -0.0020567 | -0.0022062 | -0.0035562 | -0.0067398 | -0.006548 | -0.0037786 | -0.0038442 | -0.0053374 | -0.00051776 | -0.0032424 | -0.0094841 | -0.0052435 | -0.0044518 | -0.0015488 | -0.0029884 | -0.0018901 | -0.0043604 | -0.0032078 | b14 | 0.004494 | -0.0016589 | -0.004257 | -0.0016809 | -0.0032108 | -0.00079402 | -0.0057084 | -0.0027398 | -0.0025247 | -0.001309 | -0.0054503 | -0.0018288 | -0.0029636 | -0.0033235 | 0.0032417 | 0.0038079 | 0.0016432 | -0.0028768 | 0.00059288 | -0.0058241 | -0.0016645
b15 | -0.009771 | 0.041042 | 0.038177 | 0.03392 | 0.0329 | 0.032163 | 0.032447 | 0.031024 | 0.033667 | 0.031372 | 0.033304 | 0.032538 | 0.032546 | 0.029857 | 0.031936 | 0.032429 | 0.031809 | 0.035232 | 0.030811 | 0.022279 | 0.03151 | b15 | -0.0060801 | 0.040327 | 0.033885 | 0.030065 | 0.030531 | 0.030789 | 0.032397 | 0.025918 | 0.029856 | 0.030673 | 0.03265 | 0.028575 | 0.029748 | 0.029236 | 0.032639 | 0.031432 | 0.030898 | 0.03034 | 0.033143 | 0.028019 | 0.030044
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 38.9046 | -9.5757 | -4.3382 | -6.591 | -3.8633 | -1.5506 | 4.3637 | 10.6156 | 3.5503 | -1.5707 | -0.85016 | 1.9735 | 3.0831 | 2.7972 | -1.3885 | -5.1634 | -5.5318 | -1.33 | -3.101 | -7.1602 | 0 | a1 | 37.9233 | -10.3216 | -3.8556 | -6.6582 | -6.2438 | -2.5467 | 4.6898 | 10.4839 | 3.1292 | -1.2207 | -1.0795 | -0.61159 | 2.2732 | 4.9587 | -1.5404 | -4.7525 | -3.4102 | -0.042943 | -1.329 | -6.6945 | 0
a2 | 10.2777 | 0.093547 | 1.3896 | 0.65468 | 2.4139 | 1.635 | 2.034 | 1.577 | 1.745 | 0.22462 | 1.2246 | 1.6214 | 2.0471 | 1.808 | 0.61674 | 1.7106 | 1.5251 | 1.7526 | 1.2274 | 2.062 | 1.6985 | a2 | 16.6656 | 0.33144 | 0.18601 | 0.81301 | 1.979 | 0.67094 | 0.61029 | 0.87071 | 0.93859 | 0.89016 | 0.86983 | 0.64755 | 0.039965 | -0.092125 | 0.10407 | -0.053437 | 0.89027 | 0.68437 | 0.041015 | 5.2835 | 1.2165
a3 | 4.5964 | 1.5887 | 1.6982 | 0.88932 | 1.1359 | 0.83779 | 0.74268 | 0.79192 | 0.48223 | 1.0621 | 1.5811 | 1.3032 | 1.0464 | 0.82783 | 0.66189 | 0.70396 | 1.1123 | 1.0223 | 0.93875 | -0.64048 | 1.0797 | a3 | -0.011045 | 1.4181 | 1.4614 | 0.9253 | 1.2348 | 0.87304 | 0.78206 | 0.91727 | 0.79913 | 0.88239 | 1.0312 | 0.72953 | 0.66322 | 0.88313 | 0.82541 | 0.97715 | 1.1036 | 1.2046 | 1.0296 | -0.66246 | 0.90334
a4 | 1.5559 | 0.24011 | -0.14815 | 0.55309 | 0.27568 | 0.15809 | -0.026885 | 0.12337 | 0.24244 | 0.56412 | 0.25011 | 0.0090329 | 0.18949 | 0.055888 | 0.2827 | -0.30351 | -0.17795 | -0.35351 | 0.060447 | -0.041655 | 0.15019 | a4 | 4.2356 | -0.036736 | 0.26249 | 0.31308 | 0.11291 | 0.42393 | 0.37978 | 0.4139 | 0.44573 | 0.77599 | 0.87942 | 0.60932 | 0.40627 | 0.28374 | 0.35205 | 0.62195 | 0.24722 | 0.28776 | 0.53728 | 2.1175 | 0.57684
a5 | 3.8604 | 0.35657 | 0.32225 | 0.33811 | 0.44198 | 0.32319 | 0.14127 | 0.17381 | 0.11517 | 0.5 | 0.31485 | 0.11256 | 0.082013 | 0.3493 | 0.0058204 | 0.20915 | 0.31782 | 0.15369 | 0.29283 | -0.10005 | 0.35202 | a5 | 2.2886 | 0.10023 | 0.22031 | 0.33352 | 0.39005 | 0.32171 | 0.2799 | 0.16003 | 0.41956 | 0.31687 | 0.16521 | 0.18935 | 0.13086 | 0.31002 | 0.1407 | 0.25318 | 0.036796 | 0.094665 | 0.14737 | -0.62291 | 0.25895
a6 | 0.5064 | 0.10586 | -0.013338 | 0.18654 | 0.087513 | 0.12308 | 0.053533 | 0.056676 | 0.075711 | 0.10121 | 0.13138 | 0.0052228 | 0.052438 | 0.010694 | 0.21547 | 0.14585 | 0.067895 | 0.032392 | 0.19534 | 0.19793 | 0.10693 | a6 | 0.68452 | -0.097022 | 0.025575 | 0.11675 | -0.12082 | -0.13679 | -0.14485 | 0.029127 | 0.024056 | 0.052431 | 0.15464 | 0.041841 | -0.060863 | 0.0035287 | 0.053731 | 0.031246 | 0.012156 | 0.011117 | -0.11606 | -0.32403 | 0.0043304
a7 | 1.7844 | -0.050071 | 0.002379 | -0.082032 | 0.053836 | 0.030346 | 0.055966 | 0.020052 | -0.011463 | -0.10467 | 0.030885 | 0.034378 | -0.040675 | -0.036635 | 0.041427 | 0.092625 | 0.02734 | 0.040955 | 0.026635 | 0.037419 | 0.062722 | a7 | 1.5792 | -0.027329 | 0.031892 | -0.024567 | 0.0015116 | 0.112 | 0.098486 | 0.041792 | 0.019516 | -0.014348 | 0.10093 | -0.024082 | 0.011821 | 0.099996 | 0.089026 | 0.066693 | 0.08864 | 0.02094 | 0.071858 | 0.1911 | 0.09391
a8 | 0.016239 | -0.044669 | -0.03326 | -0.03133 | -0.0017088 | 0.0062358 | 0.0023881 | -0.00567 | 0.016365 | 0.023391 | -0.022622 | 0.00060028 | 0.020875 | 0.025716 | 0.016558 | 0.035257 | -0.0072433 | -0.047282 | 0.0072555 | 0.0060998 | -0.0013457 | a8 | 0.025918 | -0.020996 | 0.00016805 | 0.10171 | 0.092089 | 0.0241 | 0.0037124 | 0.063432 | 0.048869 | 0.067132 | 0.086957 | 0.095117 | 0.051589 | 0.062077 | 0.054226 | 0.039669 | 0.063197 | 0.03702 | 0.015738 | -0.05295 | 0.045278
a9 | -0.17302 | 0.025332 | 0.031302 | 0.0014333 | 0.012968 | 0.010639 | 0.0055838 | 0.0026144 | 0.0070161 | -0.0015264 | 0.033928 | 0.0029876 | -0.026464 | 0.0024197 | -0.025867 | 0.018901 | 0.01259 | -0.0066147 | -0.018929 | -0.087521 | -0.0035165 | a9 | -0.26772 | 0.032024 | 0.042205 | 0.028349 | 0.013582 | 0.018265 | 0.012381 | 0.030098 | 0.029719 | 0.016478 | 0.0074699 | 0.0081298 | 0.019118 | 0.025573 | -0.012681 | 0.0070175 | 0.020758 | 0.0070803 | 0.0072086 | -0.077592 | 0.0057116
a10 | 0.079262 | -0.04244 | -0.048914 | 0.0079117 | -0.0037509 | 0.0023076 | -0.0059304 | -0.010526 | -0.0083906 | 0.033501 | -0.022853 | -0.0030314 | 0.0019005 | -0.0097048 | 0.044517 | -0.00028405 | 0.022256 | -0.028387 | 0.049359 | -0.039214 | 0 | a10 | -0.027318 | -0.076748 | -0.032378 | 0.003882 | 0.00090618 | 0.0027628 | -0.037359 | 0.03743 | 0.0069926 | 0.0096381 | -0.0014091 | 0.053349 | 0.030227 | 0.027652 | 0.00053753 | 0.02139 | 0.031643 | 0.026558 | -0.0010971 | -0.14543 | 0
a11 | 0.16479 | -0.066014 | -0.050583 | -0.048253 | -0.045994 | -0.044484 | -0.038195 | -0.032005 | -0.044572 | -0.046931 | -0.039444 | -0.040432 | -0.048144 | -0.034642 | -0.042017 | -0.036501 | -0.029501 | -0.034067 | -0.037665 | -0.024324 | -0.035274 | a11 | 0.17258 | -0.027873 | -0.0060944 | -0.0066975 | -0.0098721 | 0.0095286 | 0.0099609 | 0.015683 | 0.0085161 | -0.0054498 | 0.0033053 | -0.0012852 | 0.0065177 | 0.014295 | 0.00026407 | 0.0051406 | -0.0019125 | 0.0063675 | 0.00051104 | 0.021454 | 0.0070138
a12 | 0.0019201 | -0.0025692 | -0.0068748 | 0.0040668 | -0.00070898 | -0.00021464 | -0.0028499 | -0.0051965 | -0.0028161 | 0.00097301 | -0.004388 | -0.0040677 | -0.0016646 | -0.0074224 | 0.0065694 | -0.0026194 | 0.0015562 | -0.0046707 | 0.0056767 | 0.0037373 | -0.0012382 | a12 | -0.0043391 | -0.010863 | -0.005097 | 0.0019958 | -0.0083147 | -0.0074991 | -0.0092015 | -0.00021896 | -0.0029422 | -0.0050981 | -0.00061986 | 0.0013845 | 0.0031771 | -0.00082481 | -0.00028286 | 0.00084579 | 0.00021163 | -0.0007013 | -0.0057267 | -0.020036 | -0.003358
a13 | -0.023364 | 0.0011127 | 0.004218 | 0.0010951 | 0.004059 | 0.0010791 | 0.0035129 | 0.0014625 | 0.00017972 | -0.0015358 | 0.004754 | 0.001183 | -0.0026863 | 0.0017935 | -0.0015179 | 0.0021332 | 0.0021578 | 0.001451 | 0.00015516 | -0.0034348 | 0.0004548 | a13 | -0.026652 | 0.0063824 | 0.0072176 | 0.0024256 | 0.0051581 | 0.0056887 | 0.0038577 | 0.0023317 | 0.0031062 | 0.0022377 | 0.0029292 | 0.0020718 | 0.0022696 | 0.0039846 | -0.00029323 | 0.0015486 | 0.0013405 | 0.0025379 | 0.0025664 | -0.0049645 | 0.0020223
a14 | 0.013647 | -0.021484 | -0.019318 | -0.016417 | -0.015291 | -0.015489 | -0.017047 | -0.013974 | -0.014248 | -0.011196 | -0.016481 | -0.014813 | -0.013419 | -0.013213 | -0.015824 | -0.015654 | -0.015724 | -0.018006 | -0.014437 | -0.0095254 | -0.014573 | a14 | 0.013295 | -0.019411 | -0.014275 | -0.0088272 | -0.0091581 | -0.01137 | -0.012365 | -0.0073946 | -0.01058 | -0.0097351 | -0.0093004 | -0.0067403 | -0.0076766 | -0.0085332 | -0.011182 | -0.010469 | -0.0093418 | -0.0098479 | -0.010033 | -0.0031453 | -0.0094009
a15 | 0.013396 | 0.0036105 | -0.00060155 | -0.0021021 | -0.0023209 | -0.0045475 | -0.0088302 | -0.009484 | -0.0049132 | -0.0042864 | -0.007429 | -0.0011969 | -0.0054555 | -0.015248 | -0.0078564 | -0.0060904 | -0.0024276 | -0.0048643 | -0.0050928 | -0.010161 | -0.0045568 | a15 | 0.0048109 | -0.0029724 | -0.0069007 | -0.0030584 | -0.0032252 | -0.001894 | -0.0091484 | -0.0053611 | -0.0049557 | -0.0037639 | -0.0086 | -0.0035949 | -0.0034588 | -0.0048721 | 0.0024561 | 0.004024 | 0.0017694 | -0.0035423 | 0.0012663 | -0.009164 | -0.0030521
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00033683 | -0.00022349 | -0.00011924 | -6.6428e-05 | -3.1778e-05 | 0 | 0.0001274 | 0.00013311 | -1.0106e-05 | -1.6349e-05 | 9.1308e-05 | -8.1641e-05 | 1.1501e-05 | 0.00025759 | 2.5567e-05 | 1.0022e-05 | -3.0231e-05 | -4.4884e-05 | 0 | 0.00017622 | 0.00010281 | Dx (m) | -8.6693e-05 | -2.9959e-05 | 7.0951e-05 | -2.612e-05 | 0 | -2.6965e-05 | 0.00015018 | 0.00010296 | 3.9705e-05 | -2.3162e-05 | 0.00011541 | 2.5187e-05 | 5.3278e-05 | 8.1722e-05 | -0.00017439 | -0.00018525 | -0.00012616 | 2.6831e-05 | -0.00013366 | 0.00018269 | 0.00011399
Dy (m) | -0.00013499 | 7.6572e-05 | 0.00010409 | -2.2148e-05 | 0 | 0 | 2.0067e-05 | 2.9589e-05 | 1.841e-05 | -7.702e-05 | 5.618e-05 | 0 | 0 | 3.3386e-05 | -9.6937e-05 | 0 | -5.1441e-05 | 6.2208e-05 | -0.00011249 | 0.00010877 | 0.00072798 | Dy (m) | 0.00010488 | 0.00017062 | 7.3332e-05 | 0 | 0 | 0 | 8.1196e-05 | -8.5658e-05 | -1.5947e-05 | -2.1773e-05 | 0 | -0.00011949 | -6.7559e-05 | -6.2496e-05 | 0 | -4.6693e-05 | -7.1294e-05 | -5.9322e-05 | 0 | 0.00036212 | 0.00064939
Magnetic length (m) |  | 14.4223 | [19 empty] | Magnetic length (m) |  | 14.4204
Transfer funct. (mT/KA) |  | 596.1729 | [19 empty] | Transfer funct. (mT/KA) |  | 596.1794
Coil positioning (mm) |  | 0.018834 | [19 empty] | Coil positioning (mm) |  | 0.016587
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000075 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000075 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 350.9064 | 596.4561 | 596.2183 | 596.1507 | 596.0879 | 596.1229 | 596.1438 | 596.149 | 596.0961 | 596.1579 | 596.0884 | 596.1039 | 596.0638 | 596.1024 | 596.1054 | 596.1726 | 596.2155 | 596.1494 | 596.5272 | 367.9582 | 85.8631 | C1/i (mT/KA) | 347.9524 | 596.4699 | 596.1357 | 596.1462 | 596.1046 | 596.1522 | 596.2091 | 596.1682 | 596.1054 | 596.1494 | 596.0873 | 596.0939 | 596.1545 | 596.1289 | 596.1512 | 596.0944 | 596.1828 | 596.1539 | 596.5418 | 360.6786 | 85.7867
Angle (mrad) | 3.8904 | -0.95757 | -0.43382 | -0.6591 | -0.38632 | -0.15506 | 0.43637 | 1.0616 | 0.35503 | -0.15707 | -0.085016 | 0.19735 | 0.30831 | 0.27972 | -0.13885 | -0.51634 | -0.55318 | -0.133 | -0.3101 | -0.71602 | -14.6953 | Angle (mrad) | 3.7923 | -1.0322 | -0.38556 | -0.66582 | -0.62438 | -0.25467 | 0.46898 | 1.0484 | 0.31292 | -0.12207 | -0.10795 | -0.061159 | 0.22732 | 0.49587 | -0.15404 | -0.47525 | -0.34102 | -0.004294 | -0.1329 | -0.66945 | -33.7026
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 0.96052 | 0.70179 | 0.86818 | 0.21402 | 0.59831 | 0.94735 | 0.84229 | 1.2612 | 0.67934 | -0.56039 | 0.55989 | 0.81401 | 1.1192 | 0.79002 | 0.1077 | 0.73825 | -0.042429 | 0.90791 | 0.49576 | 0.19374 | 0.6106 | b2 | 1.599 | 0.60025 | 0.1817 | 0.37762 | 0.22388 | 0.1231 | 0.40133 | 0.71793 | 0.10315 | -0.23543 | 0.41395 | 0.1789 | 0.4937 | 0.05998 | -0.46133 | -0.21328 | -0.017123 | 0.096406 | -0.065497 | 1.5995 | 0.25433
b3 | 32.7919 | -4.0613 | -3.4475 | -3.6019 | -3.5326 | -5.3452 | -4.1483 | -2.9929 | -2.6782 | -4.2814 | -5.2352 | -4.0019 | -3.9485 | -6.7933 | -5.5711 | -4.6372 | -3.7443 | -3.4887 | -2.7747 | -2.2636 | -2.9356 | b3 | 33.6779 | -3.6734 | -3.4156 | -4.2133 | -3.8157 | -4.748 | -3.6429 | -4.551 | -3.6111 | -3.3819 | -4.2859 | -3.8304 | -3.1176 | -5.0002 | -4.3987 | -3.3944 | -3.9448 | -4.3018 | -2.8332 | -3.8594 | -2.7537
b4 | -0.97516 | -0.11458 | 0.17238 | 0.085877 | 0.0020324 | 0.14666 | 0.069828 | 0.29641 | 0.095792 | -0.016535 | 0.040729 | -0.0075303 | -0.013346 | -0.073833 | 0.096835 | 0.084314 | -0.06699 | 0.027797 | -0.0069976 | 0.12348 | 0.016713 | b4 | -1.3872 | -0.18517 | -0.099553 | -0.040006 | -0.018825 | -0.24224 | -0.23827 | -0.29777 | -0.26356 | -0.18592 | -0.15834 | -0.23464 | -0.2885 | -0.42734 | -0.24926 | -0.36497 | -0.37778 | -0.36031 | -0.10628 | 0.036873 | -0.25671
b5 | -6.0724 | -0.44039 | -0.58019 | -0.26301 | -0.27026 | -0.1539 | -0.16582 | -0.15621 | -0.28179 | -0.017634 | -0.0050925 | -0.072987 | -0.2182 | -0.049364 | 0.058478 | -0.42767 | -0.37629 | -0.38753 | -0.27113 | -3.4357 | -0.5089 | b5 | -6.1453 | -0.5857 | -0.42248 | -0.13394 | 0.076638 | 0.21049 | 0.28393 | 0.15541 | 0.1322 | -0.029075 | 0.14453 | -0.063111 | -0.25054 | -0.079107 | -0.00086768 | -0.40925 | -0.24249 | -0.26683 | -0.49013 | -3.2283 | -0.39139
b6 | 0.13351 | 0.063316 | 0.080556 | 0.044375 | 0.0061567 | -0.036952 | -0.0064989 | -0.079278 | 0.070191 | 0.1981 | 0.055691 | 0.1542 | 0.095475 | -0.033219 | 0.18362 | 0.175 | 0.13813 | 0.060457 | 0.1173 | 0.022074 | 0.071794 | b6 | -0.20486 | -0.0093148 | -0.077161 | -0.079665 | 0.047809 | 0.030929 | -0.04616 | 0.016862 | -0.017079 | -0.0057918 | -0.098936 | -0.034573 | -0.16825 | -0.12246 | 0.089617 | 0.057834 | 0.051182 | -0.066443 | 0.048188 | -0.32053 | -0.036314
b7 | 2.6222 | 0.92039 | 1.0531 | 1.1824 | 1.2062 | 1.2304 | 1.2174 | 1.1842 | 1.1617 | 1.2197 | 1.2378 | 1.1644 | 1.1625 | 1.1759 | 1.0748 | 1.2216 | 1.1793 | 1.1591 | 1.0949 | 0.86186 | 1.1934 | b7 | 2.6037 | 0.82195 | 0.99421 | 1.1513 | 1.1618 | 1.2218 | 1.1791 | 1.2183 | 1.154 | 1.133 | 1.1492 | 1.1346 | 1.1128 | 1.124 | 1.1268 | 1.0915 | 1.0852 | 1.1507 | 1.0433 | 0.77612 | 1.1487
b8 | 0.027688 | 0.063021 | 0.034195 | 0.036441 | 0.037423 | 0.041028 | -0.0051053 | -0.02977 | 0.043838 | 0.011876 | 0.015189 | 0.02588 | -0.0024734 | -0.056412 | -0.076593 | -0.0093033 | 0.0092754 | -0.015677 | -0.018344 | -0.047534 | 0.0047626 | b8 | -0.086617 | 0.0073846 | -0.020646 | 0.01223 | -0.045456 | 0.0050984 | -0.042429 | -0.044125 | -0.041545 | -0.030734 | -0.046624 | -0.01142 | 0.0011444 | -0.016997 | 0.0097314 | 0.023758 | 0.0061949 | -0.021672 | 0.039395 | 0.019642 | -0.013306
b9 | 0.29228 | 0.54151 | 0.51316 | 0.50755 | 0.52718 | 0.49559 | 0.50231 | 0.50722 | 0.50724 | 0.49197 | 0.51223 | 0.50968 | 0.51613 | 0.50581 | 0.49504 | 0.50643 | 0.5051 | 0.53082 | 0.51835 | 0.45852 | 0.50237 | b9 | 0.28809 | 0.52478 | 0.50926 | 0.50839 | 0.53374 | 0.50826 | 0.50811 | 0.50158 | 0.51803 | 0.52842 | 0.51212 | 0.50664 | 0.51603 | 0.51079 | 0.51093 | 0.52741 | 0.4999 | 0.52034 | 0.52672 | 0.47152 | 0.5068
b10 | 0.10365 | 0.095723 | 0.048902 | 0.030023 | 0.013821 | -0.00025639 | -0.056336 | -0.059277 | 0.0039779 | 0.009329 | -0.041908 | 0.036013 | -0.0049741 | -0.11681 | -0.0091854 | -0.0044566 | 0.014285 | 0.018588 | 0.004146 | -0.069142 | 0 | b10 | 0.040517 | 0.010604 | -0.031852 | 0.01179 | -0.0042179 | 0.012152 | -0.067105 | -0.046924 | -0.017981 | 0.010401 | -0.052041 | -0.011147 | -0.024258 | -0.037577 | 0.079021 | 0.082547 | 0.05635 | -0.012277 | 0.059817 | -0.067756 | 0
b11 | 0.59019 | 0.75135 | 0.74101 | 0.75218 | 0.74884 | 0.75316 | 0.74583 | 0.75014 | 0.75031 | 0.74937 | 0.75592 | 0.75127 | 0.75021 | 0.76682 | 0.7695 | 0.75188 | 0.75294 | 0.75111 | 0.74703 | 0.65251 | 0.74399 | b11 | 0.58557 | 0.75934 | 0.75669 | 0.76509 | 0.76531 | 0.76819 | 0.76475 | 0.76184 | 0.76647 | 0.75831 | 0.7668 | 0.75871 | 0.76568 | 0.77081 | 0.77064 | 0.75918 | 0.75847 | 0.76394 | 0.76099 | 0.66995 | 0.75504
b12 | 0.01914 | 0.013916 | 0.002849 | 0.0002815 | -0.0012934 | -0.0014109 | -0.0058092 | -0.0084929 | -0.0028206 | -0.0032589 | -0.0042591 | 0.0068525 | 0.0001346 | -0.0099577 | -0.0022968 | 0.001071 | 0.0063571 | 0.00026812 | -0.001627 | -0.0065148 | -0.00011679 | b12 | 0.0033677 | 0.0014092 | -0.0053136 | 0.00019488 | -0.0025159 | -0.001634 | -0.0072785 | -0.0039561 | 0.001755 | 0.00047163 | -0.0053128 | 0.00073265 | 0.0007393 | -0.00053915 | 0.0084068 | 0.0096878 | 0.0093756 | 0.0019871 | 0.0062757 | 0.0018915 | 0.00091716
b13 | 0.068796 | 0.063528 | 0.067723 | 0.06709 | 0.067078 | 0.064199 | 0.065466 | 0.065564 | 0.065044 | 0.063947 | 0.06272 | 0.06276 | 0.06571 | 0.061453 | 0.060358 | 0.065807 | 0.065333 | 0.06676 | 0.06584 | 0.049023 | 0.064415 | b13 | 0.079928 | 0.064603 | 0.066341 | 0.065881 | 0.064925 | 0.063905 | 0.065236 | 0.065763 | 0.062332 | 0.063128 | 0.065082 | 0.068118 | 0.071304 | 0.067304 | 0.0654 | 0.06835 | 0.065791 | 0.068987 | 0.069004 | 0.04607 | 0.065975
b14 | 0.0018038 | 0.0031107 | 0.00033835 | -0.0020567 | -0.0022062 | -0.0035562 | -0.0067398 | -0.006548 | -0.0037786 | -0.0038442 | -0.0053374 | -0.00051776 | -0.0032424 | -0.0094841 | -0.0052435 | -0.0044518 | -0.0015488 | -0.0029884 | -0.0018901 | -0.0043604 | -0.0032078 | b14 | 0.004494 | -0.0016589 | -0.004257 | -0.0016809 | -0.0032108 | -0.00079402 | -0.0057084 | -0.0027398 | -0.0025247 | -0.001309 | -0.0054503 | -0.0018288 | -0.0029636 | -0.0033235 | 0.0032417 | 0.0038079 | 0.0016432 | -0.0028768 | 0.00059288 | -0.0058241 | -0.0016645
b15 | -0.009771 | 0.041042 | 0.038177 | 0.03392 | 0.0329 | 0.032163 | 0.032447 | 0.031024 | 0.033667 | 0.031372 | 0.033304 | 0.032538 | 0.032546 | 0.029857 | 0.031936 | 0.032429 | 0.031809 | 0.035232 | 0.030811 | 0.022279 | 0.03151 | b15 | -0.0060801 | 0.040327 | 0.033885 | 0.030065 | 0.030531 | 0.030789 | 0.032397 | 0.025918 | 0.029856 | 0.030673 | 0.03265 | 0.028575 | 0.029748 | 0.029236 | 0.032639 | 0.031432 | 0.030898 | 0.03034 | 0.033143 | 0.028019 | 0.030044
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 38.9046 | -9.5757 | -4.3382 | -6.591 | -3.8633 | -1.5506 | 4.3637 | 10.6156 | 3.5503 | -1.5707 | -0.85016 | 1.9735 | 3.0831 | 2.7972 | -1.3885 | -5.1634 | -5.5318 | -1.33 | -3.101 | -7.1602 | 0 | a1 | 37.9233 | -10.3216 | -3.8556 | -6.6582 | -6.2438 | -2.5467 | 4.6898 | 10.4839 | 3.1292 | -1.2207 | -1.0795 | -0.61159 | 2.2732 | 4.9587 | -1.5404 | -4.7525 | -3.4102 | -0.042943 | -1.329 | -6.6945 | 0
a2 | 10.2777 | 0.093547 | 1.3896 | 0.65468 | 2.4139 | 1.635 | 2.034 | 1.577 | 1.745 | 0.22462 | 1.2246 | 1.6214 | 2.0471 | 1.808 | 0.61674 | 1.7106 | 1.5251 | 1.7526 | 1.2274 | 2.062 | 1.6985 | a2 | 16.6656 | 0.33144 | 0.18601 | 0.81301 | 1.979 | 0.67094 | 0.61029 | 0.87071 | 0.93859 | 0.89016 | 0.86983 | 0.64755 | 0.039965 | -0.092125 | 0.10407 | -0.053437 | 0.89027 | 0.68437 | 0.041015 | 5.2835 | 1.2165
a3 | 4.5964 | 1.5887 | 1.6982 | 0.88932 | 1.1359 | 0.83779 | 0.74268 | 0.79192 | 0.48223 | 1.0621 | 1.5811 | 1.3032 | 1.0464 | 0.82783 | 0.66189 | 0.70396 | 1.1123 | 1.0223 | 0.93875 | -0.64048 | 1.0797 | a3 | -0.011045 | 1.4181 | 1.4614 | 0.9253 | 1.2348 | 0.87304 | 0.78206 | 0.91727 | 0.79913 | 0.88239 | 1.0312 | 0.72953 | 0.66322 | 0.88313 | 0.82541 | 0.97715 | 1.1036 | 1.2046 | 1.0296 | -0.66246 | 0.90334
a4 | 1.5559 | 0.24011 | -0.14815 | 0.55309 | 0.27568 | 0.15809 | -0.026885 | 0.12337 | 0.24244 | 0.56412 | 0.25011 | 0.0090329 | 0.18949 | 0.055888 | 0.2827 | -0.30351 | -0.17795 | -0.35351 | 0.060447 | -0.041655 | 0.15019 | a4 | 4.2356 | -0.036736 | 0.26249 | 0.31308 | 0.11291 | 0.42393 | 0.37978 | 0.4139 | 0.44573 | 0.77599 | 0.87942 | 0.60932 | 0.40627 | 0.28374 | 0.35205 | 0.62195 | 0.24722 | 0.28776 | 0.53728 | 2.1175 | 0.57684
a5 | 3.8604 | 0.35657 | 0.32225 | 0.33811 | 0.44198 | 0.32319 | 0.14127 | 0.17381 | 0.11517 | 0.5 | 0.31485 | 0.11256 | 0.082013 | 0.3493 | 0.0058204 | 0.20915 | 0.31782 | 0.15369 | 0.29283 | -0.10005 | 0.35202 | a5 | 2.2886 | 0.10023 | 0.22031 | 0.33352 | 0.39005 | 0.32171 | 0.2799 | 0.16003 | 0.41956 | 0.31687 | 0.16521 | 0.18935 | 0.13086 | 0.31002 | 0.1407 | 0.25318 | 0.036796 | 0.094665 | 0.14737 | -0.62291 | 0.25895
a6 | 0.5064 | 0.10586 | -0.013338 | 0.18654 | 0.087513 | 0.12308 | 0.053533 | 0.056676 | 0.075711 | 0.10121 | 0.13138 | 0.0052228 | 0.052438 | 0.010694 | 0.21547 | 0.14585 | 0.067895 | 0.032392 | 0.19534 | 0.19793 | 0.10693 | a6 | 0.68452 | -0.097022 | 0.025575 | 0.11675 | -0.12082 | -0.13679 | -0.14485 | 0.029127 | 0.024056 | 0.052431 | 0.15464 | 0.041841 | -0.060863 | 0.0035287 | 0.053731 | 0.031246 | 0.012156 | 0.011117 | -0.11606 | -0.32403 | 0.0043304
a7 | 1.7844 | -0.050071 | 0.002379 | -0.082032 | 0.053836 | 0.030346 | 0.055966 | 0.020052 | -0.011463 | -0.10467 | 0.030885 | 0.034378 | -0.040675 | -0.036635 | 0.041427 | 0.092625 | 0.02734 | 0.040955 | 0.026635 | 0.037419 | 0.062722 | a7 | 1.5792 | -0.027329 | 0.031892 | -0.024567 | 0.0015116 | 0.112 | 0.098486 | 0.041792 | 0.019516 | -0.014348 | 0.10093 | -0.024082 | 0.011821 | 0.099996 | 0.089026 | 0.066693 | 0.08864 | 0.02094 | 0.071858 | 0.1911 | 0.09391
a8 | 0.016239 | -0.044669 | -0.03326 | -0.03133 | -0.0017088 | 0.0062358 | 0.0023881 | -0.00567 | 0.016365 | 0.023391 | -0.022622 | 0.00060028 | 0.020875 | 0.025716 | 0.016558 | 0.035257 | -0.0072433 | -0.047282 | 0.0072555 | 0.0060998 | -0.0013457 | a8 | 0.025918 | -0.020996 | 0.00016805 | 0.10171 | 0.092089 | 0.0241 | 0.0037124 | 0.063432 | 0.048869 | 0.067132 | 0.086957 | 0.095117 | 0.051589 | 0.062077 | 0.054226 | 0.039669 | 0.063197 | 0.03702 | 0.015738 | -0.05295 | 0.045278
a9 | -0.17302 | 0.025332 | 0.031302 | 0.0014333 | 0.012968 | 0.010639 | 0.0055838 | 0.0026144 | 0.0070161 | -0.0015264 | 0.033928 | 0.0029876 | -0.026464 | 0.0024197 | -0.025867 | 0.018901 | 0.01259 | -0.0066147 | -0.018929 | -0.087521 | -0.0035165 | a9 | -0.26772 | 0.032024 | 0.042205 | 0.028349 | 0.013582 | 0.018265 | 0.012381 | 0.030098 | 0.029719 | 0.016478 | 0.0074699 | 0.0081298 | 0.019118 | 0.025573 | -0.012681 | 0.0070175 | 0.020758 | 0.0070803 | 0.0072086 | -0.077592 | 0.0057116
a10 | 0.079262 | -0.04244 | -0.048914 | 0.0079117 | -0.0037509 | 0.0023076 | -0.0059304 | -0.010526 | -0.0083906 | 0.033501 | -0.022853 | -0.0030314 | 0.0019005 | -0.0097048 | 0.044517 | -0.00028405 | 0.022256 | -0.028387 | 0.049359 | -0.039214 | 0 | a10 | -0.027318 | -0.076748 | -0.032378 | 0.003882 | 0.00090618 | 0.0027628 | -0.037359 | 0.03743 | 0.0069926 | 0.0096381 | -0.0014091 | 0.053349 | 0.030227 | 0.027652 | 0.00053753 | 0.02139 | 0.031643 | 0.026558 | -0.0010971 | -0.14543 | 0
a11 | 0.16479 | -0.066014 | -0.050583 | -0.048253 | -0.045994 | -0.044484 | -0.038195 | -0.032005 | -0.044572 | -0.046931 | -0.039444 | -0.040432 | -0.048144 | -0.034642 | -0.042017 | -0.036501 | -0.029501 | -0.034067 | -0.037665 | -0.024324 | -0.035274 | a11 | 0.17258 | -0.027873 | -0.0060944 | -0.0066975 | -0.0098721 | 0.0095286 | 0.0099609 | 0.015683 | 0.0085161 | -0.0054498 | 0.0033053 | -0.0012852 | 0.0065177 | 0.014295 | 0.00026407 | 0.0051406 | -0.0019125 | 0.0063675 | 0.00051104 | 0.021454 | 0.0070138
a12 | 0.0019201 | -0.0025692 | -0.0068748 | 0.0040668 | -0.00070898 | -0.00021464 | -0.0028499 | -0.0051965 | -0.0028161 | 0.00097301 | -0.004388 | -0.0040677 | -0.0016646 | -0.0074224 | 0.0065694 | -0.0026194 | 0.0015562 | -0.0046707 | 0.0056767 | 0.0037373 | -0.0012382 | a12 | -0.0043391 | -0.010863 | -0.005097 | 0.0019958 | -0.0083147 | -0.0074991 | -0.0092015 | -0.00021896 | -0.0029422 | -0.0050981 | -0.00061986 | 0.0013845 | 0.0031771 | -0.00082481 | -0.00028286 | 0.00084579 | 0.00021163 | -0.0007013 | -0.0057267 | -0.020036 | -0.003358
a13 | -0.023364 | 0.0011127 | 0.004218 | 0.0010951 | 0.004059 | 0.0010791 | 0.0035129 | 0.0014625 | 0.00017972 | -0.0015358 | 0.004754 | 0.001183 | -0.0026863 | 0.0017935 | -0.0015179 | 0.0021332 | 0.0021578 | 0.001451 | 0.00015516 | -0.0034348 | 0.0004548 | a13 | -0.026652 | 0.0063824 | 0.0072176 | 0.0024256 | 0.0051581 | 0.0056887 | 0.0038577 | 0.0023317 | 0.0031062 | 0.0022377 | 0.0029292 | 0.0020718 | 0.0022696 | 0.0039846 | -0.00029323 | 0.0015486 | 0.0013405 | 0.0025379 | 0.0025664 | -0.0049645 | 0.0020223
a14 | 0.013647 | -0.021484 | -0.019318 | -0.016417 | -0.015291 | -0.015489 | -0.017047 | -0.013974 | -0.014248 | -0.011196 | -0.016481 | -0.014813 | -0.013419 | -0.013213 | -0.015824 | -0.015654 | -0.015724 | -0.018006 | -0.014437 | -0.0095254 | -0.014573 | a14 | 0.013295 | -0.019411 | -0.014275 | -0.0088272 | -0.0091581 | -0.01137 | -0.012365 | -0.0073946 | -0.01058 | -0.0097351 | -0.0093004 | -0.0067403 | -0.0076766 | -0.0085332 | -0.011182 | -0.010469 | -0.0093418 | -0.0098479 | -0.010033 | -0.0031453 | -0.0094009
a15 | 0.013396 | 0.0036105 | -0.00060155 | -0.0021021 | -0.0023209 | -0.0045475 | -0.0088302 | -0.009484 | -0.0049132 | -0.0042864 | -0.007429 | -0.0011969 | -0.0054555 | -0.015248 | -0.0078564 | -0.0060904 | -0.0024276 | -0.0048643 | -0.0050928 | -0.010161 | -0.0045568 | a15 | 0.0048109 | -0.0029724 | -0.0069007 | -0.0030584 | -0.0032252 | -0.001894 | -0.0091484 | -0.0053611 | -0.0049557 | -0.0037639 | -0.0086 | -0.0035949 | -0.0034588 | -0.0048721 | 0.0024561 | 0.004024 | 0.0017694 | -0.0035423 | 0.0012663 | -0.009164 | -0.0030521
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.33683 | 0.22349 | 0.11924 | 0.066428 | 0.031778 | 0 | -0.1274 | -0.13311 | 0.010106 | 0.016349 | -0.091308 | 0.081641 | -0.011501 | -0.25759 | -0.025567 | -0.010021 | 0.030231 | 0.044884 | 0 | -0.17622 |  | Dx (mm) | 0.086693 | 0.029959 | -0.070951 | 0.02612 | 0 | 0.026965 | -0.15018 | -0.10296 | -0.039705 | 0.023162 | -0.11541 | -0.025187 | -0.053278 | -0.081722 | 0.17439 | 0.18525 | 0.12616 | -0.026831 | 0.13366 | -0.18269
Dy (mm) | 0.13499 | -0.076572 | -0.10409 | 0.022148 | 0 | 0 | -0.020067 | -0.029589 | -0.01841 | 0.07702 | -0.05618 | 0 | 0 | -0.033386 | 0.096937 | 0 | 0.051441 | -0.062208 | 0.11249 | -0.10877 |  | Dy (mm) | -0.10488 | -0.17062 | -0.073332 | 0 | 0 | 0 | -0.081196 | 0.085658 | 0.015947 | 0.021773 | 0 | 0.11949 | 0.067559 | 0.062496 | 0 | 0.046693 | 0.071294 | 0.059323 | 0 | -0.36212
Mag. len. (m) | 14.4223 | [20 empty] | Mag. len. (m) | 14.4204
TF (mT/KA) | 596.1729 | [20 empty] | TF (mT/KA) | 596.1794
Coil pos. (mm) | 0.018834 | [20 empty] | Coil pos. (mm) | 0.016587
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0078625 | [20 empty] | Twist (m2 rad) | 0.0040678

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.6106 | 0.43689 | 0.60547 | 0.43358 | 1.6985 | 2.0702 | 1.4027 | 0.6333 | 0.25433 | 0.52685 | 0.16615 | 0.27224 | 1.2165 | 3.7255 | 0.5851 | 0.51418 |  | 0
3 | -2.9356 | 8.3046 | -4.1258 | 1.0607 | 1.0797 | 0.95405 | 1.0237 | 0.33812 | -2.7537 | 8.4197 | -3.8956 | 0.58171 | 0.90334 | 0.47098 | 0.98663 | 0.22293 |  | 0
4 | 0.016713 | 0.24806 | 0.047545 | 0.096407 | 0.15019 | 0.40525 | 0.10939 | 0.25393 | -0.25671 | 0.29007 | -0.22663 | 0.1139 | 0.57684 | 0.94291 | 0.40625 | 0.2246 |  | 0
5 | -0.5089 | 1.4669 | -0.22694 | 0.17441 | 0.35202 | 0.82437 | 0.25209 | 0.13278 | -0.39139 | 1.5061 | -0.11024 | 0.2577 | 0.25895 | 0.51875 | 0.22296 | 0.10928 |  | 0
6 | 0.071794 | 0.078343 | 0.073767 | 0.072163 | 0.10693 | 0.11357 | 0.091716 | 0.055647 | -0.036314 | 0.10177 | -0.021577 | 0.052105 | 0.0043304 | 0.19192 | -0.013021 | 0.081544 |  | 0
7 | 1.1934 | 0.34558 | 1.1574 | 0.079618 | 0.062722 | 0.40007 | 0.0081268 | 0.051288 | 1.1487 | 0.35604 | 1.1142 | 0.091093 | 0.09391 | 0.34662 | 0.042549 | 0.048935 |  | 0
8 | 0.0047626 | 0.036925 | 0.0056618 | 0.027145 | -0.0013457 | 0.023766 | -0.0016437 | 0.020936 | -0.013306 | 0.03105 | -0.012969 | 0.016689 | 0.045278 | 0.040151 | 0.046472 | 0.026498 |  | 0
9 | 0.50237 | 0.051048 | 0.50905 | 0.012509 | -0.0035165 | 0.04677 | 0.0052747 | 0.018187 | 0.5068 | 0.052018 | 0.51399 | 0.010521 | 0.0057116 | 0.067225 | 0.017149 | 0.013455 |  | 0
10 | 0 | 0.052486 | 6.9333e-05 | 0.00040368 | 0 | 0.032255 | -2.3985e-05 | 0.0010179 | 0 | 0.045648 | -0.00023908 | 0.00099318 | 0 | 0.044869 | 0.00031434 | 0.0012055 |  | 0
11 | 0.74399 | 0.041987 | 0.75196 | 0.006554 | -0.035274 | 0.046938 | -0.042223 | 0.0081419 | 0.75504 | 0.04413 | 0.76324 | 0.0043589 | 0.0070138 | 0.039576 | 0.0010754 | 0.010207 |  | 0
12 | -0.00011679 | 0.0070788 | -0.00053732 | 0.0026421 | -0.0012382 | 0.004036 | -0.001379 | 0.0021447 | 0.00091716 | 0.0047682 | 0.00060985 | 0.0022167 | -0.003358 | 0.00555 | -0.0033458 | 0.00272 |  | 0
13 | 0.064415 | 0.0041497 | 0.064667 | 0.002501 | 0.0004548 | 0.005895 | 0.0014545 | 0.0023415 | 0.065975 | 0.0059819 | 0.066048 | 0.0021589 | 0.0020223 | 0.0070551 | 0.0032222 | 0.0016357 |  | 0
14 | -0.0032078 | 0.0029649 | -0.0034623 | 0.00084467 | -0.014573 | 0.0069904 | -0.015867 | 0.0017208 | -0.0016645 | 0.0030241 | -0.0018483 | 0.00065676 | -0.0094009 | 0.0060932 | -0.010791 | 0.0012904 |  | 0
15 | 0.03151 | 0.010108 | 0.033177 | 0.0026879 | -0.0045568 | 0.0058094 | -0.004952 | 0.0041023 | 0.030044 | 0.0087946 | 0.031283 | 0.0029244 | -0.0030521 | 0.0040979 | -0.0031018 | 0.0035963 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.64625 | -0.98312
3 | [12 empty] | 3 | 9 | -0.51271 | -1.0213
4 | [12 empty] | 4 | 16 | -1.8549 | -1.8329
5 | [12 empty] | 5 | 25 | -1.8827 | -1.7849
6 | [12 empty] | 6 | 36 | -2.7588 | -2.7306
7 | [12 empty] | 7 | 49 | -2.7504 | -2.7066
8 | [12 empty] | 8 | 64 | -3.7364 | -3.8618
9 | [12 empty] | 9 | 81 | -4.1942 | -4.4314
10 | [12 empty] | sigma (mm) |  | 0.018834 | 0.016587
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -0.26593 | 0.79425 | 0.89574 | 0.24449 | 0.61152 | 0.94735 | 0.77836 | 1.2115 | 0.68148 | -0.54254 | 0.4932 | 0.85244 | 1.1139 | 0.58091 | 0.09849 | 0.73278 | -0.022381 | 0.91885 | 0.50818 | 0.15501 | 0.6106
3 | 32.8869 | -4.06 | -3.4484 | -3.6008 | -3.5326 | -5.3452 | -4.1467 | -2.9865 | -2.6792 | -4.2737 | -5.237 | -4.0018 | -3.9485 | -6.797 | -5.5659 | -4.6371 | -3.7456 | -3.4851 | -2.7735 | -2.259 | -2.9356
4 | -0.37129 | -0.097847 | 0.18076 | 0.09175 | 0.0040531 | 0.14666 | 0.064191 | 0.29031 | 0.095963 | -0.0074064 | 0.036458 | -0.0061283 | -0.013936 | -0.079568 | 0.09732 | 0.083306 | -0.060467 | 0.029641 | 0.00075327 | -0.016414 | 0.016713
5 | -6.0655 | -0.44693 | -0.5826 | -0.26267 | -0.27032 | -0.1539 | -0.16638 | -0.1598 | -0.28241 | -0.016294 | -0.0057677 | -0.076689 | -0.21788 | -0.051985 | 0.066002 | -0.42716 | -0.37649 | -0.38892 | -0.26467 | -3.4409 | -0.5089
6 | -0.093214 | -0.0079283 | 0.03615 | 0.016012 | -0.0073713 | -0.036952 | 0.047846 | -0.023856 | 0.066122 | 0.18821 | 0.094969 | 0.12064 | 0.10019 | 0.074116 | 0.19473 | 0.17932 | 0.12605 | 0.041195 | 0.11836 | 0.07424 | 0.071794
7 | 2.6193 | 0.916 | 1.0528 | 1.1811 | 1.2057 | 1.2304 | 1.2171 | 1.1826 | 1.1614 | 1.2203 | 1.2389 | 1.1635 | 1.1625 | 1.1696 | 1.0746 | 1.2216 | 1.179 | 1.1606 | 1.0953 | 0.85814 | 1.1934
8 | -0.029632 | 0.0051574 | 0.003867 | 0.02059 | 0.02954 | 0.041028 | 0.024958 | 0.0019653 | 0.041365 | 0.0080361 | 0.036302 | 0.006299 | 0.00031993 | 0.0048635 | -0.071816 | -0.006915 | 0.0023945 | -0.026695 | -0.019346 | -0.0050309 | 0.0047626
9 | 0.27946 | 0.5319 | 0.51277 | 0.50658 | 0.52694 | 0.49559 | 0.49858 | 0.50321 | 0.5073 | 0.49326 | 0.51089 | 0.50812 | 0.5161 | 0.49006 | 0.4972 | 0.50641 | 0.50547 | 0.53131 | 0.52129 | 0.45433 | 0.50237
10 | -0.00020656 | -7.8743e-05 | 2.4665e-05 | 2.6432e-06 | -0.00017719 | -0.00025639 | 9.2732e-06 | 1.3623e-05 | 4.4923e-08 | -3.8958e-06 | -3.8288e-06 | -6.5789e-05 | 0.00010108 | 5.521e-05 | -8.1673e-06 | -2.4272e-05 | 2.393e-06 | 3.7795e-06 | 0.0016536 | 5.14e-05 | 0
11 | 0.58619 | 0.74947 | 0.74125 | 0.75222 | 0.74886 | 0.75316 | 0.74539 | 0.74951 | 0.75036 | 0.74945 | 0.75583 | 0.75091 | 0.75021 | 0.76532 | 0.76988 | 0.75188 | 0.75287 | 0.75129 | 0.74745 | 0.65151 | 0.74399
12 | 0.0055643 | 0.038342 | -0.031611 | -0.028473 | -0.02798 | -0.014109 | 0.00028438 | -0.023632 | -0.032869 | -0.040803 | -0.0040508 | 0.032357 | 0.0066804 | 0.011739 | -0.013114 | 0.015366 | 0.050413 | -0.019108 | -0.016147 | -0.0015329 | -0.0011679 | ten times!
13 | 0.6974 | 0.64254 | 0.6923 | 0.66916 | 0.67132 | 0.64199 | 0.65071 | 0.65214 | 0.65273 | 0.63336 | 0.63055 | 0.62793 | 0.65681 | 0.59923 | 0.59082 | 0.65773 | 0.6475 | 0.67719 | 0.64598 | 0.49227 | 0.64415 | ten times!
14 | 0.060033 | -0.046706 | -0.03359 | -0.039507 | -0.030672 | -0.035562 | -0.031896 | -0.029161 | -0.039843 | -0.045385 | -0.024894 | -0.027054 | -0.029341 | -0.027314 | -0.051982 | -0.041842 | -0.024436 | -0.040415 | -0.023619 | -0.0021711 | -0.032078 | ten times!
15 | -0.09771 | 0.41042 | 0.38177 | 0.3392 | 0.329 | 0.32163 | 0.32447 | 0.31024 | 0.33667 | 0.31372 | 0.33304 | 0.32538 | 0.32546 | 0.29857 | 0.31936 | 0.32429 | 0.31809 | 0.35232 | 0.30811 | 0.22279 | 0.3151 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 9.5748 | 0.015191 | 1.3236 | 0.65711 | 2.4097 | 1.635 | 2.0353 | 1.5789 | 1.7386 | 0.26137 | 1.207 | 1.6089 | 2.0485 | 1.8064 | 0.68226 | 1.7115 | 1.5438 | 1.7217 | 1.2641 | 2.0198 | 1.6985
3 | 4.5271 | 1.5777 | 1.7045 | 0.8825 | 1.1344 | 0.83779 | 0.74233 | 0.79637 | 0.48211 | 1.0607 | 1.5855 | 1.3031 | 1.0467 | 0.82994 | 0.66151 | 0.70342 | 1.1138 | 1.0254 | 0.93888 | -0.6394 | 1.0797
4 | 1.4428 | 0.21342 | -0.1714 | 0.54918 | 0.27238 | 0.15809 | -0.023433 | 0.12773 | 0.24095 | 0.56252 | 0.25681 | 0.0068707 | 0.18971 | 0.076671 | 0.2814 | -0.30302 | -0.17565 | -0.36081 | 0.067623 | -0.13373 | 0.15019
5 | 3.8049 | 0.35104 | 0.32518 | 0.33418 | 0.44116 | 0.32319 | 0.14323 | 0.17534 | 0.11533 | 0.49503 | 0.3193 | 0.11243 | 0.082191 | 0.34978 | 0.0022056 | 0.20958 | 0.31512 | 0.15437 | 0.28895 | -0.089085 | 0.35202
6 | 0.16933 | 0.13469 | 0.025251 | 0.17922 | 0.086909 | 0.12308 | 0.064672 | 0.069985 | 0.0833 | 0.068654 | 0.15691 | 0.0042322 | 0.052273 | 0.021219 | 0.17907 | 0.14618 | 0.046192 | 0.057192 | 0.15186 | 0.23334 | 0.10693
7 | 1.7806 | -0.043974 | 0.0054777 | -0.081507 | 0.053858 | 0.030346 | 0.05605 | 0.019379 | -0.011199 | -0.10521 | 0.030386 | 0.034357 | -0.040576 | -0.034683 | 0.044659 | 0.092771 | 0.027233 | 0.041427 | 0.027484 | 0.035733 | 0.062722
8 | 0.025098 | -0.02782 | -0.0098794 | -0.036665 | -0.0019028 | 0.0062358 | 0.0074664 | 0.0015564 | 0.020726 | 0.0055717 | -0.007622 | 0.0004855 | 0.020732 | 0.033956 | -0.0063356 | 0.035346 | -0.01965 | -0.031603 | -0.020185 | 0.022312 | -0.0013457
9 | -0.19456 | 0.034234 | 0.037084 | 0.00080302 | 0.013031 | 0.010639 | 0.0045853 | 0.00094416 | 0.0070997 | -0.0021967 | 0.031577 | 0.0031186 | -0.026453 | -0.00096842 | -0.024793 | 0.018899 | 0.011845 | -0.005328 | -0.019176 | -0.095161 | -0.0035165
10 | -0.00025289 | 8.1262e-05 | 7.3798e-06 | -2.3363e-06 | -0.0028912 | 0.0023076 | 1.1109e-05 | 2.4725e-05 | -3.2363e-07 | 1.4961e-06 | 9.2098e-06 | -0.0010897 | 0.0015748 | 0.00010545 | 6.7715e-06 | -0.00049922 | -3.2605e-06 | -1.7974e-06 | -7.3778e-05 | 1.4292e-05 | 0
11 | 0.1627 | -0.064953 | -0.04986 | -0.048432 | -0.045979 | -0.044484 | -0.038505 | -0.032615 | -0.044587 | -0.046779 | -0.039858 | -0.040217 | -0.048157 | -0.036094 | -0.041764 | -0.036518 | -0.029743 | -0.033921 | -0.037547 | -0.024357 | -0.035274 | ten times!
12 | 0.0092004 | 0.0068903 | -0.022536 | 0.029666 | -0.0080003 | -0.0021464 | -0.016067 | -0.036897 | -0.019721 | -0.024859 | -0.015943 | -0.041359 | -0.016864 | -0.056481 | 0.02412 | -0.026043 | -0.0086223 | -0.017851 | 0.0044856 | 0.07074 | -0.012382 | ten times!
13 | -0.27066 | 0.049666 | 0.060063 | 0.01964 | 0.044306 | 0.010791 | 0.017487 | -0.0010804 | 0.0023664 | -0.011694 | 0.033739 | 0.021078 | -0.028043 | -0.010513 | -0.014386 | 0.020132 | 0.025822 | 0.019269 | 0.0031775 | -0.050811 | 0.004548 | ten times!
14 | 0.11018 | -0.19561 | -0.15986 | -0.16921 | -0.1523 | -0.15489 | -0.17438 | -0.14258 | -0.13696 | -0.13128 | -0.15499 | -0.14732 | -0.13471 | -0.15627 | -0.18539 | -0.15705 | -0.17012 | -0.16021 | -0.17291 | -0.090043 | -0.14573 | ten times!
15 | 0.13396 | 0.036105 | -0.0060155 | -0.021021 | -0.023209 | -0.045475 | -0.088302 | -0.09484 | -0.049132 | -0.042864 | -0.07429 | -0.011969 | -0.054555 | -0.15248 | -0.078564 | -0.060904 | -0.024276 | -0.048643 | -0.050928 | -0.10161 | -0.045568 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 1.2557 | 0.58473 | 0.14058 | 0.39057 | 0.22388 | 0.13816 | 0.32949 | 0.67205 | 0.087782 | -0.22395 | 0.35576 | 0.1778 | 0.47943 | 0.018399 | -0.37108 | -0.13393 | 0.050685 | 0.091234 | -0.020947 | 1.5448 | 0.25433
3 | 33.6207 | -3.6713 | -3.4203 | -4.2131 | -3.8157 | -4.7469 | -3.6547 | -4.5502 | -3.6117 | -3.3781 | -4.2891 | -3.8186 | -3.1155 | -5.0033 | -4.391 | -3.3773 | -3.9333 | -4.3005 | -2.8307 | -3.9935 | -2.7537
4 | -1.3183 | -0.18507 | -0.11041 | -0.039182 | -0.018825 | -0.24357 | -0.23359 | -0.29078 | -0.26075 | -0.18414 | -0.15442 | -0.22969 | -0.28956 | -0.4243 | -0.24922 | -0.34435 | -0.36996 | -0.36067 | -0.090865 | -0.048819 | -0.25671
5 | -6.1612 | -0.58075 | -0.42464 | -0.13333 | 0.076638 | 0.21024 | 0.28535 | 0.15665 | 0.13212 | -0.0287 | 0.14117 | -0.061897 | -0.25439 | -0.081985 | -0.0054643 | -0.41198 | -0.24413 | -0.26716 | -0.49202 | -3.211 | -0.39139
6 | -0.34298 | -0.01636 | -0.053089 | -0.090279 | 0.047809 | 0.019301 | 0.013519 | 0.062397 | -0.00079779 | -0.015164 | -0.052128 | -0.025503 | -0.14704 | -0.087836 | 0.020266 | -0.012431 | 0.005091 | -0.055108 | -0.0010285 | -0.29491 | -0.036314
7 | 2.6057 | 0.82333 | 0.9936 | 1.1512 | 1.1618 | 1.2217 | 1.1764 | 1.2187 | 1.1536 | 1.1339 | 1.147 | 1.1392 | 1.1142 | 1.125 | 1.1261 | 1.0904 | 1.0867 | 1.1513 | 1.0412 | 0.78549 | 1.1487
8 | -0.085157 | -0.0025852 | -0.0050983 | 0.0059813 | -0.045456 | -0.0013511 | -0.006992 | -0.018609 | -0.031643 | -0.036324 | -0.018812 | -0.0049581 | 0.01469 | 0.0033987 | -0.032199 | -0.022066 | -0.022788 | -0.014905 | 0.0062666 | 0.073401 | -0.013306
9 | 0.28774 | 0.53155 | 0.50932 | 0.50823 | 0.53374 | 0.50809 | 0.50438 | 0.50072 | 0.51772 | 0.52841 | 0.50894 | 0.50987 | 0.51643 | 0.51008 | 0.50364 | 0.51984 | 0.49734 | 0.521 | 0.52249 | 0.49284 | 0.5068
10 | 7.9068e-06 | 1.9899e-05 | -7.8736e-06 | 3.5068e-05 | -0.0042179 | -3.2567e-05 | -2.0993e-05 | 6.4138e-06 | 1.6687e-07 | -1.0481e-06 | 1.4053e-05 | 3.6951e-06 | -2.2912e-06 | 2.9497e-06 | -3.3307e-05 | -4.1102e-05 | -1.8547e-05 | 2.3081e-06 | -1.2389e-05 | -0.00033104 | 0
11 | 0.58568 | 0.76051 | 0.75669 | 0.76508 | 0.76531 | 0.76822 | 0.76453 | 0.76156 | 0.76649 | 0.75823 | 0.76641 | 0.75882 | 0.76585 | 0.77074 | 0.76969 | 0.75804 | 0.75772 | 0.76394 | 0.76044 | 0.67487 | 0.75504
12 | 0.0044965 | -0.0072572 | -0.023647 | -0.010198 | -0.025159 | -0.028503 | -0.0058381 | 0.0096446 | 0.035369 | -0.0052611 | -0.00011101 | 0.021185 | 0.035292 | 0.035191 | 0.0035619 | 0.0080093 | 0.035841 | 0.034 | -0.0023451 | 0.091014 | 0.0091716 | ten times!
13 | 0.78563 | 0.67174 | 0.6691 | 0.65915 | 0.64925 | 0.63921 | 0.65349 | 0.65063 | 0.62126 | 0.62989 | 0.64601 | 0.67467 | 0.70787 | 0.66688 | 0.64967 | 0.67437 | 0.65123 | 0.68481 | 0.68943 | 0.46127 | 0.65975 | ten times!
14 | 0.045126 | -0.022362 | -0.018604 | -0.023276 | -0.032108 | -0.014777 | -0.010899 | -0.0092033 | -0.016135 | -0.019616 | -0.023472 | -0.015898 | -0.018508 | -0.016067 | -0.014459 | -0.0083274 | -0.014631 | -0.023795 | -0.030551 | 0.011242 | -0.016645 | ten times!
15 | -0.060801 | 0.40327 | 0.33885 | 0.30065 | 0.30531 | 0.30789 | 0.32397 | 0.25918 | 0.29856 | 0.30673 | 0.3265 | 0.28575 | 0.29748 | 0.29236 | 0.32639 | 0.31432 | 0.30898 | 0.3034 | 0.33143 | 0.28019 | 0.30044 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 17.0812 | 0.2527 | 0.16874 | 0.81017 | 1.979 | 0.66817 | 0.58931 | 0.92769 | 0.9491 | 0.89642 | 0.88383 | 0.70356 | 0.068901 | -0.04687 | 0.087135 | -0.056087 | 0.90698 | 0.7182 | 0.024826 | 5.1048 | 1.2165
3 | -0.10152 | 1.4127 | 1.4634 | 0.92386 | 1.2348 | 0.87102 | 0.78871 | 0.92929 | 0.80299 | 0.87993 | 1.0491 | 0.73719 | 0.67048 | 0.89193 | 0.81457 | 0.95982 | 1.1028 | 1.2098 | 1.0169 | -0.59184 | 0.90334
4 | 4.0373 | -0.060957 | 0.25888 | 0.31103 | 0.11291 | 0.42189 | 0.3951 | 0.41464 | 0.44915 | 0.77441 | 0.88391 | 0.61221 | 0.41189 | 0.29086 | 0.34628 | 0.61541 | 0.25019 | 0.29209 | 0.53264 | 1.8156 | 0.57684
5 | 2.2648 | 0.10061 | 0.21918 | 0.33263 | 0.39005 | 0.32279 | 0.2724 | 0.16049 | 0.41992 | 0.31655 | 0.17046 | 0.19087 | 0.13325 | 0.31236 | 0.13794 | 0.25068 | 0.035272 | 0.095912 | 0.15193 | -0.67446 | 0.25895
6 | 0.73258 | -0.047235 | 0.052105 | 0.11698 | -0.12082 | -0.13786 | -0.10584 | -0.0061866 | 0.017835 | 0.043842 | 0.15875 | -0.0062252 | -0.087174 | -0.018378 | 0.048251 | 0.0088986 | -0.019098 | -0.012776 | -0.11945 | -0.21252 | 0.0043304
7 | 1.5745 | -0.026551 | 0.031273 | -0.025661 | 0.0015116 | 0.11173 | 0.097297 | 0.046037 | 0.020587 | -0.014713 | 0.10506 | -0.022534 | 0.012921 | 0.10252 | 0.085132 | 0.063211 | 0.085175 | 0.021879 | 0.070992 | 0.19004 | 0.09391
8 | 0.051059 | 0.020689 | 0.019152 | 0.10136 | 0.092089 | 0.023868 | 0.024002 | 0.044672 | 0.045537 | 0.061538 | 0.087363 | 0.066724 | 0.035663 | 0.048038 | 0.055267 | 0.027469 | 0.045193 | 0.022583 | 0.015284 | 0.02073 | 0.045278
9 | -0.26421 | 0.034199 | 0.039752 | 0.028295 | 0.013582 | 0.018226 | 0.0065265 | 0.034266 | 0.030017 | 0.01624 | 0.0073838 | 0.0095463 | 0.020838 | 0.028013 | -0.01273 | 0.0028791 | 0.016517 | 0.0078431 | 0.0072862 | -0.10465 | 0.0057116
10 | 7.1046e-06 | -4.324e-05 | 8.3208e-06 | 0.0039849 | 0.00090618 | 0.0026117 | 4.7177e-05 | -6.8589e-06 | 1.3572e-06 | 3.9019e-08 | -0.0011848 | 4.2528e-08 | 2.9622e-06 | 2.4982e-06 | 0.00051044 | -2.2307e-05 | -2.3573e-05 | 6.1162e-07 | -0.0011373 | -0.00041227 | 0
11 | 0.17305 | -0.027507 | -0.0065806 | -0.0067312 | -0.0098721 | 0.0096595 | 0.0086843 | 0.015888 | 0.0084224 | -0.0053801 | 0.003259 | -0.0013193 | 0.0065949 | 0.014273 | 0.00029599 | 0.0047466 | -0.0023623 | 0.006279 | 0.0010063 | 0.019528 | 0.0070138
12 | 0.032092 | -0.032172 | -0.013014 | 0.019511 | -0.083147 | -0.076073 | -0.050535 | -0.040259 | -0.035568 | -0.061049 | -0.0038124 | -0.043243 | -0.0013795 | -0.03564 | -0.0024677 | -0.016095 | -0.032187 | -0.03542 | -0.059689 | -0.088999 | -0.03358 | ten times!
13 | -0.27173 | 0.066107 | 0.062043 | 0.026019 | 0.051581 | 0.059231 | 0.020832 | 0.01929 | 0.028157 | 0.02432 | 0.021084 | 0.021091 | 0.0211 | 0.036102 | 0.011979 | 0.028956 | 0.021522 | 0.024664 | 0.035919 | -0.070167 | 0.020223 | ten times!
14 | 0.12427 | -0.13671 | -0.12632 | -0.087614 | -0.091581 | -0.11328 | -0.1133 | -0.096775 | -0.11135 | -0.10213 | -0.10118 | -0.096269 | -0.094834 | -0.10366 | -0.11534 | -0.12291 | -0.1134 | -0.11408 | -0.10173 | 0.038318 | -0.094009 | ten times!
15 | 0.048109 | -0.029724 | -0.069007 | -0.030584 | -0.032252 | -0.01894 | -0.091484 | -0.053611 | -0.049557 | -0.037639 | -0.086 | -0.035949 | -0.034588 | -0.048721 | 0.024561 | 0.04024 | 0.017694 | -0.035423 | 0.012663 | -0.09164 | -0.030521 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00047511 | 7.6236e-05 | -3.7154e-05 | -0.00014249 | -8.3784e-05 | -4.8727e-05 | -4.0005e-05 | -0.00012874 | -2.5077e-05 | -0.00014165 | -0.00011565 | -0.00018292 | -0.00011817 | -0.00011314 | -4.1934e-07 | 7.154e-05 | -3.9334e-05 | 0.00059437 |  | 596.1729
|  | 0.00048725 | -7.3319e-05 | -5.5707e-05 | -0.00012548 | -4.5642e-05 | 4.9799e-05 | -1.8805e-05 | -0.00012414 | -5.0339e-05 | -0.0001545 | -0.00014343 | -4.1785e-05 | -8.4725e-05 | -4.732e-05 | -0.00014259 | 5.6843e-06 | -4.2791e-05 | 0.00060785 |  | 596.1794
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3970.6022
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.133 | -4113.6223 | 5.4199 | 1.4309 | 0.29691 | -0.75654 | -0.16943 | 0.18117 | 0.2684 | -0.61899 | 0.41769 | -0.74816 | -0.48815 | -1.1608 | -0.51331 | -0.46298 | 0.66428 | 1.3839 | 0.27511 | 6.6126 | -3827.5821
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.10262 | 3.9931 | -0.85496 | -0.33121 | -0.55648 | -0.28371 | -0.052444 | 0.53898 | 1.1642 | 0.45765 | -0.054451 | 0.0176 | 0.29997 | 0.41093 | 0.38234 | -0.036238 | -0.41373 | -0.45056 | -0.03038 | -0.20749 | -0.61341
db2 | 0.61346 | 0.96052 | 0.088322 | 0.25472 | -0.39944 | -0.015155 | 0.33389 | 0.22883 | 0.64769 | 0.065877 | -1.1739 | -0.053578 | 0.20054 | 0.50574 | 0.17656 | -0.50576 | 0.12478 | -0.65589 | 0.29444 | -0.1177 | 0.19374
db3 | -4.1269 | 32.7919 | 0.06565 | 0.67943 | 0.52497 | 0.59431 | -1.2183 | -0.021427 | 1.134 | 1.4487 | -0.15446 | -1.1083 | 0.12504 | 0.17842 | -2.6664 | -1.4442 | -0.51034 | 0.38256 | 0.63818 | 1.3522 | -2.2636
db4 | 0.045491 | -0.97516 | -0.16007 | 0.12689 | 0.040386 | -0.043459 | 0.10117 | 0.024338 | 0.25092 | 0.050301 | -0.062026 | -0.0047617 | -0.053021 | -0.058837 | -0.11932 | 0.051345 | 0.038823 | -0.11248 | -0.017693 | -0.052489 | 0.12348
db5 | -0.22661 | -6.0724 | -0.21378 | -0.35358 | -0.036404 | -0.043652 | 0.072714 | 0.060786 | 0.070405 | -0.055183 | 0.20898 | 0.22152 | 0.15362 | 0.0084066 | 0.17725 | 0.28509 | -0.20106 | -0.14968 | -0.16092 | -0.044518 | -3.4357
db6 | 0.071479 | 0.13351 | -0.008162 | 0.0090777 | -0.027104 | -0.065322 | -0.10843 | -0.077977 | -0.15076 | -0.0012871 | 0.12662 | -0.015787 | 0.082717 | 0.023996 | -0.1047 | 0.11214 | 0.10352 | 0.066654 | -0.011022 | 0.045823 | 0.022074
db7 | 1.1581 | 2.6222 | -0.23771 | -0.10501 | 0.024298 | 0.048078 | 0.07226 | 0.059325 | 0.026072 | 0.0036407 | 0.061577 | 0.079699 | 0.0062877 | 0.0044257 | 0.017798 | -0.083314 | 0.06354 | 0.021173 | 0.0010057 | -0.063145 | 0.86186
db8 | 0.005805 | 0.027688 | 0.057216 | 0.02839 | 0.030636 | 0.031618 | 0.035223 | -0.01091 | -0.035575 | 0.038033 | 0.0060715 | 0.0093838 | 0.020075 | -0.0082784 | -0.062217 | -0.082398 | -0.015108 | 0.0034703 | -0.021482 | -0.024149 | -0.047534
db9 | 0.51074 | 0.29228 | 0.030769 | 0.0024173 | -0.0031937 | 0.016435 | -0.015153 | -0.0084252 | -0.0035202 | -0.003499 | -0.018766 | 0.0014946 | -0.001065 | 0.005389 | -0.0049261 | -0.015698 | -0.0043072 | -0.0056431 | 0.02008 | 0.0076102 | 0.45852
db10 | -0.0010223 | 0.10365 | 0.096745 | 0.049924 | 0.031045 | 0.014843 | 0.00076587 | -0.055313 | -0.058255 | 0.0050001 | 0.010351 | -0.040886 | 0.037036 | -0.0039518 | -0.11579 | -0.0081631 | -0.0034343 | 0.015307 | 0.019611 | 0.0051682 | -0.069142
db11 | 0.75216 | 0.59019 | -0.00080556 | -0.011151 | 1.6844e-05 | -0.003323 | 0.00099834 | -0.0063291 | -0.0020205 | -0.0018532 | -0.0027888 | 0.0037626 | -0.00088426 | -0.0019522 | 0.014657 | 0.017343 | -0.00028176 | 0.00078564 | -0.0010497 | -0.0051246 | 0.65251
db12 | -0.00052758 | 0.01914 | 0.014444 | 0.0033766 | 0.00080908 | -0.00076578 | -0.00088337 | -0.0052816 | -0.0079654 | -0.002293 | -0.0027313 | -0.0037315 | 0.0073801 | 0.00066218 | -0.0094301 | -0.0017692 | 0.0015986 | 0.0068847 | 0.0007957 | -0.0010995 | -0.0065148
db13 | 0.064799 | 0.068796 | -0.001271 | 0.002924 | 0.0022911 | 0.0022795 | -0.00060045 | 0.00066704 | 0.00076547 | 0.00024469 | -0.00085165 | -0.0020789 | -0.0020385 | 0.00091084 | -0.0033455 | -0.0044411 | 0.0010081 | 0.00053395 | 0.0019612 | 0.0010413 | 0.049023
db14 | -0.0033325 | 0.0018038 | 0.0064432 | 0.0036708 | 0.0012758 | 0.0011263 | -0.00022373 | -0.0034073 | -0.0032155 | -0.0004461 | -0.0005117 | -0.0020049 | 0.0028147 | 9.0103e-05 | -0.0061516 | -0.001911 | -0.0011194 | 0.0017837 | 0.00034414 | 0.0014424 | -0.0043604
db15 | 0.033177 | -0.009771 | 0.0078652 | 0.005001 | 0.00074372 | -0.00027602 | -0.0010131 | -0.00072944 | -0.0021521 | 0.00049092 | -0.0018042 | 0.00012768 | -0.00063871 | -0.00063044 | -0.00332 | -0.0012401 | -0.00074719 | -0.0013678 | 0.0020558 | -0.0023653 | 0.022279
da2 | 1.4056 | 10.2777 | -1.3121 | -0.015956 | -0.75093 | 1.0083 | 0.22943 | 0.62835 | 0.17136 | 0.33938 | -1.181 | -0.18101 | 0.2158 | 0.64145 | 0.4024 | -0.78887 | 0.30502 | 0.11952 | 0.34701 | -0.17825 | 2.062
da3 | 1.0237 | 4.5964 | 0.56498 | 0.67454 | -0.13437 | 0.11222 | -0.18591 | -0.28101 | -0.23177 | -0.54146 | 0.038361 | 0.55741 | 0.27951 | 0.022667 | -0.19586 | -0.3618 | -0.31973 | 0.088572 | -0.0014055 | -0.084948 | -0.64048
da4 | 0.11081 | 1.5559 | 0.1293 | -0.25896 | 0.44228 | 0.16487 | 0.047283 | -0.13769 | 0.012563 | 0.13163 | 0.45331 | 0.1393 | -0.10178 | 0.078678 | -0.054922 | 0.17189 | -0.41432 | -0.28876 | -0.46432 | -0.050363 | -0.041655
da5 | 0.2528 | 3.8604 | 0.10378 | 0.069447 | 0.085312 | 0.18918 | 0.070389 | -0.11153 | -0.078987 | -0.13763 | 0.2472 | 0.062054 | -0.14024 | -0.17079 | 0.096501 | -0.24698 | -0.043647 | 0.065018 | -0.099111 | 0.040032 | -0.10005
da6 | 0.090747 | 0.5064 | 0.015115 | -0.10409 | 0.095792 | -0.0032346 | 0.032329 | -0.037215 | -0.034071 | -0.015037 | 0.010458 | 0.040629 | -0.085525 | -0.038309 | -0.080053 | 0.12472 | 0.055104 | -0.022853 | -0.058355 | 0.10459 | 0.19793
da7 | 0.007293 | 1.7844 | -0.057364 | -0.004914 | -0.089325 | 0.046543 | 0.023052 | 0.048673 | 0.012759 | -0.018757 | -0.11196 | 0.023592 | 0.027084 | -0.047968 | -0.043928 | 0.034134 | 0.085332 | 0.020047 | 0.033662 | 0.019342 | 0.037419
da8 | -0.0021746 | 0.016239 | -0.042494 | -0.031086 | -0.029155 | 0.00046577 | 0.0084104 | 0.0045627 | -0.0034954 | 0.01854 | 0.025566 | -0.020447 | 0.0027749 | 0.023049 | 0.027891 | 0.018733 | 0.037432 | -0.0050687 | -0.045107 | 0.0094301 | 0.0060998
da9 | 0.0049063 | -0.17302 | 0.020426 | 0.026395 | -0.003473 | 0.0080616 | 0.0057326 | 0.00067756 | -0.0022918 | 0.0021098 | -0.0064327 | 0.029022 | -0.0019187 | -0.03137 | -0.0024866 | -0.030774 | 0.013994 | 0.0076842 | -0.011521 | -0.023835 | -0.087521
da10 | -0.0012477 | 0.079262 | -0.041192 | -0.047666 | 0.0091593 | -0.0025033 | 0.0035553 | -0.0046828 | -0.009278 | -0.0071429 | 0.034749 | -0.021605 | -0.0017837 | 0.0031482 | -0.0084572 | 0.045765 | 0.00096361 | 0.023504 | -0.027139 | 0.050607 | -0.039214
da11 | -0.042191 | 0.16479 | -0.023822 | -0.0083915 | -0.0060622 | -0.0038027 | -0.0022924 | 0.0039963 | 0.010187 | -0.0023806 | -0.0047395 | 0.0027475 | 0.0017593 | -0.0059533 | 0.0075495 | 0.00017462 | 0.0056902 | 0.01269 | 0.0081243 | 0.0045258 | -0.024324
da12 | -0.0015123 | 0.0019201 | -0.0010569 | -0.0053625 | 0.0055791 | 0.00080328 | 0.0012976 | -0.0013376 | -0.0036842 | -0.0013038 | 0.0024853 | -0.0028757 | -0.0025554 | -0.0001523 | -0.0059101 | 0.0080817 | -0.0011071 | 0.0030684 | -0.0031585 | 0.007189 | 0.0037373
da13 | 0.001367 | -0.023364 | -0.0002543 | 0.0028509 | -0.00027189 | 0.002692 | -0.00028791 | 0.0021459 | 9.5454e-05 | -0.0011873 | -0.0029028 | 0.003387 | -0.00018406 | -0.0040533 | 0.00042643 | -0.002885 | 0.00076615 | 0.00079072 | 8.3982e-05 | -0.0012119 | -0.0034348
da14 | -0.015669 | 0.013647 | -0.0058155 | -0.0036489 | -0.00074868 | 0.00037807 | 0.00017991 | -0.0013787 | 0.0016946 | 0.0014209 | 0.0044728 | -0.00081214 | 0.00085566 | 0.0022491 | 0.0024557 | -0.00015575 | 1.4326e-05 | -5.5864e-05 | -0.0023372 | 0.0012317 | -0.0095254
da15 | -0.004952 | 0.013396 | 0.0085625 | 0.0043504 | 0.0028499 | 0.0026311 | 0.00040451 | -0.0038782 | -0.004532 | 3.8779e-05 | 0.00066563 | -0.002477 | 0.0037551 | -0.0005035 | -0.010296 | -0.0029045 | -0.0011384 | 0.0025243 | 8.7704e-05 | -0.00014081 | -0.010161
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4056.4913
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.1386 | -4163.2299 | 5.5573 | -0.048751 | 0.12738 | -0.57044 | 0.22803 | 1.1825 | 0.49642 | -0.55702 | 0.18106 | -0.86064 | -0.74993 | 0.26661 | -0.16282 | 0.21126 | -0.74154 | 0.74133 | 0.25655 | 6.7634 | -3949.7527
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.10043 | 3.8927 | -0.93173 | -0.28513 | -0.56539 | -0.52394 | -0.15423 | 0.56941 | 1.1488 | 0.41336 | -0.021635 | -0.0075178 | 0.039274 | 0.32775 | 0.5963 | -0.053605 | -0.37482 | -0.24058 | 0.096139 | -0.032472 | -0.56902
db2 | 0.16551 | 1.599 | 0.43474 | 0.01619 | 0.21211 | 0.058362 | -0.042411 | 0.23582 | 0.55242 | -0.062363 | -0.40094 | 0.24844 | 0.013384 | 0.32819 | -0.10553 | -0.62684 | -0.37879 | -0.18264 | -0.069108 | -0.23101 | 1.5995
db3 | -3.8978 | 33.6779 | 0.22436 | 0.48214 | -0.31553 | 0.082032 | -0.85023 | 0.25486 | -0.65322 | 0.28664 | 0.51591 | -0.38807 | 0.067404 | 0.78016 | -1.1025 | -0.50089 | 0.5034 | -0.047046 | -0.40404 | 1.0646 | -3.8594
db4 | -0.22993 | -1.3872 | 0.044758 | 0.13038 | 0.18992 | 0.2111 | -0.012309 | -0.0083437 | -0.067844 | -0.033628 | 0.044009 | 0.071588 | -0.0047148 | -0.05857 | -0.19741 | -0.01933 | -0.13504 | -0.14785 | -0.13038 | 0.12365 | 0.036873
db5 | -0.10946 | -6.1453 | -0.47624 | -0.31301 | -0.024476 | 0.1861 | 0.31995 | 0.39339 | 0.26487 | 0.24167 | 0.080387 | 0.25399 | 0.046351 | -0.14108 | 0.030355 | 0.10859 | -0.29979 | -0.13303 | -0.15736 | -0.38066 | -3.2283
db6 | -0.0213 | -0.20486 | 0.011986 | -0.05586 | -0.058364 | 0.06911 | 0.052229 | -0.024859 | 0.038162 | 0.0042213 | 0.015509 | -0.077636 | -0.013273 | -0.14694 | -0.10116 | 0.11092 | 0.079134 | 0.072482 | -0.045143 | 0.069489 | -0.32053
db7 | 1.1141 | 2.6037 | -0.29213 | -0.11987 | 0.037246 | 0.047748 | 0.10771 | 0.065064 | 0.10423 | 0.039899 | 0.018906 | 0.03511 | 0.020556 | -0.0013282 | 0.0098728 | 0.012679 | -0.022623 | -0.028885 | 0.036573 | -0.07075 | 0.77612
db8 | -0.012039 | -0.086617 | 0.019424 | -0.0086061 | 0.02427 | -0.033416 | 0.017138 | -0.03039 | -0.032085 | -0.029506 | -0.018694 | -0.034585 | 0.00061919 | 0.013184 | -0.0049579 | 0.021771 | 0.035798 | 0.018234 | -0.0096329 | 0.051435 | 0.019642
db9 | 0.51508 | 0.28809 | 0.0096993 | -0.0058186 | -0.0066922 | 0.018655 | -0.0068219 | -0.0069723 | -0.013497 | 0.0029527 | 0.013341 | -0.0029624 | -0.0084419 | 0.00094973 | -0.0042884 | -0.0041485 | 0.012329 | -0.015177 | 0.0052574 | 0.011636 | 0.47152
db10 | 0.00096129 | 0.040517 | 0.0096428 | -0.032813 | 0.010829 | -0.0051792 | 0.011191 | -0.068066 | -0.047885 | -0.018943 | 0.0094394 | -0.053002 | -0.012108 | -0.025219 | -0.038538 | 0.07806 | 0.081586 | 0.055389 | -0.013238 | 0.058856 | -0.067756
db11 | 0.7634 | 0.58557 | -0.0040589 | -0.0067105 | 0.0016858 | 0.0019085 | 0.0047879 | 0.0013518 | -0.0015637 | 0.0030721 | -0.0050902 | 0.0034054 | -0.0046937 | 0.0022817 | 0.0074063 | 0.007241 | -0.0042205 | -0.0049256 | 0.00053583 | -0.0024137 | 0.66995
db12 | 0.00080476 | 0.0033677 | 0.0006044 | -0.0061184 | -0.00060988 | -0.0033207 | -0.0024387 | -0.0080832 | -0.0047608 | 0.0009502 | -0.00033313 | -0.0061176 | -7.2105e-05 | -6.546e-05 | -0.0013439 | 0.0076021 | 0.0088831 | 0.0085709 | 0.0011824 | 0.0054709 | 0.0018915
db13 | 0.066192 | 0.079928 | -0.0015889 | 0.00014882 | -0.0003104 | -0.0012667 | -0.002287 | -0.00095545 | -0.0004285 | -0.00386 | -0.0030641 | -0.0011102 | 0.0019263 | 0.0051122 | 0.0011119 | -0.00079221 | 0.0021582 | -0.00040068 | 0.002795 | 0.0028118 | 0.04607
db14 | -0.0017245 | 0.004494 | 6.5545e-05 | -0.0025325 | 4.3574e-05 | -0.0014863 | 0.00093047 | -0.0039839 | -0.0010153 | -0.00080017 | 0.0004155 | -0.0037258 | -0.00010428 | -0.0012391 | -0.001599 | 0.0049662 | 0.0055324 | 0.0033677 | -0.0011523 | 0.0023174 | -0.0058241
db15 | 0.031283 | -0.0060801 | 0.0090432 | 0.0026013 | -0.0012183 | -0.00075237 | -0.00049438 | 0.0011131 | -0.0053652 | -0.0014275 | -0.00061067 | 0.0013668 | -0.0027085 | -0.0015351 | -0.0020478 | 0.0013559 | 0.00014907 | -0.00038582 | -0.00094313 | 0.0018593 | 0.028019
da2 | 0.57898 | 16.6656 | -0.24754 | -0.39297 | 0.23403 | 1.4 | 0.091957 | 0.031309 | 0.29173 | 0.35961 | 0.31118 | 0.29085 | 0.068569 | -0.53902 | -0.67111 | -0.47491 | -0.63242 | 0.31129 | 0.10539 | -0.53797 | 5.2835
da3 | 0.9856 | -0.011045 | 0.43245 | 0.47579 | -0.060302 | 0.24923 | -0.11256 | -0.20354 | -0.068332 | -0.18648 | -0.10321 | 0.04558 | -0.25607 | -0.32238 | -0.10247 | -0.16019 | -0.0084565 | 0.11798 | 0.21902 | 0.043951 | -0.66246
da4 | 0.40645 | 4.2356 | -0.44319 | -0.14396 | -0.093366 | -0.29354 | 0.017479 | -0.026668 | 0.0074473 | 0.03928 | 0.36954 | 0.47298 | 0.20287 | -0.00017945 | -0.12271 | -0.054398 | 0.2155 | -0.15923 | -0.11869 | 0.13083 | 2.1175
da5 | 0.2228 | 2.2886 | -0.12257 | -0.0024881 | 0.11073 | 0.16726 | 0.098909 | 0.057102 | -0.062768 | 0.19676 | 0.094074 | -0.057582 | -0.033446 | -0.091935 | 0.087226 | -0.082095 | 0.030385 | -0.186 | -0.12813 | -0.075425 | -0.62291
da6 | -0.0066781 | 0.68452 | -0.090344 | 0.032253 | 0.12343 | -0.11414 | -0.13011 | -0.13817 | 0.035805 | 0.030734 | 0.059109 | 0.16132 | 0.048519 | -0.054185 | 0.010207 | 0.060409 | 0.037925 | 0.018834 | 0.017795 | -0.10938 | -0.32403
da7 | 0.042488 | 1.5792 | -0.069817 | -0.010596 | -0.067055 | -0.040976 | 0.069514 | 0.055999 | -0.00069615 | -0.022972 | -0.056836 | 0.058445 | -0.06657 | -0.030667 | 0.057509 | 0.046538 | 0.024206 | 0.046152 | -0.021548 | 0.02937 | 0.1911
da8 | 0.049211 | 0.025918 | -0.070207 | -0.049043 | 0.052494 | 0.042878 | -0.025112 | -0.045499 | 0.014221 | -0.00034224 | 0.017921 | 0.037746 | 0.045905 | 0.002378 | 0.012865 | 0.0050152 | -0.0095418 | 0.013986 | -0.012192 | -0.033473 | -0.05295
da9 | 0.017376 | -0.26772 | 0.014648 | 0.024828 | 0.010972 | -0.0037946 | 0.00088908 | -0.0049949 | 0.012721 | 0.012342 | -0.00089852 | -0.0099065 | -0.0092466 | 0.0017416 | 0.0081971 | -0.030057 | -0.010359 | 0.0033813 | -0.010296 | -0.010168 | -0.077592
da10 | 0.0057765 | -0.027318 | -0.082524 | -0.038155 | -0.0018945 | -0.0048703 | -0.0030137 | -0.043136 | 0.031654 | 0.0012161 | 0.0038617 | -0.0071856 | 0.047572 | 0.024451 | 0.021875 | -0.005239 | 0.015614 | 0.025867 | 0.020782 | -0.0068736 | -0.14543
da11 | 0.0011614 | 0.17258 | -0.029034 | -0.0072558 | -0.0078589 | -0.011034 | 0.0083672 | 0.0087995 | 0.014522 | 0.0073547 | -0.0066112 | 0.0021439 | -0.0024466 | 0.0053563 | 0.013133 | -0.00089734 | 0.0039792 | -0.0030739 | 0.0052061 | -0.00065037 | 0.021454
da12 | -0.0027653 | -0.0043391 | -0.0080978 | -0.0023317 | 0.0047611 | -0.0055494 | -0.0047338 | -0.0064362 | 0.0025463 | -0.00017688 | -0.0023328 | 0.0021454 | 0.0041498 | 0.0059424 | 0.0019405 | 0.0024824 | 0.0036111 | 0.0029769 | 0.002064 | -0.0029614 | -0.020036
da13 | 0.0031867 | -0.026652 | 0.0031957 | 0.0040309 | -0.00076111 | 0.0019714 | 0.0025019 | 0.00067096 | -0.00085501 | -8.0579e-05 | -0.00094901 | -0.00025754 | -0.0011149 | -0.00091709 | 0.00079788 | -0.00348 | -0.0016382 | -0.0018462 | -0.00064879 | -0.00062037 | -0.0049645
da14 | -0.010347 | 0.013295 | -0.0090643 | -0.0039284 | 0.0015195 | 0.0011886 | -0.0010231 | -0.0020183 | 0.0029521 | -0.00023379 | 0.00061156 | 0.0010463 | 0.0036063 | 0.0026701 | 0.0018135 | -0.00083483 | -0.00012192 | 0.0010049 | 0.00049875 | 0.00031319 | -0.0031453
da15 | -0.0031018 | 0.0048109 | 0.00012935 | -0.0037989 | 4.3367e-05 | -0.00012342 | 0.0012078 | -0.0060466 | -0.0022593 | -0.0018539 | -0.00066206 | -0.0054982 | -0.00049308 | -0.00035704 | -0.0017703 | 0.0055579 | 0.0071257 | 0.0048712 | -0.00044049 | 0.0043681 | -0.009164
Field Colinearity | 19.0074
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.8904 | -0.95757 | -0.43382 | -0.6591 | -0.38632 | -0.15506 | 0.43637 | 1.0616 | 0.35503 | -0.15707 | -0.085016 | 0.19735 | 0.30831 | 0.27972 | -0.13885 | -0.51634 | -0.55318 | -0.133 | -0.3101 | -0.71602
Twist (m2 mrad) | 29.1783 | -13.6454 | -5.8566 | -8.4035 | -4.6359 | -1.7444 | 4.5819 | 10.3502 | 3.1953 | -1.2958 | -0.63762 | 1.3321 | 1.8499 | 1.4685 | -0.62484 | -1.9363 | -1.6595 | -0.29924 | -0.46515 | -0.26851
Twist Integral (I0) (m2 rad) | 0.0078625
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.7923 | -1.0322 | -0.38556 | -0.66582 | -0.62438 | -0.25467 | 0.46898 | 1.0484 | 0.31292 | -0.12207 | -0.10795 | -0.061159 | 0.22732 | 0.49587 | -0.15404 | -0.47525 | -0.34102 | -0.004294 | -0.1329 | -0.66945
Twist (m2 mrad) | 28.4424 | -14.7083 | -5.2051 | -8.4892 | -7.4925 | -2.865 | 4.9242 | 10.2218 | 2.8163 | -1.0071 | -0.80963 | -0.41282 | 1.3639 | 2.6033 | -0.69317 | -1.7822 | -1.0231 | -0.0096615 | -0.19936 | -0.25105
Twist Integral (I0) (m2 rad) | 0.0040678

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.61346 | -4.1269 | 0.045491 | -0.22661 | 1.1581 | -1.0262 | 1.4056 | 1.0237 | 0.11081 | 18.8337 | 0.59617 | 0.16551 | -3.8978 | -0.22993 | -0.10946 | 1.1141 | -1.0043 | 0.57898 | 0.9856 | 0.40645 | 16.587 | 0.59618 | 0.6106 | -2.9356 | 0.016713 | -0.5089 | 1.1934 | 0 | 1.6985 | 1.0797 | 0.15019 | 14422.329 | 0.25433 | -2.7537 | -0.25671 | -0.39139 | 1.1487 | 0 | 1.2165 | 0.90334 | 0.57684 | 14420.447
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8