## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3082C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-10-27
Beginn der Messung / start of measuring: |  | TT.10.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000082_cc
Ende der Messung / end of measuring: |  | TT.10.JJ
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Mehler |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-10-27 |  | 2003-10-27 | 2003-10-27
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000082_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000082
Date of test Ap 1 | 2003-10-27 |  |  |  | Date of test Ap 2 | 2003-10-27
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | yellow alarm | status ok | status ok | b7 | status ok | yellow alarm | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
213 | HCMB__A001 | 3000082 |  | 3 |  |  | 3
| COMMENTS: | Warning due to yellow alarms on b9 head CS both ap. (3.58 & 4.28 sigma) and b7 pos. 2 both apertures (5.87 & 5.16 sigma).
| HOLDING POINT | OK-w | Data at CERN |  | 2003-10-27 | Analysis by |  | V. Remondino | Answer to MMS-LD |  |  | 2003-10-28 | Answer to manufacturer |  |  | 2003-10-28
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000082_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000082 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3082_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3082_cct2 |  |  |  | a even | 1
Test Operator |  | Mehler |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-10-27 |  |  | Time of test |  |  | 9:57 |  |  |  | Date of test |  | 2003-10-27 |  |  | Time of test |  |  | 9:58
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-05-02
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.59 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00022937 |  |  | Dy average |  |  | 0.00067108 |  |  |  | Dx average |  | 0.00014818 |  |  | Dy average |  |  | 0.00066344
Magnetic length |  | 14.4182 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1137 |  |  |  | Magnetic length |  | 14.4218 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0865
Phase |  | -14.5187 |  |  | Integral field (T.m) |  |  | 85.8122 |  |  |  | Phase |  | -33.9044 |  |  | Integral field (T.m) |  |  | 85.7698
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4904 | 5.9637 | 5.9608 | 5.9615 | 5.9606 | 5.9609 | 5.9614 | 5.9611 | 5.9599 | 5.9604 | 5.9607 | 5.9607 | 5.9606 | 5.9607 | 5.961 | 5.9612 | 5.9613 | 5.9604 | 5.9636 | 3.6407 | 85.8122 | C1 (T) | 3.3882 | 5.9642 | 5.9607 | 5.961 | 5.9606 | 5.961 | 5.9608 | 5.9603 | 5.9606 | 5.9604 | 5.9604 | 5.9608 | 5.9604 | 5.9604 | 5.96 | 5.9595 | 5.961 | 5.9603 | 5.9632 | 3.6915 | 85.7698
Angle (mrad) | 1.4842 | -0.97721 | -0.071604 | -0.13602 | 0.14999 | 0.018948 | 0.28937 | 0.68762 | 0.35959 | -0.22852 | -0.64323 | -0.15904 | 0.24625 | 0.19996 | 0.088852 | -0.14055 | -0.17173 | -0.13475 | -0.10969 | -0.22382 | -14.5187 | Angle (mrad) | 2.7529 | -1.2429 | -0.37401 | -0.35501 | -0.010746 | -0.14148 | 0.26999 | 0.7303 | 0.42797 | -0.083149 | -0.61484 | -0.09539 | 0.15262 | 0.23626 | 0.20205 | -0.18422 | -0.21393 | -0.059405 | 0.034357 | -0.39139 | -33.9044
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -3.8028 | 0.25781 | 0.90284 | 0.89067 | 0.64813 | 1.2774 | 0.87723 | 1.0974 | 0.4959 | 0.22318 | 0.84973 | 0.99087 | 0.85985 | 0.70296 | 0.52881 | 1.0284 | 0.55926 | 0.30529 | 0.1079 | -0.0012382 | 0.54053 | b2 | 3.2236 | 0.083894 | -0.17266 | -0.58188 | 0.21795 | -0.076991 | -0.045634 | -0.70002 | -0.31569 | -0.30961 | -0.25956 | -0.19327 | -0.93781 | -1.1517 | -0.47023 | -0.58403 | -0.70279 | -0.44151 | -0.38347 | 0.84079 | -0.24348
b3 | 33.3841 | -3.9014 | -2.7354 | -3.5509 | -2.7 | -3.8936 | -2.5732 | -2.7353 | -2.5996 | -3.3589 | -3.9174 | -2.5266 | -3.0917 | -5.1051 | -4.1524 | -3.9172 | -2.6893 | -2.3169 | -1.8345 | -3.0421 | -2.0791 | b3 | 37.287 | -3.781 | -2.7916 | -3.631 | -1.8922 | -4.2321 | -2.7513 | -3.4437 | -2.9011 | -3.8778 | -3.4984 | -2.7323 | -2.5332 | -4.7173 | -4.0283 | -4.5771 | -3.3082 | -3.3808 | -2.3382 | -3.0616 | -2.1429
b4 | -1.6073 | 0.010108 | 0.15378 | 0.45199 | 0.40592 | 0.43696 | 0.21189 | 0.19839 | 0.075539 | 0.16645 | 0.23201 | 0.48806 | 0.10703 | 0.17324 | 0.1593 | 0.27223 | 0.16228 | 0.015549 | 0.0020997 | 0.16236 | 0.15007 | b4 | -0.05887 | -0.13388 | -0.27024 | -0.028735 | 0.070469 | -0.031947 | -0.14365 | -0.2171 | -0.16742 | -0.1207 | -0.017439 | -0.15092 | -0.32596 | -0.35663 | -0.12542 | -0.17211 | -0.27601 | -0.16739 | -0.15899 | -0.28771 | -0.15665
b5 | -6.2165 | -0.7079 | -0.83871 | -0.43883 | -0.44236 | -0.12272 | -0.15954 | -0.24126 | -0.38948 | -0.11005 | -0.055983 | -0.37485 | -0.48613 | -0.28782 | -0.3719 | -0.64757 | -0.77549 | -0.65593 | -0.77512 | -3.4519 | -0.71003 | b5 | -5.5202 | -0.79285 | -0.93119 | -0.50904 | -0.5844 | -0.30307 | -0.25795 | -0.28727 | -0.39977 | -0.25885 | -0.16891 | -0.41547 | -0.50006 | -0.41445 | -0.49548 | -0.52136 | -0.76304 | -0.69432 | -0.88557 | -3.9313 | -0.76901
b6 | -0.26868 | 0.14716 | 0.067405 | 0.10807 | 0.22496 | 0.041275 | -0.055163 | 0.021376 | 0.0821 | 0.069546 | -0.020702 | -0.16833 | 0.037357 | -0.0089099 | -0.042416 | -0.028236 | 0.10575 | 0.099731 | 0.097534 | -0.055053 | 0.030611 | b6 | 0.19562 | -0.17597 | -0.14951 | -0.079905 | -0.0097685 | 0.01302 | -0.030736 | 0.092832 | -0.028228 | -0.017345 | -0.071182 | -0.11255 | 0.042186 | 0.039619 | -0.054043 | 0.053698 | 0.011661 | 0.012327 | 0.065907 | -0.068912 | -0.017176
b7 | 2.807 | 0.9337 | 1.0211 | 1.2219 | 1.221 | 1.2398 | 1.2783 | 1.2958 | 1.2127 | 1.3157 | 1.2896 | 1.28 | 1.2151 | 1.1844 | 1.1764 | 1.235 | 1.263 | 1.2748 | 1.1833 | 0.99259 | 1.255 | b7 | 3.0136 | 0.9082 | 1.0689 | 1.2397 | 1.2612 | 1.3048 | 1.2828 | 1.2587 | 1.2577 | 1.3099 | 1.2922 | 1.1981 | 1.2112 | 1.2134 | 1.246 | 1.3135 | 1.2279 | 1.2678 | 1.2048 | 0.94885 | 1.2699
b8 | -0.050076 | -0.061304 | 0.00092874 | 0.077357 | 0.021208 | -0.0099585 | -0.01418 | 0.01786 | -0.01612 | -0.082917 | -0.0071999 | 0.0021607 | 0.034909 | -0.0012813 | -0.020279 | 0.066379 | 0.031998 | 0.008959 | 0.053686 | -0.033981 | 0.0027163 | b8 | 0.022489 | -0.079241 | -0.09003 | -0.010702 | 0.016955 | 0.0070437 | -0.010901 | 0.026873 | -0.006193 | -0.019089 | -0.050525 | -0.0054224 | 0.0054342 | -0.038344 | -0.044367 | 0.011547 | -0.0031289 | 0.02506 | 0.074268 | 0.015429 | -0.0087782
b9 | 0.36829 | 0.50565 | 0.50317 | 0.48816 | 0.52489 | 0.47794 | 0.51133 | 0.50666 | 0.53266 | 0.49964 | 0.5133 | 0.49605 | 0.50248 | 0.50366 | 0.48438 | 0.50573 | 0.51505 | 0.53786 | 0.52922 | 0.45314 | 0.50167 | b9 | 0.40357 | 0.52822 | 0.51943 | 0.49814 | 0.54504 | 0.49466 | 0.52487 | 0.51237 | 0.53158 | 0.5188 | 0.51514 | 0.49871 | 0.50819 | 0.51471 | 0.51361 | 0.49807 | 0.50529 | 0.52057 | 0.52816 | 0.47665 | 0.51075
b10 | 0.033574 | -0.080066 | -0.02358 | 0.063835 | 0.09488 | -0.035379 | -0.014389 | 0.026266 | -0.019745 | -0.08461 | -0.064139 | -0.047279 | 0.032114 | -0.021363 | -0.013985 | 0.035977 | 0.077469 | 0.040901 | 0.061454 | -0.078631 | 0 | b10 | 0.079413 | -0.12094 | -0.11691 | -0.038613 | 0.022298 | 0.016263 | 0.0077361 | 0.092992 | -0.021448 | -0.061826 | -0.039635 | -0.023558 | 0.025254 | -0.0077805 | -0.034426 | 0.037904 | 0.056298 | 0.058386 | 0.11844 | -0.025102 | 0
b11 | 0.62418 | 0.74309 | 0.74336 | 0.75283 | 0.75231 | 0.75484 | 0.74644 | 0.74799 | 0.75326 | 0.74705 | 0.75757 | 0.74941 | 0.75456 | 0.76477 | 0.76125 | 0.75087 | 0.74037 | 0.74393 | 0.74008 | 0.64632 | 0.74307 | b11 | 0.61145 | 0.7627 | 0.75231 | 0.76029 | 0.75909 | 0.76761 | 0.75992 | 0.75913 | 0.76086 | 0.75676 | 0.76783 | 0.76389 | 0.76189 | 0.76923 | 0.764 | 0.75953 | 0.75515 | 0.75657 | 0.74721 | 0.657 | 0.75248
b12 | 0.005808 | -0.0034185 | -0.0043891 | 0.0036805 | 0.0054181 | -0.004275 | -0.0030111 | 0.0035282 | -0.0027621 | -0.0092182 | -0.0074633 | -0.0051474 | 0.0033881 | -0.0017645 | -0.0018352 | 0.0074581 | 0.0084166 | 0.003695 | 0.0052893 | -0.0066313 | -0.00015924 | b12 | 0.0070059 | -0.014432 | -0.014032 | -0.0080436 | 0.0027514 | -0.00089762 | 9.7939e-05 | 0.006717 | -0.0017265 | -0.0074167 | -0.0070425 | -0.005531 | 0.0039652 | 0.00040572 | -0.0044161 | 0.00047947 | 0.0039748 | 0.0057026 | 0.015427 | 0.0060823 | -0.00084791
b13 | 0.086626 | 0.068073 | 0.070735 | 0.069813 | 0.069189 | 0.065847 | 0.065916 | 0.066698 | 0.065384 | 0.068134 | 0.065854 | 0.069599 | 0.069677 | 0.064417 | 0.068536 | 0.070799 | 0.069511 | 0.072239 | 0.071261 | 0.055373 | 0.068567 | b13 | 0.091723 | 0.072072 | 0.074322 | 0.071393 | 0.071618 | 0.070213 | 0.06819 | 0.068931 | 0.068414 | 0.069331 | 0.071461 | 0.072744 | 0.07334 | 0.069654 | 0.071534 | 0.071343 | 0.071725 | 0.075909 | 0.075387 | 0.052824 | 0.071527
b14 | 0.002715 | -0.01204 | -0.0092828 | -0.0041064 | -0.0040206 | -0.0094908 | -0.0064857 | -0.0033944 | -0.0077657 | -0.01077 | -0.0090388 | -0.0085267 | -0.0036022 | -0.006932 | -0.0068679 | -0.0027108 | -0.0010183 | -0.004091 | -0.0029476 | -0.0070982 | -0.0060343 | b14 | -0.00047223 | -0.010524 | -0.0086263 | -0.005125 | -0.002379 | -0.0018529 | -0.0024841 | 0.0019567 | -0.0034252 | -0.0069247 | -0.0057446 | -0.005567 | -0.0017298 | -0.0034476 | -0.004223 | -0.00047124 | -7.6489e-05 | 0.00077493 | 0.0042705 | 7.2318e-06 | -0.0029115
b15 | -0.0074904 | 0.031937 | 0.029342 | 0.028316 | 0.031772 | 0.029398 | 0.026401 | 0.02819 | 0.028012 | 0.02507 | 0.023925 | 0.023523 | 0.028636 | 0.031143 | 0.028608 | 0.031638 | 0.030078 | 0.031113 | 0.028529 | 0.018718 | 0.027228 | b15 | -0.0052435 | 0.029959 | 0.026022 | 0.028365 | 0.029534 | 0.031699 | 0.032964 | 0.029836 | 0.03133 | 0.030731 | 0.028994 | 0.028442 | 0.030729 | 0.032785 | 0.030965 | 0.030116 | 0.031142 | 0.030409 | 0.031393 | 0.019895 | 0.028912
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 14.8417 | -9.7721 | -0.71604 | -1.3602 | 1.4999 | 0.18948 | 2.8937 | 6.8762 | 3.5959 | -2.2852 | -6.4323 | -1.5903 | 2.4625 | 1.9996 | 0.88852 | -1.4055 | -1.7173 | -1.3475 | -1.0969 | -2.2382 | 0 | a1 | 27.5287 | -12.429 | -3.7401 | -3.5501 | -0.10746 | -1.4148 | 2.6999 | 7.3029 | 4.2797 | -0.83149 | -6.1484 | -0.9539 | 1.5262 | 2.3626 | 2.0205 | -1.8422 | -2.1393 | -0.59405 | 0.34357 | -3.914 | 0
a2 | 19.8033 | 1.2327 | 1.1687 | 0.50955 | 1.4413 | 0.43846 | 0.28725 | 1.7699 | -0.30548 | 0.49643 | 0.18339 | 0.61322 | 0.61539 | 0.22939 | 0.16522 | 1.2529 | -0.55973 | 0.0099422 | 0.77393 | -8.3501 | 0.87608 | a2 | 15.6872 | -0.85058 | -1.0286 | -1.2696 | -1.0143 | -1.0312 | -0.96199 | -0.746 | -1.5762 | -1.6939 | -0.99519 | -1.3395 | -0.75678 | -1.926 | -0.53093 | -1.9948 | -2.1442 | -0.97326 | -0.79816 | -1.0664 | -0.69708
a3 | -2.0134 | 1.2027 | 1.0173 | 0.63191 | 0.53865 | 0.19101 | 0.70447 | 0.25536 | 0.56435 | 0.24085 | 1.065 | 1.2998 | 0.92501 | 1.0046 | 0.78851 | 1.3205 | 0.85518 | 1.0255 | 0.63497 | -0.032782 | 0.68071 | a3 | -2.5744 | 0.25633 | 0.56318 | 0.33395 | 0.32794 | 0.37623 | 0.64767 | 0.47576 | 0.26163 | -0.018844 | 0.52169 | 0.32344 | 0.6401 | 0.21643 | 0.30118 | 0.81387 | 0.69802 | 0.8534 | 0.68135 | -0.40503 | 0.34185
a4 | 2.4959 | -0.023828 | -0.079629 | 0.14861 | -0.44843 | -0.65673 | -0.1881 | -1.0541 | -0.51513 | -0.91684 | -0.36368 | -0.33233 | -0.35739 | -0.22054 | -0.21536 | -0.62002 | -0.24029 | -0.10293 | -0.082391 | -0.48589 | -0.26589 | a4 | 2.8529 | 0.01949 | 0.16219 | 0.38646 | 0.68577 | 0.29599 | -0.26756 | -0.19159 | 0.38736 | 0.21708 | -0.039564 | 0.22078 | 0.31836 | -0.074586 | -0.17469 | 0.17199 | 0.16277 | -0.02246 | 0.13146 | 0.47208 | 0.22427
a5 | 2.2263 | 0.19295 | 0.043739 | 0.24455 | 0.42132 | 0.2803 | 0.015986 | 0.021507 | 0.0059866 | 0.60224 | 0.37018 | -0.12351 | -0.050668 | 0.30574 | 0.19469 | 0.1532 | -0.10361 | 0.21133 | 0.26744 | -0.044839 | 0.22554 | a5 | 2.0556 | 0.11829 | 0.18334 | 0.30089 | 0.50097 | 0.36582 | 0.29624 | 0.14544 | 0.24828 | 0.40527 | 0.48403 | 0.18437 | 0.080802 | 0.33743 | 0.25751 | 0.094149 | 0.15749 | 0.16849 | 0.23206 | -0.21722 | 0.29158
a6 | 0.4028 | 0.012517 | 0.015433 | 0.10713 | 0.056685 | -9.4624e-05 | 0.053464 | -0.047122 | -0.0023442 | -0.06985 | 0.10588 | 0.016162 | -0.023376 | -0.0037536 | -0.059841 | 0.0068717 | -0.052514 | -0.0063838 | 0.067456 | -0.24927 | 0.013542 | a6 | 0.53191 | 0.0020569 | -0.041177 | 0.043935 | -0.005471 | 0.000584 | 0.099683 | 0.070897 | -0.0029897 | -0.025456 | 0.074907 | 0.019487 | -0.01663 | 0.060081 | -0.02408 | -0.063741 | -0.061025 | 0.035845 | 0.021966 | -0.19787 | 0.019214
a7 | 1.3245 | -0.037647 | 0.021437 | 0.060762 | 0.090935 | -0.0078994 | -0.049234 | -0.0011802 | 0.054051 | -0.037817 | -0.052787 | -0.041335 | 0.0135 | 0.055547 | 0.011715 | -0.082029 | 0.036152 | 0.0066628 | -0.00058103 | 0.049657 | 0.044071 | a7 | 1.6781 | -0.029423 | 0.039485 | 0.10327 | 0.16802 | 0.059332 | 0.051911 | 0.056369 | 0.12641 | 0.11222 | 0.049887 | 0.064696 | 0.12352 | 0.15759 | 0.15838 | 0.012865 | 0.062831 | 0.057966 | 0.087498 | 0.15817 | 0.13105
a8 | 0.0017814 | -0.02527 | -0.0003777 | 0.024331 | 0.016265 | 0.022346 | 0.028584 | 0.048482 | -0.0027635 | 0.064043 | 0.03958 | 0.015102 | -0.012238 | -0.032358 | -0.031036 | 0.020458 | -0.014796 | -0.0032187 | 0.014193 | -0.028433 | 0.008074 | a8 | 0.026119 | 0.0050764 | 0.014858 | 0.0056456 | 0.012115 | 0.0069206 | -0.022933 | -0.0053096 | -0.050889 | -0.044631 | -0.01135 | -0.014092 | 0.0020416 | -0.019232 | -0.021758 | -0.034243 | -0.029352 | 0.01751 | 0.0018208 | -0.036221 | -0.010182
a9 | -0.30746 | 0.0021069 | -0.0039175 | -0.028766 | -0.075501 | -0.027858 | 0.0026881 | 0.0052043 | -0.040412 | -0.052379 | 0.015398 | -2.4574e-05 | 0.0095474 | -0.030713 | -0.047411 | 0.013356 | 0.0024869 | -0.010672 | -0.005179 | -0.076356 | -0.025977 | a9 | -0.24794 | 0.060686 | 0.02723 | 0.012593 | -0.02708 | 0.018613 | 0.034748 | 0.043173 | 0.012689 | -0.016476 | 0.04051 | 0.010903 | 0.035109 | 0.010607 | -0.0026395 | 0.06127 | 0.025505 | 0.018055 | 0.00054999 | -0.063942 | 0.0096696
a10 | 0.053981 | -0.016444 | 0.014061 | 0.020608 | -0.010174 | -0.030581 | 0.016544 | -0.007692 | 0.0060624 | -0.0045325 | 0.036188 | 0.024333 | 0.0056012 | -0.021565 | -0.011816 | -0.019612 | -0.0069355 | -0.0022909 | 0.026593 | -0.081793 | 0 | a10 | 0.036523 | -0.0015375 | 0.014688 | 0.011685 | 0.03093 | -0.014026 | -0.028552 | 0.020786 | 0.0025015 | -0.020064 | -0.0066772 | 0.025946 | 0.010734 | -0.0044652 | -0.019146 | 0.0073659 | -0.011798 | 0.0040525 | -0.0019352 | -0.066596 | 0
a11 | 0.12238 | -0.083994 | -0.052859 | -0.049959 | -0.034116 | -0.045263 | -0.056917 | -0.040127 | -0.037307 | -0.038034 | -0.055146 | -0.053117 | -0.043127 | -0.045625 | -0.04967 | -0.06551 | -0.053583 | -0.04593 | -0.042976 | -0.018284 | -0.043383 | a11 | 0.19361 | -0.025964 | -0.0068235 | 0.011887 | 0.015437 | 0.013783 | 0.011499 | 0.022851 | 0.021803 | 0.017788 | 0.012306 | 0.0039485 | 0.010992 | 0.020085 | 0.015992 | -0.001509 | 0.0054796 | 0.0060802 | 0.0095275 | 0.02842 | 0.015259
a12 | -0.00032864 | -0.003361 | 0.00067287 | 0.0024165 | -0.00077811 | -0.0032152 | 0.0016344 | 0.00076125 | -0.00014527 | -0.0030438 | 0.0011924 | 0.0026246 | 0.00027986 | -0.0025835 | -0.0010959 | -0.0026883 | -0.0054265 | -0.00025014 | 0.0028945 | -0.011598 | -0.00090574 | a12 | -0.00037057 | -0.0042239 | -0.0037022 | -0.0018462 | -0.00015276 | -0.0015189 | -0.00023677 | 0.0021447 | -8.6004e-06 | -0.0054464 | 0.00029535 | 0.00043907 | 0.00040442 | 0.00019261 | -0.0015515 | -0.0014801 | -0.0078404 | -9.8612e-06 | -0.00044536 | -0.010898 | -0.0016651
a13 | -0.029851 | 0.00025751 | 0.0018174 | -0.0014761 | -0.0010843 | -0.00014086 | -0.0029273 | -0.0037687 | -0.001636 | 0.00057483 | 0.0027021 | -0.0022938 | -0.0017419 | 0.00028046 | -0.0023756 | -0.0029145 | -0.0034655 | -0.0023082 | -0.0032335 | -0.0037271 | -0.0022655 | a13 | -0.019033 | 0.0077473 | 0.0042305 | 0.0021126 | 0.0023807 | 0.0038611 | 0.0033242 | 0.0019127 | 0.0047699 | 0.004403 | 0.007308 | 0.0034029 | 0.0041001 | 0.0041649 | 0.0040189 | 0.0064552 | 0.0023462 | 0.0018075 | 0.0029466 | 0.00023468 | 0.0031594
a14 | 0.0025238 | -0.015499 | -0.012124 | -0.011148 | -0.01441 | -0.012326 | -0.011212 | -0.010461 | -0.011949 | -0.0084183 | -0.010604 | -0.010164 | -0.013958 | -0.015147 | -0.014861 | -0.015203 | -0.014456 | -0.013909 | -0.011271 | -0.0042732 | -0.011891 | a14 | 0.0088066 | -0.011676 | -0.0086985 | -0.0099641 | -0.010605 | -0.010761 | -0.0152 | -0.013385 | -0.013657 | -0.013597 | -0.011877 | -0.011386 | -0.010779 | -0.014456 | -0.01302 | -0.011832 | -0.012677 | -0.010951 | -0.012386 | -0.0037427 | -0.011164
a15 | 0.0051283 | -0.018886 | -0.013372 | -0.0054884 | -0.0059901 | -0.013462 | -0.01142 | -0.0074459 | -0.011393 | -0.018389 | -0.013236 | -0.013008 | -0.0061422 | -0.012634 | -0.012487 | -0.0073086 | -0.0037687 | -0.0076541 | -0.0045672 | -0.017981 | -0.010139 | a15 | 0.0077714 | -0.015134 | -0.014108 | -0.0080209 | -0.0026907 | -0.0030743 | -0.0029293 | 0.0031718 | -0.0051588 | -0.01062 | -0.007799 | -0.00598 | -0.0015724 | -0.0050154 | -0.0078766 | 0.00028603 | -0.00010341 | 0.00048679 | 0.0059378 | -0.0050577 | -0.004113
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00011585 | 0.00017664 | 5.5932e-05 | -0.00014044 | -0.00021504 | 7.5268e-05 | 3.5436e-05 | -6.0465e-05 | 4.5132e-05 | 0.0001916 | 0.00014909 | 0.00011062 | -7.1402e-05 | 4.4466e-05 | 2.9383e-05 | -8.4694e-05 | -0.00017818 | -9.344e-05 | -0.00013715 | 0.00020016 | 0.00022937 | Dx (m) | -0.00022995 | 0.00026943 | 0.00026477 | 8.5932e-05 | -5.1327e-05 | -3.5446e-05 | -1.6337e-05 | -0.00020951 | 4.7723e-05 | 0.00013989 | 8.7985e-05 | 5.2127e-05 | -5.6687e-05 | 1.7441e-05 | 7.7464e-05 | -8.4804e-05 | -0.00012653 | -0.00013128 | -0.00026971 | 7.2235e-05 | 0.00014818
Dy (m) | -0.00012435 | 5.7621e-05 | -2.8185e-05 | -5.5876e-05 | 1.3254e-05 | 7.3401e-05 | -3.498e-05 | 1.4239e-05 | -1.1447e-05 | 2.0132e-05 | -7.0402e-05 | -4.7381e-05 | -1.6703e-05 | 5.0592e-05 | 2.8306e-05 | 3.7036e-05 | 0 | 0 | -6.9086e-05 | 0.00022093 | 0.00067108 | Dy (m) | -2.8749e-05 | 1.2736e-05 | -3.0736e-05 | -2.7478e-05 | -6.8232e-05 | 3.1697e-05 | 6.4122e-05 | -4.029e-05 | 0 | 4.1845e-05 | 1.3377e-05 | -5.8022e-05 | -2.3135e-05 | 0 | 4.0992e-05 | -1.6648e-05 | 2.7532e-05 | 0 | 0 | 0.00016901 | 0.00066344
Magnetic length (m) |  | 14.4182 | [19 empty] | Magnetic length (m) |  | 14.4218
Transfer funct. (mT/KA) |  | 596.1137 | [19 empty] | Transfer funct. (mT/KA) |  | 596.0865
Coil positioning (mm) |  | 0.020706 | [19 empty] | Coil positioning (mm) |  | 0.017006
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000082 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000082 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 349.0362 | 596.3724 | 596.0808 | 596.1451 | 596.0564 | 596.0851 | 596.1431 | 596.1134 | 595.9912 | 596.0368 | 596.0658 | 596.073 | 596.0564 | 596.0722 | 596.1011 | 596.1208 | 596.1291 | 596.0393 | 596.3641 | 364.0676 | 85.8122 | C1/i (mT/KA) | 338.8172 | 596.4237 | 596.0699 | 596.0977 | 596.0608 | 596.0998 | 596.0773 | 596.0345 | 596.0639 | 596.0367 | 596.0357 | 596.0756 | 596.0357 | 596.0431 | 595.9956 | 595.9523 | 596.1031 | 596.032 | 596.3201 | 369.1542 | 85.7698
Angle (mrad) | 1.4842 | -0.97721 | -0.071604 | -0.13602 | 0.14999 | 0.018948 | 0.28937 | 0.68762 | 0.35959 | -0.22852 | -0.64323 | -0.15904 | 0.24625 | 0.19996 | 0.088852 | -0.14055 | -0.17173 | -0.13475 | -0.10969 | -0.22382 | -14.5187 | Angle (mrad) | 2.7529 | -1.2429 | -0.37401 | -0.35501 | -0.010746 | -0.14148 | 0.26999 | 0.7303 | 0.42797 | -0.083149 | -0.61484 | -0.09539 | 0.15262 | 0.23626 | 0.20205 | -0.18422 | -0.21393 | -0.059405 | 0.034357 | -0.39139 | -33.9044
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -3.8028 | 0.25781 | 0.90284 | 0.89067 | 0.64813 | 1.2774 | 0.87723 | 1.0974 | 0.4959 | 0.22318 | 0.84973 | 0.99087 | 0.85985 | 0.70296 | 0.52881 | 1.0284 | 0.55926 | 0.30529 | 0.1079 | -0.0012382 | 0.54053 | b2 | 3.2236 | 0.083894 | -0.17266 | -0.58188 | 0.21795 | -0.076991 | -0.045634 | -0.70002 | -0.31569 | -0.30961 | -0.25956 | -0.19327 | -0.93781 | -1.1517 | -0.47023 | -0.58403 | -0.70279 | -0.44151 | -0.38347 | 0.84079 | -0.24348
b3 | 33.3841 | -3.9014 | -2.7354 | -3.5509 | -2.7 | -3.8936 | -2.5732 | -2.7353 | -2.5996 | -3.3589 | -3.9174 | -2.5266 | -3.0917 | -5.1051 | -4.1524 | -3.9172 | -2.6893 | -2.3169 | -1.8345 | -3.0421 | -2.0791 | b3 | 37.287 | -3.781 | -2.7916 | -3.631 | -1.8922 | -4.2321 | -2.7513 | -3.4437 | -2.9011 | -3.8778 | -3.4984 | -2.7323 | -2.5332 | -4.7173 | -4.0283 | -4.5771 | -3.3082 | -3.3808 | -2.3382 | -3.0616 | -2.1429
b4 | -1.6073 | 0.010108 | 0.15378 | 0.45199 | 0.40592 | 0.43696 | 0.21189 | 0.19839 | 0.075539 | 0.16645 | 0.23201 | 0.48806 | 0.10703 | 0.17324 | 0.1593 | 0.27223 | 0.16228 | 0.015549 | 0.0020997 | 0.16236 | 0.15007 | b4 | -0.05887 | -0.13388 | -0.27024 | -0.028735 | 0.070469 | -0.031947 | -0.14365 | -0.2171 | -0.16742 | -0.1207 | -0.017439 | -0.15092 | -0.32596 | -0.35663 | -0.12542 | -0.17211 | -0.27601 | -0.16739 | -0.15899 | -0.28771 | -0.15665
b5 | -6.2165 | -0.7079 | -0.83871 | -0.43883 | -0.44236 | -0.12272 | -0.15954 | -0.24126 | -0.38948 | -0.11005 | -0.055983 | -0.37485 | -0.48613 | -0.28782 | -0.3719 | -0.64757 | -0.77549 | -0.65593 | -0.77512 | -3.4519 | -0.71003 | b5 | -5.5202 | -0.79285 | -0.93119 | -0.50904 | -0.5844 | -0.30307 | -0.25795 | -0.28727 | -0.39977 | -0.25885 | -0.16891 | -0.41547 | -0.50006 | -0.41445 | -0.49548 | -0.52136 | -0.76304 | -0.69432 | -0.88557 | -3.9313 | -0.76901
b6 | -0.26868 | 0.14716 | 0.067405 | 0.10807 | 0.22496 | 0.041275 | -0.055163 | 0.021376 | 0.0821 | 0.069546 | -0.020702 | -0.16833 | 0.037357 | -0.0089099 | -0.042416 | -0.028236 | 0.10575 | 0.099731 | 0.097534 | -0.055053 | 0.030611 | b6 | 0.19562 | -0.17597 | -0.14951 | -0.079905 | -0.0097685 | 0.01302 | -0.030736 | 0.092832 | -0.028228 | -0.017345 | -0.071182 | -0.11255 | 0.042186 | 0.039619 | -0.054043 | 0.053698 | 0.011661 | 0.012327 | 0.065907 | -0.068912 | -0.017176
b7 | 2.807 | 0.9337 | 1.0211 | 1.2219 | 1.221 | 1.2398 | 1.2783 | 1.2958 | 1.2127 | 1.3157 | 1.2896 | 1.28 | 1.2151 | 1.1844 | 1.1764 | 1.235 | 1.263 | 1.2748 | 1.1833 | 0.99259 | 1.255 | b7 | 3.0136 | 0.9082 | 1.0689 | 1.2397 | 1.2612 | 1.3048 | 1.2828 | 1.2587 | 1.2577 | 1.3099 | 1.2922 | 1.1981 | 1.2112 | 1.2134 | 1.246 | 1.3135 | 1.2279 | 1.2678 | 1.2048 | 0.94885 | 1.2699
b8 | -0.050076 | -0.061304 | 0.00092874 | 0.077357 | 0.021208 | -0.0099585 | -0.01418 | 0.01786 | -0.01612 | -0.082917 | -0.0071999 | 0.0021607 | 0.034909 | -0.0012813 | -0.020279 | 0.066379 | 0.031998 | 0.008959 | 0.053686 | -0.033981 | 0.0027163 | b8 | 0.022489 | -0.079241 | -0.09003 | -0.010702 | 0.016955 | 0.0070437 | -0.010901 | 0.026873 | -0.006193 | -0.019089 | -0.050525 | -0.0054224 | 0.0054342 | -0.038344 | -0.044367 | 0.011547 | -0.0031289 | 0.02506 | 0.074268 | 0.015429 | -0.0087782
b9 | 0.36829 | 0.50565 | 0.50317 | 0.48816 | 0.52489 | 0.47794 | 0.51133 | 0.50666 | 0.53266 | 0.49964 | 0.5133 | 0.49605 | 0.50248 | 0.50366 | 0.48438 | 0.50573 | 0.51505 | 0.53786 | 0.52922 | 0.45314 | 0.50167 | b9 | 0.40357 | 0.52822 | 0.51943 | 0.49814 | 0.54504 | 0.49466 | 0.52487 | 0.51237 | 0.53158 | 0.5188 | 0.51514 | 0.49871 | 0.50819 | 0.51471 | 0.51361 | 0.49807 | 0.50529 | 0.52057 | 0.52816 | 0.47665 | 0.51075
b10 | 0.033574 | -0.080066 | -0.02358 | 0.063835 | 0.09488 | -0.035379 | -0.014389 | 0.026266 | -0.019745 | -0.08461 | -0.064139 | -0.047279 | 0.032114 | -0.021363 | -0.013985 | 0.035977 | 0.077469 | 0.040901 | 0.061454 | -0.078631 | 0 | b10 | 0.079413 | -0.12094 | -0.11691 | -0.038613 | 0.022298 | 0.016263 | 0.0077361 | 0.092992 | -0.021448 | -0.061826 | -0.039635 | -0.023558 | 0.025254 | -0.0077805 | -0.034426 | 0.037904 | 0.056298 | 0.058386 | 0.11844 | -0.025102 | 0
b11 | 0.62418 | 0.74309 | 0.74336 | 0.75283 | 0.75231 | 0.75484 | 0.74644 | 0.74799 | 0.75326 | 0.74705 | 0.75757 | 0.74941 | 0.75456 | 0.76477 | 0.76125 | 0.75087 | 0.74037 | 0.74393 | 0.74008 | 0.64632 | 0.74307 | b11 | 0.61145 | 0.7627 | 0.75231 | 0.76029 | 0.75909 | 0.76761 | 0.75992 | 0.75913 | 0.76086 | 0.75676 | 0.76783 | 0.76389 | 0.76189 | 0.76923 | 0.764 | 0.75953 | 0.75515 | 0.75657 | 0.74721 | 0.657 | 0.75248
b12 | 0.005808 | -0.0034185 | -0.0043891 | 0.0036805 | 0.0054181 | -0.004275 | -0.0030111 | 0.0035282 | -0.0027621 | -0.0092182 | -0.0074633 | -0.0051474 | 0.0033881 | -0.0017645 | -0.0018352 | 0.0074581 | 0.0084166 | 0.003695 | 0.0052893 | -0.0066313 | -0.00015924 | b12 | 0.0070059 | -0.014432 | -0.014032 | -0.0080436 | 0.0027514 | -0.00089762 | 9.7939e-05 | 0.006717 | -0.0017265 | -0.0074167 | -0.0070425 | -0.005531 | 0.0039652 | 0.00040572 | -0.0044161 | 0.00047947 | 0.0039748 | 0.0057026 | 0.015427 | 0.0060823 | -0.00084791
b13 | 0.086626 | 0.068073 | 0.070735 | 0.069813 | 0.069189 | 0.065847 | 0.065916 | 0.066698 | 0.065384 | 0.068134 | 0.065854 | 0.069599 | 0.069677 | 0.064417 | 0.068536 | 0.070799 | 0.069511 | 0.072239 | 0.071261 | 0.055373 | 0.068567 | b13 | 0.091723 | 0.072072 | 0.074322 | 0.071393 | 0.071618 | 0.070213 | 0.06819 | 0.068931 | 0.068414 | 0.069331 | 0.071461 | 0.072744 | 0.07334 | 0.069654 | 0.071534 | 0.071343 | 0.071725 | 0.075909 | 0.075387 | 0.052824 | 0.071527
b14 | 0.002715 | -0.01204 | -0.0092828 | -0.0041064 | -0.0040206 | -0.0094908 | -0.0064857 | -0.0033944 | -0.0077657 | -0.01077 | -0.0090388 | -0.0085267 | -0.0036022 | -0.006932 | -0.0068679 | -0.0027108 | -0.0010183 | -0.004091 | -0.0029476 | -0.0070982 | -0.0060343 | b14 | -0.00047223 | -0.010524 | -0.0086263 | -0.005125 | -0.002379 | -0.0018529 | -0.0024841 | 0.0019567 | -0.0034252 | -0.0069247 | -0.0057446 | -0.005567 | -0.0017298 | -0.0034476 | -0.004223 | -0.00047124 | -7.6489e-05 | 0.00077493 | 0.0042705 | 7.2318e-06 | -0.0029115
b15 | -0.0074904 | 0.031937 | 0.029342 | 0.028316 | 0.031772 | 0.029398 | 0.026401 | 0.02819 | 0.028012 | 0.02507 | 0.023925 | 0.023523 | 0.028636 | 0.031143 | 0.028608 | 0.031638 | 0.030078 | 0.031113 | 0.028529 | 0.018718 | 0.027228 | b15 | -0.0052435 | 0.029959 | 0.026022 | 0.028365 | 0.029534 | 0.031699 | 0.032964 | 0.029836 | 0.03133 | 0.030731 | 0.028994 | 0.028442 | 0.030729 | 0.032785 | 0.030965 | 0.030116 | 0.031142 | 0.030409 | 0.031393 | 0.019895 | 0.028912
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 14.8417 | -9.7721 | -0.71604 | -1.3602 | 1.4999 | 0.18948 | 2.8937 | 6.8762 | 3.5959 | -2.2852 | -6.4323 | -1.5903 | 2.4625 | 1.9996 | 0.88852 | -1.4055 | -1.7173 | -1.3475 | -1.0969 | -2.2382 | 0 | a1 | 27.5287 | -12.429 | -3.7401 | -3.5501 | -0.10746 | -1.4148 | 2.6999 | 7.3029 | 4.2797 | -0.83149 | -6.1484 | -0.9539 | 1.5262 | 2.3626 | 2.0205 | -1.8422 | -2.1393 | -0.59405 | 0.34357 | -3.914 | 0
a2 | 19.8033 | 1.2327 | 1.1687 | 0.50955 | 1.4413 | 0.43846 | 0.28725 | 1.7699 | -0.30548 | 0.49643 | 0.18339 | 0.61322 | 0.61539 | 0.22939 | 0.16522 | 1.2529 | -0.55973 | 0.0099422 | 0.77393 | -8.3501 | 0.87608 | a2 | 15.6872 | -0.85058 | -1.0286 | -1.2696 | -1.0143 | -1.0312 | -0.96199 | -0.746 | -1.5762 | -1.6939 | -0.99519 | -1.3395 | -0.75678 | -1.926 | -0.53093 | -1.9948 | -2.1442 | -0.97326 | -0.79816 | -1.0664 | -0.69708
a3 | -2.0134 | 1.2027 | 1.0173 | 0.63191 | 0.53865 | 0.19101 | 0.70447 | 0.25536 | 0.56435 | 0.24085 | 1.065 | 1.2998 | 0.92501 | 1.0046 | 0.78851 | 1.3205 | 0.85518 | 1.0255 | 0.63497 | -0.032782 | 0.68071 | a3 | -2.5744 | 0.25633 | 0.56318 | 0.33395 | 0.32794 | 0.37623 | 0.64767 | 0.47576 | 0.26163 | -0.018844 | 0.52169 | 0.32344 | 0.6401 | 0.21643 | 0.30118 | 0.81387 | 0.69802 | 0.8534 | 0.68135 | -0.40503 | 0.34185
a4 | 2.4959 | -0.023828 | -0.079629 | 0.14861 | -0.44843 | -0.65673 | -0.1881 | -1.0541 | -0.51513 | -0.91684 | -0.36368 | -0.33233 | -0.35739 | -0.22054 | -0.21536 | -0.62002 | -0.24029 | -0.10293 | -0.082391 | -0.48589 | -0.26589 | a4 | 2.8529 | 0.01949 | 0.16219 | 0.38646 | 0.68577 | 0.29599 | -0.26756 | -0.19159 | 0.38736 | 0.21708 | -0.039564 | 0.22078 | 0.31836 | -0.074586 | -0.17469 | 0.17199 | 0.16277 | -0.02246 | 0.13146 | 0.47208 | 0.22427
a5 | 2.2263 | 0.19295 | 0.043739 | 0.24455 | 0.42132 | 0.2803 | 0.015986 | 0.021507 | 0.0059866 | 0.60224 | 0.37018 | -0.12351 | -0.050668 | 0.30574 | 0.19469 | 0.1532 | -0.10361 | 0.21133 | 0.26744 | -0.044839 | 0.22554 | a5 | 2.0556 | 0.11829 | 0.18334 | 0.30089 | 0.50097 | 0.36582 | 0.29624 | 0.14544 | 0.24828 | 0.40527 | 0.48403 | 0.18437 | 0.080802 | 0.33743 | 0.25751 | 0.094149 | 0.15749 | 0.16849 | 0.23206 | -0.21722 | 0.29158
a6 | 0.4028 | 0.012517 | 0.015433 | 0.10713 | 0.056685 | -9.4624e-05 | 0.053464 | -0.047122 | -0.0023442 | -0.06985 | 0.10588 | 0.016162 | -0.023376 | -0.0037536 | -0.059841 | 0.0068717 | -0.052514 | -0.0063838 | 0.067456 | -0.24927 | 0.013542 | a6 | 0.53191 | 0.0020569 | -0.041177 | 0.043935 | -0.005471 | 0.000584 | 0.099683 | 0.070897 | -0.0029897 | -0.025456 | 0.074907 | 0.019487 | -0.01663 | 0.060081 | -0.02408 | -0.063741 | -0.061025 | 0.035845 | 0.021966 | -0.19787 | 0.019214
a7 | 1.3245 | -0.037647 | 0.021437 | 0.060762 | 0.090935 | -0.0078994 | -0.049234 | -0.0011802 | 0.054051 | -0.037817 | -0.052787 | -0.041335 | 0.0135 | 0.055547 | 0.011715 | -0.082029 | 0.036152 | 0.0066628 | -0.00058103 | 0.049657 | 0.044071 | a7 | 1.6781 | -0.029423 | 0.039485 | 0.10327 | 0.16802 | 0.059332 | 0.051911 | 0.056369 | 0.12641 | 0.11222 | 0.049887 | 0.064696 | 0.12352 | 0.15759 | 0.15838 | 0.012865 | 0.062831 | 0.057966 | 0.087498 | 0.15817 | 0.13105
a8 | 0.0017814 | -0.02527 | -0.0003777 | 0.024331 | 0.016265 | 0.022346 | 0.028584 | 0.048482 | -0.0027635 | 0.064043 | 0.03958 | 0.015102 | -0.012238 | -0.032358 | -0.031036 | 0.020458 | -0.014796 | -0.0032187 | 0.014193 | -0.028433 | 0.008074 | a8 | 0.026119 | 0.0050764 | 0.014858 | 0.0056456 | 0.012115 | 0.0069206 | -0.022933 | -0.0053096 | -0.050889 | -0.044631 | -0.01135 | -0.014092 | 0.0020416 | -0.019232 | -0.021758 | -0.034243 | -0.029352 | 0.01751 | 0.0018208 | -0.036221 | -0.010182
a9 | -0.30746 | 0.0021069 | -0.0039175 | -0.028766 | -0.075501 | -0.027858 | 0.0026881 | 0.0052043 | -0.040412 | -0.052379 | 0.015398 | -2.4574e-05 | 0.0095474 | -0.030713 | -0.047411 | 0.013356 | 0.0024869 | -0.010672 | -0.005179 | -0.076356 | -0.025977 | a9 | -0.24794 | 0.060686 | 0.02723 | 0.012593 | -0.02708 | 0.018613 | 0.034748 | 0.043173 | 0.012689 | -0.016476 | 0.04051 | 0.010903 | 0.035109 | 0.010607 | -0.0026395 | 0.06127 | 0.025505 | 0.018055 | 0.00054999 | -0.063942 | 0.0096696
a10 | 0.053981 | -0.016444 | 0.014061 | 0.020608 | -0.010174 | -0.030581 | 0.016544 | -0.007692 | 0.0060624 | -0.0045325 | 0.036188 | 0.024333 | 0.0056012 | -0.021565 | -0.011816 | -0.019612 | -0.0069355 | -0.0022909 | 0.026593 | -0.081793 | 0 | a10 | 0.036523 | -0.0015375 | 0.014688 | 0.011685 | 0.03093 | -0.014026 | -0.028552 | 0.020786 | 0.0025015 | -0.020064 | -0.0066772 | 0.025946 | 0.010734 | -0.0044652 | -0.019146 | 0.0073659 | -0.011798 | 0.0040525 | -0.0019352 | -0.066596 | 0
a11 | 0.12238 | -0.083994 | -0.052859 | -0.049959 | -0.034116 | -0.045263 | -0.056917 | -0.040127 | -0.037307 | -0.038034 | -0.055146 | -0.053117 | -0.043127 | -0.045625 | -0.04967 | -0.06551 | -0.053583 | -0.04593 | -0.042976 | -0.018284 | -0.043383 | a11 | 0.19361 | -0.025964 | -0.0068235 | 0.011887 | 0.015437 | 0.013783 | 0.011499 | 0.022851 | 0.021803 | 0.017788 | 0.012306 | 0.0039485 | 0.010992 | 0.020085 | 0.015992 | -0.001509 | 0.0054796 | 0.0060802 | 0.0095275 | 0.02842 | 0.015259
a12 | -0.00032864 | -0.003361 | 0.00067287 | 0.0024165 | -0.00077811 | -0.0032152 | 0.0016344 | 0.00076125 | -0.00014527 | -0.0030438 | 0.0011924 | 0.0026246 | 0.00027986 | -0.0025835 | -0.0010959 | -0.0026883 | -0.0054265 | -0.00025014 | 0.0028945 | -0.011598 | -0.00090574 | a12 | -0.00037057 | -0.0042239 | -0.0037022 | -0.0018462 | -0.00015276 | -0.0015189 | -0.00023677 | 0.0021447 | -8.6004e-06 | -0.0054464 | 0.00029535 | 0.00043907 | 0.00040442 | 0.00019261 | -0.0015515 | -0.0014801 | -0.0078404 | -9.8612e-06 | -0.00044536 | -0.010898 | -0.0016651
a13 | -0.029851 | 0.00025751 | 0.0018174 | -0.0014761 | -0.0010843 | -0.00014086 | -0.0029273 | -0.0037687 | -0.001636 | 0.00057483 | 0.0027021 | -0.0022938 | -0.0017419 | 0.00028046 | -0.0023756 | -0.0029145 | -0.0034655 | -0.0023082 | -0.0032335 | -0.0037271 | -0.0022655 | a13 | -0.019033 | 0.0077473 | 0.0042305 | 0.0021126 | 0.0023807 | 0.0038611 | 0.0033242 | 0.0019127 | 0.0047699 | 0.004403 | 0.007308 | 0.0034029 | 0.0041001 | 0.0041649 | 0.0040189 | 0.0064552 | 0.0023462 | 0.0018075 | 0.0029466 | 0.00023468 | 0.0031594
a14 | 0.0025238 | -0.015499 | -0.012124 | -0.011148 | -0.01441 | -0.012326 | -0.011212 | -0.010461 | -0.011949 | -0.0084183 | -0.010604 | -0.010164 | -0.013958 | -0.015147 | -0.014861 | -0.015203 | -0.014456 | -0.013909 | -0.011271 | -0.0042732 | -0.011891 | a14 | 0.0088066 | -0.011676 | -0.0086985 | -0.0099641 | -0.010605 | -0.010761 | -0.0152 | -0.013385 | -0.013657 | -0.013597 | -0.011877 | -0.011386 | -0.010779 | -0.014456 | -0.01302 | -0.011832 | -0.012677 | -0.010951 | -0.012386 | -0.0037427 | -0.011164
a15 | 0.0051283 | -0.018886 | -0.013372 | -0.0054884 | -0.0059901 | -0.013462 | -0.01142 | -0.0074459 | -0.011393 | -0.018389 | -0.013236 | -0.013008 | -0.0061422 | -0.012634 | -0.012487 | -0.0073086 | -0.0037687 | -0.0076541 | -0.0045672 | -0.017981 | -0.010139 | a15 | 0.0077714 | -0.015134 | -0.014108 | -0.0080209 | -0.0026907 | -0.0030743 | -0.0029293 | 0.0031718 | -0.0051588 | -0.01062 | -0.007799 | -0.00598 | -0.0015724 | -0.0050154 | -0.0078766 | 0.00028603 | -0.00010341 | 0.00048679 | 0.0059378 | -0.0050577 | -0.004113
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.11585 | -0.17664 | -0.055932 | 0.14044 | 0.21504 | -0.075268 | -0.035436 | 0.060465 | -0.045132 | -0.1916 | -0.14909 | -0.11062 | 0.071402 | -0.044466 | -0.029383 | 0.084694 | 0.17818 | 0.09344 | 0.13715 | -0.20016 |  | Dx (mm) | 0.22995 | -0.26943 | -0.26477 | -0.085932 | 0.051327 | 0.035446 | 0.016337 | 0.20951 | -0.047723 | -0.13989 | -0.087985 | -0.052127 | 0.056687 | -0.017441 | -0.077464 | 0.084804 | 0.12653 | 0.13128 | 0.26971 | -0.072235
Dy (mm) | 0.12435 | -0.057621 | 0.028185 | 0.055876 | -0.013254 | -0.073401 | 0.03498 | -0.014239 | 0.011447 | -0.020132 | 0.070402 | 0.047381 | 0.016703 | -0.050592 | -0.028306 | -0.037036 | 0 | 0 | 0.069086 | -0.22093 |  | Dy (mm) | 0.028749 | -0.012736 | 0.030736 | 0.027478 | 0.068232 | -0.031697 | -0.064122 | 0.04029 | 0 | -0.041845 | -0.013377 | 0.058022 | 0.023135 | 0 | -0.040992 | 0.016648 | -0.027532 | 0 | 0 | -0.16901
Mag. len. (m) | 14.4182 | [20 empty] | Mag. len. (m) | 14.4218
TF (mT/KA) | 596.1137 | [20 empty] | TF (mT/KA) | 596.0865
Coil pos. (mm) | 0.020706 | [20 empty] | Coil pos. (mm) | 0.017006
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0022855 | [20 empty] | Twist (m2 rad) | 0.00075662

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.54053 | 1.0583 | 0.69769 | 0.34001 | 0.87608 | 4.8694 | 0.57225 | 0.61212 | -0.24348 | 0.90278 | -0.39378 | 0.33355 | -0.69708 | 3.8017 | -1.2041 | 0.47582 |  | 0
3 | -2.0791 | 8.2156 | -3.1996 | 0.82049 | 0.68071 | 0.72456 | 0.79155 | 0.35122 | -2.1429 | 9.1146 | -3.3555 | 0.77849 | 0.34185 | 0.72963 | 0.46001 | 0.23161 |  | 0
4 | 0.15007 | 0.42989 | 0.21002 | 0.15318 | -0.26589 | 0.70327 | -0.34707 | 0.3102 | -0.15665 | 0.11063 | -0.15574 | 0.11352 | 0.22427 | 0.65012 | 0.13417 | 0.24266 |  | 0
5 | -0.71003 | 1.4449 | -0.43883 | 0.25004 | 0.22554 | 0.49965 | 0.16952 | 0.19244 | -0.76901 | 1.34 | -0.51281 | 0.2282 | 0.29158 | 0.43853 | 0.25313 | 0.12705 |  | 0
6 | 0.030611 | 0.11089 | 0.040855 | 0.080322 | 0.013542 | 0.11854 | 0.0075313 | 0.042886 | -0.017176 | 0.086295 | -0.022992 | 0.038954 | 0.019214 | 0.13528 | 0.0095978 | 0.050368 |  | 0
7 | 1.255 | 0.37358 | 1.212 | 0.096396 | 0.044071 | 0.29853 | 0.0021156 | 0.045687 | 1.2699 | 0.41801 | 1.224 | 0.099984 | 0.13105 | 0.36006 | 0.081122 | 0.052243 |  | 0
8 | 0.0027163 | 0.040829 | 0.004608 | 0.023938 | 0.008074 | 0.02699 | 0.0090362 | 0.025765 | -0.0087782 | 0.038384 | -0.0097357 | 0.015923 | -0.010182 | 0.021961 | -0.010737 | 0.019694 |  | 0
9 | 0.50167 | 0.036267 | 0.50469 | 0.015835 | -0.025977 | 0.0705 | -0.015303 | 0.026541 | 0.51075 | 0.02896 | 0.51099 | 0.013165 | 0.0096696 | 0.065864 | 0.020109 | 0.023901 |  | 0
10 | 0 | 0.055162 | -5.4355e-06 | 3.3069e-05 | 0 | 0.028658 | -5.9951e-05 | 0.00031927 | 0 | 0.063452 | 8.6087e-06 | 6.3071e-05 | 0 | 0.023353 | -4.9953e-05 | 0.0017576 |  | 0
11 | 0.74307 | 0.036167 | 0.7499 | 0.007047 | -0.043383 | 0.040268 | -0.049642 | 0.011638 | 0.75248 | 0.03982 | 0.7596 | 0.0060909 | 0.015259 | 0.042699 | 0.0090843 | 0.011791 |  | 0
12 | -0.00015924 | 0.0053885 | -0.00040971 | 0.0021825 | -0.00090574 | 0.0033465 | -0.00058831 | 0.0013428 | -0.00084791 | 0.0074818 | -0.0011926 | 0.0017135 | -0.0016651 | 0.003175 | -0.0013929 | 0.0023259 |  | 0
13 | 0.068567 | 0.0055657 | 0.068198 | 0.0024926 | -0.0022655 | 0.006613 | -0.0012314 | 0.0015699 | 0.071527 | 0.006658 | 0.07115 | 0.002044 | 0.0031594 | 0.005423 | 0.0039685 | 0.0014203 |  | 0
14 | -0.0060343 | 0.0036147 | -0.0064571 | 0.00122 | -0.011891 | 0.0042728 | -0.012998 | 0.0015589 | -0.0029115 | 0.0036204 | -0.0030734 | 0.00062643 | -0.011164 | 0.005172 | -0.012702 | 0.0013995 |  | 0
15 | 0.027228 | 0.0086157 | 0.028646 | 0.0025675 | -0.010139 | 0.0058073 | -0.01037 | 0.0045099 | 0.028912 | 0.0083139 | 0.030301 | 0.0016659 | -0.004113 | 0.0059592 | -0.0044556 | 0.0055922 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.78481 | -0.92034
3 | [12 empty] | 3 | 9 | -0.62209 | -0.85654
4 | [12 empty] | 4 | 16 | -1.5233 | -1.7959
5 | [12 empty] | 5 | 25 | -1.517 | -1.7704
6 | [12 empty] | 6 | 36 | -2.8355 | -3.1169
7 | [12 empty] | 7 | 49 | -2.7126 | -2.6273
8 | [12 empty] | 8 | 64 | -3.6955 | -4.0337
9 | [12 empty] | 9 | 81 | -3.8873 | -4.032
10 | [12 empty] | sigma (mm) |  | 0.020706 | 0.017006
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -4.2873 | 0.16858 | 0.88822 | 0.9535 | 0.71559 | 1.2413 | 0.8694 | 1.1165 | 0.48285 | 0.14689 | 0.78984 | 0.96524 | 0.88764 | 0.67028 | 0.51183 | 1.0617 | 0.61563 | 0.33076 | 0.14266 | -0.072023 | 0.54053
3 | 33.4717 | -3.9008 | -2.7343 | -3.5606 | -2.7143 | -3.8792 | -2.5731 | -2.7348 | -2.6 | -3.35 | -3.9158 | -2.5199 | -3.0941 | -5.1017 | -4.1505 | -3.9172 | -2.6944 | -2.3172 | -1.8355 | -3.0175 | -2.0791
4 | -1.3727 | -0.02193 | 0.14303 | 0.4697 | 0.42699 | 0.42995 | 0.21069 | 0.20176 | 0.071419 | 0.15863 | 0.23618 | 0.47693 | 0.115 | 0.16659 | 0.15543 | 0.2838 | 0.19479 | 0.029971 | 0.031461 | 0.0021147 | 0.15007
5 | -6.1926 | -0.70047 | -0.83748 | -0.44153 | -0.45681 | -0.1218 | -0.15956 | -0.24144 | -0.3884 | -0.10572 | -0.054699 | -0.3801 | -0.48703 | -0.28788 | -0.37177 | -0.64694 | -0.78103 | -0.65867 | -0.77768 | -3.439 | -0.71003
6 | -0.32533 | 0.20613 | 0.087775 | 0.048702 | 0.13187 | 0.074416 | -0.039784 | -0.0062699 | 0.10164 | 0.15879 | 0.045843 | -0.11905 | 0.0068146 | 0.0086858 | -0.030333 | -0.064081 | 0.026322 | 0.057691 | 0.040237 | 0.011197 | 0.030611
7 | 2.8095 | 0.92984 | 1.0211 | 1.218 | 1.219 | 1.2389 | 1.2785 | 1.295 | 1.2124 | 1.3086 | 1.2903 | 1.2803 | 1.214 | 1.185 | 1.1765 | 1.2324 | 1.2607 | 1.2744 | 1.1807 | 0.99238 | 1.255
8 | -0.088145 | -0.019328 | 0.014121 | 0.044338 | -0.031437 | 0.0079326 | -0.0056092 | 0.0034088 | -0.0050247 | -0.037371 | 0.029323 | 0.027981 | 0.0181 | 0.009989 | -0.012949 | 0.045989 | -0.011188 | -0.014692 | 0.01936 | 0.01664 | 0.0027163
9 | 0.36979 | 0.49867 | 0.50268 | 0.48403 | 0.51416 | 0.47772 | 0.51137 | 0.50587 | 0.53222 | 0.4911 | 0.50959 | 0.49389 | 0.50131 | 0.50373 | 0.48433 | 0.50451 | 0.50774 | 0.53584 | 0.52573 | 0.45437 | 0.50167
10 | -1.0228e-05 | -6.2944e-06 | 1.3919e-06 | 3.3748e-07 | -1.6258e-05 | -4.2175e-06 | -4.8684e-07 | -2.0466e-06 | 8.1596e-07 | 3.8559e-05 | 1.6007e-05 | 2.8371e-06 | -2.0086e-06 | -1.6686e-06 | -1.8952e-07 | -4.2974e-06 | -0.00012999 | 1.0909e-05 | -1.2412e-06 | -0.00015621 | 0
11 | 0.62372 | 0.74282 | 0.74321 | 0.75258 | 0.75157 | 0.75479 | 0.74641 | 0.74785 | 0.75317 | 0.74594 | 0.75691 | 0.74912 | 0.75441 | 0.7648 | 0.76123 | 0.75052 | 0.7394 | 0.7437 | 0.73974 | 0.64712 | 0.74307
12 | -0.038961 | 0.050591 | -0.015602 | -0.032986 | -0.050741 | -0.0076931 | -0.014346 | 0.0071935 | -0.006923 | -0.0001123 | -0.0039854 | 0.0021028 | -0.0014427 | 0.0024741 | -0.0036621 | 0.033017 | -0.0032603 | -0.010697 | -0.017675 | 0.017735 | -0.0015924 | ten times!
13 | 0.86626 | 0.6713 | 0.70077 | 0.69778 | 0.69996 | 0.65992 | 0.65441 | 0.66969 | 0.65012 | 0.66686 | 0.64253 | 0.68509 | 0.69696 | 0.64768 | 0.68703 | 0.71405 | 0.6965 | 0.72532 | 0.70975 | 0.55008 | 0.68567 | ten times!
14 | 0.039548 | -0.064975 | -0.082416 | -0.076338 | -0.095817 | -0.068548 | -0.060443 | -0.047108 | -0.068319 | -0.065088 | -0.068687 | -0.068914 | -0.053705 | -0.052652 | -0.058846 | -0.046946 | -0.054318 | -0.064852 | -0.064299 | -0.0074129 | -0.060343 | ten times!
15 | -0.074904 | 0.31937 | 0.29342 | 0.28316 | 0.31772 | 0.29398 | 0.26401 | 0.2819 | 0.28012 | 0.2507 | 0.23925 | 0.23523 | 0.28636 | 0.31143 | 0.28608 | 0.31638 | 0.30078 | 0.31113 | 0.28529 | 0.18718 | 0.27228 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 19.3423 | 1.2312 | 1.1845 | 0.52245 | 1.4235 | 0.40653 | 0.30077 | 1.7635 | -0.29899 | 0.4939 | 0.23452 | 0.64422 | 0.6137 | 0.20426 | 0.15411 | 1.2227 | -0.57765 | -0.0013307 | 0.77859 | -8.4299 | 0.87608
3 | -2.0292 | 1.2021 | 1.0158 | 0.62377 | 0.55662 | 0.18795 | 0.70198 | 0.26711 | 0.56009 | 0.21044 | 1.0526 | 1.2893 | 0.9292 | 1.0044 | 0.78819 | 1.3316 | 0.86274 | 1.0272 | 0.63694 | -0.043615 | 0.68071
4 | 2.6171 | -0.025406 | -0.073491 | 0.1463 | -0.47113 | -0.65388 | -0.18665 | -1.0552 | -0.51402 | -0.89021 | -0.34977 | -0.33136 | -0.35462 | -0.22077 | -0.21649 | -0.62871 | -0.23595 | -0.10757 | -0.078422 | -0.66745 | -0.26589
5 | 2.2224 | 0.1961 | 0.043434 | 0.23835 | 0.41861 | 0.28119 | 0.017111 | 0.022434 | 0.0056791 | 0.59871 | 0.37525 | -0.12064 | -0.050361 | 0.30556 | 0.19382 | 0.15272 | -0.10086 | 0.21151 | 0.26273 | -0.063091 | 0.22554
6 | 0.22545 | 0.029159 | 0.0056985 | 0.080026 | 0.055495 | 0.031815 | 0.037066 | -0.040585 | -0.006383 | -0.063058 | 0.07106 | -0.0068567 | -0.03088 | 0.018266 | -0.047967 | 0.025467 | -0.054787 | -0.0066035 | 0.038631 | -0.16836 | 0.013542
7 | 1.327 | -0.040939 | 0.021418 | 0.057575 | 0.08961 | -0.0075078 | -0.048613 | -0.0022825 | 0.054076 | -0.033452 | -0.050148 | -0.04069 | 0.01362 | 0.054928 | 0.011103 | -0.08173 | 0.037238 | 0.0067866 | -0.0029098 | 0.044222 | 0.044071
8 | -0.0030079 | -0.011384 | -0.0071546 | 0.013396 | 0.027179 | 0.037868 | 0.020212 | 0.051729 | -0.0064912 | 0.064054 | 0.023654 | 0.0040407 | -0.016509 | -0.02101 | -0.025239 | 0.02874 | -0.015005 | -0.0027495 | -0.0026787 | 0.011486 | 0.008074
9 | -0.31298 | -0.0018733 | -0.0031493 | -0.032186 | -0.073677 | -0.030451 | 0.003265 | 0.0056485 | -0.040148 | -0.053741 | 0.020645 | 0.0025864 | 0.0090517 | -0.031792 | -0.047805 | 0.014941 | 0.0031411 | -0.010558 | -0.0093576 | -0.09422 | -0.025977
10 | -1.4892e-05 | 1.5227e-05 | -2.1635e-06 | -8.9231e-06 | 6.6087e-06 | 6.6404e-06 | -1.1819e-06 | 1.6919e-07 | -2.8565e-07 | 2.7692e-05 | -2.107e-05 | -1.0117e-05 | -1.0355e-06 | 7.3244e-07 | 4.3305e-07 | 9.4708e-06 | -0.0013194 | 0.00023363 | -1.5511e-05 | 4.9534e-05 | 0
11 | 0.12194 | -0.084506 | -0.052754 | -0.050311 | -0.033961 | -0.045622 | -0.056812 | -0.040124 | -0.03729 | -0.038532 | -0.054691 | -0.052771 | -0.043177 | -0.045757 | -0.049725 | -0.065184 | -0.052957 | -0.045915 | -0.043469 | -0.020734 | -0.043383 | ten times!
12 | -0.054911 | -0.0056011 | -0.0066267 | -0.0019073 | 0.00033808 | 0.0018896 | -0.00066396 | 0.015925 | -0.0072573 | -0.019978 | -0.017959 | 0.0011773 | -0.0045385 | -0.0027428 | 0.0022421 | -0.006632 | -0.049907 | -0.00097896 | -0.0026757 | -0.034888 | -0.0090574 | ten times!
13 | -0.30333 | -0.023666 | 0.014989 | -0.0010338 | 0.012446 | -0.01383 | -0.030576 | -0.03322 | -0.019805 | -0.0082441 | 0.019797 | -0.028446 | -0.0093374 | -0.0050276 | -0.028582 | -0.020066 | -0.014959 | -0.013143 | -0.018957 | -0.055803 | -0.022655 | ten times!
14 | 0.028016 | -0.16731 | -0.13421 | -0.11816 | -0.13003 | -0.11383 | -0.12305 | -0.097595 | -0.12637 | -0.10904 | -0.13617 | -0.12267 | -0.13991 | -0.14312 | -0.14496 | -0.13729 | -0.13903 | -0.1332 | -0.12378 | -0.038315 | -0.11891 | ten times!
15 | 0.051283 | -0.18886 | -0.13372 | -0.054884 | -0.059901 | -0.13462 | -0.1142 | -0.074459 | -0.11393 | -0.18389 | -0.13236 | -0.13008 | -0.061422 | -0.12634 | -0.12487 | -0.073086 | -0.037687 | -0.076541 | -0.045672 | -0.17981 | -0.10139 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 2.2061 | -0.03634 | -0.25758 | -0.6175 | 0.232 | -0.060745 | -0.045232 | -0.61288 | -0.33198 | -0.37333 | -0.2966 | -0.20782 | -0.91917 | -1.1614 | -0.50839 | -0.53677 | -0.65581 | -0.38929 | -0.30928 | 0.82283 | -0.24348
3 | 37.3039 | -3.7874 | -2.8034 | -3.6296 | -1.8846 | -4.2336 | -2.7478 | -3.4371 | -2.9025 | -3.8824 | -3.4985 | -2.7315 | -2.5286 | -4.7184 | -4.0287 | -4.574 | -3.3028 | -3.3769 | -2.3306 | -3.0794 | -2.1429
4 | 0.2537 | -0.1845 | -0.32693 | -0.037082 | 0.085569 | -0.032148 | -0.14713 | -0.20156 | -0.17191 | -0.13321 | -0.022459 | -0.1535 | -0.31885 | -0.35833 | -0.13693 | -0.16134 | -0.25432 | -0.14595 | -0.10279 | -0.34589 | -0.15665
5 | -5.5289 | -0.8068 | -0.94321 | -0.51071 | -0.58437 | -0.30321 | -0.25968 | -0.29215 | -0.40016 | -0.25925 | -0.17104 | -0.41686 | -0.50088 | -0.41424 | -0.49642 | -0.52302 | -0.76298 | -0.69479 | -0.8908 | -3.9229 | -0.76901
6 | -0.031923 | -0.089477 | -0.049195 | -0.041305 | -0.028569 | -0.0039679 | -0.039307 | 0.00055961 | -0.007044 | 0.045671 | -0.031291 | -0.089186 | 0.018962 | 0.047088 | -0.022268 | 0.014461 | -0.043785 | -0.046419 | -0.048777 | -0.054156 | -0.017176
7 | 3.0118 | 0.89938 | 1.0593 | 1.2393 | 1.2611 | 1.3046 | 1.2835 | 1.2563 | 1.2575 | 1.3096 | 1.2904 | 1.1976 | 1.2111 | 1.2131 | 1.245 | 1.3128 | 1.2284 | 1.2665 | 1.1965 | 0.95183 | 1.2699
8 | -0.024535 | -0.012631 | -0.024917 | 0.0096045 | 0.0029209 | -0.0014852 | -0.015985 | -0.022825 | 0.0057452 | 0.015388 | -0.029451 | 0.007109 | -0.0077403 | -0.03412 | -0.025593 | -0.0078501 | -0.033546 | -0.0071013 | 0.0072332 | 0.036717 | -0.0087782
9 | 0.39445 | 0.51098 | 0.50328 | 0.49655 | 0.54555 | 0.49459 | 0.52577 | 0.5025 | 0.53103 | 0.51467 | 0.51334 | 0.49886 | 0.50757 | 0.51464 | 0.51262 | 0.49643 | 0.50169 | 0.51651 | 0.51125 | 0.48165 | 0.51075
10 | -1.883e-05 | 0.0001318 | 0.00013583 | 1.0346e-05 | -1.0313e-06 | 5.1524e-07 | -4.6675e-07 | -2.2432e-05 | -8.8811e-05 | 7.2588e-06 | 8.3477e-06 | 1.9773e-08 | -1.5548e-06 | 0.0001112 | 1.6399e-06 | 1.9639e-07 | 4.5912e-06 | -3.985e-05 | -0.00010265 | -9.8758e-06 | 0
11 | 0.6104 | 0.76022 | 0.74983 | 0.75981 | 0.75899 | 0.76766 | 0.75993 | 0.75828 | 0.76081 | 0.75623 | 0.76742 | 0.76372 | 0.76175 | 0.76924 | 0.76382 | 0.75949 | 0.75496 | 0.75608 | 0.74452 | 0.65848 | 0.75248
12 | -0.082684 | -0.0079454 | -0.00049324 | -0.036721 | 0.0027131 | -0.027408 | -0.0083889 | -0.034228 | 0.0057811 | -0.0070074 | -0.026732 | -0.027149 | 0.010975 | 0.012632 | -0.0062087 | -0.037154 | -0.024768 | -0.013319 | 0.010746 | 0.087478 | -0.0084791 | ten times!
13 | 0.92 | 0.70017 | 0.72371 | 0.70847 | 0.71158 | 0.70524 | 0.68966 | 0.68206 | 0.68289 | 0.69026 | 0.71196 | 0.72017 | 0.73225 | 0.69608 | 0.71692 | 0.71223 | 0.71999 | 0.75831 | 0.74506 | 0.53308 | 0.71527 | ten times!
14 | 0.0070472 | -0.037174 | -0.033094 | -0.032992 | -0.037786 | -0.026979 | -0.027729 | -0.030858 | -0.021938 | -0.030185 | -0.035579 | -0.046318 | -0.031942 | -0.029767 | -0.019817 | -0.025706 | -0.033192 | -0.025128 | -0.027023 | 0.018947 | -0.029115 | ten times!
15 | -0.052435 | 0.29959 | 0.26022 | 0.28365 | 0.29534 | 0.31699 | 0.32964 | 0.29836 | 0.3133 | 0.30731 | 0.28994 | 0.28442 | 0.30729 | 0.32785 | 0.30965 | 0.30116 | 0.31142 | 0.30409 | 0.31393 | 0.19895 | 0.28912 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 15.6307 | -0.84812 | -1.001 | -1.2545 | -1.001 | -1.0486 | -0.98399 | -0.7414 | -1.5748 | -1.7133 | -0.99529 | -1.3189 | -0.75416 | -1.9256 | -0.54761 | -1.994 | -2.1653 | -0.98644 | -0.81978 | -1.1307 | -0.69708
3 | -2.6899 | 0.25696 | 0.57223 | 0.33995 | 0.32088 | 0.3742 | 0.64681 | 0.48439 | 0.26489 | -0.014376 | 0.52104 | 0.32701 | 0.63825 | 0.2162 | 0.29789 | 0.81181 | 0.69305 | 0.85392 | 0.67509 | -0.40759 | 0.34185
4 | 2.7791 | 0.024613 | 0.18035 | 0.39584 | 0.6891 | 0.29068 | -0.27259 | -0.19603 | 0.39015 | 0.22787 | -0.030075 | 0.22871 | 0.32 | -0.073201 | -0.17477 | 0.17216 | 0.15314 | -0.027664 | 0.11673 | 0.31206 | 0.22427
5 | 2.018 | 0.1178 | 0.18148 | 0.30265 | 0.50125 | 0.36593 | 0.29518 | 0.13997 | 0.24824 | 0.40401 | 0.48569 | 0.18659 | 0.080792 | 0.33774 | 0.25631 | 0.095476 | 0.15986 | 0.1671 | 0.23032 | -0.22485 | 0.29158
6 | 0.36514 | 0.0033414 | -0.049083 | 0.035045 | -0.038886 | 0.014439 | 0.12842 | 0.04883 | -0.00086052 | -0.00056934 | 0.082557 | -0.003858 | -0.028991 | 0.061051 | -0.0017225 | -0.071844 | -0.0519 | 0.03316 | 0.013637 | -0.13724 | 0.019214
7 | 1.6754 | -0.029276 | 0.042245 | 0.10359 | 0.16728 | 0.059322 | 0.051778 | 0.056381 | 0.12541 | 0.10932 | 0.049197 | 0.064523 | 0.12342 | 0.15745 | 0.15693 | 0.013981 | 0.064325 | 0.05702 | 0.087295 | 0.15817 | 0.13105
8 | 0.047488 | 0.015937 | 0.010738 | -0.00028652 | -0.0047316 | 0.013989 | -0.007362 | -0.019281 | -0.050604 | -0.035499 | -0.0064298 | -0.027442 | -0.0044277 | -0.019145 | -0.011946 | -0.04059 | -0.024324 | 0.016394 | 0.001751 | -0.00048539 | -0.010182
9 | -0.25359 | 0.059651 | 0.031191 | 0.013686 | -0.028726 | 0.01915 | 0.035257 | 0.038884 | 0.012752 | -0.019331 | 0.039918 | 0.012343 | 0.034477 | 0.010566 | -0.0041718 | 0.060606 | 0.027116 | 0.017774 | 0.00082631 | -0.068735 | 0.0096696
10 | -5.2603e-06 | 6.158e-05 | 2.3359e-05 | -3.0677e-06 | -3.0795e-06 | -6.0122e-07 | 4.8543e-07 | -2.1912e-05 | 0.0031136 | 2.7212e-05 | 1.5396e-06 | -5.3501e-06 | -1.4822e-06 | -0.0042591 | 4.4579e-06 | 2.98e-06 | 2.4033e-05 | 0.0035829 | -0.0034467 | -7.1442e-05 | 0
11 | 0.19354 | -0.026819 | -0.0071787 | 0.011927 | 0.01532 | 0.0138 | 0.011505 | 0.022385 | 0.021803 | 0.017094 | 0.012262 | 0.0041709 | 0.010918 | 0.020088 | 0.015797 | -0.0014329 | 0.0061924 | 0.006081 | 0.0096052 | 0.028575 | 0.015259
12 | 0.008574 | -0.021025 | -0.04524 | -0.031028 | -0.036884 | -0.00044521 | 0.028113 | -0.00098591 | 0.0015208 | -0.029638 | 0.01424 | -0.024151 | -0.0095735 | 0.0024389 | 0.007381 | -0.02705 | -0.066561 | -0.0017737 | -0.010064 | -0.045842 | -0.016651 | ten times!
13 | -0.20571 | 0.052391 | 0.02672 | 0.015656 | 0.029211 | 0.041079 | 0.033923 | 0.039969 | 0.042715 | 0.02727 | 0.064501 | 0.03196 | 0.045979 | 0.039721 | 0.031152 | 0.072285 | 0.035712 | 0.029069 | 0.055012 | 0.0002887 | 0.031594 | ten times!
14 | 0.074591 | -0.1472 | -0.12433 | -0.11174 | -0.12151 | -0.098434 | -0.1342 | -0.14923 | -0.1386 | -0.13761 | -0.12123 | -0.13002 | -0.11292 | -0.14528 | -0.12477 | -0.12265 | -0.1196 | -0.11003 | -0.13705 | -0.012745 | -0.11164 | ten times!
15 | 0.077714 | -0.15134 | -0.14108 | -0.080209 | -0.026907 | -0.030743 | -0.029293 | 0.031718 | -0.051588 | -0.1062 | -0.07799 | -0.0598 | -0.015724 | -0.050154 | -0.078766 | 0.0028603 | -0.0010341 | 0.0048679 | 0.059378 | -0.050577 | -0.04113 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00043402 | -5.5144e-05 | 5.2721e-05 | -9.6076e-05 | -4.7931e-05 | 4.9366e-05 | -4.5666e-07 | -0.00020545 | -0.00012896 | -8.0307e-05 | -6.8229e-05 | -9.6076e-05 | -6.9571e-05 | -2.109e-05 | 1.1957e-05 | 2.5881e-05 | -0.00012476 | 0.0004201 |  | 596.1137
|  | 0.00056564 | -2.7895e-05 | 1.8743e-05 | -4.3161e-05 | 2.2266e-05 | -1.5481e-05 | -8.7282e-05 | -3.7961e-05 | -8.3592e-05 | -8.5269e-05 | -1.8333e-05 | -8.5269e-05 | -7.2855e-05 | -0.00015254 | -0.00022518 | 2.7802e-05 | -9.1476e-05 | 0.00039184 |  | 596.0865
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -4018.407
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0819 | -4144.4921 | 4.8743 | -0.017615 | 1.0611 | -0.42695 | 0.054523 | 1.0275 | 0.52929 | -1.5208 | -0.75577 | -0.26926 | -0.14847 | -0.42695 | -0.16189 | 0.32294 | 0.65343 | 0.79268 | -0.71383 | 4.7351 | -3892.322
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.040654 | 1.5248 | -0.93656 | -0.03095 | -0.095368 | 0.19064 | 0.059602 | 0.33003 | 0.72828 | 0.40024 | -0.18786 | -0.60257 | -0.11838 | 0.2869 | 0.24061 | 0.12951 | -0.0999 | -0.13108 | -0.094093 | -0.06904 | -0.18316
db2 | 0.70021 | -3.8028 | -0.44239 | 0.20264 | 0.19047 | -0.052078 | 0.5772 | 0.17702 | 0.39724 | -0.20431 | -0.47703 | 0.14952 | 0.29066 | 0.15965 | 0.0027578 | -0.1714 | 0.3282 | -0.14094 | -0.39491 | -0.59231 | -0.0012382
db3 | -3.2 | 33.3841 | -0.70144 | 0.46456 | -0.35093 | 0.49998 | -0.69359 | 0.62673 | 0.46463 | 0.60038 | -0.1589 | -0.71742 | 0.67334 | 0.10825 | -1.9051 | -0.95247 | -0.7172 | 0.51068 | 0.88303 | 1.3655 | -3.0421
db4 | 0.20682 | -1.6073 | -0.19672 | -0.053046 | 0.24516 | 0.1991 | 0.23014 | 0.005061 | -0.0084293 | -0.13129 | -0.040375 | 0.025189 | 0.28124 | -0.099791 | -0.033585 | -0.047524 | 0.065407 | -0.044544 | -0.19127 | -0.20472 | 0.16236
db5 | -0.43787 | -6.2165 | -0.27003 | -0.40084 | -0.00095908 | -0.0044907 | 0.31515 | 0.27833 | 0.19661 | 0.048386 | 0.32782 | 0.38189 | 0.063023 | -0.048265 | 0.15005 | 0.065965 | -0.2097 | -0.33762 | -0.21806 | -0.33725 | -3.4519
db6 | 0.04325 | -0.26868 | 0.10391 | 0.024155 | 0.064819 | 0.18171 | -0.0019754 | -0.098413 | -0.021874 | 0.038851 | 0.026296 | -0.063952 | -0.21158 | -0.0058929 | -0.05216 | -0.085666 | -0.071486 | 0.0625 | 0.056481 | 0.054284 | -0.055053
db7 | 1.2134 | 2.807 | -0.27972 | -0.19233 | 0.0084566 | 0.0075476 | 0.026426 | 0.064877 | 0.082338 | -0.00067736 | 0.10228 | 0.076145 | 0.066538 | 0.0017006 | -0.029032 | -0.037059 | 0.021614 | 0.049626 | 0.061349 | -0.030078 | 0.99259
db8 | 0.0056781 | -0.050076 | -0.066982 | -0.0047493 | 0.071679 | 0.01553 | -0.015637 | -0.019858 | 0.012182 | -0.021798 | -0.088596 | -0.012878 | -0.0035174 | 0.029231 | -0.0069594 | -0.025957 | 0.060701 | 0.02632 | 0.0032809 | 0.048008 | -0.033981
db9 | 0.50766 | 0.36829 | -0.0020035 | -0.0044843 | -0.019494 | 0.017234 | -0.029715 | 0.0036762 | -0.0010018 | 0.025001 | -0.0080186 | 0.0056467 | -0.011612 | -0.0051818 | -0.0039993 | -0.023282 | -0.0019226 | 0.0073931 | 0.030202 | 0.021563 | 0.45314
db10 | 0.0015756 | 0.033574 | -0.081641 | -0.025155 | 0.062259 | 0.093305 | -0.036955 | -0.015964 | 0.02469 | -0.021321 | -0.086185 | -0.065714 | -0.048855 | 0.030539 | -0.022939 | -0.01556 | 0.034401 | 0.075893 | 0.039325 | 0.059878 | -0.078631
db11 | 0.75022 | 0.62418 | -0.0071307 | -0.0068596 | 0.0026064 | 0.0020931 | 0.0046234 | -0.0037805 | -0.0022287 | 0.0030357 | -0.0031719 | 0.0073532 | -0.00081482 | 0.0043392 | 0.014547 | 0.011027 | 0.00064708 | -0.009851 | -0.0062925 | -0.010143 | 0.64632
db12 | -0.0001339 | 0.005808 | -0.0032846 | -0.0042552 | 0.0038144 | 0.005552 | -0.0041411 | -0.0028772 | 0.0036621 | -0.0026282 | -0.0090843 | -0.0073294 | -0.0050135 | 0.003522 | -0.0016306 | -0.0017013 | 0.007592 | 0.0085505 | 0.0038289 | 0.0054232 | -0.0066313
db13 | 0.068427 | 0.086626 | -0.00035359 | 0.0023082 | 0.0013861 | 0.00076228 | -0.0025802 | -0.0025107 | -0.0017285 | -0.0030426 | -0.0002924 | -0.0025724 | 0.0011717 | 0.0012505 | -0.0040096 | 0.00010881 | 0.0023719 | 0.0010843 | 0.0038123 | 0.0028339 | 0.055373
db14 | -0.0062828 | 0.002715 | -0.0057568 | -0.0029999 | 0.0021764 | 0.0022622 | -0.003208 | -0.00020291 | 0.0028884 | -0.0014828 | -0.0044868 | -0.0027559 | -0.0022439 | 0.0026806 | -0.00064914 | -0.00058511 | 0.003572 | 0.0052646 | 0.0021918 | 0.0033352 | -0.0070982
db15 | 0.028646 | -0.0074904 | 0.0032907 | 0.00069547 | -0.00033061 | 0.0031255 | 0.00075209 | -0.0022454 | -0.00045592 | -0.00063448 | -0.0035756 | -0.004721 | -0.0051233 | -1.0584e-05 | 0.0024969 | -3.8024e-05 | 0.0029919 | 0.0014319 | 0.0024672 | -0.0001167 | 0.018718
da2 | 0.57347 | 19.8033 | 0.65919 | 0.59522 | -0.063918 | 0.86788 | -0.13501 | -0.28622 | 1.1965 | -0.87895 | -0.077043 | -0.39008 | 0.03975 | 0.041924 | -0.34408 | -0.40825 | 0.67942 | -1.1332 | -0.56353 | 0.20046 | -8.3501
da3 | 0.79254 | -2.0134 | 0.41015 | 0.22479 | -0.16063 | -0.25389 | -0.60153 | -0.088076 | -0.53718 | -0.2282 | -0.55169 | 0.27247 | 0.5073 | 0.13246 | 0.21207 | -0.0040374 | 0.528 | 0.062639 | 0.23292 | -0.15757 | -0.032782
da4 | -0.34828 | 2.4959 | 0.32445 | 0.26865 | 0.49689 | -0.10015 | -0.30845 | 0.16018 | -0.70578 | -0.16685 | -0.56856 | -0.015404 | 0.015954 | -0.0091054 | 0.12774 | 0.13292 | -0.27174 | 0.10799 | 0.24536 | 0.26589 | -0.48589
da5 | 0.16963 | 2.2263 | 0.023322 | -0.12589 | 0.074923 | 0.25169 | 0.11067 | -0.15365 | -0.14813 | -0.16365 | 0.43261 | 0.20055 | -0.29314 | -0.2203 | 0.13611 | 0.02506 | -0.016436 | -0.27324 | 0.041701 | 0.097803 | -0.044839
da6 | 0.0097958 | 0.4028 | 0.0027214 | 0.0056369 | 0.097338 | 0.046889 | -0.0098904 | 0.043668 | -0.056917 | -0.01214 | -0.079646 | 0.096085 | 0.0063658 | -0.033172 | -0.013549 | -0.069637 | -0.0029241 | -0.06231 | -0.01618 | 0.057661 | -0.24927
da7 | 0.0022362 | 1.3245 | -0.039883 | 0.019201 | 0.058526 | 0.088698 | -0.010136 | -0.05147 | -0.0034164 | 0.051815 | -0.040054 | -0.055023 | -0.043572 | 0.011264 | 0.053311 | 0.0094784 | -0.084265 | 0.033916 | 0.0044266 | -0.0028172 | 0.049657
da8 | 0.0095181 | 0.0017814 | -0.034788 | -0.0098958 | 0.014813 | 0.0067468 | 0.012828 | 0.019066 | 0.038964 | -0.012282 | 0.054525 | 0.030062 | 0.0055843 | -0.021756 | -0.041876 | -0.040554 | 0.01094 | -0.024315 | -0.012737 | 0.0046745 | -0.028433
da9 | -0.015114 | -0.30746 | 0.017221 | 0.011196 | -0.013652 | -0.060387 | -0.012744 | 0.017802 | 0.020318 | -0.025299 | -0.037265 | 0.030511 | 0.015089 | 0.024661 | -0.015599 | -0.032298 | 0.028469 | 0.017601 | 0.0044419 | 0.0099347 | -0.076356
da10 | 0.0010193 | 0.053981 | -0.017463 | 0.013042 | 0.019589 | -0.011194 | -0.031601 | 0.015525 | -0.0087113 | 0.0050431 | -0.0055518 | 0.035169 | 0.023314 | 0.0045819 | -0.022585 | -0.012836 | -0.020632 | -0.0079548 | -0.0033102 | 0.025574 | -0.081793
da11 | -0.049626 | 0.12238 | -0.034369 | -0.0032331 | -0.00033311 | 0.01551 | 0.0043628 | -0.0072919 | 0.0094983 | 0.012319 | 0.011591 | -0.00552 | -0.0034916 | 0.0064983 | 0.0040008 | -4.4546e-05 | -0.015885 | -0.0039573 | 0.0036955 | 0.0066496 | -0.018284
da12 | -0.00056175 | -0.00032864 | -0.0027993 | 0.0012346 | 0.0029783 | -0.00021636 | -0.0026535 | 0.0021962 | 0.001323 | 0.00041647 | -0.0024821 | 0.0017541 | 0.0031863 | 0.00084161 | -0.0020218 | -0.00053418 | -0.0021266 | -0.0048648 | 0.00031161 | 0.0034563 | -0.011598
da13 | -0.0013186 | -0.029851 | 0.0015761 | 0.003136 | -0.00015752 | 0.00023425 | 0.0011777 | -0.0016087 | -0.0024502 | -0.00031749 | 0.0018934 | 0.0040206 | -0.00097521 | -0.00042332 | 0.001599 | -0.0010571 | -0.0015959 | -0.002147 | -0.00098963 | -0.001915 | -0.0037271
da14 | -0.012618 | 0.0025238 | -0.0028813 | 0.00049365 | 0.0014698 | -0.0017925 | 0.00029176 | 0.0014061 | 0.002157 | 0.00066877 | 0.0041995 | 0.0020134 | 0.0024537 | -0.0013401 | -0.0025289 | -0.0022428 | -0.0025857 | -0.0018378 | -0.0012915 | 0.0013469 | -0.0042732
da15 | -0.01037 | 0.0051283 | -0.0085163 | -0.003002 | 0.0048813 | 0.0043796 | -0.0030928 | -0.0010508 | 0.0029237 | -0.0010237 | -0.0080197 | -0.0028665 | -0.0026385 | 0.0042275 | -0.002264 | -0.0021178 | 0.0030611 | 0.0066009 | 0.0027156 | 0.0058024 | -0.017981
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4061.1494
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0509 | -4315.6327 | 6.2552 | 0.3195 | 0.7859 | 0.16683 | 0.82113 | 0.44365 | -0.27441 | 0.21884 | -0.2375 | -0.25428 | 0.41513 | -0.25428 | -0.13013 | -0.92704 | -1.6535 | 0.8765 | -0.31635 | 4.5171 | -3806.6661
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.073418 | 2.8263 | -1.1695 | -0.3006 | -0.28159 | 0.062672 | -0.068059 | 0.34341 | 0.80371 | 0.50139 | -0.009731 | -0.54143 | -0.021972 | 0.22604 | 0.30968 | 0.27547 | -0.11081 | -0.14051 | 0.014013 | 0.10777 | -0.31798
db2 | -0.39028 | 3.2236 | 0.47417 | 0.21762 | -0.1916 | 0.60822 | 0.31329 | 0.34464 | -0.30974 | 0.074585 | 0.080673 | 0.13072 | 0.19701 | -0.54753 | -0.76143 | -0.079947 | -0.19375 | -0.31251 | -0.051232 | 0.0068118 | 0.84079
db3 | -3.3564 | 37.287 | -0.4246 | 0.5648 | -0.27459 | 1.4642 | -0.87569 | 0.60515 | -0.087318 | 0.45534 | -0.5214 | -0.14195 | 0.62408 | 0.82326 | -1.3609 | -0.67188 | -1.2207 | 0.048257 | -0.024368 | 1.0183 | -3.0616
db4 | -0.15523 | -0.05887 | 0.021345 | -0.11502 | 0.12649 | 0.22569 | 0.12328 | 0.011574 | -0.06187 | -0.012193 | 0.034527 | 0.13779 | 0.0043043 | -0.17073 | -0.2014 | 0.02981 | -0.016884 | -0.12079 | -0.012167 | -0.0037613 | -0.28771
db5 | -0.51017 | -5.5202 | -0.28268 | -0.42102 | 0.0011249 | -0.074236 | 0.2071 | 0.25222 | 0.2229 | 0.1104 | 0.25132 | 0.34126 | 0.094699 | 0.01011 | 0.095724 | 0.014693 | -0.011195 | -0.25287 | -0.18415 | -0.3754 | -3.9313
db6 | -0.022111 | 0.19562 | -0.15386 | -0.1274 | -0.057794 | 0.012343 | 0.035131 | -0.0086248 | 0.11494 | -0.0061165 | 0.0047659 | -0.049071 | -0.090443 | 0.064297 | 0.06173 | -0.031932 | 0.075809 | 0.033772 | 0.034438 | 0.088018 | -0.068912
db7 | 1.2259 | 3.0136 | -0.31773 | -0.157 | 0.013734 | 0.035231 | 0.078907 | 0.05685 | 0.032772 | 0.031731 | 0.084001 | 0.066235 | -0.027836 | -0.014726 | -0.012521 | 0.020103 | 0.087526 | 0.0019705 | 0.041914 | -0.021156 | 0.94885
db8 | -0.010598 | 0.022489 | -0.068643 | -0.079432 | -0.0001044 | 0.027553 | 0.017642 | -0.00030337 | 0.037471 | 0.004405 | -0.0084915 | -0.039927 | 0.0051756 | 0.016032 | -0.027746 | -0.033769 | 0.022145 | 0.007469 | 0.035658 | 0.084866 | 0.015429
db9 | 0.51531 | 0.40357 | 0.012909 | 0.0041179 | -0.017171 | 0.029728 | -0.020645 | 0.0095575 | -0.0029378 | 0.016268 | 0.0034955 | -0.00017298 | -0.016599 | -0.0071144 | -0.00059538 | -0.0016967 | -0.017239 | -0.010017 | 0.0052584 | 0.012854 | 0.47665
db10 | -0.0016424 | 0.079413 | -0.1193 | -0.11527 | -0.036971 | 0.02394 | 0.017905 | 0.0093785 | 0.094635 | -0.019806 | -0.060183 | -0.037992 | -0.021916 | 0.026897 | -0.006138 | -0.032783 | 0.039547 | 0.057941 | 0.060029 | 0.12009 | -0.025102
db11 | 0.76022 | 0.61145 | 0.0024804 | -0.007915 | 6.9689e-05 | -0.0011288 | 0.0073854 | -0.00029851 | -0.0010855 | 0.00064319 | -0.0034632 | 0.0076063 | 0.0036724 | 0.0016677 | 0.0090129 | 0.003777 | -0.00068851 | -0.0050711 | -0.0036551 | -0.013009 | 0.657
db12 | -0.0013343 | 0.0070059 | -0.013098 | -0.012697 | -0.0067094 | 0.0040857 | 0.00043663 | 0.0014322 | 0.0080512 | -0.00039229 | -0.0060825 | -0.0057082 | -0.0041967 | 0.0052994 | 0.00174 | -0.0030818 | 0.0018137 | 0.0053091 | 0.0070368 | 0.016761 | 0.0060823
db13 | 0.071532 | 0.091723 | 0.00053947 | 0.0027893 | -0.00013952 | 8.5819e-05 | -0.0013193 | -0.0033427 | -0.0026009 | -0.0031186 | -0.0022009 | -7.0951e-05 | 0.001212 | 0.001808 | -0.0018779 | 1.8194e-06 | -0.00018937 | 0.00019244 | 0.0043763 | 0.0038548 | 0.052824
db14 | -0.0030888 | -0.00047223 | -0.0074348 | -0.0055375 | -0.0020362 | 0.00070978 | 0.0012359 | 0.0006047 | 0.0050455 | -0.00033637 | -0.0038359 | -0.0026558 | -0.0024782 | 0.001359 | -0.00035882 | -0.0011342 | 0.0026176 | 0.0030123 | 0.0038637 | 0.0073593 | 7.2318e-06
db15 | 0.030301 | -0.0052435 | -0.00034204 | -0.0042792 | -0.0019357 | -0.00076674 | 0.0013982 | 0.002663 | -0.00046523 | 0.0010293 | 0.00043029 | -0.0013068 | -0.0018587 | 0.00042795 | 0.0024845 | 0.00066406 | -0.00018449 | 0.00084126 | 0.00010811 | 0.0010922 | 0.019895
da2 | -1.2017 | 15.6872 | 0.35116 | 0.17314 | -0.067886 | 0.18749 | 0.17053 | 0.23975 | 0.45574 | -0.37449 | -0.49213 | 0.20655 | -0.1378 | 0.44496 | -0.7243 | 0.67081 | -0.7931 | -0.94249 | 0.22848 | 0.40358 | -1.0664
da3 | 0.45963 | -2.5744 | -0.2033 | 0.10355 | -0.12568 | -0.13169 | -0.083397 | 0.18804 | 0.016133 | -0.198 | -0.47847 | 0.062061 | -0.13619 | 0.18047 | -0.2432 | -0.15845 | 0.35424 | 0.23839 | 0.39377 | 0.22172 | -0.40503
da4 | 0.13274 | 2.8529 | -0.11325 | 0.029455 | 0.25373 | 0.55303 | 0.16325 | -0.4003 | -0.32432 | 0.25462 | 0.08434 | -0.1723 | 0.088041 | 0.18562 | -0.20732 | -0.30743 | 0.039255 | 0.030037 | -0.1552 | -0.001275 | 0.47208
da5 | 0.25338 | 2.0556 | -0.13509 | -0.070043 | 0.047509 | 0.24759 | 0.11243 | 0.042856 | -0.10795 | -0.0051037 | 0.15189 | 0.23065 | -0.069013 | -0.17258 | 0.08405 | 0.0041312 | -0.15923 | -0.095888 | -0.084894 | -0.021317 | -0.21722
da6 | 0.010493 | 0.53191 | -0.008436 | -0.05167 | 0.033442 | -0.015964 | -0.0099089 | 0.089191 | 0.060404 | -0.013483 | -0.035949 | 0.064414 | 0.0089939 | -0.027123 | 0.049588 | -0.034572 | -0.074234 | -0.071518 | 0.025353 | 0.011473 | -0.19787
da7 | 0.081268 | 1.6781 | -0.11069 | -0.041782 | 0.022002 | 0.086749 | -0.021936 | -0.029356 | -0.024899 | 0.045139 | 0.030955 | -0.031381 | -0.016572 | 0.042249 | 0.076325 | 0.077109 | -0.068403 | -0.018437 | -0.023301 | 0.00623 | 0.15817
da8 | -0.010433 | 0.026119 | 0.01551 | 0.025291 | 0.016079 | 0.022548 | 0.017354 | -0.012499 | 0.0051237 | -0.040455 | -0.034197 | -0.00091651 | -0.0036586 | 0.012475 | -0.0087986 | -0.011324 | -0.023809 | -0.018919 | 0.027943 | 0.012254 | -0.036221
da9 | 0.020336 | -0.24794 | 0.04035 | 0.0068944 | -0.007743 | -0.047416 | -0.0017225 | 0.014412 | 0.022837 | -0.0076473 | -0.036812 | 0.020174 | -0.0094327 | 0.014773 | -0.0097285 | -0.022975 | 0.040934 | 0.0051695 | -0.0022805 | -0.019786 | -0.063942
da10 | 0.0011382 | 0.036523 | -0.0026757 | 0.01355 | 0.010547 | 0.029792 | -0.015164 | -0.02969 | 0.019648 | 0.0013634 | -0.021202 | -0.0078154 | 0.024808 | 0.0095954 | -0.0056034 | -0.020284 | 0.0062277 | -0.012936 | 0.0029143 | -0.0030734 | -0.066596
da11 | 0.0091757 | 0.19361 | -0.035139 | -0.015999 | 0.002711 | 0.0062612 | 0.0046076 | 0.0023228 | 0.013675 | 0.012628 | 0.0086118 | 0.0031305 | -0.0052272 | 0.0018162 | 0.01091 | 0.0068166 | -0.010685 | -0.0036961 | -0.0030955 | 0.00035175 | 0.02842
da12 | -0.0013882 | -0.00037057 | -0.0028357 | -0.002314 | -0.00045801 | 0.0012354 | -0.00013075 | 0.0011514 | 0.0035329 | 0.0013796 | -0.0040583 | 0.0016835 | 0.0018272 | 0.0017926 | 0.0015808 | -0.00016334 | -9.199e-05 | -0.0064523 | 0.0013783 | 0.00094279 | -0.010898
da13 | 0.0039607 | -0.019033 | 0.0037866 | 0.0002698 | -0.0018481 | -0.00158 | -9.9576e-05 | -0.00063644 | -0.002048 | 0.00080921 | 0.00044235 | 0.0033473 | -0.00055778 | 0.00013938 | 0.00020426 | 5.8174e-05 | 0.0024945 | -0.0016145 | -0.0021532 | -0.001014 | 0.00023468
da14 | -0.01205 | 0.0088066 | 0.00037441 | 0.0033519 | 0.0020863 | 0.0014455 | 0.0012899 | -0.0031498 | -0.001335 | -0.0016068 | -0.0015461 | 0.0001736 | 0.00066462 | 0.0012709 | -0.0024059 | -0.00096931 | 0.00021802 | -0.00062656 | 0.0010998 | -0.00033551 | -0.0037427
da15 | -0.0044556 | 0.0077714 | -0.010678 | -0.0096519 | -0.0035654 | 0.0017649 | 0.0013812 | 0.0015263 | 0.0076274 | -0.00070325 | -0.0061647 | -0.0033434 | -0.0015244 | 0.0028831 | -0.00055986 | -0.0034211 | 0.0047416 | 0.0043521 | 0.0049424 | 0.010393 | -0.0050577
Field Colinearity | 19.3857
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 1.4842 | -0.97721 | -0.071604 | -0.13602 | 0.14999 | 0.018948 | 0.28937 | 0.68762 | 0.35959 | -0.22852 | -0.64323 | -0.15904 | 0.24625 | 0.19996 | 0.088852 | -0.14055 | -0.17173 | -0.13475 | -0.10969 | -0.22382
Twist (m2 mrad) | 11.1313 | -13.9253 | -0.96665 | -1.7343 | 1.7999 | 0.21316 | 3.0384 | 6.7043 | 3.2363 | -1.8853 | -4.8242 | -1.0735 | 1.4775 | 1.0498 | 0.39983 | -0.52707 | -0.51519 | -0.30318 | -0.16454 | -0.083931
Twist Integral (I0) (m2 rad) | 0.0022855
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.7529 | -1.2429 | -0.37401 | -0.35501 | -0.010746 | -0.14148 | 0.26999 | 0.7303 | 0.42797 | -0.083149 | -0.61484 | -0.09539 | 0.15262 | 0.23626 | 0.20205 | -0.18422 | -0.21393 | -0.059405 | 0.034357 | -0.39139
Twist (m2 mrad) | 20.6465 | -17.7113 | -5.0492 | -4.5263 | -0.12895 | -1.5916 | 2.8349 | 7.1204 | 3.8518 | -0.68598 | -4.6113 | -0.64388 | 0.91574 | 1.2404 | 0.90923 | -0.69084 | -0.64177 | -0.13366 | 0.051535 | -0.14677
Twist Integral (I0) (m2 rad) | 0.00075662

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.70021 | -3.2 | 0.20682 | -0.43787 | 1.2134 | -0.40654 | 0.57347 | 0.79254 | -0.34828 | 20.706 | 0.59611 | -0.39028 | -3.3564 | -0.15523 | -0.51017 | 1.2259 | -0.73418 | -1.2017 | 0.45963 | 0.13274 | 17.0059 | 0.59609 | 0.54053 | -2.0791 | 0.15007 | -0.71003 | 1.255 | 0 | 0.87608 | 0.68071 | -0.26589 | 14418.171 | -0.24348 | -2.1429 | -0.15665 | -0.76901 | 1.2699 | 0 | -0.69708 | 0.34185 | 0.22427 | 14421.756
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8