## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3088C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-10-22
Beginn der Messung / start of measuring: |  | TT.10.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000088_cc
Ende der Messung / end of measuring: |  | TT.10.JJ
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Mahler |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-10-22 |  | 2003-10-23 | 2003-10-23
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000008_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000088
Date of test Ap 1 | 2003-10-22 |  |  |  | Date of test Ap 2 | 2003-10-22
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | yellow alarm | status ok | status ok | b7 | status ok | yellow alarm | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
208 | HCMB__A001 | 3000088 |  | 3 |  |  | 3
| COMMENTS: | Warning due to yellow alarms on b9, head CS, both ap (4.4 & 4.6 sigma) and b7, pos. 2, both apertures (5.5 & 5.8 sigma).
| HOLDING POINT | OK-w | Data at CERN |  | 2003-10-23 | Analysis by |  | V. Remondino | Answer to MMS-LD |  |  | 2003-10-23 | Answer to manufacturer |  |  | 2003-10-23
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.15 | 0.2 | -0.05 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.15 | 0.15 | 0.8 | Up | 0.8 | 0.15 | 0.15 | 0.8
[8 empty] | 0.8 | 0.15 | 0.15 | 0.8 | Down | 0.8 | 0.15 | 0.15 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000008_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000088 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3088_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3088_cct2 |  |  |  | a even | 1
Test Operator |  | Mahler |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-10-22 |  |  | Time of test |  |  | 14:29 |  |  |  | Date of test |  | 2003-10-22 |  |  | Time of test |  |  | 14:29
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-05-02
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | 4.1817e-05 |  |  | Dy average |  |  | 0.00067246 |  |  |  | Dx average |  | 0.00010048 |  |  | Dy average |  |  | 0.00065139
Magnetic length |  | 14.4206 |  |  | Tranfs. Function (mT/KA) |  |  | 595.8932 |  |  |  | Magnetic length |  | 14.4221 |  |  | Tranfs. Function (mT/KA) |  |  | 595.9372
Phase |  | -14.4306 |  |  | Integral field (T.m) |  |  | 85.8182 |  |  |  | Phase |  | -33.5732 |  |  | Integral field (T.m) |  |  | 85.7795
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4878 | 5.9618 | 5.9585 | 5.9592 | 5.9571 | 5.9581 | 5.9585 | 5.9581 | 5.9585 | 5.9586 | 5.9588 | 5.9587 | 5.9581 | 5.9572 | 5.959 | 5.9591 | 5.9596 | 5.9593 | 5.9625 | 3.6915 | 85.8182 | C1 (T) | 3.3754 | 5.9622 | 5.9588 | 5.9594 | 5.9594 | 5.9586 | 5.9587 | 5.959 | 5.9592 | 5.957 | 5.9586 | 5.9591 | 5.9586 | 5.9598 | 5.9589 | 5.9593 | 5.9602 | 5.9592 | 5.9624 | 3.7448 | 85.7795
Angle (mrad) | 2.2728 | -0.5982 | 0.06702 | -0.071456 | 0.26212 | 0.22468 | 0.46064 | 0.82044 | 0.58406 | 0.12646 | -0.75214 | -0.1263 | 0.062306 | -0.21086 | -0.15537 | -0.43273 | -0.43116 | -0.54454 | -0.38633 | -0.36796 | -14.4306 | Angle (mrad) | 3.1235 | -0.60918 | -0.128 | -0.07008 | 0.437 | 0.021095 | 0.36551 | 0.81392 | 0.6416 | -0.096258 | -0.6796 | 0.039041 | 0.20774 | 0.093781 | -0.11086 | -0.61212 | -0.45282 | -0.44103 | -0.48957 | -1.112 | -33.5732
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -5.1194 | 0.11421 | 0.61269 | 0.12435 | 0.37741 | 0.054134 | 0.33796 | 0.37973 | -0.058271 | -0.10706 | -0.39329 | 0.01497 | -0.05113 | -0.0011244 | -0.64905 | -0.27794 | 0.22433 | 0.26551 | 0.43349 | 0.42633 | -0.069312 | b2 | -0.90116 | 0.067395 | 0.082853 | -0.0048367 | -0.34808 | -0.37121 | 0.25409 | -0.17079 | -0.081414 | -0.23313 | -0.39024 | -1.024 | 0.05909 | -0.42866 | -0.42648 | -1.3012 | -1.0093 | -0.25047 | 0.064516 | -1.0926 | -0.3495
b3 | 33.6331 | -3.6968 | -3.228 | -4.0685 | -3.9518 | -4.3851 | -3.8964 | -3.1534 | -2.5089 | -2.8753 | -4.1463 | -4.306 | -4.1008 | -6.1209 | -5.5017 | -4.4687 | -2.6119 | -3.0545 | -2.8675 | -4.445 | -2.7082 | b3 | 38.1976 | -4.0601 | -3.8649 | -5.1239 | -4.5519 | -5.8028 | -3.6859 | -4.245 | -4.7071 | -4.0529 | -4.6735 | -5.531 | -4.5148 | -6.3474 | -5.4164 | -4.8565 | -3.8809 | -4.2026 | -2.7766 | -3.6033 | -3.2781
b4 | -2.2371 | -0.079263 | 0.0309 | 0.032632 | -0.169 | 0.14539 | 0.18048 | 0.09028 | -0.12439 | -0.060848 | 0.11102 | 0.23599 | -0.011829 | -0.077724 | -0.1971 | -0.081163 | 0.31307 | 0.143 | 0.08032 | 0.11631 | -0.035168 | b4 | -1.1741 | 0.074748 | 0.11376 | -0.034299 | -0.0025645 | -0.046823 | 0.099156 | 0.017266 | -0.032099 | -0.030278 | 0.16772 | -0.21633 | -0.30503 | -0.10197 | 0.016355 | -0.17566 | -0.27797 | 0.015204 | -0.058546 | 0.21211 | -0.068199
b5 | -5.7203 | -0.52373 | -0.58245 | -0.054921 | 0.037963 | 0.076226 | 0.041758 | -0.092717 | -0.2028 | 0.030749 | 0.4473 | 0.041487 | 0.14299 | -0.042013 | -0.027826 | -0.092896 | -0.22889 | -0.26614 | -0.40243 | -3.7098 | -0.38246 | b5 | -5.2608 | -0.53217 | -0.63133 | -0.27669 | -0.14787 | 0.13779 | 0.071405 | -0.041301 | -0.10335 | -0.1275 | 0.12487 | 0.11359 | -0.025132 | -0.0031114 | -0.16223 | -0.21961 | -0.3692 | -0.22707 | -0.30865 | -3.8414 | -0.42311
b6 | -0.22637 | -0.00035775 | 0.031472 | 0.1255 | 0.0060721 | 0.09853 | 0.036149 | 0.050502 | 0.087684 | 0.15055 | 0.12132 | 0.068878 | 0.021573 | 0.090116 | 0.10583 | 0.084101 | 0.065926 | 0.070737 | -0.025819 | 0.013214 | 0.055423 | b6 | 0.043832 | -0.032053 | -0.020652 | -0.054106 | -0.068673 | 0.041392 | -0.13171 | -0.025604 | -0.12572 | 0.10041 | 0.14688 | 0.0886 | -0.13798 | 0.0097184 | -0.01679 | -0.0011932 | 0.086842 | 0.02975 | -0.010911 | 0.14138 | -0.0004253
b7 | 2.6478 | 0.83337 | 0.97844 | 1.1804 | 1.1682 | 1.2193 | 1.1968 | 1.1629 | 1.1495 | 1.14 | 1.1733 | 1.1845 | 1.1217 | 1.1191 | 1.1142 | 1.1644 | 1.1394 | 1.1848 | 1.1538 | 0.88035 | 1.1705 | b7 | 2.6187 | 0.78188 | 0.9125 | 1.1245 | 1.1741 | 1.2288 | 1.1414 | 1.1046 | 1.0936 | 1.1581 | 1.1454 | 1.1473 | 1.1044 | 1.0804 | 1.0575 | 1.186 | 1.0837 | 1.1539 | 1.063 | 0.85466 | 1.1337
b8 | -0.028334 | -0.020495 | -0.026633 | 0.010924 | -0.0036821 | -0.020862 | -0.018215 | -0.0012839 | -0.035229 | 0.02552 | 0.012244 | -0.035166 | -0.067477 | -0.034857 | -0.019909 | 0.0058554 | 0.0028149 | 0.039539 | 0.022204 | -0.031688 | -0.01046 | b8 | -0.020792 | -0.019516 | -0.065623 | -0.090783 | -0.061171 | -0.0096881 | -0.022751 | -0.012194 | -0.036674 | 0.014161 | 0.024657 | -0.064397 | -0.063968 | -0.047347 | -0.030861 | 0.011372 | 0.028663 | 0.046077 | 0.015426 | 0.017393 | -0.020081
b9 | 0.36166 | 0.55564 | 0.55512 | 0.54558 | 0.5477 | 0.54242 | 0.5452 | 0.54254 | 0.5671 | 0.55075 | 0.54066 | 0.54784 | 0.55539 | 0.53134 | 0.51671 | 0.53916 | 0.54932 | 0.56256 | 0.54439 | 0.46555 | 0.53838 | b9 | 0.35041 | 0.53686 | 0.54001 | 0.512 | 0.52408 | 0.49548 | 0.52504 | 0.51205 | 0.52841 | 0.53315 | 0.50395 | 0.49042 | 0.51164 | 0.50852 | 0.49367 | 0.52233 | 0.53182 | 0.55598 | 0.54501 | 0.46486 | 0.51373
b10 | 0.017518 | -0.032365 | -0.041736 | 0.016849 | 0.0045458 | 0.034459 | -0.0094124 | 0.0020891 | -0.034929 | -0.0057877 | 0.043027 | -0.014856 | -0.040171 | -0.024186 | 0.021333 | 0.011329 | 0.032307 | 0.049106 | 0.0071401 | -0.046817 | 0 | b10 | 0.10024 | -0.014478 | -0.091699 | -0.026992 | -0.025699 | -0.0031172 | -0.031925 | -0.044864 | -0.010764 | 0.038412 | 0.063811 | -0.0013724 | -0.049539 | -0.020939 | 0.006062 | 0.037009 | 0.055061 | 0.050704 | 0.01894 | -0.0085782 | 0
b11 | 0.6222 | 0.74666 | 0.74071 | 0.75664 | 0.75122 | 0.75343 | 0.74742 | 0.74582 | 0.749 | 0.74969 | 0.76087 | 0.7541 | 0.76009 | 0.76563 | 0.76191 | 0.75774 | 0.74605 | 0.75227 | 0.73626 | 0.64321 | 0.74451 | b11 | 0.61462 | 0.75716 | 0.75533 | 0.76184 | 0.75891 | 0.76408 | 0.76139 | 0.76382 | 0.76563 | 0.75289 | 0.76654 | 0.7649 | 0.76333 | 0.76986 | 0.77093 | 0.76075 | 0.75974 | 0.7599 | 0.74552 | 0.64897 | 0.75325
b12 | 0.0091647 | -0.0045403 | -0.0048012 | -0.0014493 | -0.0011668 | 0.0017948 | -0.0025899 | -0.0019534 | -0.0067015 | -0.0035069 | 0.0028854 | -0.0038232 | -0.011551 | -0.0034401 | -0.00037475 | 0.0019855 | -4.6498e-05 | 0.0043532 | -0.00083672 | -0.0028673 | -0.0016753 | b12 | 0.0074458 | -0.00011472 | -0.0097523 | -0.0021913 | -0.0044241 | -0.0023663 | -0.010583 | -0.0050596 | 4.5754e-05 | 0.0050513 | 0.0047999 | -0.0023679 | -0.0045726 | -0.0035279 | -0.0012027 | 0.0013757 | 0.0066674 | 0.0057308 | 0.0036109 | 0.006448 | -0.00055268
b13 | 0.082833 | 0.071185 | 0.071205 | 0.068496 | 0.069542 | 0.067448 | 0.067957 | 0.070004 | 0.07039 | 0.068441 | 0.065975 | 0.069005 | 0.068276 | 0.067776 | 0.070264 | 0.070797 | 0.072705 | 0.072021 | 0.072577 | 0.051183 | 0.069475 | b13 | 0.086101 | 0.068123 | 0.072882 | 0.071057 | 0.075992 | 0.069393 | 0.069982 | 0.070188 | 0.068363 | 0.069837 | 0.072833 | 0.070645 | 0.071488 | 0.068792 | 0.071949 | 0.070351 | 0.070083 | 0.06985 | 0.067271 | 0.051051 | 0.070328
b14 | 0.0017609 | -0.004299 | -0.0032969 | -0.0011758 | -0.0011354 | -0.00036042 | -0.001608 | -0.0013635 | -0.0026863 | -0.0020099 | 0.00071125 | -0.0022116 | -0.0049755 | -0.0027313 | 0.00063481 | -0.00042183 | 0.00063577 | 0.0012688 | -0.0010065 | -0.0013124 | -0.0013441 | b14 | 0.0053541 | -0.0027247 | -0.0055401 | -0.0041858 | -0.0037769 | -0.0021599 | -0.0040286 | -0.0040401 | -0.0019089 | 0.00015797 | 0.0012887 | -0.0010771 | -0.0037242 | -0.0023212 | -0.0012353 | 3.6805e-05 | 0.0011717 | 0.00053608 | -0.0011945 | -0.00044937 | -0.0016659
b15 | -0.0091704 | 0.035552 | 0.031726 | 0.029502 | 0.027927 | 0.028921 | 0.030451 | 0.025273 | 0.027024 | 0.02961 | 0.028212 | 0.027964 | 0.02869 | 0.032583 | 0.030802 | 0.029069 | 0.029461 | 0.031083 | 0.026772 | 0.022057 | 0.028062 | b15 | 0.0023925 | 0.035442 | 0.026465 | 0.029965 | 0.025794 | 0.029108 | 0.029548 | 0.025305 | 0.029732 | 0.030475 | 0.032402 | 0.029319 | 0.029131 | 0.030814 | 0.029864 | 0.02744 | 0.029188 | 0.028249 | 0.028417 | 0.014604 | 0.027986
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 22.7277 | -5.982 | 0.6702 | -0.71456 | 2.6212 | 2.2468 | 4.6064 | 8.2044 | 5.8406 | 1.2646 | -7.5214 | -1.263 | 0.62306 | -2.1086 | -1.5537 | -4.3273 | -4.3116 | -5.4454 | -3.8634 | -3.6796 | 0 | a1 | 31.2355 | -6.0918 | -1.28 | -0.7008 | 4.37 | 0.21095 | 3.6551 | 8.1392 | 6.416 | -0.96258 | -6.796 | 0.39041 | 2.0774 | 0.93781 | -1.1086 | -6.1212 | -4.5282 | -4.4103 | -4.8957 | -11.1196 | 0
a2 | 16.1728 | 2.5922 | 1.2675 | 1.6794 | 1.9788 | 0.63743 | 1.2309 | 1.9082 | 1.9864 | 1.0507 | 1.7347 | 1.5324 | 2.0306 | 2.0342 | 1.4536 | 1.819 | 1.2465 | 1.5514 | 1.1707 | -0.013367 | 1.9975 | a2 | 12.1314 | 0.43148 | 0.12782 | 0.069639 | 0.27387 | 0.85914 | 0.62314 | 0.4856 | 0.62132 | 0.58543 | 0.46859 | 0.68478 | 0.88435 | -0.93289 | -0.38902 | 0.80547 | -0.38483 | 0.51269 | 1.059 | 1.1131 | 0.74791
a3 | -0.49653 | 1.1724 | 0.94268 | 0.50288 | 0.52873 | 0.94858 | 0.92833 | 0.79621 | 0.38156 | 0.31093 | 0.62085 | 0.38911 | 0.72747 | 0.62617 | 0.83421 | 0.78119 | 0.96248 | 0.19628 | 0.24092 | -0.66484 | 0.58259 | a3 | 1.6266 | 1.2955 | 0.64228 | 0.43605 | 0.73013 | 0.65828 | 0.60638 | 0.22685 | 0.77281 | 0.36554 | 1.0997 | 1.0813 | 0.86381 | 0.34075 | 0.16162 | 0.81423 | 0.95307 | 0.444 | 0.17218 | -0.3573 | 0.644
a4 | 2.4915 | 0.48381 | 0.067079 | 0.13678 | 0.20008 | 0.44312 | 0.18195 | -0.19879 | 0.12351 | 0.40826 | 0.5759 | 0.59468 | 0.7878 | 0.46681 | 0.40061 | 0.30311 | 0.49754 | 0.14143 | 0.31608 | 0.36666 | 0.39652 | a4 | 2.3752 | 0.0049448 | 0.071971 | -0.15584 | 0.31244 | -0.31858 | -0.006958 | 0.22436 | 0.043328 | 0.13406 | 0.47695 | 0.50601 | 0.48314 | 0.51745 | -0.42485 | 0.23491 | 0.43758 | -0.0016741 | -0.08718 | 0.48553 | 0.21372
a5 | 2.4816 | 0.089366 | 0.035501 | 0.079547 | 0.25428 | 0.12615 | 0.17034 | 0.13976 | 0.15442 | 0.26878 | 0.22971 | -0.15515 | -0.078026 | 0.25318 | 0.28323 | 0.23305 | 0.0071931 | 0.028715 | 0.15657 | -0.1529 | 0.18924 | a5 | 3.239 | 0.34729 | 0.038564 | -0.05107 | 0.026184 | 0.030853 | 0.014085 | -0.050404 | 0.22752 | 0.3143 | 0.31872 | 0.19977 | 0.11911 | 0.37749 | 0.11568 | 0.2308 | 0.2512 | -0.099301 | 0.22875 | -0.32931 | 0.22231
a6 | 0.46323 | 0.081606 | 0.09673 | 0.089267 | 0.14025 | 0.0050833 | 0.035141 | 0.024541 | 0.13394 | 0.17684 | 0.1897 | 0.17389 | -0.11276 | 0.01246 | -0.014962 | 0.096312 | 0.10032 | 0.0062216 | 0.10872 | -0.1229 | 0.080101 | a6 | -0.020907 | 0.083967 | 0.12375 | 0.021013 | 0.016382 | -0.076144 | 0.035525 | 0.092138 | 0.074707 | 0.11466 | 0.084696 | 0.15431 | 0.026097 | -0.075808 | -0.12318 | 0.04337 | 0.11642 | 0.002536 | -0.050072 | 0.013376 | 0.034431
a7 | 1.371 | -0.10258 | -0.022962 | 0.0034954 | 0.049345 | -0.01982 | -0.05119 | 0.0055801 | 0.035748 | 0.019185 | -0.027285 | -0.018792 | 0.011639 | 0.053871 | -0.0063114 | -0.092969 | 0.017382 | -0.03528 | -0.039156 | 0.0058648 | 0.030506 | a7 | 1.7173 | -0.010436 | -0.026529 | -0.0041044 | 0.02609 | 0.014567 | -0.023763 | 0.14413 | 0.16691 | 0.087767 | 0.032653 | 0.15239 | 0.083825 | 0.12798 | 0.056922 | -0.031083 | 0.028521 | -0.0037878 | 0.0085609 | 0.062794 | 0.095999
a8 | 0.04686 | 0.0013255 | 0.021759 | 0.038549 | 0.014505 | -0.010946 | -0.0049797 | 0.05161 | 0.068362 | 0.0071696 | 0.03247 | 0.018789 | 0.020767 | -0.02404 | 0.011235 | 0.042072 | 0.052021 | -0.0072794 | 0.048783 | -0.025956 | 0.020492 | a8 | -0.048243 | 0.013124 | 0.045583 | 0.022934 | 0.072507 | 0.058282 | 0.06089 | 0.040061 | 0.010516 | 0.033151 | 0.033692 | 0.053558 | 0.0018483 | -0.032749 | 0.0059921 | 0.020676 | 0.0070247 | -0.013472 | -0.018891 | -0.056334 | 0.018336
a9 | -0.27893 | 0.019541 | 0.019796 | -0.048111 | -0.053017 | -0.012373 | 0.010473 | 0.023477 | -0.00025618 | -0.022183 | 0.00072435 | -0.01036 | 0.035194 | -0.018878 | -0.030176 | 0.023084 | -0.002096 | -0.016289 | -0.037668 | -0.053051 | -0.016415 | a9 | -0.1872 | 0.063132 | 0.025139 | -0.0052886 | -0.047308 | -0.012855 | 0.025241 | 0.036247 | 0.031632 | -0.021412 | 0.016984 | 0.045083 | 0.049153 | 0.012567 | -0.032683 | 0.04013 | 0.021482 | 0.020767 | 0.0016603 | -0.050376 | 0.0068777
a10 | 0.054037 | -0.033739 | -0.016297 | -0.0097151 | 0.012887 | -0.0017054 | -0.010978 | 0.0352 | 0.034663 | 0.0043681 | 0.032903 | 0.018436 | -0.0010946 | -0.046919 | 0.0011871 | 0.0027625 | 0.014107 | -0.020726 | 0.022778 | -0.11259 | 0 | a10 | -0.062502 | -0.028087 | 0.034416 | -0.026638 | 0.041148 | -0.007401 | -0.0097722 | 0.04541 | 0.014274 | -0.012599 | -0.011995 | 0.0095762 | -0.0066571 | 0.0024713 | -0.017774 | 0.0077452 | 0.0078941 | 0.008716 | 0.0031397 | -0.029378 | 0
a11 | 0.12944 | -0.069817 | -0.04441 | -0.044693 | -0.034298 | -0.046398 | -0.044794 | -0.031086 | -0.029241 | -0.035548 | -0.047427 | -0.0532 | -0.049526 | -0.048836 | -0.047516 | -0.054447 | -0.045516 | -0.044083 | -0.046659 | -0.038388 | -0.039862 | a11 | 0.17939 | -0.0089722 | -0.0056176 | 0.00015419 | 0.00029774 | 0.0060002 | 0.011884 | 0.020063 | 0.028022 | 0.010553 | 0.011221 | 0.012939 | 0.012926 | 0.014569 | 0.004688 | -0.004439 | 0.0050847 | -0.0011664 | -0.0039 | 0.0091273 | 0.011546
a12 | 0.0059196 | -0.0028533 | -0.00038653 | -0.0014393 | 0.0021197 | -0.0012673 | -0.0045736 | 0.0030187 | 0.0042989 | 0.002507 | 0.0043399 | 0.0074277 | -0.0067803 | -0.0059308 | 0.00089345 | 0.00059987 | 0.00096646 | -0.0019056 | 0.0022181 | -0.016201 | -0.00017274 | a12 | -0.011729 | -0.0020489 | 0.0034935 | -0.0040785 | 0.00036974 | -0.0039701 | -0.0017805 | 0.0045713 | 0.0062514 | 0.0015487 | -0.00078903 | 0.0046822 | -0.0014664 | -0.0012355 | -0.0043328 | 0.00045951 | 0.0011558 | -0.0013575 | -0.0016932 | -0.0037388 | -0.00048006
a13 | -0.027493 | 0.0021693 | 0.0012647 | -0.0024792 | 2.1079e-05 | -0.0014029 | -0.00097918 | 0.00023453 | 7.5677e-06 | 0.00022278 | -0.0027266 | -0.0012926 | 0.0010642 | -0.00075939 | -4.8012e-05 | -0.0026736 | -0.0028543 | -0.0040936 | -0.0050351 | -0.0054577 | -0.0020222 | a13 | -0.018768 | 0.0077813 | 0.005505 | -0.000797 | 0.00065015 | 0.0027441 | 0.0029114 | 0.0035453 | 0.0051987 | 0.0026023 | 0.0029344 | 0.0080456 | 0.0073486 | 0.0051922 | 0.00069985 | 0.0038908 | 0.0045129 | 0.00049678 | 0.0011213 | -0.0022885 | 0.0027255
a14 | 0.008448 | -0.012692 | -0.012541 | -0.010671 | -0.011032 | -0.012372 | -0.013124 | -0.0087318 | -0.008741 | -0.012529 | -0.0094701 | -0.008794 | -0.0086725 | -0.012967 | -0.011745 | -0.010308 | -0.01063 | -0.012068 | -0.0094036 | -0.0047471 | -0.010127 | a14 | 0.011174 | -0.014838 | -0.0096379 | -0.011942 | -0.0059802 | -0.010152 | -0.010234 | -0.0083534 | -0.011838 | -0.010816 | -0.011945 | -0.008912 | -0.010354 | -0.011021 | -0.013248 | -0.0086282 | -0.010721 | -0.01134 | -0.012784 | -0.0049387 | -0.0098736
a15 | 0.0081968 | -0.00926 | -0.0076288 | -0.0026879 | -0.0037768 | -0.00093878 | -0.003471 | -0.0023565 | -0.0053045 | -0.0049761 | 0.00081288 | -0.0036092 | -0.008144 | -0.0063427 | -0.0017046 | -0.001359 | -0.0011163 | -0.00014761 | -0.0037819 | -0.0065727 | -0.0033881 | a15 | 0.008437 | -0.0053093 | -0.0093013 | -0.0072714 | -0.005329 | -0.0032498 | -0.006742 | -0.0047795 | -0.0047482 | -0.001441 | 0.002256 | -0.001522 | -0.0063098 | -0.0041213 | -0.0039675 | 0.00077197 | 0.0016005 | 0.001864 | -0.0020148 | -0.0027127 | -0.0029456
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -7.5251e-05 | 6.5918e-05 | 9.3221e-05 | -3.9009e-05 | 0 | -7.7695e-05 | 1.9839e-05 | 0 | 8.2212e-05 | 1.3563e-05 | -9.1179e-05 | 3.6233e-05 | 8.9336e-05 | 4.684e-05 | -4.7248e-05 | -2.4843e-05 | -7.1387e-05 | -0.00011334 | -1.3101e-05 | 0.00010518 | 4.1817e-05 | Dx (m) | -0.00020864 | 3.1752e-05 | 0.00020701 | 6.0234e-05 | 5.753e-05 | 0 | 7.162e-05 | 9.7103e-05 | 2.2706e-05 | -8.6332e-05 | -0.00014111 | 0 | 0.00011056 | 4.6121e-05 | -1.3128e-05 | -8.2597e-05 | -0.00012335 | -0.00011342 | -4.3153e-05 | 2.3561e-05 | 0.00010048
Dy (m) | -0.00013204 | 8.2995e-05 | 4.3002e-05 | 1.9523e-05 | -2.9571e-05 | 0 | 2.6158e-05 | -8.0289e-05 | -7.5496e-05 | 0 | -7.9209e-05 | -3.9007e-05 | 0 | 0.00010716 | 0 | 0 | -3.6497e-05 | 4.0212e-05 | -5.3422e-05 | 0.00030336 | 0.00067246 | Dy (m) | 0.00023396 | 6.3436e-05 | -7.6041e-05 | 5.9432e-05 | -9.2215e-05 | 1.6411e-05 | 2.0716e-05 | -0.00010365 | -3.2524e-05 | 2.9661e-05 | 2.868e-05 | -2.1328e-05 | 1.2962e-05 | 0 | 3.9272e-05 | -1.779e-05 | -1.6851e-05 | -1.9671e-05 | 0 | 7.6605e-05 | 0.00065139
Magnetic length (m) |  | 14.4206 | [19 empty] | Magnetic length (m) |  | 14.4221
Transfer funct. (mT/KA) |  | 595.8932 | [19 empty] | Transfer funct. (mT/KA) |  | 595.9372
Coil positioning (mm) |  | 0.017787 | [19 empty] | Coil positioning (mm) |  | 0.021253
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000088 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000088 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 348.7814 | 596.1773 | 595.8526 | 595.9193 | 595.7149 | 595.8074 | 595.8476 | 595.8091 | 595.8478 | 595.8629 | 595.8835 | 595.8725 | 595.8128 | 595.7198 | 595.8978 | 595.9104 | 595.9554 | 595.9326 | 596.2542 | 369.1482 | 85.8182 | C1/i (mT/KA) | 337.545 | 596.2212 | 595.8805 | 595.9404 | 595.9411 | 595.8632 | 595.87 | 595.9046 | 595.9194 | 595.698 | 595.8639 | 595.9133 | 595.8638 | 595.9833 | 595.8937 | 595.9258 | 596.0239 | 595.9204 | 596.2435 | 374.4801 | 85.7795
Angle (mrad) | 2.2728 | -0.5982 | 0.06702 | -0.071456 | 0.26212 | 0.22468 | 0.46064 | 0.82044 | 0.58406 | 0.12646 | -0.75214 | -0.1263 | 0.062306 | -0.21086 | -0.15537 | -0.43273 | -0.43116 | -0.54454 | -0.38633 | -0.36796 | -14.4306 | Angle (mrad) | 3.1235 | -0.60918 | -0.128 | -0.07008 | 0.437 | 0.021095 | 0.36551 | 0.81392 | 0.6416 | -0.096258 | -0.6796 | 0.039041 | 0.20774 | 0.093781 | -0.11086 | -0.61212 | -0.45282 | -0.44103 | -0.48957 | -1.112 | -33.5732
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -5.1194 | 0.11421 | 0.61269 | 0.12435 | 0.37741 | 0.054134 | 0.33796 | 0.37973 | -0.058271 | -0.10706 | -0.39329 | 0.01497 | -0.05113 | -0.0011244 | -0.64905 | -0.27794 | 0.22433 | 0.26551 | 0.43349 | 0.42633 | -0.069312 | b2 | -0.90116 | 0.067395 | 0.082853 | -0.0048367 | -0.34808 | -0.37121 | 0.25409 | -0.17079 | -0.081414 | -0.23313 | -0.39024 | -1.024 | 0.05909 | -0.42866 | -0.42648 | -1.3012 | -1.0093 | -0.25047 | 0.064516 | -1.0926 | -0.3495
b3 | 33.6331 | -3.6968 | -3.228 | -4.0685 | -3.9518 | -4.3851 | -3.8964 | -3.1534 | -2.5089 | -2.8753 | -4.1463 | -4.306 | -4.1008 | -6.1209 | -5.5017 | -4.4687 | -2.6119 | -3.0545 | -2.8675 | -4.445 | -2.7082 | b3 | 38.1976 | -4.0601 | -3.8649 | -5.1239 | -4.5519 | -5.8028 | -3.6859 | -4.245 | -4.7071 | -4.0529 | -4.6735 | -5.531 | -4.5148 | -6.3474 | -5.4164 | -4.8565 | -3.8809 | -4.2026 | -2.7766 | -3.6033 | -3.2781
b4 | -2.2371 | -0.079263 | 0.0309 | 0.032632 | -0.169 | 0.14539 | 0.18048 | 0.09028 | -0.12439 | -0.060848 | 0.11102 | 0.23599 | -0.011829 | -0.077724 | -0.1971 | -0.081163 | 0.31307 | 0.143 | 0.08032 | 0.11631 | -0.035168 | b4 | -1.1741 | 0.074748 | 0.11376 | -0.034299 | -0.0025645 | -0.046823 | 0.099156 | 0.017266 | -0.032099 | -0.030278 | 0.16772 | -0.21633 | -0.30503 | -0.10197 | 0.016355 | -0.17566 | -0.27797 | 0.015204 | -0.058546 | 0.21211 | -0.068199
b5 | -5.7203 | -0.52373 | -0.58245 | -0.054921 | 0.037963 | 0.076226 | 0.041758 | -0.092717 | -0.2028 | 0.030749 | 0.4473 | 0.041487 | 0.14299 | -0.042013 | -0.027826 | -0.092896 | -0.22889 | -0.26614 | -0.40243 | -3.7098 | -0.38246 | b5 | -5.2608 | -0.53217 | -0.63133 | -0.27669 | -0.14787 | 0.13779 | 0.071405 | -0.041301 | -0.10335 | -0.1275 | 0.12487 | 0.11359 | -0.025132 | -0.0031114 | -0.16223 | -0.21961 | -0.3692 | -0.22707 | -0.30865 | -3.8414 | -0.42311
b6 | -0.22637 | -0.00035775 | 0.031472 | 0.1255 | 0.0060721 | 0.09853 | 0.036149 | 0.050502 | 0.087684 | 0.15055 | 0.12132 | 0.068878 | 0.021573 | 0.090116 | 0.10583 | 0.084101 | 0.065926 | 0.070737 | -0.025819 | 0.013214 | 0.055423 | b6 | 0.043832 | -0.032053 | -0.020652 | -0.054106 | -0.068673 | 0.041392 | -0.13171 | -0.025604 | -0.12572 | 0.10041 | 0.14688 | 0.0886 | -0.13798 | 0.0097184 | -0.01679 | -0.0011932 | 0.086842 | 0.02975 | -0.010911 | 0.14138 | -0.0004253
b7 | 2.6478 | 0.83337 | 0.97844 | 1.1804 | 1.1682 | 1.2193 | 1.1968 | 1.1629 | 1.1495 | 1.14 | 1.1733 | 1.1845 | 1.1217 | 1.1191 | 1.1142 | 1.1644 | 1.1394 | 1.1848 | 1.1538 | 0.88035 | 1.1705 | b7 | 2.6187 | 0.78188 | 0.9125 | 1.1245 | 1.1741 | 1.2288 | 1.1414 | 1.1046 | 1.0936 | 1.1581 | 1.1454 | 1.1473 | 1.1044 | 1.0804 | 1.0575 | 1.186 | 1.0837 | 1.1539 | 1.063 | 0.85466 | 1.1337
b8 | -0.028334 | -0.020495 | -0.026633 | 0.010924 | -0.0036821 | -0.020862 | -0.018215 | -0.0012839 | -0.035229 | 0.02552 | 0.012244 | -0.035166 | -0.067477 | -0.034857 | -0.019909 | 0.0058554 | 0.0028149 | 0.039539 | 0.022204 | -0.031688 | -0.01046 | b8 | -0.020792 | -0.019516 | -0.065623 | -0.090783 | -0.061171 | -0.0096881 | -0.022751 | -0.012194 | -0.036674 | 0.014161 | 0.024657 | -0.064397 | -0.063968 | -0.047347 | -0.030861 | 0.011372 | 0.028663 | 0.046077 | 0.015426 | 0.017393 | -0.020081
b9 | 0.36166 | 0.55564 | 0.55512 | 0.54558 | 0.5477 | 0.54242 | 0.5452 | 0.54254 | 0.5671 | 0.55075 | 0.54066 | 0.54784 | 0.55539 | 0.53134 | 0.51671 | 0.53916 | 0.54932 | 0.56256 | 0.54439 | 0.46555 | 0.53838 | b9 | 0.35041 | 0.53686 | 0.54001 | 0.512 | 0.52408 | 0.49548 | 0.52504 | 0.51205 | 0.52841 | 0.53315 | 0.50395 | 0.49042 | 0.51164 | 0.50852 | 0.49367 | 0.52233 | 0.53182 | 0.55598 | 0.54501 | 0.46486 | 0.51373
b10 | 0.017518 | -0.032365 | -0.041736 | 0.016849 | 0.0045458 | 0.034459 | -0.0094124 | 0.0020891 | -0.034929 | -0.0057877 | 0.043027 | -0.014856 | -0.040171 | -0.024186 | 0.021333 | 0.011329 | 0.032307 | 0.049106 | 0.0071401 | -0.046817 | 0 | b10 | 0.10024 | -0.014478 | -0.091699 | -0.026992 | -0.025699 | -0.0031172 | -0.031925 | -0.044864 | -0.010764 | 0.038412 | 0.063811 | -0.0013724 | -0.049539 | -0.020939 | 0.006062 | 0.037009 | 0.055061 | 0.050704 | 0.01894 | -0.0085782 | 0
b11 | 0.6222 | 0.74666 | 0.74071 | 0.75664 | 0.75122 | 0.75343 | 0.74742 | 0.74582 | 0.749 | 0.74969 | 0.76087 | 0.7541 | 0.76009 | 0.76563 | 0.76191 | 0.75774 | 0.74605 | 0.75227 | 0.73626 | 0.64321 | 0.74451 | b11 | 0.61462 | 0.75716 | 0.75533 | 0.76184 | 0.75891 | 0.76408 | 0.76139 | 0.76382 | 0.76563 | 0.75289 | 0.76654 | 0.7649 | 0.76333 | 0.76986 | 0.77093 | 0.76075 | 0.75974 | 0.7599 | 0.74552 | 0.64897 | 0.75325
b12 | 0.0091647 | -0.0045403 | -0.0048012 | -0.0014493 | -0.0011668 | 0.0017948 | -0.0025899 | -0.0019534 | -0.0067015 | -0.0035069 | 0.0028854 | -0.0038232 | -0.011551 | -0.0034401 | -0.00037475 | 0.0019855 | -4.6498e-05 | 0.0043532 | -0.00083672 | -0.0028673 | -0.0016753 | b12 | 0.0074458 | -0.00011472 | -0.0097523 | -0.0021913 | -0.0044241 | -0.0023663 | -0.010583 | -0.0050596 | 4.5754e-05 | 0.0050513 | 0.0047999 | -0.0023679 | -0.0045726 | -0.0035279 | -0.0012027 | 0.0013757 | 0.0066674 | 0.0057308 | 0.0036109 | 0.006448 | -0.00055268
b13 | 0.082833 | 0.071185 | 0.071205 | 0.068496 | 0.069542 | 0.067448 | 0.067957 | 0.070004 | 0.07039 | 0.068441 | 0.065975 | 0.069005 | 0.068276 | 0.067776 | 0.070264 | 0.070797 | 0.072705 | 0.072021 | 0.072577 | 0.051183 | 0.069475 | b13 | 0.086101 | 0.068123 | 0.072882 | 0.071057 | 0.075992 | 0.069393 | 0.069982 | 0.070188 | 0.068363 | 0.069837 | 0.072833 | 0.070645 | 0.071488 | 0.068792 | 0.071949 | 0.070351 | 0.070083 | 0.06985 | 0.067271 | 0.051051 | 0.070328
b14 | 0.0017609 | -0.004299 | -0.0032969 | -0.0011758 | -0.0011354 | -0.00036042 | -0.001608 | -0.0013635 | -0.0026863 | -0.0020099 | 0.00071125 | -0.0022116 | -0.0049755 | -0.0027313 | 0.00063481 | -0.00042183 | 0.00063577 | 0.0012688 | -0.0010065 | -0.0013124 | -0.0013441 | b14 | 0.0053541 | -0.0027247 | -0.0055401 | -0.0041858 | -0.0037769 | -0.0021599 | -0.0040286 | -0.0040401 | -0.0019089 | 0.00015797 | 0.0012887 | -0.0010771 | -0.0037242 | -0.0023212 | -0.0012353 | 3.6805e-05 | 0.0011717 | 0.00053608 | -0.0011945 | -0.00044937 | -0.0016659
b15 | -0.0091704 | 0.035552 | 0.031726 | 0.029502 | 0.027927 | 0.028921 | 0.030451 | 0.025273 | 0.027024 | 0.02961 | 0.028212 | 0.027964 | 0.02869 | 0.032583 | 0.030802 | 0.029069 | 0.029461 | 0.031083 | 0.026772 | 0.022057 | 0.028062 | b15 | 0.0023925 | 0.035442 | 0.026465 | 0.029965 | 0.025794 | 0.029108 | 0.029548 | 0.025305 | 0.029732 | 0.030475 | 0.032402 | 0.029319 | 0.029131 | 0.030814 | 0.029864 | 0.02744 | 0.029188 | 0.028249 | 0.028417 | 0.014604 | 0.027986
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 22.7277 | -5.982 | 0.6702 | -0.71456 | 2.6212 | 2.2468 | 4.6064 | 8.2044 | 5.8406 | 1.2646 | -7.5214 | -1.263 | 0.62306 | -2.1086 | -1.5537 | -4.3273 | -4.3116 | -5.4454 | -3.8634 | -3.6796 | 0 | a1 | 31.2355 | -6.0918 | -1.28 | -0.7008 | 4.37 | 0.21095 | 3.6551 | 8.1392 | 6.416 | -0.96258 | -6.796 | 0.39041 | 2.0774 | 0.93781 | -1.1086 | -6.1212 | -4.5282 | -4.4103 | -4.8957 | -11.1196 | 0
a2 | 16.1728 | 2.5922 | 1.2675 | 1.6794 | 1.9788 | 0.63743 | 1.2309 | 1.9082 | 1.9864 | 1.0507 | 1.7347 | 1.5324 | 2.0306 | 2.0342 | 1.4536 | 1.819 | 1.2465 | 1.5514 | 1.1707 | -0.013367 | 1.9975 | a2 | 12.1314 | 0.43148 | 0.12782 | 0.069639 | 0.27387 | 0.85914 | 0.62314 | 0.4856 | 0.62132 | 0.58543 | 0.46859 | 0.68478 | 0.88435 | -0.93289 | -0.38902 | 0.80547 | -0.38483 | 0.51269 | 1.059 | 1.1131 | 0.74791
a3 | -0.49653 | 1.1724 | 0.94268 | 0.50288 | 0.52873 | 0.94858 | 0.92833 | 0.79621 | 0.38156 | 0.31093 | 0.62085 | 0.38911 | 0.72747 | 0.62617 | 0.83421 | 0.78119 | 0.96248 | 0.19628 | 0.24092 | -0.66484 | 0.58259 | a3 | 1.6266 | 1.2955 | 0.64228 | 0.43605 | 0.73013 | 0.65828 | 0.60638 | 0.22685 | 0.77281 | 0.36554 | 1.0997 | 1.0813 | 0.86381 | 0.34075 | 0.16162 | 0.81423 | 0.95307 | 0.444 | 0.17218 | -0.3573 | 0.644
a4 | 2.4915 | 0.48381 | 0.067079 | 0.13678 | 0.20008 | 0.44312 | 0.18195 | -0.19879 | 0.12351 | 0.40826 | 0.5759 | 0.59468 | 0.7878 | 0.46681 | 0.40061 | 0.30311 | 0.49754 | 0.14143 | 0.31608 | 0.36666 | 0.39652 | a4 | 2.3752 | 0.0049448 | 0.071971 | -0.15584 | 0.31244 | -0.31858 | -0.006958 | 0.22436 | 0.043328 | 0.13406 | 0.47695 | 0.50601 | 0.48314 | 0.51745 | -0.42485 | 0.23491 | 0.43758 | -0.0016741 | -0.08718 | 0.48553 | 0.21372
a5 | 2.4816 | 0.089366 | 0.035501 | 0.079547 | 0.25428 | 0.12615 | 0.17034 | 0.13976 | 0.15442 | 0.26878 | 0.22971 | -0.15515 | -0.078026 | 0.25318 | 0.28323 | 0.23305 | 0.0071931 | 0.028715 | 0.15657 | -0.1529 | 0.18924 | a5 | 3.239 | 0.34729 | 0.038564 | -0.05107 | 0.026184 | 0.030853 | 0.014085 | -0.050404 | 0.22752 | 0.3143 | 0.31872 | 0.19977 | 0.11911 | 0.37749 | 0.11568 | 0.2308 | 0.2512 | -0.099301 | 0.22875 | -0.32931 | 0.22231
a6 | 0.46323 | 0.081606 | 0.09673 | 0.089267 | 0.14025 | 0.0050833 | 0.035141 | 0.024541 | 0.13394 | 0.17684 | 0.1897 | 0.17389 | -0.11276 | 0.01246 | -0.014962 | 0.096312 | 0.10032 | 0.0062216 | 0.10872 | -0.1229 | 0.080101 | a6 | -0.020907 | 0.083967 | 0.12375 | 0.021013 | 0.016382 | -0.076144 | 0.035525 | 0.092138 | 0.074707 | 0.11466 | 0.084696 | 0.15431 | 0.026097 | -0.075808 | -0.12318 | 0.04337 | 0.11642 | 0.002536 | -0.050072 | 0.013376 | 0.034431
a7 | 1.371 | -0.10258 | -0.022962 | 0.0034954 | 0.049345 | -0.01982 | -0.05119 | 0.0055801 | 0.035748 | 0.019185 | -0.027285 | -0.018792 | 0.011639 | 0.053871 | -0.0063114 | -0.092969 | 0.017382 | -0.03528 | -0.039156 | 0.0058648 | 0.030506 | a7 | 1.7173 | -0.010436 | -0.026529 | -0.0041044 | 0.02609 | 0.014567 | -0.023763 | 0.14413 | 0.16691 | 0.087767 | 0.032653 | 0.15239 | 0.083825 | 0.12798 | 0.056922 | -0.031083 | 0.028521 | -0.0037878 | 0.0085609 | 0.062794 | 0.095999
a8 | 0.04686 | 0.0013255 | 0.021759 | 0.038549 | 0.014505 | -0.010946 | -0.0049797 | 0.05161 | 0.068362 | 0.0071696 | 0.03247 | 0.018789 | 0.020767 | -0.02404 | 0.011235 | 0.042072 | 0.052021 | -0.0072794 | 0.048783 | -0.025956 | 0.020492 | a8 | -0.048243 | 0.013124 | 0.045583 | 0.022934 | 0.072507 | 0.058282 | 0.06089 | 0.040061 | 0.010516 | 0.033151 | 0.033692 | 0.053558 | 0.0018483 | -0.032749 | 0.0059921 | 0.020676 | 0.0070247 | -0.013472 | -0.018891 | -0.056334 | 0.018336
a9 | -0.27893 | 0.019541 | 0.019796 | -0.048111 | -0.053017 | -0.012373 | 0.010473 | 0.023477 | -0.00025618 | -0.022183 | 0.00072435 | -0.01036 | 0.035194 | -0.018878 | -0.030176 | 0.023084 | -0.002096 | -0.016289 | -0.037668 | -0.053051 | -0.016415 | a9 | -0.1872 | 0.063132 | 0.025139 | -0.0052886 | -0.047308 | -0.012855 | 0.025241 | 0.036247 | 0.031632 | -0.021412 | 0.016984 | 0.045083 | 0.049153 | 0.012567 | -0.032683 | 0.04013 | 0.021482 | 0.020767 | 0.0016603 | -0.050376 | 0.0068777
a10 | 0.054037 | -0.033739 | -0.016297 | -0.0097151 | 0.012887 | -0.0017054 | -0.010978 | 0.0352 | 0.034663 | 0.0043681 | 0.032903 | 0.018436 | -0.0010946 | -0.046919 | 0.0011871 | 0.0027625 | 0.014107 | -0.020726 | 0.022778 | -0.11259 | 0 | a10 | -0.062502 | -0.028087 | 0.034416 | -0.026638 | 0.041148 | -0.007401 | -0.0097722 | 0.04541 | 0.014274 | -0.012599 | -0.011995 | 0.0095762 | -0.0066571 | 0.0024713 | -0.017774 | 0.0077452 | 0.0078941 | 0.008716 | 0.0031397 | -0.029378 | 0
a11 | 0.12944 | -0.069817 | -0.04441 | -0.044693 | -0.034298 | -0.046398 | -0.044794 | -0.031086 | -0.029241 | -0.035548 | -0.047427 | -0.0532 | -0.049526 | -0.048836 | -0.047516 | -0.054447 | -0.045516 | -0.044083 | -0.046659 | -0.038388 | -0.039862 | a11 | 0.17939 | -0.0089722 | -0.0056176 | 0.00015419 | 0.00029774 | 0.0060002 | 0.011884 | 0.020063 | 0.028022 | 0.010553 | 0.011221 | 0.012939 | 0.012926 | 0.014569 | 0.004688 | -0.004439 | 0.0050847 | -0.0011664 | -0.0039 | 0.0091273 | 0.011546
a12 | 0.0059196 | -0.0028533 | -0.00038653 | -0.0014393 | 0.0021197 | -0.0012673 | -0.0045736 | 0.0030187 | 0.0042989 | 0.002507 | 0.0043399 | 0.0074277 | -0.0067803 | -0.0059308 | 0.00089345 | 0.00059987 | 0.00096646 | -0.0019056 | 0.0022181 | -0.016201 | -0.00017274 | a12 | -0.011729 | -0.0020489 | 0.0034935 | -0.0040785 | 0.00036974 | -0.0039701 | -0.0017805 | 0.0045713 | 0.0062514 | 0.0015487 | -0.00078903 | 0.0046822 | -0.0014664 | -0.0012355 | -0.0043328 | 0.00045951 | 0.0011558 | -0.0013575 | -0.0016932 | -0.0037388 | -0.00048006
a13 | -0.027493 | 0.0021693 | 0.0012647 | -0.0024792 | 2.1079e-05 | -0.0014029 | -0.00097918 | 0.00023453 | 7.5677e-06 | 0.00022278 | -0.0027266 | -0.0012926 | 0.0010642 | -0.00075939 | -4.8012e-05 | -0.0026736 | -0.0028543 | -0.0040936 | -0.0050351 | -0.0054577 | -0.0020222 | a13 | -0.018768 | 0.0077813 | 0.005505 | -0.000797 | 0.00065015 | 0.0027441 | 0.0029114 | 0.0035453 | 0.0051987 | 0.0026023 | 0.0029344 | 0.0080456 | 0.0073486 | 0.0051922 | 0.00069985 | 0.0038908 | 0.0045129 | 0.00049678 | 0.0011213 | -0.0022885 | 0.0027255
a14 | 0.008448 | -0.012692 | -0.012541 | -0.010671 | -0.011032 | -0.012372 | -0.013124 | -0.0087318 | -0.008741 | -0.012529 | -0.0094701 | -0.008794 | -0.0086725 | -0.012967 | -0.011745 | -0.010308 | -0.01063 | -0.012068 | -0.0094036 | -0.0047471 | -0.010127 | a14 | 0.011174 | -0.014838 | -0.0096379 | -0.011942 | -0.0059802 | -0.010152 | -0.010234 | -0.0083534 | -0.011838 | -0.010816 | -0.011945 | -0.008912 | -0.010354 | -0.011021 | -0.013248 | -0.0086282 | -0.010721 | -0.01134 | -0.012784 | -0.0049387 | -0.0098736
a15 | 0.0081968 | -0.00926 | -0.0076288 | -0.0026879 | -0.0037768 | -0.00093878 | -0.003471 | -0.0023565 | -0.0053045 | -0.0049761 | 0.00081288 | -0.0036092 | -0.008144 | -0.0063427 | -0.0017046 | -0.001359 | -0.0011163 | -0.00014761 | -0.0037819 | -0.0065727 | -0.0033881 | a15 | 0.008437 | -0.0053093 | -0.0093013 | -0.0072714 | -0.005329 | -0.0032498 | -0.006742 | -0.0047795 | -0.0047482 | -0.001441 | 0.002256 | -0.001522 | -0.0063098 | -0.0041213 | -0.0039675 | 0.00077197 | 0.0016005 | 0.001864 | -0.0020148 | -0.0027127 | -0.0029456
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.075251 | -0.065918 | -0.093221 | 0.039009 | 0 | 0.077695 | -0.019839 | 0 | -0.082212 | -0.013563 | 0.091179 | -0.036233 | -0.089336 | -0.04684 | 0.047248 | 0.024843 | 0.071387 | 0.11334 | 0.013101 | -0.10518 |  | Dx (mm) | 0.20864 | -0.031752 | -0.20701 | -0.060234 | -0.05753 | 0 | -0.07162 | -0.097103 | -0.022706 | 0.086332 | 0.14111 | 0 | -0.11056 | -0.046121 | 0.013128 | 0.082597 | 0.12335 | 0.11342 | 0.043153 | -0.023561
Dy (mm) | 0.13204 | -0.082995 | -0.043002 | -0.019523 | 0.029571 | 0 | -0.026158 | 0.080289 | 0.075496 | 0 | 0.079209 | 0.039007 | 0 | -0.10716 | 0 | 0 | 0.036497 | -0.040212 | 0.053422 | -0.30336 |  | Dy (mm) | -0.23396 | -0.063436 | 0.076041 | -0.059432 | 0.092215 | -0.016411 | -0.020716 | 0.10365 | 0.032524 | -0.029661 | -0.02868 | 0.021328 | -0.012962 | 0 | -0.039272 | 0.01779 | 0.016851 | 0.019671 | 0 | -0.076605
Mag. len. (m) | 14.4206 | [20 empty] | Mag. len. (m) | 14.4221
TF (mT/KA) | 595.8932 | [20 empty] | TF (mT/KA) | 595.9372
Coil pos. (mm) | 0.017787 | [20 empty] | Coil pos. (mm) | 0.021253
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.015404 | [20 empty] | Twist (m2 rad) | 0.018726

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -0.069312 | 1.206 | 0.078587 | 0.30678 | 1.9975 | 3.3251 | 1.6053 | 0.45985 | -0.3495 | 0.45669 | -0.30936 | 0.40208 | 0.74791 | 2.6695 | 0.37962 | 0.51655 |  | 0
3 | -2.7082 | 8.4346 | -3.8307 | 0.96292 | 0.58259 | 0.46753 | 0.6603 | 0.28317 | -3.2781 | 9.5889 | -4.5708 | 0.85723 | 0.644 | 0.45143 | 0.64791 | 0.33332 |  | 0
4 | -0.035168 | 0.5261 | 0.030253 | 0.14362 | 0.39652 | 0.53175 | 0.32812 | 0.23349 | -0.068199 | 0.28988 | -0.044557 | 0.12638 | 0.21372 | 0.57052 | 0.13564 | 0.28527 |  | 0
5 | -0.38246 | 1.4777 | -0.094176 | 0.24699 | 0.18924 | 0.54624 | 0.12625 | 0.1236 | -0.42311 | 1.389 | -0.15256 | 0.21731 | 0.22231 | 0.71927 | 0.14689 | 0.14888 |  | 0
6 | 0.055423 | 0.079757 | 0.064513 | 0.044868 | 0.080101 | 0.12422 | 0.072314 | 0.07262 | -0.0004253 | 0.083915 | -0.0054123 | 0.061877 | 0.034431 | 0.075019 | 0.034594 | 0.074473 |  | 0
7 | 1.1705 | 0.35682 | 1.1324 | 0.090719 | 0.030506 | 0.31178 | -0.012293 | 0.043143 | 1.1337 | 0.36096 | 1.0958 | 0.10397 | 0.095999 | 0.37863 | 0.046588 | 0.065402 |  | 0
8 | -0.01046 | 0.026011 | -0.0095235 | 0.019019 | 0.020492 | 0.02705 | 0.020437 | 0.017763 | -0.020081 | 0.038069 | -0.020013 | 0.02541 | 0.018336 | 0.03611 | 0.021686 | 0.028237 |  | 0
9 | 0.53838 | 0.045563 | 0.54633 | 0.01097 | -0.016415 | 0.065888 | -0.0069506 | 0.025474 | 0.51373 | 0.043101 | 0.51919 | 0.017828 | 0.0068777 | 0.054512 | 0.015623 | 0.029493 |  | 0
10 | 0 | 0.029452 | 0.00027533 | 0.00093299 | 0 | 0.035985 | 0.0003701 | 0.0016635 | 0 | 0.045315 | -0.00024651 | 0.00078426 | 0 | 0.025607 | 0.00033321 | 0.00099151 |  | 0
11 | 0.74451 | 0.037576 | 0.75193 | 0.0076856 | -0.039862 | 0.040026 | -0.04545 | 0.0094355 | 0.75325 | 0.040641 | 0.76105 | 0.0061804 | 0.011546 | 0.039752 | 0.0064227 | 0.0097541 |  | 0
12 | -0.0016753 | 0.0043457 | -0.0020918 | 0.001674 | -0.00017274 | 0.0052919 | 3.6365e-05 | 0.0024847 | -0.00055268 | 0.0053295 | -0.00072217 | 0.0021856 | -0.00048006 | 0.0040063 | -0.00028076 | 0.0023314 |  | 0
13 | 0.069475 | 0.005497 | 0.069612 | 0.0019202 | -0.0020222 | 0.0062093 | -0.0010736 | 0.0015595 | 0.070328 | 0.0060343 | 0.070368 | 0.0019146 | 0.0027255 | 0.0056648 | 0.0035903 | 0.0024824 |  | 0
14 | -0.0013441 | 0.0017943 | -0.0014274 | 0.00057093 | -0.010127 | 0.004735 | -0.011053 | 0.00096006 | -0.0016659 | 0.0025328 | -0.0018714 | 0.00053521 | -0.0098736 | 0.0053498 | -0.011026 | 0.0014208 |  | 0
15 | 0.028062 | 0.0089972 | 0.029479 | 0.0023679 | -0.0033881 | 0.0038518 | -0.0036552 | 0.0028378 | 0.027986 | 0.0070489 | 0.029259 | 0.0023477 | -0.0029456 | 0.0041258 | -0.0033119 | 0.0033664 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.97924 | -0.78584
3 | [12 empty] | 3 | 9 | -0.64974 | -0.62635
4 | [12 empty] | 4 | 16 | -1.6976 | -1.6614
5 | [12 empty] | 5 | 25 | -1.7445 | -1.7155
6 | [12 empty] | 6 | 36 | -2.8633 | -2.69
7 | [12 empty] | 7 | 49 | -2.7716 | -2.4954
8 | [12 empty] | 8 | 64 | -3.9965 | -3.6199
9 | [12 empty] | 9 | 81 | -4.0914 | -3.7753
10 | [12 empty] | sigma (mm) |  | 0.017787 | 0.021253
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -5.4249 | 0.07409 | 0.57252 | 0.14187 | 0.37925 | 0.094216 | 0.32601 | 0.38725 | -0.079149 | -0.11165 | -0.34302 | -0.0015998 | -0.09423 | -0.042748 | -0.61847 | -0.26488 | 0.2504 | 0.3053 | 0.43942 | 0.39505 | -0.069312
3 | 33.7208 | -3.7048 | -3.228 | -4.0692 | -3.9507 | -4.3871 | -3.8966 | -3.1562 | -2.5091 | -2.8755 | -4.14 | -4.3004 | -4.101 | -6.1304 | -5.5 | -4.4683 | -2.6127 | -3.0584 | -2.8647 | -4.4625 | -2.7082
4 | -2.0587 | -0.089131 | 0.017765 | 0.032771 | -0.16723 | 0.144 | 0.17963 | 0.09292 | -0.12557 | -0.06075 | 0.10571 | 0.23492 | -0.0088233 | -0.084571 | -0.19679 | -0.08062 | 0.31698 | 0.14983 | 0.083528 | 0.035421 | -0.035168
5 | -5.6973 | -0.52572 | -0.58281 | -0.056873 | 0.039183 | 0.073974 | 0.041699 | -0.092137 | -0.19771 | 0.031349 | 0.44846 | 0.044216 | 0.14356 | -0.041164 | -0.029297 | -0.09351 | -0.2292 | -0.26857 | -0.40062 | -3.6984 | -0.38246
6 | -0.2328 | 0.022036 | 0.064013 | 0.10922 | 0.0065871 | 0.065096 | 0.045001 | 0.05066 | 0.12199 | 0.15601 | 0.082799 | 0.083767 | 0.056942 | 0.10658 | 0.087249 | 0.073892 | 0.037443 | 0.023845 | -0.031892 | 0.045266 | 0.055423
7 | 2.6513 | 0.83277 | 0.97703 | 1.1799 | 1.1684 | 1.2199 | 1.1967 | 1.1647 | 1.1504 | 1.1401 | 1.1739 | 1.1843 | 1.1193 | 1.1195 | 1.1146 | 1.1643 | 1.1401 | 1.1831 | 1.1547 | 0.88222 | 1.1705
8 | -0.058472 | -0.0040223 | -0.0026816 | 0.0013507 | -0.0044199 | -0.040695 | -0.013254 | -0.00039691 | -0.013298 | 0.029035 | -0.010928 | -0.026015 | -0.044128 | -0.022193 | -0.031398 | -0.00044778 | -0.015675 | 0.0098436 | 0.0179 | -0.0010719 | -0.01046
9 | 0.36474 | 0.55599 | 0.55343 | 0.54533 | 0.5479 | 0.54101 | 0.54526 | 0.54404 | 0.56697 | 0.5507 | 0.53996 | 0.54793 | 0.55349 | 0.5334 | 0.51618 | 0.53901 | 0.54837 | 0.56006 | 0.54498 | 0.48102 | 0.53838
10 | 2.9818e-05 | -4.6576e-06 | 4.7872e-06 | 7.6864e-07 | 0.0039492 | 2.4796e-05 | -9.4737e-07 | 0.00062096 | -6.1213e-06 | 0.00019363 | 8.1588e-06 | -3.7766e-06 | -0.00022837 | -1.28e-05 | 0.00015697 | 0.00025555 | 1.3088e-06 | -3.3556e-06 | -2.23e-07 | -0.00017255 | 0
11 | 0.62226 | 0.74662 | 0.74043 | 0.75669 | 0.75126 | 0.75334 | 0.74746 | 0.74598 | 0.74885 | 0.74966 | 0.76092 | 0.7542 | 0.75942 | 0.76593 | 0.76192 | 0.75771 | 0.74608 | 0.752 | 0.73635 | 0.64619 | 0.74451
12 | 0.022024 | -0.013551 | -0.0015407 | -0.033012 | -0.011663 | -0.019043 | -0.016202 | -0.019401 | -0.026162 | -0.028517 | -0.015133 | -0.020939 | -0.072456 | -0.011418 | -0.027182 | 0.00744 | -0.037837 | -0.012924 | -0.016978 | 0.021014 | -0.016753 | ten times!
13 | 0.83584 | 0.71773 | 0.71382 | 0.6869 | 0.69293 | 0.6747 | 0.68195 | 0.69468 | 0.69716 | 0.6842 | 0.65351 | 0.68681 | 0.67936 | 0.6874 | 0.70241 | 0.70805 | 0.72374 | 0.72282 | 0.72203 | 0.52178 | 0.69475 | ten times!
14 | 0.032205 | -0.017362 | -0.0059112 | -0.020804 | -0.012273 | -0.022109 | -0.010357 | -0.015193 | -0.011865 | -0.016791 | -0.013541 | -0.014931 | -0.028647 | -0.0091467 | -0.0056369 | -0.010165 | -0.011298 | -0.016275 | -0.014617 | 0.022402 | -0.013441 | ten times!
15 | -0.091704 | 0.35552 | 0.31726 | 0.29502 | 0.27927 | 0.28921 | 0.30451 | 0.25273 | 0.27024 | 0.2961 | 0.28212 | 0.27964 | 0.2869 | 0.32583 | 0.30802 | 0.29069 | 0.29461 | 0.31083 | 0.26772 | 0.22057 | 0.28062 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 15.6547 | 2.5652 | 1.2615 | 1.6677 | 1.9926 | 0.62876 | 1.2211 | 1.938 | 2.0124 | 1.0512 | 1.7667 | 1.5539 | 2.0383 | 1.9605 | 1.449 | 1.8168 | 1.2496 | 1.5343 | 1.1884 | -0.18024 | 1.9975
3 | -0.47749 | 1.1769 | 0.94402 | 0.50205 | 0.52962 | 0.94251 | 0.9298 | 0.79493 | 0.38501 | 0.31191 | 0.61003 | 0.39129 | 0.73989 | 0.62856 | 0.83087 | 0.77986 | 0.9542 | 0.19447 | 0.23944 | -0.65181 | 0.58259
4 | 2.6252 | 0.47497 | 0.061964 | 0.1358 | 0.19981 | 0.44081 | 0.183 | -0.19704 | 0.1301 | 0.40911 | 0.56263 | 0.59297 | 0.78616 | 0.46854 | 0.39746 | 0.30174 | 0.49939 | 0.13814 | 0.32066 | 0.098068 | 0.39652
5 | 2.4802 | 0.09094 | 0.038551 | 0.079243 | 0.25422 | 0.12604 | 0.17082 | 0.13857 | 0.15571 | 0.26949 | 0.2218 | -0.15409 | -0.080989 | 0.25619 | 0.28344 | 0.23235 | 0.0043792 | 0.029344 | 0.15656 | -0.15552 | 0.18924
6 | 0.30343 | 0.10363 | 0.11082 | 0.097352 | 0.12806 | 0.0056268 | 0.045832 | -0.0084134 | 0.10435 | 0.17693 | 0.15778 | 0.15734 | -0.11239 | 0.055677 | -0.014857 | 0.097127 | 0.085201 | 0.024448 | 0.087144 | -0.028424 | 0.080101
7 | 1.3711 | -0.10324 | -0.022598 | 0.002964 | 0.049389 | -0.01947 | -0.051427 | 0.0056225 | 0.039158 | 0.019225 | -0.028903 | -0.017947 | 0.012403 | 0.051869 | -0.0065299 | -0.093399 | 0.015811 | -0.034285 | -0.039908 | 0.00078248 | 0.030506
8 | 0.034266 | 0.023633 | 0.033861 | 0.044445 | 0.0068829 | -0.010494 | 0.0018294 | 0.031112 | 0.048204 | 0.007028 | 0.012286 | 0.0085561 | 0.022247 | 0.0023393 | 0.011905 | 0.041802 | 0.042656 | 0.0042348 | 0.035329 | 0.037879 | 0.020492
9 | -0.28231 | 0.016941 | 0.018042 | -0.047737 | -0.053088 | -0.012303 | 0.010228 | 0.023388 | 0.0026486 | -0.022152 | -0.0026682 | -0.0096995 | 0.035142 | -0.021414 | -0.030206 | 0.023047 | -0.0032534 | -0.014 | -0.038028 | -0.066839 | -0.016415
10 | -1.7284e-05 | 6.5471e-06 | 4.4934e-06 | -3.6155e-07 | -0.00017994 | 0.0004151 | 2.1798e-07 | -2.3806e-05 | -1.3237e-05 | 0.0040845 | -5.1362e-06 | -1.4838e-06 | -0.0036971 | -2.3553e-06 | 0.0025077 | 0.0035581 | 1.006e-06 | 6.6673e-06 | 9.7115e-07 | -0.00018074 | 0
11 | 0.12837 | -0.070183 | -0.044567 | -0.044675 | -0.034275 | -0.046334 | -0.044896 | -0.030984 | -0.028685 | -0.035526 | -0.047831 | -0.05293 | -0.049918 | -0.049255 | -0.047543 | -0.054456 | -0.04556 | -0.04383 | -0.046649 | -0.040054 | -0.039862 | ten times!
12 | -0.0034013 | 0.01418 | 0.01858 | -0.0042712 | 0.006681 | -0.011903 | -0.033325 | -0.0094874 | 0.0054815 | 0.025091 | 0.0082662 | 0.054946 | -0.067132 | -0.0082914 | 0.0089506 | 0.0064675 | -0.0076281 | 0.0046615 | -0.0047215 | -0.056457 | -0.0017274 | ten times!
13 | -0.28157 | 0.012567 | 0.0026226 | -0.021785 | 0.00046753 | -0.0066783 | -0.012104 | 0.0031824 | -0.0038687 | 0.00092836 | -0.021094 | -0.014703 | 0.0047171 | -0.014477 | 0.0037633 | -0.024778 | -0.022918 | -0.030086 | -0.048998 | -0.06144 | -0.020222 | ten times!
14 | 0.089372 | -0.10765 | -0.12003 | -0.1011 | -0.11712 | -0.12312 | -0.12525 | -0.10403 | -0.1078 | -0.12584 | -0.11371 | -0.098 | -0.092717 | -0.10336 | -0.11678 | -0.1028 | -0.1145 | -0.11025 | -0.10541 | 0.0019409 | -0.10127 | ten times!
15 | 0.081968 | -0.0926 | -0.076288 | -0.026879 | -0.037768 | -0.0093878 | -0.03471 | -0.023565 | -0.053045 | -0.049761 | 0.0081288 | -0.036092 | -0.08144 | -0.063427 | -0.017046 | -0.01359 | -0.011163 | -0.0014761 | -0.037819 | -0.065727 | -0.033881 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.8835 | 0.04256 | -0.0055278 | -0.044195 | -0.37097 | -0.37248 | 0.22156 | -0.21652 | -0.091031 | -0.19324 | -0.31636 | -1.0213 | -0.00095109 | -0.4631 | -0.41886 | -1.2523 | -0.95108 | -0.19337 | 0.078613 | -1.0993 | -0.3495
3 | 38.1428 | -4.0597 | -3.8597 | -5.1226 | -4.5468 | -5.8019 | -3.6846 | -4.2406 | -4.707 | -4.0531 | -4.6801 | -5.5291 | -4.5218 | -6.3483 | -5.4135 | -4.8532 | -3.8735 | -4.2029 | -2.7762 | -3.6089 | -3.2781
4 | -1.0941 | 0.065589 | 0.083697 | -0.037506 | -0.0039979 | -0.046942 | 0.10029 | 0.015093 | -0.03091 | -0.029882 | 0.16142 | -0.21532 | -0.30605 | -0.102 | 0.015787 | -0.17042 | -0.26626 | 0.020804 | -0.055412 | 0.19675 | -0.068199
5 | -5.2621 | -0.53404 | -0.62982 | -0.27802 | -0.14858 | 0.13816 | 0.068415 | -0.039223 | -0.10347 | -0.13105 | 0.11805 | 0.11456 | -0.029718 | -0.0029796 | -0.16074 | -0.21935 | -0.37177 | -0.22805 | -0.30851 | -3.8408 | -0.42311
6 | -0.2908 | -0.023057 | 0.045307 | -0.030113 | -0.043983 | 0.041308 | -0.10268 | 0.017526 | -0.11504 | 0.064203 | 0.089501 | 0.089747 | -0.095264 | 0.027305 | -0.022479 | -0.035964 | 0.039832 | -0.016469 | -0.0271 | 0.14679 | -0.0004253
7 | 2.6251 | 0.78128 | 0.90834 | 1.1217 | 1.1755 | 1.2284 | 1.1402 | 1.1059 | 1.0934 | 1.1572 | 1.1436 | 1.1478 | 1.1015 | 1.0795 | 1.0576 | 1.1858 | 1.0823 | 1.1516 | 1.0627 | 0.85661 | 1.1337
8 | -0.034585 | -0.013379 | -0.012117 | -0.076122 | -0.049035 | -0.0095889 | -0.0053012 | 0.012972 | -0.030544 | -0.0072006 | -0.0090377 | -0.063944 | -0.037648 | -0.03631 | -0.033306 | -0.0085944 | -0.0020359 | 0.016593 | 0.0043587 | 0.024363 | -0.020081
9 | 0.34708 | 0.53756 | 0.53135 | 0.51198 | 0.5253 | 0.49554 | 0.52394 | 0.51223 | 0.52852 | 0.53159 | 0.49937 | 0.49053 | 0.50879 | 0.50801 | 0.494 | 0.52079 | 0.52829 | 0.55303 | 0.54458 | 0.46595 | 0.51373
10 | 0.00012194 | -5.9013e-07 | 2.8402e-05 | -4.1961e-06 | -7.8078e-07 | -0.0031751 | 7.6444e-06 | -1.2108e-05 | -1.5745e-06 | -6.45e-06 | -7.9686e-06 | -0.00121 | 5.7083e-06 | -5.2999e-05 | 3.4732e-07 | 1.5274e-08 | -1.2632e-05 | -1.0629e-05 | 1.5711e-05 | 4.9021e-06 | 0
11 | 0.61539 | 0.75724 | 0.7542 | 0.76191 | 0.75876 | 0.76412 | 0.76092 | 0.76381 | 0.76576 | 0.75257 | 0.76612 | 0.76496 | 0.76302 | 0.76975 | 0.77105 | 0.76068 | 0.75922 | 0.75946 | 0.74542 | 0.64925 | 0.75325
12 | -0.021351 | 0.010637 | 0.011933 | 0.0086336 | -0.012958 | -0.02398 | -0.070877 | 0.00010676 | 0.012608 | 0.0074097 | -0.025144 | -0.022468 | 0.0093926 | -0.012883 | -0.018888 | -0.026772 | 0.0061904 | 0.0014524 | 0.015618 | 0.074208 | -0.0055268 | ten times!
13 | 0.83248 | 0.68777 | 0.71445 | 0.71407 | 0.75404 | 0.69521 | 0.69923 | 0.69226 | 0.68036 | 0.70072 | 0.72956 | 0.705 | 0.71276 | 0.6871 | 0.72359 | 0.70232 | 0.69835 | 0.69633 | 0.67311 | 0.51332 | 0.70328 | ten times!
14 | 0.033175 | -0.015205 | -0.016107 | -0.023435 | -0.029595 | -0.02116 | -0.021708 | -0.024244 | -0.014802 | -0.019735 | -0.025301 | -0.011038 | -0.010045 | -0.011508 | -0.014299 | -0.018184 | -0.01771 | -0.020724 | -0.022043 | 5.1367e-05 | -0.016659 | ten times!
15 | 0.023925 | 0.35442 | 0.26465 | 0.29965 | 0.25794 | 0.29108 | 0.29548 | 0.25305 | 0.29732 | 0.30475 | 0.32402 | 0.29319 | 0.29131 | 0.30814 | 0.29864 | 0.2744 | 0.29188 | 0.28249 | 0.28417 | 0.14604 | 0.27986 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 13.1429 | 0.40602 | 0.17803 | 0.036903 | 0.32819 | 0.84794 | 0.61927 | 0.53995 | 0.6414 | 0.56757 | 0.43457 | 0.69866 | 0.88871 | -0.93104 | -0.41429 | 0.80772 | -0.39097 | 0.51649 | 1.0581 | 1.0796 | 0.74791
3 | 1.4907 | 1.2964 | 0.64339 | 0.43403 | 0.73335 | 0.65814 | 0.60666 | 0.23038 | 0.77317 | 0.36334 | 1.0887 | 1.0821 | 0.87254 | 0.34496 | 0.16271 | 0.81135 | 0.94437 | 0.44398 | 0.17284 | -0.35242 | 0.644
4 | 1.9266 | -0.00040368 | 0.085145 | -0.16043 | 0.316 | -0.31805 | -0.0063726 | 0.22421 | 0.045335 | 0.12679 | 0.46721 | 0.50543 | 0.48616 | 0.52155 | -0.4267 | 0.23135 | 0.43175 | 0.0020271 | -0.089502 | 0.41446 | 0.21372
5 | 3.2433 | 0.34748 | 0.046561 | -0.051644 | 0.028324 | 0.031053 | 0.014031 | -0.046992 | 0.22922 | 0.31226 | 0.31644 | 0.19922 | 0.11944 | 0.37647 | 0.11597 | 0.22975 | 0.24655 | -0.099558 | 0.22939 | -0.32603 | 0.22231
6 | 0.068867 | 0.10136 | 0.09732 | 0.044514 | -0.021303 | -0.069027 | 0.043269 | 0.056669 | 0.063491 | 0.12411 | 0.094664 | 0.14568 | 0.034421 | -0.073724 | -0.10878 | 0.036829 | 0.10874 | -0.0053235 | -0.050203 | 0.037006 | 0.034431
7 | 1.7195 | -0.010775 | -0.020589 | -0.0057573 | 0.03013 | 0.014501 | -0.022161 | 0.14625 | 0.1675 | 0.086762 | 0.030986 | 0.15296 | 0.083567 | 0.12736 | 0.056391 | -0.031869 | 0.027966 | -0.0035318 | 0.0088966 | 0.062796 | 0.095999
8 | 0.0087159 | 0.030094 | 0.028708 | 0.037104 | 0.048484 | 0.062109 | 0.06686 | 0.016742 | 0.0027665 | 0.041463 | 0.039366 | 0.048636 | 0.0075267 | -0.032476 | 0.015317 | 0.014743 | 0.0015604 | -0.019728 | -0.018925 | -0.040135 | 0.018336
9 | -0.16788 | 0.062174 | 0.032602 | -0.0069874 | -0.0448 | -0.012882 | 0.02452 | 0.041044 | 0.031989 | -0.020233 | 0.018849 | 0.045098 | 0.048424 | 0.012627 | -0.032434 | 0.039443 | 0.020476 | 0.019716 | 0.0015885 | -0.05109 | 0.0068777
10 | 6.6313e-05 | -7.7909e-07 | -5.3936e-05 | 1.655e-06 | -8.0828e-06 | -2.4986e-05 | 6.5625e-06 | -1.2202e-05 | 7.232e-07 | 1.0464e-06 | 5.0813e-06 | -2.0142e-05 | 3.9218e-06 | 0.0028666 | -1.2395e-06 | -1.2558e-07 | -5.7949e-06 | 7.4265e-07 | 0.0032387 | -7.5133e-06 | 0
11 | 0.1821 | -0.009019 | -0.0046698 | -8.9038e-05 | 0.00057549 | 0.0059751 | 0.01166 | 0.02069 | 0.028113 | 0.010564 | 0.011382 | 0.012972 | 0.012783 | 0.014532 | 0.0046942 | -0.0044794 | 0.0049198 | -0.0011397 | -0.0038527 | 0.0093899 | 0.011546
12 | 0.052547 | 0.011759 | 0.0038592 | -0.011314 | -0.045504 | -0.031662 | -0.0060999 | -0.0032075 | 0.047652 | 0.028523 | 0.0039315 | 0.036186 | -0.0023881 | -0.010665 | -0.023448 | -0.0065078 | -0.00070745 | -0.023672 | -0.017273 | -0.010163 | -0.0048006 | ten times!
13 | -0.19593 | 0.072889 | 0.043015 | -0.015373 | 0.0065339 | 0.02717 | 0.022871 | 0.032452 | 0.050406 | 0.033199 | 0.042516 | 0.080631 | 0.064364 | 0.048035 | 0.0079574 | 0.044352 | 0.05509 | 0.014723 | 0.015431 | -0.024039 | 0.027255 | ten times!
14 | 0.10185 | -0.13125 | -0.12881 | -0.10836 | -0.081915 | -0.09759 | -0.10127 | -0.10895 | -0.12723 | -0.099695 | -0.11442 | -0.094269 | -0.10617 | -0.11178 | -0.12239 | -0.090828 | -0.11288 | -0.11972 | -0.12712 | -0.040699 | -0.098736 | ten times!
15 | 0.08437 | -0.053093 | -0.093013 | -0.072714 | -0.05329 | -0.032498 | -0.06742 | -0.047795 | -0.047482 | -0.01441 | 0.02256 | -0.01522 | -0.063098 | -0.041213 | -0.039675 | 0.0077197 | 0.016005 | 0.01864 | -0.020148 | -0.027127 | -0.029456 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00047674 | -6.8161e-05 | 4.3772e-05 | -0.00029924 | -0.00014401 | -7.6552e-05 | -0.00014116 | -7.6216e-05 | -5.0876e-05 | -1.6306e-05 | -3.4766e-05 | -0.00013495 | -0.00029102 | 7.6915e-06 | 2.8836e-05 | 0.00010435 | 6.6091e-05 | 0.00060579 |  | 595.8932
|  | 0.00047652 | -9.5182e-05 | 5.3324e-06 | 6.507e-06 | -0.00012421 | -0.0001128 | -5.4741e-05 | -2.9906e-05 | -0.00040142 | -0.00012304 | -4.0142e-05 | -0.0001232 | 7.732e-05 | -7.3032e-05 | -1.9167e-05 | 0.00014545 | -2.8228e-05 | 0.00051394 |  | 595.9372
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3975.6138
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0287 | -4146.5184 | 5.4443 | -0.0050348 | 1.1144 | -2.316 | -0.76361 | -0.088948 | -0.73508 | -0.085592 | 0.16783 | 0.51355 | 0.32894 | -0.67298 | -2.2338 | 0.75354 | 0.965 | 1.7202 | 1.3376 | 6.7349 | -3804.7092
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.061188 | 2.334 | -0.53701 | 0.12821 | -0.010268 | 0.32331 | 0.28587 | 0.52183 | 0.88162 | 0.64525 | 0.18764 | -0.69095 | -0.065112 | 0.12349 | -0.14967 | -0.094184 | -0.37154 | -0.36998 | -0.48335 | -0.32515 | -0.30677
db2 | 0.077828 | -5.1194 | 0.036378 | 0.53487 | 0.046526 | 0.29958 | -0.023694 | 0.26013 | 0.3019 | -0.1361 | -0.18489 | -0.47111 | -0.062858 | -0.12896 | -0.078952 | -0.72688 | -0.35577 | 0.1465 | 0.18768 | 0.35566 | 0.42633
db3 | -2.7401 | 33.6331 | 0.13332 | 0.60214 | -0.23833 | -0.12162 | -0.55493 | -0.066266 | 0.67672 | 1.3212 | 0.9548 | -0.31614 | -0.4759 | -0.27069 | -2.2907 | -1.6715 | -0.63855 | 1.2182 | 0.77562 | 0.96265 | -4.445
db4 | 0.03121 | -2.2371 | -0.11047 | -0.00030923 | 0.0014225 | -0.20021 | 0.11418 | 0.14927 | 0.05907 | -0.1556 | -0.092058 | 0.079813 | 0.20478 | -0.043039 | -0.10893 | -0.22831 | -0.11237 | 0.28186 | 0.11179 | 0.04911 | 0.11631
db5 | -0.29435 | -5.7203 | -0.42937 | -0.4881 | 0.039431 | 0.13232 | 0.17058 | 0.13611 | 0.0016352 | -0.10845 | 0.1251 | 0.54165 | 0.13584 | 0.23735 | 0.052339 | 0.066526 | 0.0014562 | -0.13454 | -0.17179 | -0.30808 | -3.7098
db6 | 0.066042 | -0.22637 | -0.0664 | -0.03457 | 0.059453 | -0.05997 | 0.032488 | -0.029893 | -0.01554 | 0.021642 | 0.084506 | 0.055278 | 0.0028361 | -0.044469 | 0.024074 | 0.039787 | 0.018059 | -0.00011571 | 0.0046948 | -0.091861 | 0.013214
db7 | 1.2074 | 2.6478 | -0.29908 | -0.15401 | 0.047928 | 0.035799 | 0.086805 | 0.064354 | 0.030501 | 0.017031 | 0.0075589 | 0.040846 | 0.052076 | -0.010715 | -0.013345 | -0.018273 | 0.031929 | 0.0069159 | 0.052342 | 0.021337 | 0.88035
db8 | -0.0091505 | -0.028334 | -0.011345 | -0.017483 | 0.020074 | 0.0054684 | -0.011712 | -0.009064 | 0.0078666 | -0.026079 | 0.03467 | 0.021394 | -0.026015 | -0.058327 | -0.025707 | -0.010759 | 0.015006 | 0.011965 | 0.04869 | 0.031354 | -0.031688
db9 | 0.51864 | 0.36166 | 0.0090007 | 0.0084822 | -0.0010546 | 0.0010666 | -0.004211 | -0.0014325 | -0.0040962 | 0.020469 | 0.0041111 | -0.0059769 | 0.001203 | 0.0087571 | -0.015293 | -0.029923 | -0.007472 | 0.0026876 | 0.015928 | -0.0022457 | 0.46555
db10 | 0.0010412 | 0.017518 | -0.033406 | -0.042777 | 0.015808 | 0.0035047 | 0.033418 | -0.010454 | 0.0010479 | -0.03597 | -0.0068289 | 0.041986 | -0.015897 | -0.041213 | -0.025228 | 0.020291 | 0.010287 | 0.031266 | 0.048065 | 0.006099 | -0.046817
db11 | 0.75612 | 0.6222 | -0.0053138 | -0.011261 | 0.0046631 | -0.00075445 | 0.0014538 | -0.0045553 | -0.006153 | -0.0029744 | -0.0022809 | 0.0089006 | 0.0021281 | 0.0081182 | 0.013655 | 0.009939 | 0.0057668 | -0.0059193 | 0.00029575 | -0.015709 | 0.64321
db12 | -0.0019868 | 0.0091647 | -0.0025535 | -0.0028144 | 0.00053752 | 0.00082005 | 0.0037816 | -0.0006031 | 3.3453e-05 | -0.0047147 | -0.0015201 | 0.0048722 | -0.0018364 | -0.0095643 | -0.0014533 | 0.0016121 | 0.0039723 | 0.0019403 | 0.00634 | 0.0011501 | -0.0028673
db13 | 0.06967 | 0.082833 | 0.0015144 | 0.0015347 | -0.0011741 | -0.00012797 | -0.002222 | -0.001713 | 0.00033375 | 0.00071953 | -0.0012293 | -0.0036956 | -0.00066549 | -0.0013938 | -0.0018946 | 0.00059424 | 0.0011265 | 0.0030353 | 0.0023511 | 0.0029065 | 0.051183
db14 | -0.0014462 | 0.0017609 | -0.0028529 | -0.0018507 | 0.00027034 | 0.00031082 | 0.0010858 | -0.00016178 | 8.2708e-05 | -0.0012402 | -0.00056371 | 0.0021574 | -0.00076543 | -0.0035293 | -0.0012851 | 0.002081 | 0.0010244 | 0.002082 | 0.002715 | 0.00043972 | -0.0013124
db15 | 0.029479 | -0.0091704 | 0.0060733 | 0.0022468 | 2.2824e-05 | -0.0015522 | -0.00055858 | 0.0009723 | -0.0042061 | -0.0024554 | 0.00013106 | -0.001267 | -0.0015152 | -0.00078865 | 0.0031042 | 0.0013229 | -0.00041003 | -1.7786e-05 | 0.0016042 | -0.0027067 | 0.022057
da2 | 1.6058 | 16.1728 | 0.98633 | -0.33829 | 0.073561 | 0.37299 | -0.96839 | -0.37489 | 0.30238 | 0.38059 | -0.55511 | 0.12886 | -0.073392 | 0.42481 | 0.4284 | -0.15219 | 0.21322 | -0.35934 | -0.054466 | -0.43508 | -0.013367
da3 | 0.66061 | -0.49653 | 0.51182 | 0.28207 | -0.15773 | -0.13188 | 0.28797 | 0.26772 | 0.1356 | -0.27906 | -0.34968 | -0.039764 | -0.2715 | 0.066862 | -0.034442 | 0.1736 | 0.12058 | 0.30187 | -0.46433 | -0.41969 | -0.66484
da4 | 0.32943 | 2.4915 | 0.15438 | -0.26235 | -0.19265 | -0.12935 | 0.11369 | -0.14748 | -0.52822 | -0.20592 | 0.078827 | 0.24647 | 0.26525 | 0.45837 | 0.13738 | 0.071176 | -0.026324 | 0.16811 | -0.188 | -0.01335 | 0.36666
da5 | 0.12648 | 2.4816 | -0.037113 | -0.090977 | -0.046932 | 0.1278 | -0.00032751 | 0.04386 | 0.013282 | 0.027941 | 0.1423 | 0.10323 | -0.28163 | -0.20451 | 0.1267 | 0.15675 | 0.10657 | -0.11929 | -0.097763 | 0.030092 | -0.1529
da6 | 0.074627 | 0.46323 | 0.0069784 | 0.022103 | 0.014639 | 0.065628 | -0.069544 | -0.039486 | -0.050086 | 0.059311 | 0.10221 | 0.11507 | 0.099259 | -0.18738 | -0.062168 | -0.08959 | 0.021684 | 0.025688 | -0.068406 | 0.03409 | -0.1229
da7 | -0.012228 | 1.371 | -0.09035 | -0.010734 | 0.015723 | 0.061572 | -0.0075922 | -0.038962 | 0.017808 | 0.047976 | 0.031413 | -0.015057 | -0.0065648 | 0.023866 | 0.066098 | 0.0059163 | -0.080741 | 0.02961 | -0.023052 | -0.026929 | 0.0058648
da8 | 0.021232 | 0.04686 | -0.019906 | 0.00052692 | 0.017318 | -0.0067272 | -0.032178 | -0.026212 | 0.030379 | 0.04713 | -0.014062 | 0.011238 | -0.0024426 | -0.00046445 | -0.045272 | -0.0099972 | 0.02084 | 0.030789 | -0.028511 | 0.027551 | -0.025956
da9 | -0.0066177 | -0.27893 | 0.026158 | 0.026414 | -0.041494 | -0.046399 | -0.0057552 | 0.017091 | 0.030094 | 0.0063616 | -0.015565 | 0.0073421 | -0.0037422 | 0.041812 | -0.012261 | -0.023558 | 0.029701 | 0.0045217 | -0.0096716 | -0.03105 | -0.053051
da10 | 0.0021177 | 0.054037 | -0.035856 | -0.018414 | -0.011833 | 0.01077 | -0.0038231 | -0.013095 | 0.033082 | 0.032546 | 0.0022504 | 0.030785 | 0.016318 | -0.0032123 | -0.049037 | -0.00093059 | 0.00064476 | 0.011989 | -0.022844 | 0.020661 | -0.11259
da11 | -0.045416 | 0.12944 | -0.024401 | 0.001006 | 0.00072318 | 0.011119 | -0.00098139 | 0.00062241 | 0.014331 | 0.016175 | 0.0098684 | -0.0020105 | -0.0077839 | -0.0041093 | -0.0034199 | -0.0020997 | -0.0090305 | -9.9819e-05 | 0.0013336 | -0.0012425 | -0.038388
da12 | 0.00018072 | 0.0059196 | -0.003034 | -0.00056725 | -0.00162 | 0.001939 | -0.001448 | -0.0047543 | 0.002838 | 0.0041182 | 0.0023262 | 0.0041592 | 0.007247 | -0.006961 | -0.0061116 | 0.00071274 | 0.00041915 | 0.00078574 | -0.0020864 | 0.0020374 | -0.016201
da13 | -0.0010756 | -0.027493 | 0.0032449 | 0.0023402 | -0.0014036 | 0.0010967 | -0.00032732 | 9.6406e-05 | 0.0013101 | 0.0010831 | 0.0012984 | -0.001651 | -0.000217 | 0.0021398 | 0.00031619 | 0.0010276 | -0.001598 | -0.0017787 | -0.003018 | -0.0039596 | -0.0054577
da14 | -0.010916 | 0.008448 | -0.0017758 | -0.0016246 | 0.00024546 | -0.00011539 | -0.0014559 | -0.002208 | 0.0021844 | 0.0021752 | -0.0016124 | 0.0014461 | 0.0021221 | 0.0022436 | -0.0020509 | -0.00082856 | 0.0006084 | 0.00028597 | -0.0011523 | 0.0015126 | -0.0047471
da15 | -0.0036552 | 0.0081968 | -0.0056048 | -0.0039737 | 0.00096732 | -0.00012161 | 0.0027164 | 0.00018417 | 0.0012986 | -0.0016494 | -0.001321 | 0.0044681 | 4.5936e-05 | -0.0044889 | -0.0026876 | 0.0019506 | 0.0022961 | 0.0025388 | 0.0035076 | -0.00012678 | -0.0065727
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4025.636
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0761 | -4335.546 | 5.3846 | -0.33279 | 0.67241 | 0.68415 | -0.62311 | -0.509 | 0.071635 | 0.32 | -3.3954 | -0.61136 | 0.21763 | -0.61304 | 1.3923 | -0.11128 | 0.4274 | 2.0736 | 0.33678 | 5.7588 | -3715.7259
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.059435 | 3.183 | -0.54974 | -0.068569 | -0.010645 | 0.49643 | 0.08053 | 0.42495 | 0.87336 | 0.70103 | -0.036823 | -0.62017 | 0.098476 | 0.26718 | 0.15322 | -0.051429 | -0.55268 | -0.39338 | -0.38159 | -0.43014 | -1.0525
db2 | -0.30621 | -0.90116 | 0.37361 | 0.38907 | 0.30138 | -0.04187 | -0.064999 | 0.56031 | 0.13543 | 0.2248 | 0.073082 | -0.084029 | -0.71781 | 0.3653 | -0.12245 | -0.12027 | -0.99494 | -0.70307 | 0.055739 | 0.37073 | -1.0926
db3 | -3.4819 | 38.1976 | 0.51182 | 0.70704 | -0.55196 | 0.020031 | -1.2309 | 0.88598 | 0.32691 | -0.13524 | 0.51901 | -0.1016 | -0.9591 | 0.057103 | -1.7756 | -0.84454 | -0.28461 | 0.69104 | 0.3693 | 1.7953 | -3.6033
db4 | -0.043187 | -1.1741 | 0.11793 | 0.15694 | 0.0088882 | 0.040622 | -0.0036365 | 0.14234 | 0.060452 | 0.011088 | 0.012909 | 0.21091 | -0.17314 | -0.26185 | -0.058785 | 0.059541 | -0.13247 | -0.23479 | 0.05839 | -0.015359 | 0.21211
db5 | -0.35153 | -5.2608 | -0.38064 | -0.4798 | -0.12516 | 0.0036641 | 0.28932 | 0.22294 | 0.11023 | 0.048182 | 0.024033 | 0.2764 | 0.26512 | 0.1264 | 0.14842 | -0.010696 | -0.068075 | -0.21767 | -0.075542 | -0.15712 | -3.8414
db6 | -0.0067663 | 0.043832 | -0.025286 | -0.013885 | -0.047339 | -0.061907 | 0.048159 | -0.12494 | -0.018838 | -0.11896 | 0.10718 | 0.15364 | 0.095366 | -0.13121 | 0.016485 | -0.010023 | 0.0055731 | 0.093608 | 0.036517 | -0.0041442 | 0.14138
db7 | 1.1717 | 2.6187 | -0.31487 | -0.18424 | 0.027771 | 0.077405 | 0.13209 | 0.044665 | 0.007902 | -0.003143 | 0.061357 | 0.048654 | 0.050544 | 0.007698 | -0.016333 | -0.039201 | 0.089302 | -0.013006 | 0.05715 | -0.033749 | 0.85466
db8 | -0.021367 | -0.020792 | 0.0018512 | -0.044256 | -0.069415 | -0.039803 | 0.011679 | -0.0013835 | 0.0091737 | -0.015307 | 0.035528 | 0.046025 | -0.043029 | -0.0426 | -0.02598 | -0.0094931 | 0.03274 | 0.050031 | 0.067445 | 0.036794 | 0.017393
db9 | 0.49258 | 0.35041 | 0.016278 | 0.019432 | -0.0085789 | 0.0034998 | -0.0251 | 0.0044652 | -0.0085336 | 0.0078277 | 0.012569 | -0.016626 | -0.030156 | -0.0089384 | -0.012059 | -0.026911 | 0.0017532 | 0.01124 | 0.035404 | 0.024433 | 0.46486
db10 | -0.0028549 | 0.10024 | -0.011623 | -0.088844 | -0.024137 | -0.022844 | -0.00026221 | -0.02907 | -0.042009 | -0.0079087 | 0.041267 | 0.066666 | 0.0014826 | -0.046684 | -0.018084 | 0.0089169 | 0.039864 | 0.057916 | 0.053559 | 0.021795 | -0.0085782
db11 | 0.7654 | 0.61462 | -0.0040882 | -0.0059195 | 0.00059201 | -0.002345 | 0.0028266 | 0.00014071 | 0.0025661 | 0.0043785 | -0.0083647 | 0.0052892 | 0.0036481 | 0.0020813 | 0.0086095 | 0.0096746 | -0.00049879 | -0.0015112 | -0.0013485 | -0.015731 | 0.64897
db12 | -0.0010489 | 0.0074458 | 0.00093421 | -0.0087033 | -0.0011424 | -0.0033752 | -0.0013173 | -0.0095341 | -0.0040107 | 0.0010947 | 0.0061002 | 0.0058488 | -0.001319 | -0.0035236 | -0.002479 | -0.00015377 | 0.0024246 | 0.0077163 | 0.0067798 | 0.0046599 | 0.006448
db13 | 0.070504 | 0.086101 | -0.0023814 | 0.0023775 | 0.00055298 | 0.0054871 | -0.001111 | -0.00052277 | -0.00031627 | -0.0021412 | -0.00066781 | 0.0023284 | 0.00014063 | 0.00098368 | -0.0017122 | 0.0014442 | -0.0001531 | -0.00042106 | -0.00065445 | -0.0032333 | 0.051051
db14 | -0.0019292 | 0.0053541 | -0.00079545 | -0.0036109 | -0.0022566 | -0.0018476 | -0.00023069 | -0.0020994 | -0.0021108 | 2.0291e-05 | 0.0020872 | 0.0032179 | 0.00085211 | -0.0017949 | -0.00039194 | 0.00069389 | 0.001966 | 0.0031009 | 0.0024653 | 0.00073477 | -0.00044937
db15 | 0.029259 | 0.0023925 | 0.0061828 | -0.0027936 | 0.0007065 | -0.0034645 | -0.00015074 | 0.00028915 | -0.0039536 | 0.00047285 | 0.0012157 | 0.0031436 | 6.0166e-05 | -0.00012811 | 0.0015557 | 0.00060487 | -0.0018189 | -7.0474e-05 | -0.0010099 | -0.0008415 | 0.014604
da2 | 0.37697 | 12.1314 | 0.054502 | -0.24916 | -0.30734 | -0.1031 | 0.48216 | 0.24617 | 0.10862 | 0.24435 | 0.20845 | 0.091617 | 0.30781 | 0.50738 | -1.3099 | -0.76599 | 0.4285 | -0.7618 | 0.13571 | 0.68199 | 1.1131
da3 | 0.64803 | 1.6266 | 0.64749 | -0.0057458 | -0.21198 | 0.082105 | 0.010247 | -0.041648 | -0.42118 | 0.12478 | -0.28249 | 0.45172 | 0.43326 | 0.21578 | -0.30728 | -0.48641 | 0.1662 | 0.30504 | -0.20403 | -0.47585 | -0.3573
da4 | 0.13623 | 2.3752 | -0.13128 | -0.064254 | -0.29206 | 0.17621 | -0.45481 | -0.14318 | 0.088134 | -0.092897 | -0.0021643 | 0.34072 | 0.36978 | 0.34691 | 0.38123 | -0.56107 | 0.098689 | 0.30136 | -0.1379 | -0.2234 | 0.48553
da5 | 0.14664 | 3.239 | 0.20065 | -0.10808 | -0.19771 | -0.12046 | -0.11579 | -0.13256 | -0.19705 | 0.080878 | 0.16766 | 0.17207 | 0.053132 | -0.027528 | 0.23085 | -0.030959 | 0.084155 | 0.10456 | -0.24594 | 0.08211 | -0.32931
da6 | 0.03691 | -0.020907 | 0.047058 | 0.086839 | -0.015897 | -0.020528 | -0.11305 | -0.001385 | 0.055229 | 0.037797 | 0.077748 | 0.047786 | 0.1174 | -0.010813 | -0.11272 | -0.16009 | 0.0064607 | 0.079513 | -0.034374 | -0.086982 | 0.013376
da7 | 0.046145 | 1.7173 | -0.056581 | -0.072674 | -0.050249 | -0.020055 | -0.031578 | -0.069908 | 0.097983 | 0.12076 | 0.041622 | -0.013492 | 0.10625 | 0.03768 | 0.081835 | 0.010777 | -0.077228 | -0.017624 | -0.049933 | -0.037584 | 0.062794
da8 | 0.02304 | -0.048243 | -0.0099161 | 0.022542 | -0.00010621 | 0.049467 | 0.035242 | 0.03785 | 0.01702 | -0.012524 | 0.010111 | 0.010651 | 0.030518 | -0.021192 | -0.05579 | -0.017048 | -0.0023643 | -0.016016 | -0.036513 | -0.041931 | -0.056334
da9 | 0.014982 | -0.1872 | 0.04815 | 0.010157 | -0.02027 | -0.062289 | -0.027837 | 0.010259 | 0.021266 | 0.016651 | -0.036393 | 0.0020021 | 0.030101 | 0.034171 | -0.0024152 | -0.047665 | 0.025149 | 0.0065007 | 0.0057853 | -0.013322 | -0.050376
da10 | 0.0029926 | -0.062502 | -0.031079 | 0.031423 | -0.029631 | 0.038155 | -0.010394 | -0.012765 | 0.042418 | 0.011282 | -0.015592 | -0.014988 | 0.0065836 | -0.0096497 | -0.00052129 | -0.020767 | 0.0047526 | 0.0049015 | 0.0057233 | 0.00014709 | -0.029378
da11 | 0.0063504 | 0.17939 | -0.015323 | -0.011968 | -0.0061962 | -0.0060526 | -0.00035014 | 0.0055335 | 0.013713 | 0.021671 | 0.0042029 | 0.0048706 | 0.0065889 | 0.0065754 | 0.0082186 | -0.0016624 | -0.010789 | -0.0012657 | -0.0075168 | -0.01025 | 0.0091273
da12 | -1.2237e-05 | -0.011729 | -0.0020366 | 0.0035057 | -0.0040662 | 0.00038197 | -0.0039578 | -0.0017683 | 0.0045836 | 0.0062636 | 0.0015609 | -0.00077679 | 0.0046944 | -0.0014542 | -0.0012233 | -0.0043206 | 0.00047175 | 0.0011681 | -0.0013453 | -0.001681 | -0.0037388
da13 | 0.0035769 | -0.018768 | 0.0042044 | 0.0019282 | -0.0043739 | -0.0029267 | -0.00083277 | -0.00066551 | -3.1559e-05 | 0.0016218 | -0.00097458 | -0.00064248 | 0.0044687 | 0.0037718 | 0.0016153 | -0.002877 | 0.0003139 | 0.00093603 | -0.0030801 | -0.0024556 | -0.0022885
da14 | -0.010708 | 0.011174 | -0.0041297 | 0.0010701 | -0.0012341 | 0.0047278 | 0.0005556 | 0.00047449 | 0.0023546 | -0.00113 | -0.00010834 | -0.0012367 | 0.001796 | 0.00035414 | -0.00031308 | -0.0025399 | 0.0020798 | -1.2734e-05 | -0.00063192 | -0.0020761 | -0.0049387
da15 | -0.0033119 | 0.008437 | -0.0019974 | -0.0059894 | -0.0039595 | -0.0020171 | 6.2112e-05 | -0.0034301 | -0.0014675 | -0.0014363 | 0.0018709 | 0.0055679 | 0.0017899 | -0.0029979 | -0.00080939 | -0.00065559 | 0.0040839 | 0.0049124 | 0.0051759 | 0.0012971 | -0.0027127
Field Colinearity | 19.1426
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.2728 | -0.5982 | 0.06702 | -0.071456 | 0.26212 | 0.22468 | 0.46064 | 0.82044 | 0.58406 | 0.12646 | -0.75214 | -0.1263 | 0.062306 | -0.21086 | -0.15537 | -0.43273 | -0.43116 | -0.54454 | -0.38633 | -0.36796
Twist (m2 mrad) | 17.0458 | -8.5244 | 0.90477 | -0.91106 | 3.1454 | 2.5276 | 4.8368 | 7.9992 | 5.2565 | 1.0433 | -5.6411 | -0.85252 | 0.37384 | -1.107 | -0.69917 | -1.6227 | -1.2935 | -1.2252 | -0.5795 | -0.13798
Twist Integral (I0) (m2 rad) | 0.015404
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.1235 | -0.60918 | -0.128 | -0.07008 | 0.437 | 0.021095 | 0.36551 | 0.81392 | 0.6416 | -0.096258 | -0.6796 | 0.039041 | 0.20774 | 0.093781 | -0.11086 | -0.61212 | -0.45282 | -0.44103 | -0.48957 | -1.112
Twist (m2 mrad) | 23.4266 | -8.6807 | -1.7281 | -0.89352 | 5.244 | 0.23732 | 3.8379 | 7.9357 | 5.7744 | -0.79413 | -5.097 | 0.26353 | 1.2465 | 0.49235 | -0.49889 | -2.2954 | -1.3584 | -0.99232 | -0.73436 | -0.41698
Twist Integral (I0) (m2 rad) | 0.018726

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.077828 | -3.8301 | 0.03121 | -0.094352 | 1.1324 | -0.61188 | 1.6058 | 0.66061 | 0.32943 | 17.787 | 0.59589 | -0.30621 | -4.5719 | -0.043187 | -0.15153 | 1.0967 | -0.59435 | 0.37697 | 0.64803 | 0.13623 | 21.2527 | 0.59594 | -0.069312 | -2.7082 | -0.035168 | -0.38246 | 1.1705 | 0 | 1.9975 | 0.58259 | 0.39652 | 14420.576 | -0.3495 | -3.2781 | -0.068199 | -0.42311 | 1.1337 | 0 | 0.74791 | 0.644 | 0.21372 | 14422.08
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.15 | 0.15 | 0.8 | 0.8 | 0.15 | 0.15 | 0.8 | 0.8 | 0.15 | 0.15 | 0.8 | 0.8 | 0.15 | 0.15 | 0.8 | 0.8