## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3090C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-10-25
Beginn der Messung / start of measuring: |  | TT.10.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000090_cc
Ende der Messung / end of measuring: |  | TT.10.JJ
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Mehler |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-10-25 |  | 2003-10-27 | 2003-10-27
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000090_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000090
Date of test Ap 1 | 2003-10-25 |  |  |  | Date of test Ap 2 | 2003-10-25
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | yellow alarm | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | yellow alarm | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | yellow alarm
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
212 | HCMB__A001 | 3000090 |  | 3 |  |  | 3
| COMMENTS: | Yellow alarms in ap.1 on b9 pos. 1 (5 sigma) and in ap. 2 on b3 pos. 14, b7 pos. 2, b9 pos.1, (sigmas < =5) and a2 pos. 20 (7 sigma).
| HOLDING POINT | OK-w | Data at CERN |  | 2003-10-27 | Analysis by |  | V. Remondino | Answer to MMS-LD |  |  | 2003-10-27 | Answer to manufacturer |  |  | 2003-10-28
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000090_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000090 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3090_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3090_cct2 |  |  |  | a even | 1
Test Operator |  | Mehler |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-10-25 |  |  | Time of test |  |  | 6:49 |  |  |  | Date of test |  | 2003-10-25 |  |  | Time of test |  |  | 6:49
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-05-02
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 21 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 21
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00014515 |  |  | Dy average |  |  | 0.00076136 |  |  |  | Dx average |  | 0.00017761 |  |  | Dy average |  |  | 0.00075249
Magnetic length |  | 14.4193 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1404 |  |  |  | Magnetic length |  | 14.4212 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1376
Phase |  | -13.9642 |  |  | Integral field (T.m) |  |  | 85.8347 |  |  |  | Phase |  | -33.5994 |  |  | Integral field (T.m) |  |  | 85.7977
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4644 | 5.9645 | 5.961 | 5.9611 | 5.9595 | 5.9609 | 5.9611 | 5.9614 | 5.9608 | 5.961 | 5.9608 | 5.9617 | 5.9608 | 5.9604 | 5.9603 | 5.9612 | 5.9619 | 5.9617 | 5.9651 | 3.6921 | 85.8347 | C1 (T) | 3.4189 | 5.9636 | 5.9601 | 5.9608 | 5.9602 | 5.9608 | 5.9608 | 5.9611 | 5.9609 | 5.9609 | 5.9613 | 5.9621 | 5.9615 | 5.9608 | 5.9606 | 5.9613 | 5.9611 | 5.9615 | 5.9653 | 3.6886 | 85.7977
Angle (mrad) | 2.1911 | -0.60803 | -0.044153 | 0.011131 | -0.1707 | 0.018936 | 0.52368 | 0.85698 | 0.47191 | -0.51792 | -0.88115 | -0.32586 | 0.10329 | 0.21401 | -0.10065 | -0.1001 | -0.2695 | -0.13817 | -0.33179 | 0.024615 | -13.9642 | Angle (mrad) | 4.5962 | -0.48804 | -0.22112 | -0.19274 | 0.13627 | -0.002202 | 0.2732 | 0.54102 | 0.67164 | -0.75034 | -1.0537 | -0.51025 | 0.04996 | 0.21488 | -0.10336 | -0.087147 | -0.23635 | -0.086977 | -0.054922 | -1.1887 | -33.5994
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.0815 | 0.025759 | -0.31827 | -0.86137 | -0.61862 | 0.098373 | -0.39466 | -1.387 | -1.0768 | -0.96921 | -0.80679 | -0.56764 | -1.0567 | -1.3308 | -0.62016 | -0.4169 | -1.0611 | -1.8016 | -2.0087 | -1.569 | -0.8081 | b2 | -3.7208 | 0.73455 | 0.83909 | 0.35919 | 0.60421 | 1.2554 | 1.5206 | 1.0421 | 1.5208 | 0.51409 | 0.099529 | 0.16057 | 0.60412 | 0.79135 | 0.64947 | 0.40316 | 0.46437 | 0.80446 | 1.2759 | 0.62725 | 0.61991
b3 | 38.0119 | -3.7143 | -3.4196 | -4.8184 | -2.7231 | -4.263 | -2.4628 | -3.613 | -2.3349 | -3.3663 | -3.7473 | -3.674 | -3.331 | -4.7718 | -4.887 | -3.7209 | -2.8029 | -3.3659 | -2.3252 | -2.8349 | -2.2399 | b3 | 37.9504 | -3.3457 | -3.2559 | -4.088 | -3.6777 | -3.9966 | -2.8598 | -4.1731 | -3.1052 | -3.6731 | -3.6096 | -4.0403 | -3.4464 | -6.0211 | -4.212 | -3.0132 | -2.6015 | -3.198 | -2.1334 | -2.3838 | -2.301
b4 | -1.4749 | 0.17301 | 0.023436 | 0.22769 | -0.2317 | -0.024517 | -0.22189 | 0.17577 | 0.12044 | -0.0072668 | 0.14299 | 0.088159 | 0.039088 | 0.093655 | 0.070396 | 0.18245 | 0.17218 | 0.21903 | 0.076359 | 0.0065883 | 0.024294 | b4 | -1.8344 | -0.16993 | 0.021094 | -0.17867 | -0.20967 | -0.079787 | 0.18562 | 0.21547 | 0.28008 | -0.10321 | -0.44426 | -0.24346 | -0.098056 | 0.07067 | -0.02547 | -0.098063 | -0.061283 | 0.02494 | 0.048287 | -0.04401 | -0.10134
b5 | -5.8369 | -0.68176 | -0.90556 | -0.31508 | -0.61689 | -0.36873 | -0.53708 | -0.19889 | -0.66276 | -0.29618 | -0.2618 | -0.45886 | -0.60907 | -0.6189 | -0.41509 | -0.75925 | -0.65653 | -0.57508 | -0.55882 | -4.0582 | -0.80215 | b5 | -5.7145 | -0.69199 | -0.86749 | -0.31993 | -0.61362 | -0.43722 | -0.74775 | -0.39804 | -0.59142 | -0.44676 | -0.19211 | -0.35225 | -0.64524 | -0.3178 | -0.35832 | -0.70695 | -0.71142 | -0.58604 | -0.53281 | -3.8689 | -0.79136
b6 | -0.19586 | -0.028996 | -0.026315 | -0.051665 | 0.079658 | 0.052853 | -0.0094079 | -0.010162 | -0.062451 | -0.06854 | 0.11783 | -0.071962 | 0.052238 | -0.028609 | -0.026847 | 0.11827 | 0.15145 | 0.18966 | 0.081452 | 0.0051018 | 0.018115 | b6 | 0.029014 | 0.033157 | -0.035303 | 0.050129 | 0.040291 | -0.038149 | -0.14889 | -0.088984 | -0.13663 | -0.060622 | 0.041582 | 0.055984 | -0.021601 | -0.11955 | -0.093241 | 0.016435 | 0.026332 | -0.045645 | -0.040229 | -0.055015 | -0.030339
b7 | 2.7987 | 0.89334 | 0.93428 | 1.1231 | 1.0857 | 1.1303 | 1.0287 | 1.0909 | 1.0665 | 1.0902 | 1.0959 | 1.0393 | 1.0715 | 1.0241 | 1.0756 | 1.1098 | 1.0779 | 1.164 | 1.1012 | 0.92682 | 1.1147 | b7 | 2.6727 | 0.87299 | 0.9638 | 1.191 | 1.1893 | 1.2114 | 1.2192 | 1.2313 | 1.192 | 1.2188 | 1.2272 | 1.1841 | 1.152 | 1.1561 | 1.1219 | 1.1239 | 1.1341 | 1.2233 | 1.122 | 0.88489 | 1.1887
b8 | -0.030424 | 0.0408 | -0.017341 | -0.037009 | 0.014579 | -0.0027879 | -0.01089 | -0.0049655 | 0.001592 | 0.014289 | 0.059874 | 0.00084819 | -0.035197 | -0.019531 | -0.010794 | -0.056749 | 0.002757 | 0.029011 | 0.044835 | -0.027842 | -0.0011224 | b8 | -0.005837 | 0.034446 | -0.010091 | 0.025843 | 0.02005 | -0.013068 | -0.0074769 | 0.0066177 | 0.01575 | 0.035182 | 0.072126 | -0.038874 | -0.022192 | 0.012441 | -0.020122 | -0.030921 | 0.0067228 | 0.005717 | 0.02959 | -0.029902 | 0.0052061
b9 | 0.32825 | 0.48552 | 0.46374 | 0.45781 | 0.51098 | 0.47964 | 0.51368 | 0.47218 | 0.52668 | 0.47829 | 0.48922 | 0.47226 | 0.49271 | 0.48928 | 0.46774 | 0.49163 | 0.48403 | 0.50934 | 0.48304 | 0.41568 | 0.47999 | b9 | 0.32546 | 0.50112 | 0.47959 | 0.44987 | 0.47314 | 0.46094 | 0.47743 | 0.46368 | 0.49029 | 0.47599 | 0.47515 | 0.46429 | 0.48905 | 0.46051 | 0.46846 | 0.49915 | 0.48967 | 0.49295 | 0.47517 | 0.4198 | 0.47065
b10 | 0.045409 | 0.017805 | -0.021007 | 0.055973 | 0.018992 | -0.010596 | -0.013725 | 0.024964 | -0.016254 | 0.014249 | 0.083324 | -0.040067 | -0.097436 | -0.054228 | -0.073812 | -0.051895 | 0.047339 | 0.072421 | 0.06039 | -0.069266 | 0 | b10 | 0.12518 | 0.0025258 | -0.061958 | 0.047824 | 0.034441 | -0.030604 | -0.035099 | 0.00023349 | -0.01824 | -0.0041485 | 0.060636 | -0.040084 | -0.046652 | 0.0039595 | -0.03312 | -0.034044 | 0.021083 | 0.035393 | 0.036156 | -0.016367 | 0
b11 | 0.60254 | 0.73462 | 0.74141 | 0.74972 | 0.73982 | 0.74488 | 0.74157 | 0.75218 | 0.7455 | 0.74913 | 0.75537 | 0.75438 | 0.75163 | 0.75942 | 0.75261 | 0.74671 | 0.7417 | 0.74024 | 0.72763 | 0.63304 | 0.73804 | b11 | 0.58193 | 0.74278 | 0.74325 | 0.75005 | 0.74702 | 0.75138 | 0.73825 | 0.7502 | 0.74957 | 0.74489 | 0.75804 | 0.7552 | 0.75335 | 0.76414 | 0.7521 | 0.74538 | 0.74187 | 0.74394 | 0.73712 | 0.64258 | 0.73987
b12 | 0.011138 | 0.0054873 | -0.0037046 | -0.0056262 | -0.0012368 | -0.003483 | -0.0024729 | -0.0018938 | -0.0056832 | -0.0021979 | 0.0072131 | -0.0054431 | -0.0096184 | -0.0059644 | -0.0088088 | -0.007483 | 0.0018 | 0.0053649 | 0.004657 | -0.01037 | -0.0020329 | b12 | 0.010688 | 0.0087521 | -0.0048833 | 0.0096938 | 0.0082714 | -0.0012056 | -0.0052088 | 0.0036801 | 0.00164 | 0.00071973 | 0.0075399 | -0.0036139 | -0.0044894 | -0.0038075 | -0.0034149 | -0.0041399 | 0.0031937 | 0.0061622 | 0.0046938 | -0.0027194 | 0.0014608
b13 | 0.080338 | 0.064423 | 0.067983 | 0.064101 | 0.062173 | 0.058808 | 0.05902 | 0.061376 | 0.060855 | 0.061141 | 0.059838 | 0.063056 | 0.063893 | 0.062096 | 0.061665 | 0.063866 | 0.063276 | 0.062222 | 0.063723 | 0.048863 | 0.062523 | b13 | 0.084632 | 0.063879 | 0.068175 | 0.066298 | 0.066859 | 0.064612 | 0.070124 | 0.066809 | 0.065697 | 0.06325 | 0.061823 | 0.062028 | 0.064592 | 0.065525 | 0.06187 | 0.062915 | 0.064324 | 0.066853 | 0.066448 | 0.048792 | 0.065173
b14 | 0.0010157 | -0.0042862 | -0.0076353 | -0.0038791 | -0.0027856 | -0.0046299 | -0.0069238 | -0.0039221 | -0.0066263 | -0.0054032 | 0.00032907 | -0.0069918 | -0.009445 | -0.0082911 | -0.010137 | -0.0075665 | -0.0023812 | -0.0023089 | -0.0017915 | -0.0054112 | -0.0050745 | b14 | 0.0032671 | -0.0030972 | -0.0079384 | 7.043e-05 | -0.00055006 | -0.004354 | -0.0080905 | -0.0021195 | -0.0046753 | -0.0047125 | -0.00055738 | -0.0065183 | -0.0060689 | -0.0054407 | -0.0072081 | -0.0068118 | -0.0031147 | -0.0022563 | -0.0013844 | -0.0031184 | -0.0039016
b15 | -0.0093393 | 0.035127 | 0.03659 | 0.031816 | 0.031871 | 0.031248 | 0.033004 | 0.033784 | 0.033109 | 0.035375 | 0.034831 | 0.034088 | 0.027657 | 0.032421 | 0.032458 | 0.032622 | 0.036093 | 0.03701 | 0.031093 | 0.026048 | 0.031817 | b15 | 0.0033875 | 0.032029 | 0.033112 | 0.028117 | 0.028899 | 0.028541 | 0.027226 | 0.027572 | 0.030284 | 0.032119 | 0.036496 | 0.031071 | 0.028982 | 0.031545 | 0.033598 | 0.02933 | 0.030024 | 0.033164 | 0.02968 | 0.030752 | 0.029843
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 21.9113 | -6.0803 | -0.44153 | 0.11131 | -1.707 | 0.18936 | 5.2368 | 8.5698 | 4.7191 | -5.1792 | -8.8115 | -3.2586 | 1.0329 | 2.1401 | -1.0065 | -1.001 | -2.695 | -1.3817 | -3.3179 | 0.24615 | 0 | a1 | 45.9624 | -4.8804 | -2.2112 | -1.9274 | 1.3627 | -0.022024 | 2.732 | 5.4102 | 6.7164 | -7.5034 | -10.5367 | -5.1025 | 0.4996 | 2.1488 | -1.0336 | -0.87147 | -2.3635 | -0.86977 | -0.54922 | -11.8871 | 0
a2 | 12.591 | 2.0907 | 0.81622 | -0.068417 | 1.0527 | 1.3094 | 1.3968 | 0.78767 | 0.87503 | -0.41691 | -0.17751 | 0.91459 | 0.99979 | 1.2728 | 0.24549 | 1.2773 | 0.71554 | 0.45385 | 0.027663 | 1.844 | 1.1475 | a2 | 6.8532 | -1.9284 | -2.4735 | -1.8425 | -2.2208 | -3.0028 | -2.8308 | -3.4564 | -2.3204 | -3.4013 | -2.6556 | -1.8578 | -1.4532 | -2.2936 | -2.6003 | -2.4758 | -2.0058 | -2.1813 | -1.7596 | 19.7868 | -1.3853
a3 | -0.68961 | 0.85443 | 1.0936 | 1.203 | 1.0681 | 0.2767 | -0.032745 | 0.047781 | 0.10941 | 0.40326 | 0.15661 | 0.063531 | 0.3807 | 0.68656 | 0.7208 | 0.37754 | 0.41989 | 0.13577 | 0.40489 | 0.2208 | 0.42217 | a3 | 3.5986 | 1.1685 | 1.2409 | 1.0827 | 0.85891 | 0.22441 | -0.19745 | 0.72186 | 0.90186 | 0.48146 | 0.25369 | 0.16414 | 0.23113 | 0.74464 | 0.6895 | 0.7448 | 0.7166 | 1.2276 | 1.0202 | -0.16777 | 0.74175
a4 | 2.5888 | 0.022954 | 0.093556 | 0.53825 | -0.15518 | -0.050263 | -0.17443 | 0.0057102 | -0.0031471 | 0.32103 | 0.0026672 | -0.19021 | -0.21405 | -0.43539 | 0.066344 | -0.31804 | 0.025073 | -0.16414 | 0.014064 | 0.27516 | 0.055193 | a4 | 1.5286 | -0.30772 | -0.15725 | -0.50028 | -0.49166 | -0.50641 | -0.26395 | -0.45699 | -0.60567 | 0.17362 | 0.18796 | -0.19172 | -0.017307 | -0.18632 | -0.0011554 | 0.069758 | 0.1151 | 0.04612 | 0.038293 | 1.9633 | -0.050233
a5 | 2.9798 | 0.029532 | 0.094292 | 0.012568 | 0.22995 | -0.10356 | -0.16814 | -0.038996 | 0.15596 | 0.17621 | 0.11555 | -0.09504 | -0.11213 | 0.13495 | -0.004854 | -0.047916 | 0.057827 | 0.077971 | 0.4056 | -0.063265 | 0.13606 | a5 | 3.002 | 0.10541 | -0.018712 | 0.13785 | 0.26296 | 0.0071786 | 0.0417 | 0.15908 | 0.022408 | 0.12642 | -0.035665 | 0.011832 | -0.2778 | 0.14093 | 0.17303 | 0.1131 | 0.24143 | 0.077953 | 0.2532 | -0.19618 | 0.16374
a6 | 0.53022 | 0.1042 | 0.098212 | 0.22235 | 0.11084 | 0.23509 | 0.18819 | 0.10326 | 0.045405 | 0.027653 | 0.14375 | 0.15401 | 0.12873 | 0.11314 | 0.17216 | 0.16768 | 0.19109 | 0.18625 | 0.18833 | -0.092471 | 0.14746 | a6 | -0.049979 | 0.14712 | 0.037453 | -0.03284 | -0.095348 | -0.17957 | -0.23357 | -0.111 | -0.10896 | 0.034803 | 0.14378 | 0.13727 | -0.092003 | -0.097809 | -0.01624 | -0.02077 | -0.059034 | 0.024508 | 0.12783 | 0.10306 | -0.018707
a7 | 1.6659 | -0.085627 | 0.01731 | -0.0036148 | 0.085283 | -0.023196 | -0.013782 | 0.078847 | 0.0047617 | 0.022332 | 0.024614 | -0.0069458 | 0.008031 | 0.028789 | 0.03566 | 0.049079 | 0.02165 | -0.033842 | -0.0092452 | 0.019074 | 0.061452 | a7 | 1.8814 | -0.0030491 | 0.0057168 | 0.02064 | 0.099406 | -0.017247 | -0.086038 | 0.087605 | 0.12345 | 0.022313 | 0.0075667 | -0.017955 | -0.0010082 | 0.14327 | 0.11594 | 0.11061 | 0.052562 | 0.099066 | 0.065419 | 0.044978 | 0.10082
a8 | -0.017062 | -0.02037 | -0.027331 | 0.00031201 | 0.010569 | 0.048035 | 0.0033026 | 0.0027363 | 0.017525 | 0.0075095 | -0.0028899 | 0.022718 | 0.049879 | 0.042147 | 0.018621 | 0.047302 | 0.012581 | -0.021393 | 0.017627 | -0.062024 | 0.0094022 | a8 | -0.13054 | -0.0040839 | 0.00094514 | 0.040105 | 2.5396e-05 | -0.047049 | -0.0446 | -0.025584 | -0.024565 | -0.049638 | -0.020469 | -0.008831 | -0.037107 | -0.029362 | -0.036887 | -0.034518 | -0.037775 | -0.024709 | 0.015997 | -0.027752 | -0.023974
a9 | -0.24773 | 0.004919 | -0.0025641 | 0.0014939 | -0.025836 | -0.047156 | -0.024492 | -0.014122 | -0.012616 | -0.035147 | -0.047688 | -0.02335 | -0.021642 | -0.040067 | -0.037724 | -0.034472 | 0.00015218 | -0.013759 | -0.024393 | -0.043823 | -0.029663 | a9 | -0.2272 | 0.054759 | 0.04728 | 0.023284 | -0.0024942 | -0.012924 | -0.0094635 | 0.040894 | 0.037031 | -0.017484 | -0.021599 | 0.017074 | 0.024135 | 9.6e-06 | -0.026222 | 0.013357 | 0.031842 | 0.048018 | 0.016829 | -0.059463 | 0.005068
a10 | 0.077943 | 0.0046519 | -0.023282 | 0.029423 | -0.00076691 | 0.012116 | 0.0099132 | 0.0014899 | -0.011108 | -0.012326 | 0.004415 | 0.0070824 | 0.039179 | -0.015474 | -0.00020454 | -0.014237 | -0.013112 | -0.027558 | 0.042877 | -0.12659 | 0 | a10 | -0.049939 | 0.036088 | -0.008789 | 0.037329 | 0.02559 | -0.014017 | 0.016275 | 0.032049 | 0.001894 | -0.022881 | -0.031214 | -0.011012 | 0.0073144 | -0.0090082 | -0.02198 | 0.0187 | 0.0032755 | -0.0032645 | 0.050488 | -0.12646 | 0
a11 | 0.15343 | -0.065556 | -0.048974 | -0.057788 | -0.049077 | -0.051602 | -0.045979 | -0.031006 | -0.038362 | -0.045878 | -0.044013 | -0.042903 | -0.038097 | -0.047866 | -0.053355 | -0.038909 | -0.035002 | -0.0316 | -0.041712 | -0.032346 | -0.038465 | a11 | 0.19706 | -0.014931 | -0.010393 | -0.0030377 | 0.00075619 | -0.0092303 | 0.00080901 | 0.01373 | 0.013231 | -0.006335 | -0.0056902 | -0.0043379 | -0.00028195 | 0.00084993 | 2.9523e-05 | 0.002151 | 0.0074727 | 0.014263 | 0.0042125 | 0.0029067 | 0.0061522
a12 | 0.012046 | 0.002199 | -0.0013327 | 0.0070911 | 0.00068581 | 0.0069813 | 0.0063375 | -0.00081885 | -0.0038963 | -0.0017299 | 0.0035761 | 0.0081231 | 0.0082357 | 0.0015567 | 0.0053925 | 0.0037177 | 0.00090195 | 0.00066649 | 0.0072999 | -0.011716 | 0.0028507 | a12 | -0.0045816 | 0.0072438 | -0.003341 | -4.61e-05 | -0.0034709 | -0.0095009 | -0.0044274 | -0.003645 | -0.0046928 | -0.0071037 | -0.0039376 | -0.0041072 | -0.0038413 | -0.0073944 | -0.0047277 | -0.0013505 | -0.0029692 | -0.0037744 | 0.0069579 | -0.0078521 | -0.0032097
a13 | -0.023641 | -0.0015152 | 0.00017382 | -0.0026265 | -0.0023286 | -0.005036 | -0.0047474 | -0.0035845 | -0.0031357 | -0.0043423 | -0.0063118 | -0.0038634 | -0.0021467 | -0.0029683 | -0.0036875 | -0.0042346 | -0.0029797 | -0.0033188 | -0.003029 | -0.0038225 | -0.0039471 | a13 | -0.020707 | 0.0069228 | 0.0058918 | 0.0017191 | 0.0012747 | -0.00018516 | -0.001116 | 0.0048694 | 0.0048321 | 0.0013438 | -0.00051225 | 0.0011596 | 0.00065129 | 0.0026692 | 0.00016357 | 0.0036839 | 0.002483 | 0.0032273 | 0.0035929 | -0.00097798 | 0.0015732
a14 | 0.009823 | -0.015899 | -0.017686 | -0.012847 | -0.013496 | -0.013763 | -0.016662 | -0.015753 | -0.014853 | -0.014004 | -0.016628 | -0.015216 | -0.010528 | -0.013954 | -0.014877 | -0.014577 | -0.016488 | -0.017183 | -0.012222 | -0.0051273 | -0.013755 | a14 | 0.015498 | -0.014424 | -0.013904 | -0.0098939 | -0.010742 | -0.012716 | -0.010423 | -0.012562 | -0.014617 | -0.014862 | -0.017291 | -0.013759 | -0.012062 | -0.013231 | -0.014585 | -0.012752 | -0.01244 | -0.015442 | -0.012042 | -0.0068122 | -0.012144
a15 | 0.0065293 | -0.0081537 | -0.011871 | -0.0053056 | -0.0065295 | -0.010661 | -0.011626 | -0.0067196 | -0.010642 | -0.0091352 | -0.0018375 | -0.013747 | -0.016407 | -0.014134 | -0.016065 | -0.016355 | -0.0062826 | -0.0034333 | -0.0020837 | -0.014874 | -0.0091872 | a15 | 0.0046008 | -0.0055195 | -0.012079 | -0.00076473 | -0.0020327 | -0.009603 | -0.013055 | -0.0022196 | -0.0061768 | -0.0065837 | -0.0012577 | -0.011072 | -0.011608 | -0.007619 | -0.01097 | -0.0094112 | -0.0054635 | -0.0047479 | -0.0025421 | -0.012478 | -0.0066592
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00017265 | -3.9925e-05 | 4.4444e-05 | -0.00012099 | -4.3565e-05 | 2.597e-05 | 3.2744e-05 | -5.6183e-05 | 3.5665e-05 | -3.3919e-05 | -0.00018637 | 9.0914e-05 | 0.00022432 | 0.00011877 | 0.00016595 | 0.00011619 | -0.00010968 | -0.00016876 | -0.00013484 | 0.00016734 | 0.00014515 | Dx (m) | -0.0002838 | 0 | 0.00014142 | -0.00010808 | -7.8463e-05 | 6.8824e-05 | 8.0805e-05 | 0 | 4.1277e-05 | 0 | -0.00013651 | 9.0086e-05 | 0.0001053 | 0 | 7.4855e-05 | 7.7531e-05 | -4.8386e-05 | -8.0716e-05 | -8.3995e-05 | 4.4635e-05 | 0.00017761
Dy (m) | -0.00017616 | -1.4332e-05 | 5.6326e-05 | -7.6007e-05 | 0 | -2.5854e-05 | -2.0698e-05 | 0 | 2.7171e-05 | 2.5891e-05 | -2.0863e-05 | -1.0821e-05 | -7.7471e-05 | 4.2132e-05 | 1.2174e-05 | 3.8468e-05 | 2.4871e-05 | 5.6091e-05 | -0.00010796 | 0.00034861 | 0.00076136 | Dy (m) | 0.00024245 | -8.266e-05 | 2.2111e-05 | -8.5106e-05 | -5.818e-05 | 3.2573e-05 | -3.7566e-05 | -7.2603e-05 | 0 | 5.2289e-05 | 6.9039e-05 | 2.5322e-05 | -1.6452e-05 | 2.0049e-05 | 4.9692e-05 | -4.2879e-05 | 0 | 0 | -0.00011597 | 0.00033418 | 0.00075249
Magnetic length (m) |  | 14.4193 | [19 empty] | Magnetic length (m) |  | 14.4212
Transfer funct. (mT/KA) |  | 596.1404 | [19 empty] | Transfer funct. (mT/KA) |  | 596.1376
Coil positioning (mm) |  | 0.019114 | [19 empty] | Coil positioning (mm) |  | 0.020573
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000090 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000090 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 346.4437 | 596.4453 | 596.1023 | 596.114 | 595.9504 | 596.0863 | 596.1114 | 596.1423 | 596.0793 | 596.1049 | 596.0757 | 596.1718 | 596.084 | 596.0391 | 596.0321 | 596.1172 | 596.1897 | 596.1711 | 596.5095 | 369.208 | 85.8347 | C1/i (mT/KA) | 341.8893 | 596.3576 | 596.0131 | 596.0849 | 596.0214 | 596.0772 | 596.0759 | 596.1115 | 596.0946 | 596.0946 | 596.1268 | 596.2053 | 596.1484 | 596.0771 | 596.0625 | 596.1289 | 596.1145 | 596.1529 | 596.5287 | 368.8552 | 85.7977
Angle (mrad) | 2.1911 | -0.60803 | -0.044153 | 0.011131 | -0.1707 | 0.018936 | 0.52368 | 0.85698 | 0.47191 | -0.51792 | -0.88115 | -0.32586 | 0.10329 | 0.21401 | -0.10065 | -0.1001 | -0.2695 | -0.13817 | -0.33179 | 0.024615 | -13.9642 | Angle (mrad) | 4.5962 | -0.48804 | -0.22112 | -0.19274 | 0.13627 | -0.002202 | 0.2732 | 0.54102 | 0.67164 | -0.75034 | -1.0537 | -0.51025 | 0.04996 | 0.21488 | -0.10336 | -0.087147 | -0.23635 | -0.086977 | -0.054922 | -1.1887 | -33.5994
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 1.0815 | 0.025759 | -0.31827 | -0.86137 | -0.61862 | 0.098373 | -0.39466 | -1.387 | -1.0768 | -0.96921 | -0.80679 | -0.56764 | -1.0567 | -1.3308 | -0.62016 | -0.4169 | -1.0611 | -1.8016 | -2.0087 | -1.569 | -0.8081 | b2 | -3.7208 | 0.73455 | 0.83909 | 0.35919 | 0.60421 | 1.2554 | 1.5206 | 1.0421 | 1.5208 | 0.51409 | 0.099529 | 0.16057 | 0.60412 | 0.79135 | 0.64947 | 0.40316 | 0.46437 | 0.80446 | 1.2759 | 0.62725 | 0.61991
b3 | 38.0119 | -3.7143 | -3.4196 | -4.8184 | -2.7231 | -4.263 | -2.4628 | -3.613 | -2.3349 | -3.3663 | -3.7473 | -3.674 | -3.331 | -4.7718 | -4.887 | -3.7209 | -2.8029 | -3.3659 | -2.3252 | -2.8349 | -2.2399 | b3 | 37.9504 | -3.3457 | -3.2559 | -4.088 | -3.6777 | -3.9966 | -2.8598 | -4.1731 | -3.1052 | -3.6731 | -3.6096 | -4.0403 | -3.4464 | -6.0211 | -4.212 | -3.0132 | -2.6015 | -3.198 | -2.1334 | -2.3838 | -2.301
b4 | -1.4749 | 0.17301 | 0.023436 | 0.22769 | -0.2317 | -0.024517 | -0.22189 | 0.17577 | 0.12044 | -0.0072668 | 0.14299 | 0.088159 | 0.039088 | 0.093655 | 0.070396 | 0.18245 | 0.17218 | 0.21903 | 0.076359 | 0.0065883 | 0.024294 | b4 | -1.8344 | -0.16993 | 0.021094 | -0.17867 | -0.20967 | -0.079787 | 0.18562 | 0.21547 | 0.28008 | -0.10321 | -0.44426 | -0.24346 | -0.098056 | 0.07067 | -0.02547 | -0.098063 | -0.061283 | 0.02494 | 0.048287 | -0.04401 | -0.10134
b5 | -5.8369 | -0.68176 | -0.90556 | -0.31508 | -0.61689 | -0.36873 | -0.53708 | -0.19889 | -0.66276 | -0.29618 | -0.2618 | -0.45886 | -0.60907 | -0.6189 | -0.41509 | -0.75925 | -0.65653 | -0.57508 | -0.55882 | -4.0582 | -0.80215 | b5 | -5.7145 | -0.69199 | -0.86749 | -0.31993 | -0.61362 | -0.43722 | -0.74775 | -0.39804 | -0.59142 | -0.44676 | -0.19211 | -0.35225 | -0.64524 | -0.3178 | -0.35832 | -0.70695 | -0.71142 | -0.58604 | -0.53281 | -3.8689 | -0.79136
b6 | -0.19586 | -0.028996 | -0.026315 | -0.051665 | 0.079658 | 0.052853 | -0.0094079 | -0.010162 | -0.062451 | -0.06854 | 0.11783 | -0.071962 | 0.052238 | -0.028609 | -0.026847 | 0.11827 | 0.15145 | 0.18966 | 0.081452 | 0.0051018 | 0.018115 | b6 | 0.029014 | 0.033157 | -0.035303 | 0.050129 | 0.040291 | -0.038149 | -0.14889 | -0.088984 | -0.13663 | -0.060622 | 0.041582 | 0.055984 | -0.021601 | -0.11955 | -0.093241 | 0.016435 | 0.026332 | -0.045645 | -0.040229 | -0.055015 | -0.030339
b7 | 2.7987 | 0.89334 | 0.93428 | 1.1231 | 1.0857 | 1.1303 | 1.0287 | 1.0909 | 1.0665 | 1.0902 | 1.0959 | 1.0393 | 1.0715 | 1.0241 | 1.0756 | 1.1098 | 1.0779 | 1.164 | 1.1012 | 0.92682 | 1.1147 | b7 | 2.6727 | 0.87299 | 0.9638 | 1.191 | 1.1893 | 1.2114 | 1.2192 | 1.2313 | 1.192 | 1.2188 | 1.2272 | 1.1841 | 1.152 | 1.1561 | 1.1219 | 1.1239 | 1.1341 | 1.2233 | 1.122 | 0.88489 | 1.1887
b8 | -0.030424 | 0.0408 | -0.017341 | -0.037009 | 0.014579 | -0.0027879 | -0.01089 | -0.0049655 | 0.001592 | 0.014289 | 0.059874 | 0.00084819 | -0.035197 | -0.019531 | -0.010794 | -0.056749 | 0.002757 | 0.029011 | 0.044835 | -0.027842 | -0.0011224 | b8 | -0.005837 | 0.034446 | -0.010091 | 0.025843 | 0.02005 | -0.013068 | -0.0074769 | 0.0066177 | 0.01575 | 0.035182 | 0.072126 | -0.038874 | -0.022192 | 0.012441 | -0.020122 | -0.030921 | 0.0067228 | 0.005717 | 0.02959 | -0.029902 | 0.0052061
b9 | 0.32825 | 0.48552 | 0.46374 | 0.45781 | 0.51098 | 0.47964 | 0.51368 | 0.47218 | 0.52668 | 0.47829 | 0.48922 | 0.47226 | 0.49271 | 0.48928 | 0.46774 | 0.49163 | 0.48403 | 0.50934 | 0.48304 | 0.41568 | 0.47999 | b9 | 0.32546 | 0.50112 | 0.47959 | 0.44987 | 0.47314 | 0.46094 | 0.47743 | 0.46368 | 0.49029 | 0.47599 | 0.47515 | 0.46429 | 0.48905 | 0.46051 | 0.46846 | 0.49915 | 0.48967 | 0.49295 | 0.47517 | 0.4198 | 0.47065
b10 | 0.045409 | 0.017805 | -0.021007 | 0.055973 | 0.018992 | -0.010596 | -0.013725 | 0.024964 | -0.016254 | 0.014249 | 0.083324 | -0.040067 | -0.097436 | -0.054228 | -0.073812 | -0.051895 | 0.047339 | 0.072421 | 0.06039 | -0.069266 | 0 | b10 | 0.12518 | 0.0025258 | -0.061958 | 0.047824 | 0.034441 | -0.030604 | -0.035099 | 0.00023349 | -0.01824 | -0.0041485 | 0.060636 | -0.040084 | -0.046652 | 0.0039595 | -0.03312 | -0.034044 | 0.021083 | 0.035393 | 0.036156 | -0.016367 | 0
b11 | 0.60254 | 0.73462 | 0.74141 | 0.74972 | 0.73982 | 0.74488 | 0.74157 | 0.75218 | 0.7455 | 0.74913 | 0.75537 | 0.75438 | 0.75163 | 0.75942 | 0.75261 | 0.74671 | 0.7417 | 0.74024 | 0.72763 | 0.63304 | 0.73804 | b11 | 0.58193 | 0.74278 | 0.74325 | 0.75005 | 0.74702 | 0.75138 | 0.73825 | 0.7502 | 0.74957 | 0.74489 | 0.75804 | 0.7552 | 0.75335 | 0.76414 | 0.7521 | 0.74538 | 0.74187 | 0.74394 | 0.73712 | 0.64258 | 0.73987
b12 | 0.011138 | 0.0054873 | -0.0037046 | -0.0056262 | -0.0012368 | -0.003483 | -0.0024729 | -0.0018938 | -0.0056832 | -0.0021979 | 0.0072131 | -0.0054431 | -0.0096184 | -0.0059644 | -0.0088088 | -0.007483 | 0.0018 | 0.0053649 | 0.004657 | -0.01037 | -0.0020329 | b12 | 0.010688 | 0.0087521 | -0.0048833 | 0.0096938 | 0.0082714 | -0.0012056 | -0.0052088 | 0.0036801 | 0.00164 | 0.00071973 | 0.0075399 | -0.0036139 | -0.0044894 | -0.0038075 | -0.0034149 | -0.0041399 | 0.0031937 | 0.0061622 | 0.0046938 | -0.0027194 | 0.0014608
b13 | 0.080338 | 0.064423 | 0.067983 | 0.064101 | 0.062173 | 0.058808 | 0.05902 | 0.061376 | 0.060855 | 0.061141 | 0.059838 | 0.063056 | 0.063893 | 0.062096 | 0.061665 | 0.063866 | 0.063276 | 0.062222 | 0.063723 | 0.048863 | 0.062523 | b13 | 0.084632 | 0.063879 | 0.068175 | 0.066298 | 0.066859 | 0.064612 | 0.070124 | 0.066809 | 0.065697 | 0.06325 | 0.061823 | 0.062028 | 0.064592 | 0.065525 | 0.06187 | 0.062915 | 0.064324 | 0.066853 | 0.066448 | 0.048792 | 0.065173
b14 | 0.0010157 | -0.0042862 | -0.0076353 | -0.0038791 | -0.0027856 | -0.0046299 | -0.0069238 | -0.0039221 | -0.0066263 | -0.0054032 | 0.00032907 | -0.0069918 | -0.009445 | -0.0082911 | -0.010137 | -0.0075665 | -0.0023812 | -0.0023089 | -0.0017915 | -0.0054112 | -0.0050745 | b14 | 0.0032671 | -0.0030972 | -0.0079384 | 7.043e-05 | -0.00055006 | -0.004354 | -0.0080905 | -0.0021195 | -0.0046753 | -0.0047125 | -0.00055738 | -0.0065183 | -0.0060689 | -0.0054407 | -0.0072081 | -0.0068118 | -0.0031147 | -0.0022563 | -0.0013844 | -0.0031184 | -0.0039016
b15 | -0.0093393 | 0.035127 | 0.03659 | 0.031816 | 0.031871 | 0.031248 | 0.033004 | 0.033784 | 0.033109 | 0.035375 | 0.034831 | 0.034088 | 0.027657 | 0.032421 | 0.032458 | 0.032622 | 0.036093 | 0.03701 | 0.031093 | 0.026048 | 0.031817 | b15 | 0.0033875 | 0.032029 | 0.033112 | 0.028117 | 0.028899 | 0.028541 | 0.027226 | 0.027572 | 0.030284 | 0.032119 | 0.036496 | 0.031071 | 0.028982 | 0.031545 | 0.033598 | 0.02933 | 0.030024 | 0.033164 | 0.02968 | 0.030752 | 0.029843
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 21.9113 | -6.0803 | -0.44153 | 0.11131 | -1.707 | 0.18936 | 5.2368 | 8.5698 | 4.7191 | -5.1792 | -8.8115 | -3.2586 | 1.0329 | 2.1401 | -1.0065 | -1.001 | -2.695 | -1.3817 | -3.3179 | 0.24615 | 0 | a1 | 45.9624 | -4.8804 | -2.2112 | -1.9274 | 1.3627 | -0.022024 | 2.732 | 5.4102 | 6.7164 | -7.5034 | -10.5367 | -5.1025 | 0.4996 | 2.1488 | -1.0336 | -0.87147 | -2.3635 | -0.86977 | -0.54922 | -11.8871 | 0
a2 | 12.591 | 2.0907 | 0.81622 | -0.068417 | 1.0527 | 1.3094 | 1.3968 | 0.78767 | 0.87503 | -0.41691 | -0.17751 | 0.91459 | 0.99979 | 1.2728 | 0.24549 | 1.2773 | 0.71554 | 0.45385 | 0.027663 | 1.844 | 1.1475 | a2 | 6.8532 | -1.9284 | -2.4735 | -1.8425 | -2.2208 | -3.0028 | -2.8308 | -3.4564 | -2.3204 | -3.4013 | -2.6556 | -1.8578 | -1.4532 | -2.2936 | -2.6003 | -2.4758 | -2.0058 | -2.1813 | -1.7596 | 19.7868 | -1.3853
a3 | -0.68961 | 0.85443 | 1.0936 | 1.203 | 1.0681 | 0.2767 | -0.032745 | 0.047781 | 0.10941 | 0.40326 | 0.15661 | 0.063531 | 0.3807 | 0.68656 | 0.7208 | 0.37754 | 0.41989 | 0.13577 | 0.40489 | 0.2208 | 0.42217 | a3 | 3.5986 | 1.1685 | 1.2409 | 1.0827 | 0.85891 | 0.22441 | -0.19745 | 0.72186 | 0.90186 | 0.48146 | 0.25369 | 0.16414 | 0.23113 | 0.74464 | 0.6895 | 0.7448 | 0.7166 | 1.2276 | 1.0202 | -0.16777 | 0.74175
a4 | 2.5888 | 0.022954 | 0.093556 | 0.53825 | -0.15518 | -0.050263 | -0.17443 | 0.0057102 | -0.0031471 | 0.32103 | 0.0026672 | -0.19021 | -0.21405 | -0.43539 | 0.066344 | -0.31804 | 0.025073 | -0.16414 | 0.014064 | 0.27516 | 0.055193 | a4 | 1.5286 | -0.30772 | -0.15725 | -0.50028 | -0.49166 | -0.50641 | -0.26395 | -0.45699 | -0.60567 | 0.17362 | 0.18796 | -0.19172 | -0.017307 | -0.18632 | -0.0011554 | 0.069758 | 0.1151 | 0.04612 | 0.038293 | 1.9633 | -0.050233
a5 | 2.9798 | 0.029532 | 0.094292 | 0.012568 | 0.22995 | -0.10356 | -0.16814 | -0.038996 | 0.15596 | 0.17621 | 0.11555 | -0.09504 | -0.11213 | 0.13495 | -0.004854 | -0.047916 | 0.057827 | 0.077971 | 0.4056 | -0.063265 | 0.13606 | a5 | 3.002 | 0.10541 | -0.018712 | 0.13785 | 0.26296 | 0.0071786 | 0.0417 | 0.15908 | 0.022408 | 0.12642 | -0.035665 | 0.011832 | -0.2778 | 0.14093 | 0.17303 | 0.1131 | 0.24143 | 0.077953 | 0.2532 | -0.19618 | 0.16374
a6 | 0.53022 | 0.1042 | 0.098212 | 0.22235 | 0.11084 | 0.23509 | 0.18819 | 0.10326 | 0.045405 | 0.027653 | 0.14375 | 0.15401 | 0.12873 | 0.11314 | 0.17216 | 0.16768 | 0.19109 | 0.18625 | 0.18833 | -0.092471 | 0.14746 | a6 | -0.049979 | 0.14712 | 0.037453 | -0.03284 | -0.095348 | -0.17957 | -0.23357 | -0.111 | -0.10896 | 0.034803 | 0.14378 | 0.13727 | -0.092003 | -0.097809 | -0.01624 | -0.02077 | -0.059034 | 0.024508 | 0.12783 | 0.10306 | -0.018707
a7 | 1.6659 | -0.085627 | 0.01731 | -0.0036148 | 0.085283 | -0.023196 | -0.013782 | 0.078847 | 0.0047617 | 0.022332 | 0.024614 | -0.0069458 | 0.008031 | 0.028789 | 0.03566 | 0.049079 | 0.02165 | -0.033842 | -0.0092452 | 0.019074 | 0.061452 | a7 | 1.8814 | -0.0030491 | 0.0057168 | 0.02064 | 0.099406 | -0.017247 | -0.086038 | 0.087605 | 0.12345 | 0.022313 | 0.0075667 | -0.017955 | -0.0010082 | 0.14327 | 0.11594 | 0.11061 | 0.052562 | 0.099066 | 0.065419 | 0.044978 | 0.10082
a8 | -0.017062 | -0.02037 | -0.027331 | 0.00031201 | 0.010569 | 0.048035 | 0.0033026 | 0.0027363 | 0.017525 | 0.0075095 | -0.0028899 | 0.022718 | 0.049879 | 0.042147 | 0.018621 | 0.047302 | 0.012581 | -0.021393 | 0.017627 | -0.062024 | 0.0094022 | a8 | -0.13054 | -0.0040839 | 0.00094514 | 0.040105 | 2.5396e-05 | -0.047049 | -0.0446 | -0.025584 | -0.024565 | -0.049638 | -0.020469 | -0.008831 | -0.037107 | -0.029362 | -0.036887 | -0.034518 | -0.037775 | -0.024709 | 0.015997 | -0.027752 | -0.023974
a9 | -0.24773 | 0.004919 | -0.0025641 | 0.0014939 | -0.025836 | -0.047156 | -0.024492 | -0.014122 | -0.012616 | -0.035147 | -0.047688 | -0.02335 | -0.021642 | -0.040067 | -0.037724 | -0.034472 | 0.00015218 | -0.013759 | -0.024393 | -0.043823 | -0.029663 | a9 | -0.2272 | 0.054759 | 0.04728 | 0.023284 | -0.0024942 | -0.012924 | -0.0094635 | 0.040894 | 0.037031 | -0.017484 | -0.021599 | 0.017074 | 0.024135 | 9.6e-06 | -0.026222 | 0.013357 | 0.031842 | 0.048018 | 0.016829 | -0.059463 | 0.005068
a10 | 0.077943 | 0.0046519 | -0.023282 | 0.029423 | -0.00076691 | 0.012116 | 0.0099132 | 0.0014899 | -0.011108 | -0.012326 | 0.004415 | 0.0070824 | 0.039179 | -0.015474 | -0.00020454 | -0.014237 | -0.013112 | -0.027558 | 0.042877 | -0.12659 | 0 | a10 | -0.049939 | 0.036088 | -0.008789 | 0.037329 | 0.02559 | -0.014017 | 0.016275 | 0.032049 | 0.001894 | -0.022881 | -0.031214 | -0.011012 | 0.0073144 | -0.0090082 | -0.02198 | 0.0187 | 0.0032755 | -0.0032645 | 0.050488 | -0.12646 | 0
a11 | 0.15343 | -0.065556 | -0.048974 | -0.057788 | -0.049077 | -0.051602 | -0.045979 | -0.031006 | -0.038362 | -0.045878 | -0.044013 | -0.042903 | -0.038097 | -0.047866 | -0.053355 | -0.038909 | -0.035002 | -0.0316 | -0.041712 | -0.032346 | -0.038465 | a11 | 0.19706 | -0.014931 | -0.010393 | -0.0030377 | 0.00075619 | -0.0092303 | 0.00080901 | 0.01373 | 0.013231 | -0.006335 | -0.0056902 | -0.0043379 | -0.00028195 | 0.00084993 | 2.9523e-05 | 0.002151 | 0.0074727 | 0.014263 | 0.0042125 | 0.0029067 | 0.0061522
a12 | 0.012046 | 0.002199 | -0.0013327 | 0.0070911 | 0.00068581 | 0.0069813 | 0.0063375 | -0.00081885 | -0.0038963 | -0.0017299 | 0.0035761 | 0.0081231 | 0.0082357 | 0.0015567 | 0.0053925 | 0.0037177 | 0.00090195 | 0.00066649 | 0.0072999 | -0.011716 | 0.0028507 | a12 | -0.0045816 | 0.0072438 | -0.003341 | -4.61e-05 | -0.0034709 | -0.0095009 | -0.0044274 | -0.003645 | -0.0046928 | -0.0071037 | -0.0039376 | -0.0041072 | -0.0038413 | -0.0073944 | -0.0047277 | -0.0013505 | -0.0029692 | -0.0037744 | 0.0069579 | -0.0078521 | -0.0032097
a13 | -0.023641 | -0.0015152 | 0.00017382 | -0.0026265 | -0.0023286 | -0.005036 | -0.0047474 | -0.0035845 | -0.0031357 | -0.0043423 | -0.0063118 | -0.0038634 | -0.0021467 | -0.0029683 | -0.0036875 | -0.0042346 | -0.0029797 | -0.0033188 | -0.003029 | -0.0038225 | -0.0039471 | a13 | -0.020707 | 0.0069228 | 0.0058918 | 0.0017191 | 0.0012747 | -0.00018516 | -0.001116 | 0.0048694 | 0.0048321 | 0.0013438 | -0.00051225 | 0.0011596 | 0.00065129 | 0.0026692 | 0.00016357 | 0.0036839 | 0.002483 | 0.0032273 | 0.0035929 | -0.00097798 | 0.0015732
a14 | 0.009823 | -0.015899 | -0.017686 | -0.012847 | -0.013496 | -0.013763 | -0.016662 | -0.015753 | -0.014853 | -0.014004 | -0.016628 | -0.015216 | -0.010528 | -0.013954 | -0.014877 | -0.014577 | -0.016488 | -0.017183 | -0.012222 | -0.0051273 | -0.013755 | a14 | 0.015498 | -0.014424 | -0.013904 | -0.0098939 | -0.010742 | -0.012716 | -0.010423 | -0.012562 | -0.014617 | -0.014862 | -0.017291 | -0.013759 | -0.012062 | -0.013231 | -0.014585 | -0.012752 | -0.01244 | -0.015442 | -0.012042 | -0.0068122 | -0.012144
a15 | 0.0065293 | -0.0081537 | -0.011871 | -0.0053056 | -0.0065295 | -0.010661 | -0.011626 | -0.0067196 | -0.010642 | -0.0091352 | -0.0018375 | -0.013747 | -0.016407 | -0.014134 | -0.016065 | -0.016355 | -0.0062826 | -0.0034333 | -0.0020837 | -0.014874 | -0.0091872 | a15 | 0.0046008 | -0.0055195 | -0.012079 | -0.00076473 | -0.0020327 | -0.009603 | -0.013055 | -0.0022196 | -0.0061768 | -0.0065837 | -0.0012577 | -0.011072 | -0.011608 | -0.007619 | -0.01097 | -0.0094112 | -0.0054635 | -0.0047479 | -0.0025421 | -0.012478 | -0.0066592
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.17265 | 0.039925 | -0.044444 | 0.12099 | 0.043565 | -0.02597 | -0.032744 | 0.056183 | -0.035665 | 0.033919 | 0.18637 | -0.090914 | -0.22432 | -0.11877 | -0.16595 | -0.11619 | 0.10968 | 0.16876 | 0.13484 | -0.16734 |  | Dx (mm) | 0.2838 | 0 | -0.14142 | 0.10808 | 0.078463 | -0.068824 | -0.080805 | 0 | -0.041278 | 0 | 0.13651 | -0.090086 | -0.1053 | 0 | -0.074855 | -0.077531 | 0.048386 | 0.080716 | 0.083995 | -0.044635
Dy (mm) | 0.17616 | 0.014332 | -0.056326 | 0.076007 | 0 | 0.025854 | 0.020698 | 0 | -0.027171 | -0.025891 | 0.020863 | 0.010821 | 0.077471 | -0.042132 | -0.012174 | -0.038468 | -0.024871 | -0.056091 | 0.10796 | -0.34861 |  | Dy (mm) | -0.24245 | 0.08266 | -0.022111 | 0.085106 | 0.05818 | -0.032573 | 0.037566 | 0.072603 | 0 | -0.052289 | -0.069039 | -0.025322 | 0.016452 | -0.020049 | -0.049692 | 0.042879 | 0 | 0 | 0.11597 | -0.33418
Mag. len. (m) | 14.4193 | [20 empty] | Mag. len. (m) | 14.4212
TF (mT/KA) | 596.1404 | [20 empty] | TF (mT/KA) | 596.1376
Coil pos. (mm) | 0.019114 | [20 empty] | Coil pos. (mm) | 0.020573
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0073554 | [20 empty] | Twist (m2 rad) | 0.014654

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -0.8081 | 0.7112 | -0.84434 | 0.55388 | 1.1475 | 2.7143 | 0.75535 | 0.64215 | 0.61991 | 1.0776 | 0.75613 | 0.41685 | -1.3853 | 5.2926 | -2.3721 | 0.55562 |  | 0
3 | -2.2399 | 9.311 | -3.5192 | 0.80417 | 0.42217 | 0.4494 | 0.46312 | 0.38516 | -2.301 | 9.3098 | -3.5811 | 0.83395 | 0.74175 | 0.79158 | 0.68111 | 0.41486 |  | 0
4 | 0.024294 | 0.36747 | 0.072177 | 0.1355 | 0.055193 | 0.62361 | -0.036114 | 0.22377 | -0.10134 | 0.43287 | -0.049671 | 0.17406 | -0.050233 | 0.64294 | -0.16842 | 0.25919 |  | 0
5 | -0.80215 | 1.402 | -0.52821 | 0.18863 | 0.13606 | 0.67026 | 0.050715 | 0.1386 | -0.79136 | 1.3571 | -0.52946 | 0.18392 | 0.16374 | 0.66956 | 0.084868 | 0.12929 |  | 0
6 | 0.018115 | 0.091707 | 0.023819 | 0.075807 | 0.14746 | 0.11671 | 0.14192 | 0.052922 | -0.030339 | 0.064789 | -0.028593 | 0.063162 | -0.018707 | 0.11036 | -0.02722 | 0.11652 |  | 0
7 | 1.1147 | 0.39513 | 1.066 | 0.06577 | 0.061452 | 0.37185 | 0.01169 | 0.039895 | 1.1887 | 0.35958 | 1.1514 | 0.094247 | 0.10082 | 0.41469 | 0.045542 | 0.06287 |  | 0
8 | -0.0011224 | 0.029831 | 0.00017491 | 0.022866 | 0.0094022 | 0.028358 | 0.011692 | 0.02256 | 0.0052061 | 0.02748 | 0.0090118 | 0.01758 | -0.023974 | 0.033315 | -0.023386 | 0.019984 |  | 0
9 | 0.47999 | 0.041967 | 0.48451 | 0.018718 | -0.029663 | 0.052762 | -0.02226 | 0.016893 | 0.47065 | 0.038102 | 0.47619 | 0.014455 | 0.005068 | 0.06084 | 0.014241 | 0.025549 |  | 0
10 | 0 | 0.052694 | 1.2866e-05 | 2.8921e-05 | 0 | 0.038972 | 0.00015759 | 0.00060048 | 0 | 0.044956 | 0.00014599 | 0.0014301 | 0 | 0.038609 | 7.6262e-05 | 0.0013223 |  | 0
11 | 0.73804 | 0.040486 | 0.7458 | 0.0078004 | -0.038465 | 0.045085 | -0.044818 | 0.0090375 | 0.73987 | 0.043496 | 0.7481 | 0.0068992 | 0.0061522 | 0.044686 | 1.7904e-05 | 0.0083546 |  | 0
12 | -0.0020329 | 0.0060314 | -0.0022634 | 0.0028865 | 0.0028507 | 0.0053359 | 0.0029088 | 0.0026176 | 0.0014608 | 0.0055421 | 0.0017362 | 0.0028296 | -0.0032097 | 0.0041669 | -0.0036391 | 0.0029319 |  | 0
13 | 0.062523 | 0.0055537 | 0.062192 | 0.0023354 | -0.0039471 | 0.0047387 | -0.0032426 | 0.0017448 | 0.065173 | 0.0062231 | 0.064858 | 0.002068 | 0.0015732 | 0.0056069 | 0.002299 | 0.0023993 |  | 0
14 | -0.0050745 | 0.0030613 | -0.0054131 | 0.00097723 | -0.013755 | 0.0060728 | -0.015273 | 0.0012473 | -0.0039016 | 0.0030065 | -0.0039908 | 0.0011461 | -0.012144 | 0.0067388 | -0.013764 | 0.0013307 |  | 0
15 | 0.031817 | 0.0098426 | 0.033344 | 0.0023064 | -0.0091872 | 0.0059258 | -0.0094993 | 0.0047716 | 0.029843 | 0.0065213 | 0.030655 | 0.0024424 | -0.0066592 | 0.0048063 | -0.0068181 | 0.0040449 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.51687 | -0.73134
3 | [12 empty] | 3 | 9 | -0.58603 | -0.5307
4 | [12 empty] | 4 | 16 | -1.748 | -1.5493
5 | [12 empty] | 5 | 25 | -1.8221 | -1.8695
6 | [12 empty] | 6 | 36 | -2.7592 | -2.4559
7 | [12 empty] | 7 | 49 | -2.9716 | -2.5643
8 | [12 empty] | 8 | 64 | -3.7848 | -3.9769
9 | [12 empty] | 9 | 81 | -4.0296 | -3.9519
10 | [12 empty] | sigma (mm) |  | 0.019114 | 0.020573
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | 0.29511 | 0.044646 | -0.34339 | -0.78203 | -0.60466 | 0.08619 | -0.40423 | -1.3631 | -1.087 | -0.957 | -0.72425 | -0.60686 | -1.1411 | -1.4008 | -0.7166 | -0.46947 | -1.0262 | -1.7356 | -1.9667 | -1.6339 | -0.8081
3 | 38.1373 | -3.7154 | -3.4204 | -4.816 | -2.7213 | -4.2633 | -2.4648 | -3.6148 | -2.3341 | -3.3677 | -3.752 | -3.673 | -3.3324 | -4.7666 | -4.8851 | -3.715 | -2.8063 | -3.3708 | -2.3267 | -2.8516 | -2.2399
4 | -1.1142 | 0.17952 | 0.012717 | 0.23688 | -0.22538 | -0.0274 | -0.22684 | 0.17839 | 0.11388 | -0.0059764 | 0.15503 | 0.078102 | 0.0048972 | 0.075022 | 0.054202 | 0.16213 | 0.18878 | 0.24083 | 0.10439 | -0.14801 | 0.024294
5 | -5.7995 | -0.68098 | -0.90753 | -0.30827 | -0.61791 | -0.36654 | -0.53602 | -0.19872 | -0.66378 | -0.29571 | -0.26737 | -0.4603 | -0.60269 | -0.6213 | -0.41702 | -0.75711 | -0.66282 | -0.58756 | -0.55607 | -4.0485 | -0.80215
6 | -0.26283 | -0.042017 | -0.012004 | -0.099718 | 0.062964 | 0.063002 | 0.0023794 | -0.031794 | -0.049072 | -0.081795 | 0.045923 | -0.03864 | 0.13729 | 0.013892 | 0.035998 | 0.16311 | 0.10953 | 0.121 | 0.028692 | 0.057495 | 0.018115
7 | 2.7997 | 0.89255 | 0.93459 | 1.1249 | 1.0855 | 1.1308 | 1.0285 | 1.091 | 1.0663 | 1.0899 | 1.0913 | 1.0394 | 1.0698 | 1.0224 | 1.0748 | 1.1063 | 1.0776 | 1.1625 | 1.0995 | 0.9338 | 1.1147
8 | -0.077631 | 0.031711 | -0.007574 | -0.063021 | 0.0041033 | 0.0025 | -0.0032134 | -0.01745 | 0.010593 | 0.0070825 | 0.016501 | 0.020934 | 0.016025 | 0.0086095 | 0.025949 | -0.029245 | -0.022228 | -0.011075 | 0.012945 | 0.012082 | -0.0011224
9 | 0.33137 | 0.48518 | 0.46394 | 0.45541 | 0.51054 | 0.47966 | 0.51355 | 0.47144 | 0.52653 | 0.4782 | 0.48105 | 0.47037 | 0.48275 | 0.48621 | 0.46126 | 0.48873 | 0.48146 | 0.50369 | 0.48118 | 0.43291 | 0.47999
10 | 0.00011365 | -8.4931e-07 | -1.0921e-06 | 3.3688e-05 | 3.3263e-05 | -1.6852e-06 | -1.303e-06 | 0.00010513 | -7.817e-07 | 6.8651e-07 | -2.4496e-05 | 3.4133e-06 | 7.1121e-06 | 1.3816e-05 | 3.7505e-05 | 2.0143e-05 | -1.7879e-06 | -1.9976e-05 | 2.8797e-05 | -0.00031902 | 0
11 | 0.60267 | 0.7345 | 0.74135 | 0.75051 | 0.73985 | 0.74494 | 0.7416 | 0.75225 | 0.74543 | 0.74921 | 0.75454 | 0.75412 | 0.75065 | 0.75892 | 0.75163 | 0.74605 | 0.74156 | 0.73963 | 0.72773 | 0.63456 | 0.73804
12 | -0.01593 | 0.036564 | -0.015787 | -0.11241 | -0.031488 | -0.024969 | -0.011781 | -0.043279 | -0.04091 | -0.035824 | -0.0075168 | -0.01426 | 0.0038113 | -0.0067023 | -0.015538 | -0.021301 | -0.030466 | -0.019158 | -0.016389 | -0.03658 | -0.020329 | ten times!
13 | 0.81527 | 0.64379 | 0.68485 | 0.63713 | 0.62266 | 0.58444 | 0.58583 | 0.61544 | 0.60983 | 0.61558 | 0.59526 | 0.62444 | 0.61649 | 0.61793 | 0.60517 | 0.63622 | 0.63789 | 0.63257 | 0.62898 | 0.49537 | 0.62523 | ten times!
14 | 0.032909 | -0.055374 | -0.057454 | -0.073812 | -0.03929 | -0.041886 | -0.06232 | -0.054852 | -0.054157 | -0.061965 | -0.050483 | -0.045621 | -0.053827 | -0.046297 | -0.055403 | -0.039269 | -0.055126 | -0.072938 | -0.054294 | 0.024485 | -0.050745 | ten times!
15 | -0.093393 | 0.35127 | 0.3659 | 0.31816 | 0.31871 | 0.31248 | 0.33004 | 0.33784 | 0.33109 | 0.35375 | 0.34831 | 0.34088 | 0.27657 | 0.32421 | 0.32458 | 0.32622 | 0.36093 | 0.3701 | 0.31093 | 0.26048 | 0.31817 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 11.8172 | 2.0929 | 0.79928 | -0.042454 | 1.0472 | 1.3232 | 1.4026 | 0.78735 | 0.86802 | -0.42878 | -0.17174 | 0.91994 | 1.0402 | 1.2587 | 0.25256 | 1.2656 | 0.70192 | 0.42894 | 0.050771 | 1.732 | 1.1475
3 | -0.72263 | 0.85383 | 1.0946 | 1.1885 | 1.0693 | 0.27658 | -0.032942 | 0.047725 | 0.10997 | 0.4013 | 0.156 | 0.060311 | 0.3717 | 0.67813 | 0.72289 | 0.37226 | 0.42016 | 0.14283 | 0.4031 | 0.22934 | 0.42217
4 | 2.7097 | 0.024976 | 0.082541 | 0.54353 | -0.15754 | -0.048653 | -0.17311 | 0.0062257 | -0.0060755 | 0.31782 | -0.0011145 | -0.19107 | -0.20886 | -0.43776 | 0.064966 | -0.32622 | 0.019738 | -0.17483 | 0.01539 | -0.060214 | 0.055193
5 | 2.9631 | 0.028431 | 0.09514 | 0.0058107 | 0.22853 | -0.10217 | -0.16627 | -0.040702 | 0.15593 | 0.17541 | 0.10695 | -0.090693 | -0.10482 | 0.13854 | 0.0034528 | -0.040848 | 0.052771 | 0.071855 | 0.39555 | -0.067293 | 0.13606
6 | 0.25469 | 0.10088 | 0.11706 | 0.19238 | 0.10953 | 0.22456 | 0.18051 | 0.1017 | 0.055692 | 0.037348 | 0.13407 | 0.14981 | 0.10006 | 0.12958 | 0.17887 | 0.18476 | 0.19971 | 0.21131 | 0.14681 | 0.02269 | 0.14746
7 | 1.6693 | -0.085532 | 0.016407 | -0.0024721 | 0.085093 | -0.022652 | -0.013645 | 0.078784 | 0.0050368 | 0.02238 | 0.024321 | -0.0060991 | 0.013761 | 0.030511 | 0.036878 | 0.050443 | 0.02111 | -0.031685 | -0.012217 | 0.010804 | 0.061452
8 | -0.024147 | -0.023737 | -0.015093 | -0.016148 | 0.011098 | 0.041623 | -0.0020782 | 0.0031097 | 0.024047 | 0.013898 | -0.0035106 | 0.019314 | 0.029631 | 0.049608 | 0.018355 | 0.054317 | 0.018238 | -0.0068557 | -0.005365 | 0.002718 | 0.0094022
9 | -0.25909 | 0.0046856 | -0.0037383 | -0.002643 | -0.025819 | -0.046845 | -0.02417 | -0.014167 | -0.01306 | -0.03473 | -0.049044 | -0.02278 | -0.012993 | -0.04225 | -0.038217 | -0.036404 | 0.0015368 | -0.0091465 | -0.030905 | -0.067821 | -0.029663
10 | -7.7483e-05 | -1.588e-06 | 2.7562e-06 | 1.6091e-05 | 0.00049078 | -8.1067e-07 | -1.2614e-06 | 0.0025146 | 2.2627e-06 | -7.6242e-07 | -2.9926e-05 | -1.3756e-05 | -0.00010031 | 3.1055e-06 | -2.333e-05 | 1.4039e-07 | -2.7406e-06 | 2.7859e-06 | -2.1384e-05 | 4.0317e-05 | 0
11 | 0.15082 | -0.065664 | -0.049147 | -0.058066 | -0.049097 | -0.051427 | -0.045811 | -0.030976 | -0.038552 | -0.045877 | -0.044542 | -0.042387 | -0.036419 | -0.047909 | -0.052846 | -0.038816 | -0.035037 | -0.031478 | -0.042674 | -0.035954 | -0.038465 | ten times!
12 | 0.049375 | 0.015899 | 0.013757 | 0.038762 | 0.0075742 | 0.058157 | 0.053655 | -0.0067669 | -0.028081 | -0.0050854 | 0.035252 | 0.073936 | 0.044017 | 0.031546 | 0.054904 | 0.051046 | 0.022435 | 0.035254 | 0.027323 | -0.0014369 | 0.028507 | ten times!
13 | -0.25075 | -0.0098282 | -0.0075614 | -0.012125 | -0.01879 | -0.052178 | -0.05055 | -0.029077 | -0.036784 | -0.04086 | -0.039472 | -0.048635 | -0.033932 | -0.045028 | -0.056698 | -0.057523 | -0.016421 | -0.012004 | -0.016208 | -0.059211 | -0.039471 | ten times!
14 | 0.1025 | -0.16045 | -0.16423 | -0.1431 | -0.13262 | -0.14657 | -0.17538 | -0.15442 | -0.14425 | -0.12995 | -0.16945 | -0.16549 | -0.15324 | -0.14211 | -0.16747 | -0.15108 | -0.15181 | -0.14996 | -0.14755 | 0.0030093 | -0.13755 | ten times!
15 | 0.065294 | -0.081537 | -0.11871 | -0.053056 | -0.065295 | -0.10661 | -0.11626 | -0.067196 | -0.10642 | -0.091352 | -0.018375 | -0.13747 | -0.16407 | -0.14134 | -0.16065 | -0.16354 | -0.062826 | -0.034333 | -0.020837 | -0.14874 | -0.091872 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -5.0905 | 0.74592 | 0.78169 | 0.42201 | 0.64403 | 1.2222 | 1.4926 | 1.0483 | 1.5057 | 0.51113 | 0.15544 | 0.11726 | 0.56187 | 0.78959 | 0.60835 | 0.37944 | 0.47918 | 0.83483 | 1.3109 | 0.62133 | 0.61991
3 | 37.9768 | -3.3502 | -3.2548 | -4.0921 | -3.6799 | -3.9947 | -2.8589 | -4.179 | -3.1031 | -3.6747 | -3.6011 | -4.0434 | -3.4483 | -6.0204 | -4.2123 | -3.014 | -2.601 | -3.1984 | -2.1334 | -2.5 | -2.301
4 | -1.624 | -0.16788 | -0.0076748 | -0.16777 | -0.19475 | -0.086922 | 0.17177 | 0.21819 | 0.27434 | -0.10476 | -0.43751 | -0.25099 | -0.11512 | 0.070005 | -0.033804 | -0.10982 | -0.053183 | 0.03607 | 0.065727 | -0.069217 | -0.10134
5 | -5.7133 | -0.68841 | -0.8692 | -0.32234 | -0.61618 | -0.43628 | -0.75387 | -0.40041 | -0.59307 | -0.44729 | -0.19669 | -0.35179 | -0.64636 | -0.31722 | -0.36014 | -0.70684 | -0.7118 | -0.58495 | -0.52745 | -3.8798 | -0.79136
6 | -0.39969 | 0.033068 | 0.012758 | 0.0053181 | 0.0093964 | -0.0085246 | -0.11526 | -0.086739 | -0.11926 | -0.061034 | -0.01773 | 0.093795 | 0.021208 | -0.12056 | -0.065633 | 0.048863 | 0.0069651 | -0.080495 | -0.070813 | -0.04638 | -0.030339
7 | 2.6864 | 0.87285 | 0.96321 | 1.1913 | 1.1887 | 1.2117 | 1.2183 | 1.2306 | 1.1923 | 1.2198 | 1.2237 | 1.1828 | 1.1508 | 1.1564 | 1.1221 | 1.1223 | 1.1339 | 1.2231 | 1.1217 | 0.88816 | 1.1887
8 | -0.023381 | 0.036576 | 0.021334 | 0.0038954 | 0.002512 | 0.0020585 | 0.010511 | 0.0080149 | 0.025274 | 0.035613 | 0.042304 | -0.019395 | 0.002228 | 0.012441 | -0.0030075 | -0.012439 | -0.004427 | -0.013007 | 0.011726 | -0.011733 | 0.0052061
9 | 0.31306 | 0.5027 | 0.47506 | 0.44882 | 0.4725 | 0.46006 | 0.47625 | 0.46491 | 0.48989 | 0.47662 | 0.47191 | 0.46253 | 0.48651 | 0.46061 | 0.46772 | 0.49818 | 0.48913 | 0.49144 | 0.47667 | 0.44179 | 0.47065
10 | -6.8481e-05 | 0.0017998 | 6.494e-06 | -1.1745e-05 | -1.1379e-05 | -7.5808e-06 | 9.6387e-06 | 0.00081986 | -3.9429e-05 | -0.0039536 | -4.2053e-06 | 5.1934e-08 | 8.6104e-06 | 0.0039495 | -4.2807e-06 | 4.776e-06 | -3.2315e-05 | 7.0726e-05 | 2.3051e-05 | -6.7113e-05 | 0
11 | 0.58068 | 0.74317 | 0.74285 | 0.74937 | 0.74646 | 0.75153 | 0.73787 | 0.75003 | 0.74961 | 0.74513 | 0.75755 | 0.75506 | 0.75301 | 0.76423 | 0.75209 | 0.74513 | 0.74177 | 0.74362 | 0.73738 | 0.6442 | 0.73987
12 | -0.027224 | 0.091561 | 0.018303 | 0.047393 | 0.046207 | 0.019376 | -0.012386 | 0.039296 | 0.035542 | 0.0067013 | 0.016075 | 0.0030977 | 0.0031924 | -0.038453 | -0.001515 | -0.0058518 | 0.0099675 | 0.023531 | 0.010482 | -0.0095137 | 0.014608 | ten times!
13 | 0.8105 | 0.62967 | 0.67552 | 0.65648 | 0.66414 | 0.64699 | 0.69325 | 0.66111 | 0.6555 | 0.63845 | 0.62794 | 0.61846 | 0.63952 | 0.65728 | 0.62011 | 0.62093 | 0.64439 | 0.66993 | 0.65469 | 0.50427 | 0.65173 | ten times!
14 | 0.015568 | -0.034729 | -0.038622 | -0.024857 | -0.025148 | -0.024787 | -0.066826 | -0.022522 | -0.036458 | -0.04429 | -0.045888 | -0.039823 | -0.037129 | -0.053149 | -0.04688 | -0.052714 | -0.043111 | -0.044608 | -0.036802 | 0.014461 | -0.039016 | ten times!
15 | 0.033875 | 0.32029 | 0.33112 | 0.28117 | 0.28899 | 0.28541 | 0.27226 | 0.27572 | 0.30284 | 0.32119 | 0.36496 | 0.31071 | 0.28982 | 0.31545 | 0.33598 | 0.2933 | 0.30024 | 0.33164 | 0.2968 | 0.30752 | 0.29843 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 7.8155 | -1.8958 | -2.4613 | -1.8153 | -2.2036 | -3.0163 | -2.8201 | -3.4208 | -2.316 | -3.4239 | -2.689 | -1.8681 | -1.4437 | -2.3078 | -2.6189 | -2.4538 | -2.0099 | -2.1929 | -1.7406 | 19.6922 | -1.3853
3 | 3.4436 | 1.171 | 1.237 | 1.095 | 0.86787 | 0.2178 | -0.20244 | 0.7191 | 0.89745 | 0.48051 | 0.24375 | 0.16 | 0.23109 | 0.74489 | 0.68926 | 0.74649 | 0.71562 | 1.227 | 1.0186 | -0.1549 | 0.74175
4 | 1.0022 | -0.29426 | -0.16238 | -0.49738 | -0.48811 | -0.50965 | -0.25655 | -0.45019 | -0.60546 | 0.16813 | 0.18598 | -0.19357 | -0.021692 | -0.18782 | -0.0022974 | 0.078953 | 0.11235 | 0.044639 | 0.047828 | 1.657 | -0.050233
5 | 3.0082 | 0.1046 | -0.017384 | 0.13764 | 0.26447 | 0.0031781 | 0.037794 | 0.16098 | 0.021085 | 0.12549 | -0.040593 | 0.015886 | -0.28054 | 0.14022 | 0.17131 | 0.11242 | 0.24227 | 0.077371 | 0.25142 | -0.20023 | 0.16374
6 | -0.0097341 | 0.12166 | 0.04526 | -0.069403 | -0.12252 | -0.16606 | -0.25218 | -0.14255 | -0.10716 | 0.057296 | 0.17332 | 0.14728 | -0.09873 | -0.089628 | 0.0065006 | -0.034752 | -0.059932 | 0.021686 | 0.079967 | 0.20814 | -0.018707
7 | 1.8961 | -0.0042215 | 0.00568 | 0.01795 | 0.098925 | -0.018755 | -0.087406 | 0.087407 | 0.12303 | 0.02307 | 0.010768 | -0.018688 | -0.0024668 | 0.14337 | 0.11439 | 0.11006 | 0.053314 | 0.099887 | 0.063452 | 0.040354 | 0.10082
8 | -0.063062 | -0.023577 | 0.0090819 | 0.020904 | -0.012837 | -0.040402 | -0.0534 | -0.041426 | -0.023845 | -0.037925 | -0.0036439 | -0.0025747 | -0.039697 | -0.025017 | -0.026856 | -0.044103 | -0.0385 | -0.026533 | -0.010601 | 0.037018 | -0.023974
9 | -0.20363 | 0.054649 | 0.045896 | 0.018993 | -0.004618 | -0.013962 | -0.0080692 | 0.040885 | 0.037072 | -0.017599 | -0.017126 | 0.016012 | 0.024949 | 5.1627e-05 | -0.027964 | 0.014898 | 0.031758 | 0.048157 | 0.012364 | -0.065347 | 0.005068
10 | 0.00015691 | -2.9297e-05 | 1.3528e-05 | -2.6966e-05 | -9.939e-06 | 6.4427e-06 | 4.4561e-07 | 9.9524e-06 | 0.0022153 | 3.0277e-05 | 2.764e-05 | 7.2852e-06 | 6.4867e-06 | 3.6905e-06 | 5.6827e-06 | -2.2151e-06 | 0.0030628 | -0.0039417 | -6.7257e-06 | -7.1843e-05 | 0
11 | 0.19958 | -0.015399 | -0.010768 | -0.0035683 | 0.00062103 | -0.0096788 | 0.00070413 | 0.013557 | 0.013106 | -0.0063106 | -0.0050056 | -0.0046365 | -0.00049588 | 0.00080054 | -0.00030926 | 0.0021981 | 0.0075657 | 0.01446 | 0.0034821 | 0.0020919 | 0.0061522
12 | 0.14051 | 0.035165 | -0.016887 | -0.041601 | -0.062873 | -0.080243 | -0.063505 | -0.070689 | -0.04552 | -0.047691 | -0.0087534 | -0.029248 | -0.04543 | -0.064671 | -0.025488 | -0.030531 | -0.03054 | -0.039583 | 0.013052 | 0.036268 | -0.032097 | ten times!
13 | -0.23464 | 0.071186 | 0.04254 | 0.025322 | 0.019437 | -0.0096287 | -0.015276 | 0.04987 | 0.043707 | 0.011554 | 0.012634 | 0.00085523 | -0.002436 | 0.025858 | -0.0094519 | 0.031512 | 0.029433 | 0.041805 | 0.044892 | -0.020074 | 0.015732 | ten times!
14 | 0.15099 | -0.16604 | -0.14708 | -0.11797 | -0.11996 | -0.12495 | -0.12134 | -0.1421 | -0.14827 | -0.13479 | -0.15075 | -0.13933 | -0.13461 | -0.1271 | -0.13886 | -0.14389 | -0.12222 | -0.15126 | -0.147 | 0.011923 | -0.12144 | ten times!
15 | 0.046008 | -0.055195 | -0.12079 | -0.0076473 | -0.020327 | -0.09603 | -0.13055 | -0.022196 | -0.061768 | -0.065837 | -0.012577 | -0.11072 | -0.11608 | -0.07619 | -0.1097 | -0.094112 | -0.054635 | -0.047479 | -0.025421 | -0.12478 | -0.066592 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00051153 | -6.3837e-05 | -4.421e-05 | -0.00031864 | -9.0676e-05 | -4.8572e-05 | 3.2617e-06 | -0.00010242 | -5.9475e-05 | -0.00010846 | 5.2747e-05 | -9.4534e-05 | -0.00016985 | -0.00018159 | -3.8842e-05 | 8.2773e-05 | 5.1572e-05 | 0.00061922 |  | 596.1404
|  | 0.00036913 | -0.00020876 | -8.8319e-05 | -0.00019484 | -0.00010124 | -0.00010342 | -4.3698e-05 | -7.2047e-05 | -7.2047e-05 | -1.8033e-05 | 0.00011365 | 1.82e-05 | -0.0001014 | -0.00012589 | -1.451e-05 | -3.8666e-05 | 2.5749e-05 | 0.00065614 |  | 596.1376
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3997.1999
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0982 | -4188.1441 | 5.8224 | 0.068361 | 0.26464 | -2.4799 | -0.20005 | 0.22102 | 0.73939 | -0.31748 | 0.11198 | -0.37787 | 1.2343 | -0.23864 | -0.99187 | -1.1093 | 0.31832 | 1.5346 | 1.2225 | 6.8995 | -3806.2557
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.071559 | 2.2627 | -0.53647 | 0.027406 | 0.08269 | -0.09914 | 0.090495 | 0.59524 | 0.92854 | 0.54347 | -0.44636 | -0.80959 | -0.2543 | 0.17485 | 0.28557 | -0.029087 | -0.028541 | -0.19794 | -0.066607 | -0.26023 | 0.096174
db2 | -0.84289 | 1.0815 | 0.86865 | 0.52463 | -0.018472 | 0.22427 | 0.94127 | 0.44823 | -0.5441 | -0.23392 | -0.12631 | 0.0361 | 0.27525 | -0.2138 | -0.48787 | 0.22274 | 0.426 | -0.21822 | -0.95866 | -1.1658 | -1.569
db3 | -3.519 | 38.0119 | -0.19529 | 0.099337 | -1.2994 | 0.79591 | -0.74403 | 1.0561 | -0.094068 | 1.1841 | 0.15266 | -0.22836 | -0.15505 | 0.18794 | -1.2528 | -1.368 | -0.20197 | 0.71611 | 0.15304 | 1.1938 | -2.8349
db4 | 0.073293 | -1.4749 | 0.099719 | -0.049857 | 0.1544 | -0.305 | -0.09781 | -0.29518 | 0.10247 | 0.047149 | -0.08056 | 0.069692 | 0.014866 | -0.034205 | 0.020362 | -0.0028976 | 0.10916 | 0.098886 | 0.14573 | 0.003066 | 0.0065883
db5 | -0.52757 | -5.8369 | -0.15418 | -0.37798 | 0.21249 | -0.089316 | 0.15884 | -0.0095055 | 0.32868 | -0.13519 | 0.23139 | 0.26578 | 0.068712 | -0.081497 | -0.091323 | 0.11248 | -0.23168 | -0.12896 | -0.047503 | -0.031241 | -4.0582
db6 | 0.025469 | -0.19586 | -0.054465 | -0.051784 | -0.077134 | 0.054189 | 0.027384 | -0.034877 | -0.035631 | -0.08792 | -0.09401 | 0.092358 | -0.097432 | 0.026769 | -0.054078 | -0.052316 | 0.092797 | 0.12598 | 0.16419 | 0.055983 | 0.0051018
db7 | 1.0668 | 2.7987 | -0.17346 | -0.13252 | 0.056277 | 0.018949 | 0.063523 | -0.038128 | 0.024129 | -0.00032868 | 0.023375 | 0.029135 | -0.027487 | 0.0046553 | -0.042679 | 0.0088143 | 0.043014 | 0.011078 | 0.09721 | 0.034436 | 0.92682
db8 | 0.00073996 | -0.030424 | 0.04006 | -0.018081 | -0.037749 | 0.013839 | -0.0035279 | -0.01163 | -0.0057055 | 0.00085207 | 0.013549 | 0.059134 | 0.00010822 | -0.035937 | -0.020271 | -0.011534 | -0.057489 | 0.002017 | 0.028271 | 0.044095 | -0.027842
db9 | 0.4871 | 0.32825 | -0.0015739 | -0.023363 | -0.029289 | 0.023879 | -0.0074603 | 0.026585 | -0.014915 | 0.03958 | -0.0088065 | 0.0021198 | -0.014838 | 0.005613 | 0.0021766 | -0.019359 | 0.0045337 | -0.0030642 | 0.022239 | -0.004058 | 0.41568
db10 | 0.00091318 | 0.045409 | 0.016891 | -0.02192 | 0.05506 | 0.018079 | -0.011509 | -0.014638 | 0.024051 | -0.017167 | 0.013336 | 0.082411 | -0.04098 | -0.098349 | -0.055141 | -0.074725 | -0.052808 | 0.046426 | 0.071507 | 0.059477 | -0.069266
db11 | 0.74603 | 0.60254 | -0.011406 | -0.0046163 | 0.003688 | -0.0062138 | -0.0011461 | -0.0044574 | 0.0061537 | -0.00053245 | 0.0031007 | 0.0093373 | 0.0083536 | 0.0056008 | 0.013391 | 0.006585 | 0.00067635 | -0.0043287 | -0.0057859 | -0.0184 | 0.63304
db12 | -0.0021719 | 0.011138 | 0.0076592 | -0.0015327 | -0.0034543 | 0.00093505 | -0.0013111 | -0.00030098 | 0.00027804 | -0.0035113 | -2.601e-05 | 0.009385 | -0.0032712 | -0.0074465 | -0.0037925 | -0.0066369 | -0.0053111 | 0.0039719 | 0.0075367 | 0.0068289 | -0.01037
db13 | 0.062417 | 0.080338 | 0.0020051 | 0.0055655 | 0.0016839 | -0.00024469 | -0.0036093 | -0.0033971 | -0.0010418 | -0.001562 | -0.0012768 | -0.0025793 | 0.00063883 | 0.0014752 | -0.00032132 | -0.00075261 | 0.0014483 | 0.00085857 | -0.00019564 | 0.0013052 | 0.048863
db14 | -0.0052597 | 0.0010157 | 0.00097357 | -0.0023756 | 0.0013806 | 0.0024742 | 0.00062986 | -0.001664 | 0.0013377 | -0.0013665 | -0.00014341 | 0.0055888 | -0.0017321 | -0.0041853 | -0.0030314 | -0.0048774 | -0.0023067 | 0.0028786 | 0.0029509 | 0.0034682 | -0.0054112
db15 | 0.033344 | -0.0093393 | 0.0017828 | 0.0032457 | -0.0015281 | -0.0014737 | -0.002096 | -0.0003399 | 0.00043956 | -0.00023527 | 0.0020304 | 0.0014872 | 0.00074384 | -0.0056876 | -0.00092313 | -0.00088628 | -0.00072184 | 0.0027486 | 0.0036653 | -0.0022516 | 0.026048
da2 | 0.75403 | 12.591 | 1.3366 | 0.062184 | -0.82245 | 0.29868 | 0.55534 | 0.64272 | 0.033636 | 0.12099 | -1.1709 | -0.93154 | 0.16055 | 0.24576 | 0.51875 | -0.50855 | 0.52327 | -0.038489 | -0.30018 | -0.72637 | 1.844
da3 | 0.46499 | -0.68961 | 0.38944 | 0.6286 | 0.73805 | 0.60309 | -0.18829 | -0.49774 | -0.41721 | -0.35558 | -0.061733 | -0.30838 | -0.40146 | -0.084288 | 0.22157 | 0.25581 | -0.087453 | -0.0451 | -0.32922 | -0.0601 | 0.2208
da4 | -0.034178 | 2.5888 | 0.057132 | 0.12773 | 0.57243 | -0.121 | -0.016086 | -0.14025 | 0.039888 | 0.031031 | 0.35521 | 0.036845 | -0.15603 | -0.17987 | -0.40121 | 0.10052 | -0.28386 | 0.05925 | -0.12997 | 0.048242 | 0.27516
da5 | 0.051098 | 2.9798 | -0.021566 | 0.043194 | -0.03853 | 0.17886 | -0.15466 | -0.21924 | -0.090094 | 0.10486 | 0.12511 | 0.064449 | -0.14614 | -0.16322 | 0.083848 | -0.055952 | -0.099015 | 0.0067292 | 0.026872 | 0.3545 | -0.063265
da6 | 0.14335 | 0.53022 | -0.039156 | -0.045139 | 0.079002 | -0.032515 | 0.091734 | 0.044837 | -0.040092 | -0.097946 | -0.1157 | 0.00040264 | 0.010654 | -0.014626 | -0.030212 | 0.028811 | 0.024326 | 0.047739 | 0.042898 | 0.044979 | -0.092471
da7 | 0.011117 | 1.6659 | -0.096743 | 0.0061929 | -0.014732 | 0.074166 | -0.034312 | -0.024899 | 0.06773 | -0.0063552 | 0.011215 | 0.013497 | -0.018063 | -0.0030859 | 0.017672 | 0.024543 | 0.037962 | 0.010533 | -0.044959 | -0.020362 | 0.019074
da8 | 0.012716 | -0.017062 | -0.033086 | -0.040047 | -0.012404 | -0.0021469 | 0.03532 | -0.0094129 | -0.0099792 | 0.004809 | -0.0052061 | -0.015605 | 0.010002 | 0.037163 | 0.029431 | 0.0059057 | 0.034587 | -0.00013471 | -0.034108 | 0.0049113 | -0.062024
da9 | -0.022137 | -0.24773 | 0.027056 | 0.019573 | 0.023631 | -0.0036993 | -0.025019 | -0.0023552 | 0.0080146 | 0.0095208 | -0.01301 | -0.025552 | -0.0012133 | 0.00049489 | -0.01793 | -0.015587 | -0.012335 | 0.022289 | 0.0083777 | -0.0022562 | -0.043823
da10 | 0.0018377 | 0.077943 | 0.0028142 | -0.02512 | 0.027586 | -0.0026046 | 0.010278 | 0.0080755 | -0.0003478 | -0.012946 | -0.014163 | 0.0025772 | 0.0052447 | 0.037342 | -0.017312 | -0.0020423 | -0.016075 | -0.01495 | -0.029396 | 0.041039 | -0.12659
da11 | -0.044871 | 0.15343 | -0.020685 | -0.0041026 | -0.012917 | -0.0042062 | -0.0067312 | -0.0011077 | 0.013865 | 0.0065091 | -0.001007 | 0.00085761 | 0.001968 | 0.0067741 | -0.0029945 | -0.0084841 | 0.0059619 | 0.0098692 | 0.013271 | 0.0031593 | -0.032346
da12 | 0.0030548 | 0.012046 | -0.00085588 | -0.0043875 | 0.0040362 | -0.002369 | 0.0039264 | 0.0032827 | -0.0038737 | -0.0069511 | -0.0047848 | 0.00052125 | 0.0050683 | 0.0051808 | -0.0014981 | 0.0023376 | 0.00066288 | -0.0021529 | -0.0023883 | 0.0042451 | -0.011716
da13 | -0.0033157 | -0.023641 | 0.0018005 | 0.0034895 | 0.00068915 | 0.0009871 | -0.0017203 | -0.0014317 | -0.00026884 | 0.00018002 | -0.0010266 | -0.0029961 | -0.00054776 | 0.0011689 | 0.00034735 | -0.00037186 | -0.0009189 | 0.00033599 | -3.1322e-06 | 0.00028672 | -0.0038225
da14 | -0.014813 | 0.009823 | -0.0010854 | -0.002873 | 0.0019659 | 0.0013167 | 0.00105 | -0.0018491 | -0.00093933 | -4.0299e-05 | 0.00080902 | -0.0018151 | -0.00040307 | 0.0042848 | 0.00085967 | -6.3559e-05 | 0.00023615 | -0.0016747 | -0.0023696 | 0.002591 | -0.0051273
da15 | -0.0094993 | 0.0065293 | 0.0013456 | -0.0023713 | 0.0041937 | 0.0029699 | -0.0011619 | -0.002127 | 0.0027798 | -0.0011431 | 0.00036409 | 0.0076618 | -0.0042477 | -0.0069074 | -0.0046342 | -0.0065654 | -0.0068552 | 0.0032168 | 0.0060661 | 0.0074157 | -0.014874
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4038.3721
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0993 | -4264.5584 | 4.3323 | -1.4469 | -0.24241 | -1.3077 | -0.37158 | -0.39339 | 0.20383 | -0.079685 | -0.079685 | 0.46049 | 1.7774 | 0.82285 | -0.37326 | -0.61819 | 0.49572 | 0.25415 | 0.89834 | 7.2027 | -3812.1858
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.10556 | 4.7018 | -0.38247 | -0.11555 | -0.08718 | 0.24183 | 0.10336 | 0.37876 | 0.64658 | 0.77721 | -0.64478 | -0.9481 | -0.40469 | 0.15552 | 0.32044 | 0.0022032 | 0.018417 | -0.13078 | 0.018587 | 0.050642 | -1.0831
db2 | 0.75795 | -3.7208 | -0.023394 | 0.081141 | -0.39875 | -0.15374 | 0.49744 | 0.76269 | 0.2842 | 0.76286 | -0.24386 | -0.65842 | -0.59738 | -0.15383 | 0.033399 | -0.10848 | -0.35478 | -0.29358 | 0.046516 | 0.51795 | 0.62725
db3 | -3.5806 | 37.9504 | 0.23493 | 0.32471 | -0.50738 | -0.097147 | -0.41603 | 0.72079 | -0.59252 | 0.47542 | -0.092527 | -0.028948 | -0.45975 | 0.1342 | -2.4405 | -0.63141 | 0.56737 | 0.97906 | 0.38256 | 1.4472 | -2.3838
db4 | -0.048094 | -1.8344 | -0.12184 | 0.069188 | -0.13057 | -0.16158 | -0.031693 | 0.23371 | 0.26356 | 0.32818 | -0.055113 | -0.39617 | -0.19536 | -0.049962 | 0.11876 | 0.022624 | -0.049969 | -0.013189 | 0.073034 | 0.096381 | -0.04401
db5 | -0.52873 | -5.7145 | -0.16326 | -0.33876 | 0.2088 | -0.084887 | 0.091506 | -0.21902 | 0.13069 | -0.062684 | 0.081972 | 0.33663 | 0.17648 | -0.11651 | 0.21093 | 0.17041 | -0.17822 | -0.18269 | -0.057307 | -0.0040783 | -3.8689
db6 | -0.031385 | 0.029014 | 0.064542 | -0.0039181 | 0.081514 | 0.071676 | -0.0067646 | -0.1175 | -0.057599 | -0.10524 | -0.029237 | 0.072967 | 0.087369 | 0.0097835 | -0.088163 | -0.061856 | 0.047819 | 0.057717 | -0.01426 | -0.0088447 | -0.055015
db7 | 1.1519 | 2.6727 | -0.27893 | -0.18812 | 0.039102 | 0.037423 | 0.059492 | 0.06729 | 0.079424 | 0.040079 | 0.066842 | 0.075293 | 0.032226 | 0.00012262 | 0.0041986 | -0.029985 | -0.028025 | -0.017864 | 0.071398 | -0.029955 | 0.88489
db8 | 0.0067634 | -0.005837 | 0.027683 | -0.016854 | 0.01908 | 0.013287 | -0.019832 | -0.01424 | -0.00014577 | 0.0089867 | 0.028419 | 0.065363 | -0.045638 | -0.028956 | 0.0056774 | -0.026886 | -0.037684 | -4.0597e-05 | -0.0010464 | 0.022827 | -0.029902
db9 | 0.47702 | 0.32546 | 0.024097 | 0.0025667 | -0.027153 | -0.0038851 | -0.016089 | 0.00040408 | -0.013347 | 0.013261 | -0.0010339 | -0.0018779 | -0.012735 | 0.012026 | -0.016515 | -0.0085673 | 0.022125 | 0.012646 | 0.015927 | -0.00185 | 0.4198
db10 | -0.0034276 | 0.12518 | 0.0059534 | -0.05853 | 0.051252 | 0.037869 | -0.027176 | -0.031671 | 0.0036611 | -0.014812 | -0.0007209 | 0.064063 | -0.036657 | -0.043224 | 0.0073871 | -0.029693 | -0.030616 | 0.024511 | 0.038821 | 0.039583 | -0.016367
db11 | 0.74825 | 0.58193 | -0.0054688 | -0.0050044 | 0.0017999 | -0.001236 | 0.0031278 | -0.010006 | 0.0019514 | 0.0013186 | -0.0033587 | 0.009786 | 0.0069498 | 0.0051012 | 0.015885 | 0.0038515 | -0.0028716 | -0.006385 | -0.0043069 | -0.011135 | 0.64258
db12 | 0.0013102 | 0.010688 | 0.0074419 | -0.0061935 | 0.0083836 | 0.0069612 | -0.0025158 | -0.006519 | 0.0023699 | 0.00032981 | -0.00059046 | 0.0062297 | -0.0049241 | -0.0057995 | -0.0051177 | -0.0047251 | -0.0054501 | 0.0018835 | 0.004852 | 0.0033836 | -0.0027194
db13 | 0.065116 | 0.084632 | -0.0012368 | 0.0030598 | 0.0011825 | 0.0017436 | -0.00050404 | 0.0050086 | 0.0016932 | 0.00058177 | -0.0018653 | -0.0032926 | -0.0030874 | -0.00052369 | 0.00040908 | -0.0032458 | -0.0022006 | -0.00079211 | 0.0017378 | 0.0013321 | 0.048792
db14 | -0.0041571 | 0.0032671 | 0.0010599 | -0.0037813 | 0.0042275 | 0.003607 | -0.00019687 | -0.0039334 | 0.0020376 | -0.00051817 | -0.00055542 | 0.0035997 | -0.0023612 | -0.0019118 | -0.0012837 | -0.003051 | -0.0026547 | 0.0010424 | 0.0019008 | 0.0027727 | -0.0031184
db15 | 0.030655 | 0.0033875 | 0.001374 | 0.0024574 | -0.0025379 | -0.0017561 | -0.0021142 | -0.0034291 | -0.0030832 | -0.00037101 | 0.0014641 | 0.0058411 | 0.00041647 | -0.0016729 | 0.00088981 | 0.0029434 | -0.0013244 | -0.00063107 | 0.0025087 | -0.0009751 | 0.030752
da2 | -2.3755 | 6.8532 | 0.44719 | -0.097908 | 0.53304 | 0.15475 | -0.62721 | -0.45526 | -1.0809 | 0.055174 | -1.0258 | -0.28007 | 0.51776 | 0.9223 | 0.081917 | -0.22475 | -0.10021 | 0.36974 | 0.1943 | 0.61594 | 19.7868
da3 | 0.68197 | 3.5986 | 0.48657 | 0.55888 | 0.40077 | 0.17694 | -0.45756 | -0.87942 | 0.039888 | 0.21989 | -0.20051 | -0.42828 | -0.51784 | -0.45084 | 0.062668 | 0.007528 | 0.062822 | 0.03463 | 0.54567 | 0.33821 | -0.16777
da4 | -0.16976 | 1.5286 | -0.13797 | 0.012507 | -0.33052 | -0.3219 | -0.33666 | -0.094196 | -0.28724 | -0.43592 | 0.34338 | 0.35771 | -0.021964 | 0.15245 | -0.016567 | 0.1686 | 0.23951 | 0.28485 | 0.21588 | 0.20805 | 1.9633
da5 | 0.085684 | 3.002 | 0.019726 | -0.1044 | 0.052168 | 0.17727 | -0.078506 | -0.043984 | 0.0734 | -0.063276 | 0.040739 | -0.12135 | -0.073853 | -0.36348 | 0.055243 | 0.08735 | 0.027412 | 0.15575 | -0.0077315 | 0.16752 | -0.19618
da6 | -0.021909 | -0.049979 | 0.16903 | 0.059362 | -0.010931 | -0.073439 | -0.15766 | -0.21166 | -0.089088 | -0.087052 | 0.056713 | 0.16569 | 0.15918 | -0.070094 | -0.0759 | 0.0056696 | 0.0011387 | -0.037125 | 0.046418 | 0.14974 | 0.10306
da7 | 0.046014 | 1.8814 | -0.049064 | -0.040298 | -0.025374 | 0.053391 | -0.063261 | -0.13205 | 0.041591 | 0.077431 | -0.023702 | -0.038448 | -0.06397 | -0.047023 | 0.097251 | 0.069927 | 0.064597 | 0.0065471 | 0.053051 | 0.019404 | 0.044978
da8 | -0.02045 | -0.13054 | 0.016366 | 0.021395 | 0.060555 | 0.020476 | -0.026598 | -0.02415 | -0.0051337 | -0.0041143 | -0.029188 | -1.8233e-05 | 0.011619 | -0.016657 | -0.0089118 | -0.016437 | -0.014068 | -0.017325 | -0.0042585 | 0.036447 | -0.027752
da9 | 0.014685 | -0.2272 | 0.040075 | 0.032595 | 0.0085992 | -0.017179 | -0.027609 | -0.024148 | 0.02621 | 0.022346 | -0.032169 | -0.036283 | 0.0023895 | 0.00945 | -0.014675 | -0.040907 | -0.0013273 | 0.017157 | 0.033333 | 0.0021437 | -0.059463
da10 | 0.0059354 | -0.049939 | 0.030152 | -0.014724 | 0.031394 | 0.019655 | -0.019952 | 0.01034 | 0.026114 | -0.0040414 | -0.028817 | -0.03715 | -0.016947 | 0.001379 | -0.014944 | -0.027916 | 0.012765 | -0.0026599 | -0.0091999 | 0.044552 | -0.12646
da11 | 0.00018157 | 0.19706 | -0.015113 | -0.010574 | -0.0032193 | 0.00057462 | -0.0094119 | 0.00062744 | 0.013548 | 0.01305 | -0.0065165 | -0.0058718 | -0.0045194 | -0.00046352 | 0.00066836 | -0.00015205 | 0.0019694 | 0.0072912 | 0.014082 | 0.0040309 | 0.0029067
da12 | -0.0030071 | -0.0045816 | 0.010251 | -0.00033383 | 0.002961 | -0.00046378 | -0.0064938 | -0.0014202 | -0.00063792 | -0.0016857 | -0.0040966 | -0.00093044 | -0.0011001 | -0.00083418 | -0.0043873 | -0.0017205 | 0.0016567 | 3.7899e-05 | -0.0007673 | 0.009965 | -0.0078521
da13 | 0.0023706 | -0.020707 | 0.0045522 | 0.0035212 | -0.00065152 | -0.0010959 | -0.0025558 | -0.0034866 | 0.0024987 | 0.0024615 | -0.0010268 | -0.0028829 | -0.001211 | -0.0017193 | 0.00029855 | -0.002207 | 0.0013133 | 0.00011236 | 0.00085671 | 0.0012223 | -0.00097798
da14 | -0.013208 | 0.015498 | -0.0012156 | -0.00069589 | 0.0033142 | 0.0024659 | 0.0004921 | 0.0027856 | 0.00064653 | -0.0014088 | -0.0016543 | -0.0040832 | -0.00055107 | 0.0011464 | -2.2378e-05 | -0.0013766 | 0.0004561 | 0.00076844 | -0.0022339 | 0.0011665 | -0.0068122
da15 | -0.0068181 | 0.0046008 | 0.0012986 | -0.0052604 | 0.0060534 | 0.0047854 | -0.002785 | -0.0062369 | 0.0045985 | 0.00064132 | 0.00023443 | 0.0055604 | -0.0042538 | -0.0047904 | -0.00080088 | -0.0041521 | -0.0025932 | 0.0013545 | 0.0020702 | 0.004276 | -0.012478
Field Colinearity | 19.6352
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.1911 | -0.60803 | -0.044153 | 0.011131 | -0.1707 | 0.018936 | 0.52368 | 0.85698 | 0.47191 | -0.51792 | -0.88115 | -0.32586 | 0.10329 | 0.21401 | -0.10065 | -0.1001 | -0.2695 | -0.13817 | -0.33179 | 0.024615
Twist (m2 mrad) | 16.4334 | -8.6644 | -0.59607 | 0.14192 | -2.0484 | 0.21303 | 5.4987 | 8.3556 | 4.2472 | -4.2728 | -6.6086 | -2.1995 | 0.61975 | 1.1235 | -0.45291 | -0.37538 | -0.8085 | -0.31087 | -0.49768 | 0.0092306
Twist Integral (I0) (m2 rad) | 0.0073554
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.5962 | -0.48804 | -0.22112 | -0.19274 | 0.13627 | -0.002202 | 0.2732 | 0.54102 | 0.67164 | -0.75034 | -1.0537 | -0.51025 | 0.04996 | 0.21488 | -0.10336 | -0.087147 | -0.23635 | -0.086977 | -0.054922 | -1.1887
Twist (m2 mrad) | 34.4715 | -6.9545 | -2.9851 | -2.4575 | 1.6352 | -0.024772 | 2.8686 | 5.2749 | 6.0448 | -6.1903 | -7.9025 | -3.4442 | 0.29976 | 1.1281 | -0.46512 | -0.3268 | -0.70904 | -0.1957 | -0.082383 | -0.44576
Twist Integral (I0) (m2 rad) | 0.014654

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.84289 | -3.519 | 0.073293 | -0.52757 | 1.0668 | -0.71559 | 0.75403 | 0.46499 | -0.034178 | 19.1142 | 0.59614 | 0.75795 | -3.5806 | -0.048094 | -0.52873 | 1.1519 | -1.0556 | -2.3755 | 0.68197 | -0.16976 | 20.573 | 0.59614 | -0.8081 | -2.2399 | 0.024294 | -0.80215 | 1.1147 | 0 | 1.1475 | 0.42217 | 0.055193 | 14419.336 | 0.61991 | -2.301 | -0.10134 | -0.79136 | 1.1887 | 0 | -1.3853 | 0.74175 | -0.050233 | 14421.168
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8