## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3095C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-11-04
Beginn der Messung / start of measuring: |  | TT.11.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000095_cc
Ende der Messung / end of measuring: |  | TT.11.JJ
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Mehler |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-11-04 |  | 2003-11-04 | 2003-11-04
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000095_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000095
Date of test Ap 1 | 2003-11-04 |  |  |  | Date of test Ap 2 | 2003-11-04
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | yellow alarm | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | yellow alarm
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
220 | HCMB__A001 | 3000095 |  | 3 |  |  | 3
| COMMENTS: | Yellows alarms in ap.1 on b9 pos. 1 and in ap. 2 on b3 pos. 14 & on b9 pos.1.
| HOLDING POINT | OK-w | Data at CERN |  | 2003-11-04 | Analysis by |  | E. Wildner | Answer to MMS-LD |  |  | 2003-11-05 | Answer to manufacturer |  |  | 2003-11-05
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000095_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000095 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3095_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3095_cct2 |  |  |  | a even | 1
Test Operator |  | Mehler |  |  | Test Controller |  |  | E. Wildner
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-11-04 |  |  | Time of test |  |  | 7:06 |  |  |  | Date of test |  | 2003-11-04 |  |  | Time of test |  |  | 7:06
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-02-05
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 24 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 24
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | 7.2543e-05 |  |  | Dy average |  |  | 0.00086166 |  |  |  | Dx average |  | 0.0002183 |  |  | Dy average |  |  | 0.00067491
Magnetic length |  | 14.419 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0109 |  |  |  | Magnetic length |  | 14.4199 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0506
Phase |  | -14.1837 |  |  | Integral field (T.m) |  |  | 85.8143 |  |  |  | Phase |  | -33.8621 |  |  | Integral field (T.m) |  |  | 85.7714
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4678 | 5.9628 | 5.9589 | 5.9591 | 5.9596 | 5.9593 | 5.9603 | 5.9605 | 5.9598 | 5.9599 | 5.9602 | 5.9596 | 5.9588 | 5.9592 | 5.9597 | 5.9606 | 5.9603 | 5.96 | 5.9634 | 3.6848 | 85.8143 | C1 (T) | 3.4266 | 5.9624 | 5.9593 | 5.9599 | 5.9603 | 5.9603 | 5.9604 | 5.9598 | 5.9606 | 5.9601 | 5.9604 | 5.9608 | 5.9601 | 5.9607 | 5.9603 | 5.9602 | 5.961 | 5.9595 | 5.9631 | 3.6617 | 85.7714
Angle (mrad) | 0.56889 | -1.3016 | -0.61493 | -0.26193 | -0.26792 | -0.15142 | 0.37705 | 0.81166 | 0.58944 | 0.099261 | -0.42283 | -0.14159 | 0.33508 | 0.76688 | 0.49306 | -0.030798 | -0.37643 | -0.03561 | 0.067249 | -0.42877 | -14.1837 | Angle (mrad) | 2.9106 | -1.52 | -0.58988 | -0.27883 | -0.26065 | -0.48182 | 0.2319 | 0.85468 | 0.64319 | 0.16269 | -0.35423 | -0.015795 | 0.13168 | 0.53317 | 0.43703 | -0.069097 | -0.51868 | -0.16551 | 0.1072 | -0.84629 | -33.8621
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.2917 | -0.46018 | 0.12788 | -0.62341 | -0.052965 | 0.34279 | -0.028677 | -0.25491 | 0.39955 | 0.30647 | 0.13438 | -0.22924 | 0.12585 | 0.086515 | 0.50903 | 0.67702 | 0.7392 | 0.29999 | 0.75049 | 0.41568 | 0.092379 | b2 | -1.1213 | -0.86106 | -0.14374 | -0.41653 | 0.85797 | -0.29428 | -0.12563 | 0.28218 | 0.014852 | 0.25132 | 0.024027 | -0.013122 | -0.71149 | 0.016058 | 0.14113 | 0.20541 | -0.085872 | 0.17708 | -0.66882 | -0.45903 | -0.1187
b3 | 42.8096 | -3.3938 | -2.7015 | -3.1143 | -2.1666 | -3.2502 | -2.2506 | -3.0746 | -1.9807 | -2.99 | -2.9643 | -3.37 | -2.8422 | -4.2203 | -3.8523 | -3.4074 | -2.4091 | -2.731 | -1.835 | -2.2261 | -1.5116 | b3 | 43.1642 | -3.3793 | -2.9041 | -3.0933 | -2.4372 | -4.0362 | -2.0654 | -2.1566 | -1.9896 | -2.7668 | -3.5 | -2.6188 | -2.3761 | -5.5033 | -3.73 | -2.5635 | -1.7019 | -2.2789 | -2.2575 | -2.4028 | -1.4602
b4 | -1.0714 | -0.078132 | -0.024812 | -0.14404 | -0.081079 | 0.24346 | 0.071604 | 0.14944 | 0.074763 | -0.093276 | 0.10957 | -0.086739 | 0.03965 | -0.024941 | 0.11394 | 0.18676 | -0.046898 | 0.035293 | 0.17431 | -0.070103 | -0.0024857 | b4 | -0.97049 | 0.13888 | 0.0061742 | -0.15232 | 0.082642 | 0.016766 | -0.064104 | -0.024922 | -0.014591 | 0.12703 | -0.022578 | -0.15721 | -0.096666 | 0.091333 | -0.036992 | -0.15409 | 0.025543 | -0.079755 | -0.053826 | 0.028473 | -0.047373
b5 | -5.0911 | -0.79674 | -0.96037 | -0.58735 | -0.54451 | -0.46823 | -0.36287 | -0.3178 | -0.36989 | -0.34812 | -0.34956 | -0.52068 | -0.63445 | -0.64063 | -0.58711 | -0.56926 | -0.72223 | -0.62219 | -0.66927 | -4.0119 | -0.80801 | b5 | -4.9043 | -1.0724 | -1.1203 | -0.74684 | -0.68606 | -0.67954 | -0.67359 | -0.66537 | -0.64939 | -0.61001 | -0.54112 | -0.81747 | -0.83932 | -0.54571 | -0.88771 | -1.0214 | -1.0735 | -0.94193 | -1.046 | -4.5695 | -1.055
b6 | 0.072293 | 0.11328 | 0.05959 | 0.12184 | 0.088944 | 0.013987 | -0.035969 | 0.025578 | -0.049623 | 0.020738 | -0.025678 | 0.079801 | 0.010609 | 0.0021035 | -0.022527 | -0.075818 | -0.0040077 | 0.068068 | 0.10119 | -0.0030074 | 0.027727 | b6 | -0.077281 | 0.089733 | 0.044933 | 0.029223 | -0.22551 | -0.087319 | 0.015226 | 0.042898 | 0.068093 | -0.058519 | -0.024781 | -0.088679 | -0.051166 | -0.14177 | -0.094673 | 0.0058462 | -0.010193 | 0.00774 | -0.019455 | -0.10951 | -0.031792
b7 | 2.8928 | 0.97237 | 1.0561 | 1.1794 | 1.1515 | 1.2143 | 1.1342 | 1.1499 | 1.0877 | 1.1211 | 1.1187 | 1.0922 | 1.0516 | 1.1327 | 1.1193 | 1.1033 | 1.093 | 1.1471 | 1.1441 | 0.88657 | 1.1614 | b7 | 2.9901 | 0.95318 | 1.0077 | 1.1337 | 1.1414 | 1.2082 | 1.2029 | 1.1998 | 1.1587 | 1.1644 | 1.257 | 1.2004 | 1.1289 | 1.1658 | 1.1103 | 1.1253 | 1.1268 | 1.1628 | 1.1381 | 0.84754 | 1.1895
b8 | 0.058757 | 0.061356 | 0.016798 | -0.0066228 | 0.0071371 | -0.009797 | -0.029018 | -0.002059 | -0.023003 | -0.018657 | -0.0341 | -0.0066275 | 0.027758 | -0.027606 | -0.016613 | -0.020609 | 0.0080487 | 0.032607 | 0.006028 | -0.016947 | -0.00058567 | b8 | -0.041611 | 0.053159 | -0.013194 | -0.042999 | 0.0071099 | -0.040628 | -0.011714 | 0.025825 | -0.020651 | -0.0024872 | 0.013843 | 0.055959 | 0.023498 | -0.016627 | -0.02914 | -0.0026356 | -0.0072239 | 0.0098383 | 0.042852 | -0.055548 | -0.00068897
b9 | 0.32723 | 0.47545 | 0.46555 | 0.45909 | 0.46995 | 0.45423 | 0.46506 | 0.44524 | 0.47406 | 0.45094 | 0.45456 | 0.44575 | 0.45711 | 0.4607 | 0.44378 | 0.45485 | 0.45259 | 0.47874 | 0.47124 | 0.41336 | 0.45442 | b9 | 0.37635 | 0.50529 | 0.48032 | 0.45231 | 0.46536 | 0.44662 | 0.49912 | 0.48459 | 0.50009 | 0.4676 | 0.45664 | 0.47065 | 0.4627 | 0.46463 | 0.45923 | 0.50178 | 0.48165 | 0.48233 | 0.4786 | 0.40526 | 0.47031
b10 | 0.080641 | 0.084551 | 0.025326 | 0.014366 | 0.0089201 | -0.023553 | -0.068181 | -0.0092361 | -0.031329 | -0.041129 | -0.031943 | 0.017222 | -0.00078076 | -0.04567 | -0.041341 | -0.0047735 | 0.019198 | 0.030961 | 0.041057 | 0.01513 | 0 | b10 | 0.061494 | 0.10923 | -0.015031 | -0.032632 | -0.032116 | -0.053645 | -0.026686 | 0.037691 | 0.0019233 | 9.2058e-05 | -0.0065985 | 0.047246 | 0.02095 | -0.018174 | -0.034082 | -0.011732 | -0.014986 | -0.011293 | 0.038922 | -0.056143 | 0
b11 | 0.60331 | 0.72782 | 0.71922 | 0.73006 | 0.72509 | 0.72938 | 0.7298 | 0.73493 | 0.73091 | 0.73134 | 0.73838 | 0.73064 | 0.73332 | 0.73311 | 0.73674 | 0.7343 | 0.72509 | 0.73346 | 0.72064 | 0.63234 | 0.72324 | b11 | 0.59185 | 0.74142 | 0.73814 | 0.74493 | 0.74363 | 0.7455 | 0.73792 | 0.74178 | 0.75022 | 0.74272 | 0.75077 | 0.744 | 0.74537 | 0.75873 | 0.7498 | 0.74135 | 0.73851 | 0.73913 | 0.72976 | 0.63944 | 0.73566
b12 | 0.015898 | 0.0060942 | 0.00025751 | -0.0024787 | -0.0020657 | -0.0041331 | -0.011407 | -0.00401 | -0.0063919 | -0.006917 | -0.0078165 | 0.001004 | -0.0016925 | -0.0083656 | -0.0070234 | -0.0037463 | 0.0028113 | 0.0036637 | 0.0024361 | 0.0014948 | -0.0020625 | b12 | 0.013211 | 0.0084785 | -0.0015097 | 0.00010991 | -0.0038479 | -0.0027947 | -0.0014481 | 0.003893 | 0.001312 | -0.0014268 | 0.00037714 | 0.006864 | 0.0039592 | -0.0019352 | -0.00041294 | 0.00618 | 0.0014791 | 0.0026282 | 0.0056913 | -0.001132 | 0.0017979
b13 | 0.085482 | 0.067595 | 0.069082 | 0.064198 | 0.062976 | 0.063402 | 0.061748 | 0.061963 | 0.061263 | 0.061292 | 0.065667 | 0.063705 | 0.065837 | 0.06487 | 0.066341 | 0.06509 | 0.06557 | 0.065797 | 0.06653 | 0.050473 | 0.064784 | b13 | 0.088368 | 0.07129 | 0.070721 | 0.06836 | 0.06943 | 0.069892 | 0.065852 | 0.070224 | 0.069463 | 0.06842 | 0.072321 | 0.069589 | 0.070863 | 0.066028 | 0.069586 | 0.065899 | 0.070174 | 0.07146 | 0.071168 | 0.054657 | 0.069576
b14 | 0.0061579 | 0.001936 | -0.0018942 | -0.0040846 | -0.0032018 | -0.0054799 | -0.0095148 | -0.0047974 | -0.0054024 | -0.0067173 | -0.0058688 | -0.00075766 | -0.0042221 | -0.007324 | -0.0069293 | -0.0050052 | -0.0023677 | -0.0019368 | 0.00016237 | 0.00060947 | -0.0036167 | b14 | 0.0061101 | 0.00098025 | -0.0053126 | -0.0044213 | -0.0063618 | -0.0078668 | -0.0049425 | -0.0012305 | -0.0037014 | -0.004669 | -0.0049355 | -0.0024232 | -0.0038712 | -0.0063849 | -0.0057997 | -0.0040434 | -0.0043398 | -0.0052129 | -0.002597 | -0.0043326 | -0.0039751
b15 | -0.0016122 | 0.045799 | 0.039108 | 0.035608 | 0.034536 | 0.036901 | 0.038632 | 0.038458 | 0.038004 | 0.038734 | 0.035656 | 0.037404 | 0.038946 | 0.038983 | 0.036175 | 0.03775 | 0.034582 | 0.042708 | 0.035034 | 0.021605 | 0.036221 | b15 | -0.0065662 | 0.031097 | 0.027985 | 0.025569 | 0.024061 | 0.024293 | 0.023366 | 0.024261 | 0.024701 | 0.024205 | 0.026327 | 0.024469 | 0.028317 | 0.02679 | 0.026457 | 0.027967 | 0.028094 | 0.030965 | 0.028901 | 0.026174 | 0.025542
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 5.6889 | -13.0161 | -6.1493 | -2.6193 | -2.6792 | -1.5142 | 3.7705 | 8.1166 | 5.8944 | 0.99261 | -4.2283 | -1.4159 | 3.3508 | 7.6688 | 4.9306 | -0.30798 | -3.7643 | -0.3561 | 0.67249 | -4.2877 | 0 | a1 | 29.1057 | -15.2001 | -5.8988 | -2.7883 | -2.6065 | -4.8182 | 2.319 | 8.5468 | 6.4318 | 1.6269 | -3.5423 | -0.15795 | 1.3168 | 5.3317 | 4.3703 | -0.69097 | -5.1868 | -1.6551 | 1.072 | -8.4629 | 0
a2 | 10.5206 | 1.7057 | 0.93659 | 0.37049 | 0.27757 | 0.73803 | 0.96357 | 0.7958 | 1.1997 | 0.5963 | 0.63032 | 0.15145 | 0.92424 | 1.1352 | 0.84983 | 1.4834 | 1.3928 | 1.0697 | -0.031404 | -4.3018 | 0.97139 | a2 | 5.9193 | -0.24467 | 0.12946 | -0.053691 | 0.46639 | -0.50087 | -1.9214 | -0.92363 | -1.3545 | -0.59541 | -1.1561 | -1.7077 | -0.73709 | -1.0268 | -0.28547 | -0.3406 | -0.67328 | -0.27022 | -1.4299 | 0.85576 | -0.45325
a3 | 0.0076977 | 0.82155 | 0.36295 | 0.50349 | 0.25271 | 0.51044 | 0.23253 | -0.13802 | 0.50021 | 0.097894 | 0.19165 | 0.57794 | 0.14629 | 0.15492 | 0.07936 | 0.39241 | 0.59799 | 0.37933 | 0.61059 | -0.19691 | 0.32069 | a3 | 0.82328 | 0.30697 | 0.25704 | 0.067993 | -0.011124 | -0.22219 | -0.0025971 | -0.12012 | 0.28968 | -0.092234 | 0.63123 | 0.49111 | -0.017039 | -0.12892 | -0.027902 | 0.14078 | 0.42063 | 0.16125 | 0.54665 | 0.18452 | 0.17083
a4 | 2.6653 | 0.011158 | 0.15395 | 0.081577 | -0.082052 | -0.074651 | -0.45694 | 0.021311 | -0.034296 | 0.18943 | -0.28147 | -0.34061 | -0.32947 | -0.30358 | -0.077585 | -0.23753 | -0.32031 | -0.3659 | 0.085648 | -0.57401 | -0.060697 | a4 | 1.6229 | -0.064452 | -0.35188 | -0.44451 | -0.12957 | -0.015753 | 0.3582 | 0.2493 | 0.35242 | -0.12207 | 0.2131 | -0.25202 | -0.38459 | -0.10238 | 0.0026689 | 0.38255 | -0.012594 | -0.081201 | 0.08072 | 0.47115 | 0.046925
a5 | 2.9268 | -0.15787 | 0.041543 | 0.0071212 | -0.080298 | -0.3487 | -0.14973 | -0.22101 | 0.014729 | 0.006077 | 0.018293 | 0.21265 | -0.17117 | -0.03975 | -0.075647 | -0.046405 | 0.0048629 | -0.04533 | 0.19236 | -0.048116 | 0.043485 | a5 | 2.7125 | 0.0080802 | -0.10832 | 0.05403 | -0.014349 | -0.062308 | -0.1863 | -0.013429 | 0.06637 | 0.060949 | 0.17848 | -0.029735 | -0.18703 | -0.048096 | -0.20715 | -0.025955 | -0.063658 | -0.035455 | 0.21959 | -0.15881 | 0.055638
a6 | 0.61022 | -0.1464 | -0.058901 | -0.023489 | 0.076281 | 0.05489 | 0.015145 | -0.010718 | 0.0093195 | 0.057886 | 0.055672 | 0.0076411 | -0.039183 | -0.064788 | -0.059633 | -0.10784 | 0.012529 | -0.026241 | 0.054683 | -0.098417 | 0.0052611 | a6 | 0.34408 | -0.055972 | -0.073239 | 0.022242 | 0.0038866 | 0.068744 | 0.045172 | 0.067528 | -0.03353 | 0.052269 | -0.069984 | -0.027572 | -0.15226 | -0.20616 | -0.1519 | -0.079694 | -0.072512 | -0.060058 | -0.0039164 | -0.23156 | -0.034989
a7 | 1.5255 | -0.078735 | 0.0095776 | -0.01517 | -0.054016 | -0.084932 | -0.068148 | -0.031631 | 0.049949 | -0.020549 | -0.0034651 | -0.0058345 | -0.023093 | 0.026273 | -0.013593 | -0.046601 | 0.026061 | 0.06983 | 0.043195 | 0.0048096 | 0.034877 | a7 | 1.8027 | 0.016234 | 0.011981 | -0.023333 | 0.052282 | -0.07567 | 0.024604 | 0.10247 | 0.056719 | 0.064939 | 0.010511 | 0.034364 | -0.023953 | 0.05193 | 0.054132 | 0.055947 | 0.067486 | 0.03057 | 0.070253 | 0.11616 | 0.088027
a8 | 0.055266 | 0.019421 | 0.0030897 | 0.0037123 | 0.01907 | -0.012151 | 0.0035421 | -0.036046 | 0.0087662 | 0.0072948 | 0.036917 | 0.0048588 | -0.0021397 | 0.014131 | 0.031133 | 0.0052184 | 0.023526 | -0.01327 | 0.014254 | -0.022263 | 0.0077983 | a8 | 0.0068092 | 0.0028122 | 0.0018472 | 0.016562 | 0.053297 | 0.039363 | 0.0091324 | 0.01691 | -0.028641 | 0.0054098 | -0.028724 | -0.036398 | -0.0446 | -0.0387 | -0.080386 | -0.027869 | -0.026065 | 0.0049402 | 0.010717 | -0.080387 | -0.010207
a9 | -0.28063 | -0.0058266 | -0.013064 | -0.023784 | -0.08105 | -0.058142 | -0.032889 | -0.011612 | -0.0023489 | -0.057172 | -0.017965 | 0.01176 | -0.027013 | -0.03921 | -0.03975 | -0.0089429 | -0.0095819 | -0.005983 | -0.012612 | -0.07686 | -0.033643 | a9 | -0.23319 | 0.010005 | 0.0035402 | -0.0296 | -0.07907 | -0.052721 | -0.0023116 | 0.026467 | 0.024176 | -0.027105 | 0.014646 | 0.011942 | -0.0083854 | -0.023221 | -0.044807 | 0.021957 | 0.0090242 | -0.0044896 | -0.0029833 | -0.038955 | -0.016202
a10 | 0.063711 | -0.032753 | 0.0057382 | 0.014657 | 0.027809 | 0.0018541 | -0.0077407 | -0.016522 | 0.0041969 | 0.00464 | 0.012123 | 0.0027942 | -0.012187 | -0.021415 | 0.0033025 | -0.010692 | 0.02043 | -0.039393 | 0.024596 | -0.029928 | 0 | a10 | -0.0018558 | 0.0028519 | 0.0016651 | -0.0064852 | 0.024377 | 0.003438 | 0.041627 | 0.054425 | 0.040892 | 0.025366 | -0.014948 | 0.022685 | -0.031911 | -0.015479 | -0.027112 | 0.0066411 | -0.014825 | -0.021816 | 0.0069687 | -0.15838 | 0
a11 | 0.13231 | -0.073174 | -0.051664 | -0.049438 | -0.059835 | -0.065915 | -0.055549 | -0.046886 | -0.041002 | -0.048017 | -0.044523 | -0.04183 | -0.049954 | -0.041569 | -0.045457 | -0.052822 | -0.045616 | -0.038711 | -0.039197 | -0.033266 | -0.043476 | a11 | 0.18124 | -0.020849 | -0.0098851 | 0.0017307 | -0.0055119 | -0.011416 | -0.0010328 | 0.011259 | 0.016222 | -0.0013113 | 0.0012689 | 0.0051143 | -0.0046658 | 0.0038737 | -0.0072968 | 0.0022111 | 0.0070256 | 0.0041006 | 0.014622 | 0.020786 | 0.0063797
a12 | 0.0046056 | -0.008269 | -0.0018375 | -0.0053365 | -8.8525e-05 | -0.0018001 | -0.0055462 | -0.011051 | -0.0057585 | -0.005292 | -3.5046e-05 | -0.0053279 | -0.0082954 | -0.0062235 | -0.0040964 | -0.0056434 | -0.0026605 | -0.0084674 | 0.0016219 | -0.006609 | -0.0044537 | a12 | -8.7616e-05 | 0.00032488 | -0.0014258 | -0.0053288 | -0.00029301 | -0.0017505 | -0.00054588 | 0.0052155 | 0.00022144 | 0.0037172 | -0.0052702 | -0.0011046 | -0.011988 | -0.011923 | -0.0090435 | -0.004568 | -0.004253 | -0.0053314 | 0.00020139 | -0.016154 | -0.0032892
a13 | -0.031103 | -0.0021964 | -0.0024337 | -0.0050231 | -0.0084853 | -0.0083208 | -0.0041702 | -0.0056166 | -0.0011549 | -0.0043912 | -0.0036131 | -0.0012427 | -0.005785 | -0.0034524 | -0.0048589 | -0.003974 | -0.0034775 | -0.0025269 | -0.0013607 | -0.0055844 | -0.0048765 | a13 | -0.025526 | 0.00058949 | 0.0020518 | -0.0024824 | -0.00076297 | -0.0048622 | 0.00030624 | 0.004145 | 0.0023496 | -0.00023395 | 0.0019164 | 0.0016598 | -0.0015762 | -0.00021846 | -0.0001451 | 0.0023059 | 0.0033661 | 0.0019733 | 0.0032866 | -0.00066543 | -7.3625e-05
a14 | 0.013285 | -0.019474 | -0.016738 | -0.016244 | -0.014874 | -0.018009 | -0.02067 | -0.021153 | -0.019139 | -0.017407 | -0.016108 | -0.017891 | -0.019324 | -0.018882 | -0.015757 | -0.017194 | -0.0164 | -0.020232 | -0.015411 | -0.010144 | -0.016638 | a14 | 0.012908 | -0.010653 | -0.010509 | -0.010642 | -0.0088607 | -0.0087396 | -0.0091119 | -0.0078789 | -0.0090887 | -0.0086423 | -0.0104 | -0.010673 | -0.013585 | -0.011391 | -0.012184 | -0.011585 | -0.011851 | -0.012423 | -0.010532 | -0.0056529 | -0.0096305
a15 | 0.0074049 | 0.0028675 | -0.0049734 | -0.0053651 | -0.0061557 | -0.0088651 | -0.01463 | -0.007666 | -0.0098426 | -0.011645 | -0.010608 | -0.0040476 | -0.0070853 | -0.011699 | -0.012522 | -0.0082892 | -0.0043187 | -0.0032381 | -0.0013988 | -0.0015182 | -0.0065679 | a15 | 0.00055144 | -0.00077071 | -0.0090575 | -0.010502 | -0.010245 | -0.013038 | -0.008776 | -0.0028242 | -0.0055969 | -0.0071642 | -0.0080299 | -0.00302 | -0.0062202 | -0.010735 | -0.010891 | -0.0071202 | -0.0078331 | -0.0081412 | -0.0035154 | -0.017403 | -0.0074964
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00025458 | -0.00020312 | -5.8588e-05 | -3.1006e-05 | -1.5434e-05 | 5.4843e-05 | 0.00015661 | 1.8841e-05 | 7.3199e-05 | 9.5923e-05 | 7.4971e-05 | -3.9572e-05 | 0 | 0.00010275 | 9.5521e-05 | 0 | -4.1837e-05 | -7.6324e-05 | -9.342e-05 | -4.4762e-05 | 7.2543e-05 | Dx (m) | -0.00016011 | -0.00025019 | 3.4662e-05 | 7.4496e-05 | 7.3836e-05 | 0.00012243 | 6.1623e-05 | -8.8216e-05 | 0 | 0 | 1.4997e-05 | -0.00010834 | -4.8199e-05 | 4.0883e-05 | 7.6634e-05 | 2.6857e-05 | 3.4813e-05 | 2.6249e-05 | -9.0974e-05 | 0.00016209 | 0.0002183
Dy (m) | -0.00012413 | 5.6235e-05 | -1.7768e-05 | -3.6227e-05 | -6.6473e-05 | 0 | 3.0044e-05 | 3.9417e-05 | 0 | 0 | -2.3401e-05 | 0 | 2.8242e-05 | 5.5538e-05 | 0 | 2.5415e-05 | -5.0535e-05 | 8.7285e-05 | -6.3111e-05 | 7.8095e-05 | 0.00086166 | Dy (m) | 5.4433e-05 | -1.3589e-05 | 0 | 1.4637e-05 | -5.519e-05 | 0 | -9.582e-05 | -0.00012341 | -9.253e-05 | -5.8057e-05 | 3.3819e-05 | -5.1105e-05 | 7.2476e-05 | 3.4478e-05 | 6.222e-05 | -1.5305e-05 | 3.3794e-05 | 5.0023e-05 | -1.4415e-05 | 0.00041492 | 0.00067491
Magnetic length (m) |  | 14.419 | [19 empty] | Magnetic length (m) |  | 14.4199
Transfer funct. (mT/KA) |  | 596.0109 | [19 empty] | Transfer funct. (mT/KA) |  | 596.0506
Coil positioning (mm) |  | 0.015163 | [19 empty] | Coil positioning (mm) |  | 0.018692
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000095 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000095 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 346.7788 | 596.2757 | 595.8891 | 595.9106 | 595.9598 | 595.9268 | 596.0341 | 596.0505 | 595.9777 | 595.9929 | 596.0165 | 595.9618 | 595.8817 | 595.9201 | 595.9721 | 596.064 | 596.0254 | 595.9977 | 596.3394 | 368.4831 | 85.8143 | C1/i (mT/KA) | 342.6644 | 596.2368 | 595.9281 | 595.988 | 596.0347 | 596.0313 | 596.0376 | 595.9807 | 596.0577 | 596.0118 | 596.0358 | 596.0818 | 596.0058 | 596.0738 | 596.0328 | 596.0215 | 596.0986 | 595.9477 | 596.3061 | 366.1748 | 85.7714
Angle (mrad) | 0.56889 | -1.3016 | -0.61493 | -0.26193 | -0.26792 | -0.15142 | 0.37705 | 0.81166 | 0.58944 | 0.099261 | -0.42283 | -0.14159 | 0.33508 | 0.76688 | 0.49306 | -0.030798 | -0.37643 | -0.03561 | 0.067249 | -0.42877 | -14.1837 | Angle (mrad) | 2.9106 | -1.52 | -0.58988 | -0.27883 | -0.26065 | -0.48182 | 0.2319 | 0.85468 | 0.64319 | 0.16269 | -0.35423 | -0.015795 | 0.13168 | 0.53317 | 0.43703 | -0.069097 | -0.51868 | -0.16551 | 0.1072 | -0.84629 | -33.8621
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.2917 | -0.46018 | 0.12788 | -0.62341 | -0.052965 | 0.34279 | -0.028677 | -0.25491 | 0.39955 | 0.30647 | 0.13438 | -0.22924 | 0.12585 | 0.086515 | 0.50903 | 0.67702 | 0.7392 | 0.29999 | 0.75049 | 0.41568 | 0.092379 | b2 | -1.1213 | -0.86106 | -0.14374 | -0.41653 | 0.85797 | -0.29428 | -0.12563 | 0.28218 | 0.014852 | 0.25132 | 0.024027 | -0.013122 | -0.71149 | 0.016058 | 0.14113 | 0.20541 | -0.085872 | 0.17708 | -0.66882 | -0.45903 | -0.1187
b3 | 42.8096 | -3.3938 | -2.7015 | -3.1143 | -2.1666 | -3.2502 | -2.2506 | -3.0746 | -1.9807 | -2.99 | -2.9643 | -3.37 | -2.8422 | -4.2203 | -3.8523 | -3.4074 | -2.4091 | -2.731 | -1.835 | -2.2261 | -1.5116 | b3 | 43.1642 | -3.3793 | -2.9041 | -3.0933 | -2.4372 | -4.0362 | -2.0654 | -2.1566 | -1.9896 | -2.7668 | -3.5 | -2.6188 | -2.3761 | -5.5033 | -3.73 | -2.5635 | -1.7019 | -2.2789 | -2.2575 | -2.4028 | -1.4602
b4 | -1.0714 | -0.078132 | -0.024812 | -0.14404 | -0.081079 | 0.24346 | 0.071604 | 0.14944 | 0.074763 | -0.093276 | 0.10957 | -0.086739 | 0.03965 | -0.024941 | 0.11394 | 0.18676 | -0.046898 | 0.035293 | 0.17431 | -0.070103 | -0.0024857 | b4 | -0.97049 | 0.13888 | 0.0061742 | -0.15232 | 0.082642 | 0.016766 | -0.064104 | -0.024922 | -0.014591 | 0.12703 | -0.022578 | -0.15721 | -0.096666 | 0.091333 | -0.036992 | -0.15409 | 0.025543 | -0.079755 | -0.053826 | 0.028473 | -0.047373
b5 | -5.0911 | -0.79674 | -0.96037 | -0.58735 | -0.54451 | -0.46823 | -0.36287 | -0.3178 | -0.36989 | -0.34812 | -0.34956 | -0.52068 | -0.63445 | -0.64063 | -0.58711 | -0.56926 | -0.72223 | -0.62219 | -0.66927 | -4.0119 | -0.80801 | b5 | -4.9043 | -1.0724 | -1.1203 | -0.74684 | -0.68606 | -0.67954 | -0.67359 | -0.66537 | -0.64939 | -0.61001 | -0.54112 | -0.81747 | -0.83932 | -0.54571 | -0.88771 | -1.0214 | -1.0735 | -0.94193 | -1.046 | -4.5695 | -1.055
b6 | 0.072293 | 0.11328 | 0.05959 | 0.12184 | 0.088944 | 0.013987 | -0.035969 | 0.025578 | -0.049623 | 0.020738 | -0.025678 | 0.079801 | 0.010609 | 0.0021035 | -0.022527 | -0.075818 | -0.0040077 | 0.068068 | 0.10119 | -0.0030074 | 0.027727 | b6 | -0.077281 | 0.089733 | 0.044933 | 0.029223 | -0.22551 | -0.087319 | 0.015226 | 0.042898 | 0.068093 | -0.058519 | -0.024781 | -0.088679 | -0.051166 | -0.14177 | -0.094673 | 0.0058462 | -0.010193 | 0.00774 | -0.019455 | -0.10951 | -0.031792
b7 | 2.8928 | 0.97237 | 1.0561 | 1.1794 | 1.1515 | 1.2143 | 1.1342 | 1.1499 | 1.0877 | 1.1211 | 1.1187 | 1.0922 | 1.0516 | 1.1327 | 1.1193 | 1.1033 | 1.093 | 1.1471 | 1.1441 | 0.88657 | 1.1614 | b7 | 2.9901 | 0.95318 | 1.0077 | 1.1337 | 1.1414 | 1.2082 | 1.2029 | 1.1998 | 1.1587 | 1.1644 | 1.257 | 1.2004 | 1.1289 | 1.1658 | 1.1103 | 1.1253 | 1.1268 | 1.1628 | 1.1381 | 0.84754 | 1.1895
b8 | 0.058757 | 0.061356 | 0.016798 | -0.0066228 | 0.0071371 | -0.009797 | -0.029018 | -0.002059 | -0.023003 | -0.018657 | -0.0341 | -0.0066275 | 0.027758 | -0.027606 | -0.016613 | -0.020609 | 0.0080487 | 0.032607 | 0.006028 | -0.016947 | -0.00058567 | b8 | -0.041611 | 0.053159 | -0.013194 | -0.042999 | 0.0071099 | -0.040628 | -0.011714 | 0.025825 | -0.020651 | -0.0024872 | 0.013843 | 0.055959 | 0.023498 | -0.016627 | -0.02914 | -0.0026356 | -0.0072239 | 0.0098383 | 0.042852 | -0.055548 | -0.00068897
b9 | 0.32723 | 0.47545 | 0.46555 | 0.45909 | 0.46995 | 0.45423 | 0.46506 | 0.44524 | 0.47406 | 0.45094 | 0.45456 | 0.44575 | 0.45711 | 0.4607 | 0.44378 | 0.45485 | 0.45259 | 0.47874 | 0.47124 | 0.41336 | 0.45442 | b9 | 0.37635 | 0.50529 | 0.48032 | 0.45231 | 0.46536 | 0.44662 | 0.49912 | 0.48459 | 0.50009 | 0.4676 | 0.45664 | 0.47065 | 0.4627 | 0.46463 | 0.45923 | 0.50178 | 0.48165 | 0.48233 | 0.4786 | 0.40526 | 0.47031
b10 | 0.080641 | 0.084551 | 0.025326 | 0.014366 | 0.0089201 | -0.023553 | -0.068181 | -0.0092361 | -0.031329 | -0.041129 | -0.031943 | 0.017222 | -0.00078076 | -0.04567 | -0.041341 | -0.0047735 | 0.019198 | 0.030961 | 0.041057 | 0.01513 | 0 | b10 | 0.061494 | 0.10923 | -0.015031 | -0.032632 | -0.032116 | -0.053645 | -0.026686 | 0.037691 | 0.0019233 | 9.2058e-05 | -0.0065985 | 0.047246 | 0.02095 | -0.018174 | -0.034082 | -0.011732 | -0.014986 | -0.011293 | 0.038922 | -0.056143 | 0
b11 | 0.60331 | 0.72782 | 0.71922 | 0.73006 | 0.72509 | 0.72938 | 0.7298 | 0.73493 | 0.73091 | 0.73134 | 0.73838 | 0.73064 | 0.73332 | 0.73311 | 0.73674 | 0.7343 | 0.72509 | 0.73346 | 0.72064 | 0.63234 | 0.72324 | b11 | 0.59185 | 0.74142 | 0.73814 | 0.74493 | 0.74363 | 0.7455 | 0.73792 | 0.74178 | 0.75022 | 0.74272 | 0.75077 | 0.744 | 0.74537 | 0.75873 | 0.7498 | 0.74135 | 0.73851 | 0.73913 | 0.72976 | 0.63944 | 0.73566
b12 | 0.015898 | 0.0060942 | 0.00025751 | -0.0024787 | -0.0020657 | -0.0041331 | -0.011407 | -0.00401 | -0.0063919 | -0.006917 | -0.0078165 | 0.001004 | -0.0016925 | -0.0083656 | -0.0070234 | -0.0037463 | 0.0028113 | 0.0036637 | 0.0024361 | 0.0014948 | -0.0020625 | b12 | 0.013211 | 0.0084785 | -0.0015097 | 0.00010991 | -0.0038479 | -0.0027947 | -0.0014481 | 0.003893 | 0.001312 | -0.0014268 | 0.00037714 | 0.006864 | 0.0039592 | -0.0019352 | -0.00041294 | 0.00618 | 0.0014791 | 0.0026282 | 0.0056913 | -0.001132 | 0.0017979
b13 | 0.085482 | 0.067595 | 0.069082 | 0.064198 | 0.062976 | 0.063402 | 0.061748 | 0.061963 | 0.061263 | 0.061292 | 0.065667 | 0.063705 | 0.065837 | 0.06487 | 0.066341 | 0.06509 | 0.06557 | 0.065797 | 0.06653 | 0.050473 | 0.064784 | b13 | 0.088368 | 0.07129 | 0.070721 | 0.06836 | 0.06943 | 0.069892 | 0.065852 | 0.070224 | 0.069463 | 0.06842 | 0.072321 | 0.069589 | 0.070863 | 0.066028 | 0.069586 | 0.065899 | 0.070174 | 0.07146 | 0.071168 | 0.054657 | 0.069576
b14 | 0.0061579 | 0.001936 | -0.0018942 | -0.0040846 | -0.0032018 | -0.0054799 | -0.0095148 | -0.0047974 | -0.0054024 | -0.0067173 | -0.0058688 | -0.00075766 | -0.0042221 | -0.007324 | -0.0069293 | -0.0050052 | -0.0023677 | -0.0019368 | 0.00016237 | 0.00060947 | -0.0036167 | b14 | 0.0061101 | 0.00098025 | -0.0053126 | -0.0044213 | -0.0063618 | -0.0078668 | -0.0049425 | -0.0012305 | -0.0037014 | -0.004669 | -0.0049355 | -0.0024232 | -0.0038712 | -0.0063849 | -0.0057997 | -0.0040434 | -0.0043398 | -0.0052129 | -0.002597 | -0.0043326 | -0.0039751
b15 | -0.0016122 | 0.045799 | 0.039108 | 0.035608 | 0.034536 | 0.036901 | 0.038632 | 0.038458 | 0.038004 | 0.038734 | 0.035656 | 0.037404 | 0.038946 | 0.038983 | 0.036175 | 0.03775 | 0.034582 | 0.042708 | 0.035034 | 0.021605 | 0.036221 | b15 | -0.0065662 | 0.031097 | 0.027985 | 0.025569 | 0.024061 | 0.024293 | 0.023366 | 0.024261 | 0.024701 | 0.024205 | 0.026327 | 0.024469 | 0.028317 | 0.02679 | 0.026457 | 0.027967 | 0.028094 | 0.030965 | 0.028901 | 0.026174 | 0.025542
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 5.6889 | -13.0161 | -6.1493 | -2.6193 | -2.6792 | -1.5142 | 3.7705 | 8.1166 | 5.8944 | 0.99261 | -4.2283 | -1.4159 | 3.3508 | 7.6688 | 4.9306 | -0.30798 | -3.7643 | -0.3561 | 0.67249 | -4.2877 | 0 | a1 | 29.1057 | -15.2001 | -5.8988 | -2.7883 | -2.6065 | -4.8182 | 2.319 | 8.5468 | 6.4318 | 1.6269 | -3.5423 | -0.15795 | 1.3168 | 5.3317 | 4.3703 | -0.69097 | -5.1868 | -1.6551 | 1.072 | -8.4629 | 0
a2 | 10.5206 | 1.7057 | 0.93659 | 0.37049 | 0.27757 | 0.73803 | 0.96357 | 0.7958 | 1.1997 | 0.5963 | 0.63032 | 0.15145 | 0.92424 | 1.1352 | 0.84983 | 1.4834 | 1.3928 | 1.0697 | -0.031404 | -4.3018 | 0.97139 | a2 | 5.9193 | -0.24467 | 0.12946 | -0.053691 | 0.46639 | -0.50087 | -1.9214 | -0.92363 | -1.3545 | -0.59541 | -1.1561 | -1.7077 | -0.73709 | -1.0268 | -0.28547 | -0.3406 | -0.67328 | -0.27022 | -1.4299 | 0.85576 | -0.45325
a3 | 0.0076977 | 0.82155 | 0.36295 | 0.50349 | 0.25271 | 0.51044 | 0.23253 | -0.13802 | 0.50021 | 0.097894 | 0.19165 | 0.57794 | 0.14629 | 0.15492 | 0.07936 | 0.39241 | 0.59799 | 0.37933 | 0.61059 | -0.19691 | 0.32069 | a3 | 0.82328 | 0.30697 | 0.25704 | 0.067993 | -0.011124 | -0.22219 | -0.0025971 | -0.12012 | 0.28968 | -0.092234 | 0.63123 | 0.49111 | -0.017039 | -0.12892 | -0.027902 | 0.14078 | 0.42063 | 0.16125 | 0.54665 | 0.18452 | 0.17083
a4 | 2.6653 | 0.011158 | 0.15395 | 0.081577 | -0.082052 | -0.074651 | -0.45694 | 0.021311 | -0.034296 | 0.18943 | -0.28147 | -0.34061 | -0.32947 | -0.30358 | -0.077585 | -0.23753 | -0.32031 | -0.3659 | 0.085648 | -0.57401 | -0.060697 | a4 | 1.6229 | -0.064452 | -0.35188 | -0.44451 | -0.12957 | -0.015753 | 0.3582 | 0.2493 | 0.35242 | -0.12207 | 0.2131 | -0.25202 | -0.38459 | -0.10238 | 0.0026689 | 0.38255 | -0.012594 | -0.081201 | 0.08072 | 0.47115 | 0.046925
a5 | 2.9268 | -0.15787 | 0.041543 | 0.0071212 | -0.080298 | -0.3487 | -0.14973 | -0.22101 | 0.014729 | 0.006077 | 0.018293 | 0.21265 | -0.17117 | -0.03975 | -0.075647 | -0.046405 | 0.0048629 | -0.04533 | 0.19236 | -0.048116 | 0.043485 | a5 | 2.7125 | 0.0080802 | -0.10832 | 0.05403 | -0.014349 | -0.062308 | -0.1863 | -0.013429 | 0.06637 | 0.060949 | 0.17848 | -0.029735 | -0.18703 | -0.048096 | -0.20715 | -0.025955 | -0.063658 | -0.035455 | 0.21959 | -0.15881 | 0.055638
a6 | 0.61022 | -0.1464 | -0.058901 | -0.023489 | 0.076281 | 0.05489 | 0.015145 | -0.010718 | 0.0093195 | 0.057886 | 0.055672 | 0.0076411 | -0.039183 | -0.064788 | -0.059633 | -0.10784 | 0.012529 | -0.026241 | 0.054683 | -0.098417 | 0.0052611 | a6 | 0.34408 | -0.055972 | -0.073239 | 0.022242 | 0.0038866 | 0.068744 | 0.045172 | 0.067528 | -0.03353 | 0.052269 | -0.069984 | -0.027572 | -0.15226 | -0.20616 | -0.1519 | -0.079694 | -0.072512 | -0.060058 | -0.0039164 | -0.23156 | -0.034989
a7 | 1.5255 | -0.078735 | 0.0095776 | -0.01517 | -0.054016 | -0.084932 | -0.068148 | -0.031631 | 0.049949 | -0.020549 | -0.0034651 | -0.0058345 | -0.023093 | 0.026273 | -0.013593 | -0.046601 | 0.026061 | 0.06983 | 0.043195 | 0.0048096 | 0.034877 | a7 | 1.8027 | 0.016234 | 0.011981 | -0.023333 | 0.052282 | -0.07567 | 0.024604 | 0.10247 | 0.056719 | 0.064939 | 0.010511 | 0.034364 | -0.023953 | 0.05193 | 0.054132 | 0.055947 | 0.067486 | 0.03057 | 0.070253 | 0.11616 | 0.088027
a8 | 0.055266 | 0.019421 | 0.0030897 | 0.0037123 | 0.01907 | -0.012151 | 0.0035421 | -0.036046 | 0.0087662 | 0.0072948 | 0.036917 | 0.0048588 | -0.0021397 | 0.014131 | 0.031133 | 0.0052184 | 0.023526 | -0.01327 | 0.014254 | -0.022263 | 0.0077983 | a8 | 0.0068092 | 0.0028122 | 0.0018472 | 0.016562 | 0.053297 | 0.039363 | 0.0091324 | 0.01691 | -0.028641 | 0.0054098 | -0.028724 | -0.036398 | -0.0446 | -0.0387 | -0.080386 | -0.027869 | -0.026065 | 0.0049402 | 0.010717 | -0.080387 | -0.010207
a9 | -0.28063 | -0.0058266 | -0.013064 | -0.023784 | -0.08105 | -0.058142 | -0.032889 | -0.011612 | -0.0023489 | -0.057172 | -0.017965 | 0.01176 | -0.027013 | -0.03921 | -0.03975 | -0.0089429 | -0.0095819 | -0.005983 | -0.012612 | -0.07686 | -0.033643 | a9 | -0.23319 | 0.010005 | 0.0035402 | -0.0296 | -0.07907 | -0.052721 | -0.0023116 | 0.026467 | 0.024176 | -0.027105 | 0.014646 | 0.011942 | -0.0083854 | -0.023221 | -0.044807 | 0.021957 | 0.0090242 | -0.0044896 | -0.0029833 | -0.038955 | -0.016202
a10 | 0.063711 | -0.032753 | 0.0057382 | 0.014657 | 0.027809 | 0.0018541 | -0.0077407 | -0.016522 | 0.0041969 | 0.00464 | 0.012123 | 0.0027942 | -0.012187 | -0.021415 | 0.0033025 | -0.010692 | 0.02043 | -0.039393 | 0.024596 | -0.029928 | 0 | a10 | -0.0018558 | 0.0028519 | 0.0016651 | -0.0064852 | 0.024377 | 0.003438 | 0.041627 | 0.054425 | 0.040892 | 0.025366 | -0.014948 | 0.022685 | -0.031911 | -0.015479 | -0.027112 | 0.0066411 | -0.014825 | -0.021816 | 0.0069687 | -0.15838 | 0
a11 | 0.13231 | -0.073174 | -0.051664 | -0.049438 | -0.059835 | -0.065915 | -0.055549 | -0.046886 | -0.041002 | -0.048017 | -0.044523 | -0.04183 | -0.049954 | -0.041569 | -0.045457 | -0.052822 | -0.045616 | -0.038711 | -0.039197 | -0.033266 | -0.043476 | a11 | 0.18124 | -0.020849 | -0.0098851 | 0.0017307 | -0.0055119 | -0.011416 | -0.0010328 | 0.011259 | 0.016222 | -0.0013113 | 0.0012689 | 0.0051143 | -0.0046658 | 0.0038737 | -0.0072968 | 0.0022111 | 0.0070256 | 0.0041006 | 0.014622 | 0.020786 | 0.0063797
a12 | 0.0046056 | -0.008269 | -0.0018375 | -0.0053365 | -8.8525e-05 | -0.0018001 | -0.0055462 | -0.011051 | -0.0057585 | -0.005292 | -3.5046e-05 | -0.0053279 | -0.0082954 | -0.0062235 | -0.0040964 | -0.0056434 | -0.0026605 | -0.0084674 | 0.0016219 | -0.006609 | -0.0044537 | a12 | -8.7616e-05 | 0.00032488 | -0.0014258 | -0.0053288 | -0.00029301 | -0.0017505 | -0.00054588 | 0.0052155 | 0.00022144 | 0.0037172 | -0.0052702 | -0.0011046 | -0.011988 | -0.011923 | -0.0090435 | -0.004568 | -0.004253 | -0.0053314 | 0.00020139 | -0.016154 | -0.0032892
a13 | -0.031103 | -0.0021964 | -0.0024337 | -0.0050231 | -0.0084853 | -0.0083208 | -0.0041702 | -0.0056166 | -0.0011549 | -0.0043912 | -0.0036131 | -0.0012427 | -0.005785 | -0.0034524 | -0.0048589 | -0.003974 | -0.0034775 | -0.0025269 | -0.0013607 | -0.0055844 | -0.0048765 | a13 | -0.025526 | 0.00058949 | 0.0020518 | -0.0024824 | -0.00076297 | -0.0048622 | 0.00030624 | 0.004145 | 0.0023496 | -0.00023395 | 0.0019164 | 0.0016598 | -0.0015762 | -0.00021846 | -0.0001451 | 0.0023059 | 0.0033661 | 0.0019733 | 0.0032866 | -0.00066543 | -7.3625e-05
a14 | 0.013285 | -0.019474 | -0.016738 | -0.016244 | -0.014874 | -0.018009 | -0.02067 | -0.021153 | -0.019139 | -0.017407 | -0.016108 | -0.017891 | -0.019324 | -0.018882 | -0.015757 | -0.017194 | -0.0164 | -0.020232 | -0.015411 | -0.010144 | -0.016638 | a14 | 0.012908 | -0.010653 | -0.010509 | -0.010642 | -0.0088607 | -0.0087396 | -0.0091119 | -0.0078789 | -0.0090887 | -0.0086423 | -0.0104 | -0.010673 | -0.013585 | -0.011391 | -0.012184 | -0.011585 | -0.011851 | -0.012423 | -0.010532 | -0.0056529 | -0.0096305
a15 | 0.0074049 | 0.0028675 | -0.0049734 | -0.0053651 | -0.0061557 | -0.0088651 | -0.01463 | -0.007666 | -0.0098426 | -0.011645 | -0.010608 | -0.0040476 | -0.0070853 | -0.011699 | -0.012522 | -0.0082892 | -0.0043187 | -0.0032381 | -0.0013988 | -0.0015182 | -0.0065679 | a15 | 0.00055144 | -0.00077071 | -0.0090575 | -0.010502 | -0.010245 | -0.013038 | -0.008776 | -0.0028242 | -0.0055969 | -0.0071642 | -0.0080299 | -0.00302 | -0.0062202 | -0.010735 | -0.010891 | -0.0071202 | -0.0078331 | -0.0081412 | -0.0035154 | -0.017403 | -0.0074964
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.25458 | 0.20312 | 0.058588 | 0.031006 | 0.015434 | -0.054843 | -0.15661 | -0.018841 | -0.073199 | -0.095923 | -0.074971 | 0.039572 | 0 | -0.10275 | -0.095521 | 0 | 0.041837 | 0.076324 | 0.09342 | 0.044762 |  | Dx (mm) | 0.16011 | 0.25019 | -0.034662 | -0.074496 | -0.073836 | -0.12243 | -0.061623 | 0.088216 | 0 | 0 | -0.014997 | 0.10834 | 0.048199 | -0.040883 | -0.076634 | -0.026857 | -0.034813 | -0.026249 | 0.090974 | -0.16209
Dy (mm) | 0.12413 | -0.056235 | 0.017768 | 0.036227 | 0.066473 | 0 | -0.030044 | -0.039417 | 0 | 0 | 0.023401 | 0 | -0.028242 | -0.055538 | 0 | -0.025415 | 0.050535 | -0.087285 | 0.063111 | -0.078095 |  | Dy (mm) | -0.054433 | 0.013589 | 0 | -0.014637 | 0.05519 | 0 | 0.09582 | 0.12341 | 0.09253 | 0.058057 | -0.033819 | 0.051105 | -0.072476 | -0.034478 | -0.06222 | 0.015305 | -0.033794 | -0.050023 | 0.014415 | -0.41492
Mag. len. (m) | 14.419 | [20 empty] | Mag. len. (m) | 14.4199
TF (mT/KA) | 596.0109 | [20 empty] | TF (mT/KA) | 596.0506
Coil pos. (mm) | 0.015163 | [20 empty] | Coil pos. (mm) | 0.018692
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | -0.0075755 | [20 empty] | Twist (m2 rad) | -0.0011627

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.092379 | 0.66667 | 0.15544 | 0.38558 | 0.97139 | 2.5417 | 0.83991 | 0.45485 | -0.1187 | 0.45946 | -0.076575 | 0.39736 | -0.45325 | 1.6216 | -0.70071 | 0.64001 |  | 0
3 | -1.5116 | 10.236 | -2.9186 | 0.63311 | 0.32069 | 0.26712 | 0.34784 | 0.24075 | -1.4602 | 10.3217 | -2.8523 | 0.9201 | 0.17083 | 0.28358 | 0.14952 | 0.25491 |  | 0
4 | -0.0024857 | 0.26927 | 0.036358 | 0.11048 | -0.060697 | 0.66574 | -0.13219 | 0.20297 | -0.047373 | 0.23004 | -0.019054 | 0.097054 | 0.046925 | 0.44708 | -0.01695 | 0.25779 |  | 0
5 | -0.80801 | 1.2518 | -0.56005 | 0.17283 | 0.043485 | 0.67707 | -0.046149 | 0.1343 | -1.055 | 1.2232 | -0.81216 | 0.19477 | 0.055638 | 0.62319 | -0.022013 | 0.11429 |  | 0
6 | 0.027727 | 0.057097 | 0.030577 | 0.042072 | 0.0052611 | 0.15269 | -0.0096203 | 0.052566 | -0.031792 | 0.077719 | -0.026345 | 0.072955 | -0.034989 | 0.1229 | -0.045853 | 0.063354 |  | 0
7 | 1.1614 | 0.40665 | 1.1143 | 0.054481 | 0.034877 | 0.34624 | -0.012151 | 0.044631 | 1.1895 | 0.42688 | 1.143 | 0.07147 | 0.088027 | 0.39733 | 0.032274 | 0.041422 |  | 0
8 | -0.00058567 | 0.027437 | -0.00082689 | 0.012854 | 0.0077983 | 0.020614 | 0.0076315 | 0.015269 | -0.00068897 | 0.032003 | 0.0022926 | 0.015855 | -0.010207 | 0.035235 | -0.011884 | 0.029183 |  | 0
9 | 0.45442 | 0.032576 | 0.45867 | 0.010453 | -0.033643 | 0.06204 | -0.024132 | 0.024019 | 0.47031 | 0.031408 | 0.47456 | 0.017512 | -0.016202 | 0.057341 | -0.0089183 | 0.02847 |  | 0
10 | 0 | 0.040289 | -0.00018631 | 0.00094988 | 0 | 0.023923 | 0.00050064 | 0.0010859 | 0 | 0.041302 | 0.00015876 | 0.00066107 | 0 | 0.04358 | 0.00022854 | 0.00068806 |  | 0
11 | 0.72324 | 0.035256 | 0.73006 | 0.0051455 | -0.043476 | 0.041569 | -0.049534 | 0.0091229 | 0.73566 | 0.040561 | 0.74349 | 0.0064728 | 0.0063797 | 0.041434 | 0.00020731 | 0.0093176 |  | 0
12 | -0.0020625 | 0.0061887 | -0.0025548 | 0.0014871 | -0.0044537 | 0.0038226 | -0.0045648 | 0.0020696 | 0.0017979 | 0.0043615 | 0.0015253 | 0.0024883 | -0.0032892 | 0.0054258 | -0.0035821 | 0.0027987 |  | 0
13 | 0.064784 | 0.0061416 | 0.064508 | 0.0024055 | -0.0048765 | 0.0063805 | -0.0041256 | 0.0027945 | 0.069576 | 0.0057978 | 0.069319 | 0.0021426 | -7.3625e-05 | 0.0062504 | 0.00077556 | 0.0026256 |  | 0
14 | -0.0036167 | 0.0036803 | -0.0039218 | 0.00087122 | -0.016638 | 0.0073007 | -0.017972 | 0.001334 | -0.0039751 | 0.0030213 | -0.0043702 | 0.00080994 | -0.0096305 | 0.005481 | -0.0109 | 0.00065329 |  | 0
15 | 0.036221 | 0.0097641 | 0.037946 | 0.0028266 | -0.0065679 | 0.0053775 | -0.0071935 | 0.0043702 | 0.025542 | 0.0077349 | 0.026546 | 0.0023796 | -0.0074964 | 0.0042645 | -0.0074156 | 0.0032767 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.8704 | -0.68458
3 | [12 empty] | 3 | 9 | -0.94055 | -0.72507
4 | [12 empty] | 4 | 16 | -1.8988 | -1.8441
5 | [12 empty] | 5 | 25 | -1.8816 | -1.9025
6 | [12 empty] | 6 | 36 | -3.057 | -2.6885
7 | [12 empty] | 7 | 49 | -3.0096 | -2.9112
8 | [12 empty] | 8 | 64 | -4.268 | -3.8392
9 | [12 empty] | 9 | 81 | -4.1449 | -3.8019
10 | [12 empty] | sigma (mm) |  | 0.015163 | 0.018692
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -3.5737 | -0.38452 | 0.14726 | -0.6099 | -0.047055 | 0.32182 | -0.070965 | -0.26108 | 0.3825 | 0.27272 | 0.10876 | -0.21356 | 0.12536 | 0.034487 | 0.46574 | 0.67585 | 0.75462 | 0.32061 | 0.77519 | 0.42921 | 0.092379
3 | 42.9161 | -3.3911 | -2.7007 | -3.113 | -2.1674 | -3.2479 | -2.2462 | -3.0742 | -1.9797 | -2.9916 | -2.964 | -3.3694 | -2.8406 | -4.2178 | -3.8503 | -3.4064 | -2.4116 | -2.7259 | -1.837 | -2.2177 | -1.5116
4 | -0.68093 | -0.037965 | -0.011399 | -0.13969 | -0.080357 | 0.23742 | 0.059292 | 0.15008 | 0.068392 | -0.10113 | 0.1035 | -0.081891 | 0.040787 | -0.03991 | 0.10074 | 0.18704 | -0.039731 | 0.047398 | 0.19187 | -0.026965 | -0.0024857
5 | -5.0742 | -0.80109 | -0.96171 | -0.58871 | -0.54342 | -0.468 | -0.36466 | -0.31753 | -0.37096 | -0.34753 | -0.34974 | -0.52161 | -0.63413 | -0.63951 | -0.58774 | -0.56846 | -0.72199 | -0.62304 | -0.67104 | -4.0096 | -0.80801
6 | -0.1208 | 0.045136 | 0.037812 | 0.10874 | 0.081404 | 0.037491 | 0.027444 | 0.033665 | -0.021523 | 0.058694 | 0.0038944 | 0.064547 | 0.010839 | 0.042666 | 0.015209 | -0.0754 | -0.019682 | 0.035015 | 0.064427 | -0.017146 | 0.027727
7 | 2.8895 | 0.96679 | 1.0557 | 1.1795 | 1.152 | 1.214 | 1.1323 | 1.1505 | 1.087 | 1.1204 | 1.118 | 1.0923 | 1.0516 | 1.1312 | 1.1187 | 1.1032 | 1.0933 | 1.1466 | 1.1442 | 0.88759 | 1.1614
8 | 0.0031599 | 0.016065 | 0.003853 | -0.013727 | 0.0011886 | 0.001926 | 0.0057205 | 0.0021041 | -0.0066736 | 0.0016985 | -0.018261 | -0.014928 | 0.028117 | -0.0043044 | 0.0033352 | -0.020502 | -0.0010898 | 0.015658 | -0.015064 | -0.022829 | -0.00058567
9 | 0.32055 | 0.46733 | 0.46482 | 0.45914 | 0.47085 | 0.45354 | 0.45953 | 0.4455 | 0.47284 | 0.44885 | 0.45344 | 0.44539 | 0.45729 | 0.45885 | 0.44169 | 0.45499 | 0.45272 | 0.47931 | 0.47003 | 0.41424 | 0.45442
10 | -4.4772e-05 | 1.0726e-05 | -4.1132e-07 | -2.7021e-06 | -2.3351e-06 | -2.285e-05 | 3.1462e-05 | -3.807e-06 | 0.00014245 | 0.00013676 | 6.7496e-06 | 0.0002148 | 4.9112e-05 | -2.2853e-06 | 5.561e-05 | -0.0039838 | -2.4123e-06 | 1.8791e-05 | 6.2318e-07 | 8.3093e-06 | 0
11 | 0.60106 | 0.72732 | 0.71919 | 0.72999 | 0.72511 | 0.72924 | 0.72875 | 0.73516 | 0.73061 | 0.73091 | 0.738 | 0.73061 | 0.73347 | 0.73278 | 0.73631 | 0.7344 | 0.72493 | 0.73376 | 0.72056 | 0.63263 | 0.72324
12 | -0.021892 | -0.035102 | -0.0263 | -0.040123 | -0.031499 | -0.016786 | -0.044922 | -0.030297 | -0.032265 | -0.027669 | -0.044011 | -0.0077551 | -0.015771 | -0.035252 | -0.025503 | -0.036751 | 0.0075083 | 0.0027446 | -0.020117 | 0.0020791 | -0.020625 | ten times!
13 | 0.85545 | 0.68132 | 0.6894 | 0.63845 | 0.62258 | 0.63172 | 0.61083 | 0.62532 | 0.6096 | 0.60799 | 0.65042 | 0.63728 | 0.66255 | 0.65097 | 0.65835 | 0.65424 | 0.65012 | 0.67261 | 0.65774 | 0.51058 | 0.64784 | ten times!
14 | 0.072529 | -0.058579 | -0.038539 | -0.051539 | -0.039778 | -0.038133 | -0.041704 | -0.039518 | -0.031115 | -0.036574 | -0.038718 | -0.019766 | -0.040573 | -0.034903 | -0.040837 | -0.048317 | -0.037389 | -0.043884 | -0.026057 | -0.00089309 | -0.036167 | ten times!
15 | -0.016122 | 0.45799 | 0.39108 | 0.35608 | 0.34536 | 0.36901 | 0.38632 | 0.38458 | 0.38004 | 0.38734 | 0.35656 | 0.37404 | 0.38946 | 0.38983 | 0.36175 | 0.3775 | 0.34582 | 0.42708 | 0.35034 | 0.21605 | 0.36221 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 9.8952 | 1.6636 | 0.93974 | 0.38193 | 0.29406 | 0.74133 | 0.9599 | 0.78123 | 1.204 | 0.59741 | 0.64018 | 0.14876 | 0.9148 | 1.1095 | 0.85072 | 1.4732 | 1.4042 | 1.0383 | -0.02449 | -4.3212 | 0.97139
3 | -0.088575 | 0.82037 | 0.36143 | 0.50397 | 0.25388 | 0.50972 | 0.22028 | -0.13691 | 0.49977 | 0.1011 | 0.18747 | 0.58032 | 0.14649 | 0.14917 | 0.078052 | 0.39325 | 0.60078 | 0.3848 | 0.60724 | -0.19334 | 0.32069
4 | 2.6387 | 0.0081605 | 0.15739 | 0.086532 | -0.073244 | -0.079151 | -0.46503 | 0.017383 | -0.034042 | 0.18956 | -0.27922 | -0.34259 | -0.33369 | -0.31291 | -0.079285 | -0.24094 | -0.31177 | -0.37787 | 0.091358 | -0.64722 | -0.060697
5 | 2.8785 | -0.14725 | 0.042247 | 0.0060371 | -0.082383 | -0.34782 | -0.14935 | -0.22077 | 0.01493 | 0.0077101 | 0.019698 | 0.21257 | -0.17108 | -0.041674 | -0.077322 | -0.046972 | 0.0047683 | -0.042994 | 0.18898 | -0.046889 | 0.043485
6 | 0.34641 | -0.12146 | -0.065722 | -0.038403 | 0.049561 | 0.053246 | 0.023405 | 0.0050691 | 0.01061 | 0.05719 | 0.046341 | 0.0077226 | -0.028701 | -0.041632 | -0.060091 | -0.097943 | -0.0073506 | 0.0072179 | 0.027776 | -0.074057 | 0.0052611
7 | 1.5167 | -0.078939 | 0.0093801 | -0.015118 | -0.054332 | -0.085206 | -0.068279 | -0.031944 | 0.050213 | -0.02026 | -0.0019969 | -0.0059137 | -0.02277 | 0.02624 | -0.012369 | -0.046817 | 0.025488 | 0.071419 | 0.04249 | 0.004675 | 0.034877
8 | 0.069771 | 0.03256 | -0.0004428 | -0.0037673 | 0.0049586 | -0.013652 | 0.0076936 | -0.02789 | 0.0086853 | 0.004714 | 0.031277 | 0.0046399 | 0.0039355 | 0.024276 | 0.029346 | 0.010658 | 0.012952 | 0.006609 | 0.00081259 | -0.0054525 | 0.0077983
9 | -0.29452 | 0.00021264 | -0.013481 | -0.0243 | -0.081591 | -0.058089 | -0.034615 | -0.011969 | -0.0021863 | -0.056937 | -0.017088 | 0.011701 | -0.027024 | -0.041717 | -0.039583 | -0.0090071 | -0.010548 | -0.0029605 | -0.0152 | -0.075525 | -0.033643
10 | -0.00015609 | 6.5337e-05 | 1.5227e-06 | 8.8584e-07 | 1.2283e-07 | -0.00027235 | 4e-05 | -1.3537e-06 | 0.0024314 | 0.0019306 | -4.0128e-06 | 0.0037679 | -4.5573e-06 | 2.2431e-05 | 0.00074829 | 0.00028627 | -2.2307e-06 | 4.0075e-06 | -2.7421e-06 | -2.9669e-06 | 0
11 | 0.13028 | -0.071866 | -0.051597 | -0.049273 | -0.059746 | -0.065979 | -0.056333 | -0.047123 | -0.041274 | -0.048345 | -0.044406 | -0.041693 | -0.049985 | -0.042283 | -0.04571 | -0.052884 | -0.045636 | -0.038086 | -0.039394 | -0.032999 | -0.043476 | ten times!
12 | 0.027046 | -0.052709 | -0.026033 | -0.068682 | -0.029511 | -0.021222 | -0.046977 | -0.094014 | -0.058182 | -0.055893 | -0.013109 | -0.052931 | -0.069829 | -0.039308 | -0.04424 | -0.044757 | -0.048968 | -0.042773 | -0.012522 | -0.036502 | -0.044537 | ten times!
13 | -0.34274 | 0.0091171 | -0.01658 | -0.045248 | -0.08147 | -0.090761 | -0.068642 | -0.060659 | -0.022262 | -0.056681 | -0.044316 | -0.0070134 | -0.058762 | -0.052471 | -0.060099 | -0.040713 | -0.028614 | -0.014753 | -0.0026763 | -0.052008 | -0.048765 | ten times!
14 | 0.11898 | -0.17833 | -0.1707 | -0.1717 | -0.16686 | -0.18409 | -0.21601 | -0.20024 | -0.19732 | -0.18327 | -0.1745 | -0.17759 | -0.18419 | -0.18089 | -0.16742 | -0.16404 | -0.1769 | -0.16959 | -0.17124 | -0.086984 | -0.16638 | ten times!
15 | 0.074049 | 0.028675 | -0.049734 | -0.053651 | -0.061557 | -0.088651 | -0.1463 | -0.07666 | -0.098426 | -0.11645 | -0.10608 | -0.040476 | -0.070853 | -0.11699 | -0.12522 | -0.082892 | -0.043187 | -0.032381 | -0.013988 | -0.015182 | -0.065679 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.9396 | -0.76111 | -0.15558 | -0.44376 | 0.83673 | -0.35241 | -0.14063 | 0.30282 | 0.018005 | 0.25069 | 0.01534 | 0.023209 | -0.69787 | -0.0098889 | 0.1077 | 0.19756 | -0.094514 | 0.16909 | -0.64373 | -0.51386 | -0.1187
3 | 43.176 | -3.3856 | -2.904 | -3.0942 | -2.4374 | -4.0359 | -2.06 | -2.1508 | -1.9838 | -2.768 | -3.5013 | -2.6181 | -2.3703 | -5.502 | -3.7305 | -2.5632 | -1.7016 | -2.2785 | -2.2565 | -2.4365 | -1.4602
4 | -0.82047 | 0.20203 | -0.0029626 | -0.1656 | 0.070536 | -0.0028096 | -0.078071 | -0.011501 | -0.013146 | 0.12786 | -0.025907 | -0.13673 | -0.083957 | 0.086474 | -0.049966 | -0.16064 | 0.017256 | -0.085155 | -0.030691 | -0.13029 | -0.047373
5 | -4.9062 | -1.0792 | -1.1198 | -0.7463 | -0.69089 | -0.68269 | -0.67204 | -0.66403 | -0.65031 | -0.60911 | -0.54053 | -0.81506 | -0.83535 | -0.54532 | -0.88706 | -1.0217 | -1.0729 | -0.94099 | -1.0455 | -4.5465 | -1.055
6 | -0.28088 | 0.0056423 | 0.057261 | 0.059151 | -0.19475 | -0.035113 | 0.042219 | 0.010005 | 0.069946 | -0.057188 | -0.018253 | -0.13396 | -0.069757 | -0.12558 | -0.065831 | 0.016815 | 0.0028465 | 0.017973 | -0.05564 | -0.078035 | -0.031792
7 | 2.9927 | 0.94772 | 1.0075 | 1.1322 | 1.1428 | 1.2062 | 1.2029 | 1.1997 | 1.1576 | 1.1645 | 1.2575 | 1.1972 | 1.1298 | 1.166 | 1.1115 | 1.1251 | 1.127 | 1.1628 | 1.1365 | 0.85757 | 1.1895
8 | -0.063994 | -0.006268 | -0.0053594 | -0.026938 | 0.021226 | -0.014897 | 0.0026561 | 0.0072451 | -0.019598 | -0.0032277 | 0.016832 | 0.032251 | 0.013289 | -0.0073106 | -0.011266 | 0.0038644 | 0.00052337 | 0.015902 | 0.022342 | -0.01703 | -0.00068897
9 | 0.37119 | 0.49084 | 0.48005 | 0.45107 | 0.46482 | 0.44314 | 0.50036 | 0.48639 | 0.50209 | 0.46838 | 0.45685 | 0.46856 | 0.46339 | 0.46452 | 0.45874 | 0.50166 | 0.48164 | 0.48275 | 0.47678 | 0.43523 | 0.47031
10 | -5.0744e-05 | -5.0823e-05 | 1.9869e-05 | -3.4715e-06 | 3.0879e-06 | 4.2782e-05 | 4.0536e-06 | 1.6434e-05 | 0.0028062 | 4.7276e-05 | -7.9637e-07 | -1.3053e-05 | 1.6084e-05 | -5.9888e-06 | -1.4894e-05 | 1.3475e-07 | -1.6673e-06 | -1.1868e-06 | -6.4701e-06 | -0.0002487 | 0
11 | 0.59049 | 0.74005 | 0.73811 | 0.74499 | 0.74344 | 0.74528 | 0.73783 | 0.74197 | 0.75023 | 0.74286 | 0.75089 | 0.74348 | 0.74581 | 0.75894 | 0.75014 | 0.74141 | 0.73864 | 0.73935 | 0.72943 | 0.64366 | 0.73566
12 | 0.042046 | -0.041059 | 0.0022063 | 0.037303 | -0.0025899 | 0.032452 | 0.014371 | -0.001188 | 0.014654 | -0.014364 | 0.01097 | 0.016023 | 0.016289 | -0.0002435 | 0.033577 | 0.074542 | 0.031232 | 0.038827 | 0.011546 | 0.053164 | 0.017979 | ten times!
13 | 0.87082 | 0.70992 | 0.70581 | 0.68227 | 0.68697 | 0.69155 | 0.64952 | 0.69563 | 0.6882 | 0.68036 | 0.72533 | 0.69373 | 0.71758 | 0.66129 | 0.69825 | 0.6568 | 0.70364 | 0.71831 | 0.71232 | 0.55914 | 0.69576 | ten times!
14 | 0.069512 | -0.054355 | -0.045137 | -0.027261 | -0.053644 | -0.054175 | -0.044492 | -0.0328 | -0.041279 | -0.050115 | -0.043867 | -0.047333 | -0.046239 | -0.051781 | -0.035719 | -0.035146 | -0.033163 | -0.042081 | -0.04804 | 0.051076 | -0.039751 | ten times!
15 | -0.065662 | 0.31097 | 0.27985 | 0.25569 | 0.24061 | 0.24293 | 0.23366 | 0.24261 | 0.24701 | 0.24205 | 0.26327 | 0.24469 | 0.28317 | 0.2679 | 0.26457 | 0.27967 | 0.28094 | 0.30965 | 0.28901 | 0.26174 | 0.25542 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 6.1802 | -0.2483 | 0.13051 | -0.058422 | 0.48212 | -0.50407 | -1.8981 | -0.89107 | -1.3329 | -0.57651 | -1.169 | -1.6982 | -0.75725 | -1.0497 | -0.31302 | -0.33554 | -0.67832 | -0.28313 | -1.432 | 0.74199 | -0.45325
3 | 0.76811 | 0.30948 | 0.25489 | 0.061756 | -0.013618 | -0.22253 | 0.0023822 | -0.12346 | 0.28992 | -0.093535 | 0.63166 | 0.49734 | -0.015004 | -0.1291 | -0.028272 | 0.143 | 0.42071 | 0.16017 | 0.54549 | 0.20008 | 0.17083
4 | 1.4579 | -0.061499 | -0.35276 | -0.44614 | -0.12091 | -0.017548 | 0.37069 | 0.2689 | 0.36655 | -0.11373 | 0.20942 | -0.24143 | -0.39678 | -0.10727 | -0.014062 | 0.38606 | -0.021652 | -0.092507 | 0.079567 | 0.018983 | 0.046925
5 | 2.6951 | 0.01184 | -0.10907 | 0.054643 | -0.010604 | -0.059833 | -0.18591 | -0.016738 | 0.064517 | 0.061948 | 0.17793 | -0.027523 | -0.18596 | -0.052013 | -0.21231 | -0.026611 | -0.064502 | -0.035804 | 0.21978 | -0.18321 | 0.055638
6 | 0.29965 | -0.061977 | -0.073092 | 0.027485 | -0.016984 | 0.065474 | 0.0050284 | 0.012079 | -0.07137 | 0.02841 | -0.054924 | -0.050539 | -0.12298 | -0.19122 | -0.12605 | -0.085242 | -0.058244 | -0.039244 | -0.011963 | -0.1008 | -0.034989
7 | 1.8013 | 0.015647 | 0.012007 | -0.023084 | 0.053741 | -0.073686 | 0.025298 | 0.10055 | 0.057506 | 0.064998 | 0.010526 | 0.03481 | -0.022367 | 0.051043 | 0.050849 | 0.055655 | 0.067011 | 0.030826 | 0.069598 | 0.10131 | 0.088027
8 | 0.034019 | -0.001597 | 0.0019049 | 0.01864 | 0.038463 | 0.036326 | -0.013441 | -0.012331 | -0.050417 | -0.0073654 | -0.021353 | -0.048326 | -0.028629 | -0.031608 | -0.068555 | -0.031205 | -0.018257 | 0.016239 | 0.007598 | -0.0042288 | -0.010207
9 | -0.23126 | 0.0088412 | 0.0035707 | -0.030109 | -0.077179 | -0.052498 | 0.00040017 | 0.021463 | 0.024082 | -0.027108 | 0.01441 | 0.0093622 | -0.0067673 | -0.023887 | -0.04703 | 0.022146 | 0.0084828 | -0.0050918 | -0.0036159 | -0.064878 | -0.016202
10 | 2.5215e-05 | -6.174e-06 | 0.0014635 | 4.6698e-06 | -3.5636e-06 | 0.0026159 | -3.1727e-06 | -6.221e-06 | 5.8255e-05 | 1.7035e-06 | -9.3759e-07 | -7.3201e-06 | -1.7684e-06 | 1.7442e-06 | 1.5269e-06 | 1.477e-06 | -8.6566e-07 | -3.0722e-06 | -1.9153e-06 | -0.00032751 | 0
11 | 0.18172 | -0.020977 | -0.0099171 | 0.0014749 | -0.0053885 | -0.011554 | -0.00096475 | 0.01065 | 0.016143 | -0.0012577 | 0.001226 | 0.0049648 | -0.0041063 | 0.0035151 | -0.0077619 | 0.0020705 | 0.0069621 | 0.0040951 | 0.014557 | 0.018788 | 0.0063797
12 | 0.061926 | -0.0046305 | -0.013756 | -0.04753 | -0.030376 | -0.021706 | -0.049867 | -0.011598 | -0.043155 | 0.0091326 | -0.035235 | -0.037419 | -0.083086 | -0.10323 | -0.059951 | -0.052362 | -0.024963 | -0.027716 | -0.0073384 | -0.0022209 | -0.032892 | ten times!
13 | -0.26852 | 0.026174 | 0.017732 | -0.031382 | -0.0099479 | -0.056804 | 0.0023902 | 0.047926 | 0.026115 | -0.0002667 | 0.016695 | 0.026386 | -0.0129 | -0.0074292 | -0.011351 | 0.021153 | 0.029384 | 0.015245 | 0.040479 | -0.027408 | -0.00073625 | ten times!
14 | 0.12541 | -0.10842 | -0.10767 | -0.10978 | -0.10577 | -0.10054 | -0.11401 | -0.10139 | -0.10971 | -0.097996 | -0.09766 | -0.11433 | -0.11648 | -0.10991 | -0.11516 | -0.12095 | -0.11294 | -0.11323 | -0.10612 | 0.0096782 | -0.096305 | ten times!
15 | 0.0055144 | -0.0077071 | -0.090575 | -0.10502 | -0.10245 | -0.13038 | -0.08776 | -0.028242 | -0.055969 | -0.071642 | -0.080299 | -0.0302 | -0.062202 | -0.10735 | -0.10891 | -0.071202 | -0.078331 | -0.081412 | -0.035154 | -0.17403 | -0.074964 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00044432 | -0.00020433 | -0.00016826 | -8.5709e-05 | -0.00014108 | 3.8953e-05 | 6.647e-05 | -5.5676e-05 | -3.0173e-05 | 9.4238e-06 | -8.2353e-05 | -0.00021675 | -0.00015232 | -6.5072e-05 | 8.912e-05 | 2.4356e-05 | -2.2119e-05 | 0.00055119 |  | 596.0109
|  | 0.00031241 | -0.0002055 | -0.00010501 | -2.6657e-05 | -3.2361e-05 | -2.1792e-05 | -0.00011725 | 1.193e-05 | -6.5077e-05 | -2.4811e-05 | 5.2363e-05 | -7.5143e-05 | 3.8942e-05 | -2.9845e-05 | -4.8803e-05 | 8.0549e-05 | -0.00017262 | 0.00042867 |  | 596.0506
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3999.2169
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.9738 | -4181.308 | 5.0657 | -1.4212 | -1.0604 | -0.23491 | -0.78863 | 1.0118 | 1.287 | 0.065439 | 0.32048 | 0.71647 | -0.20135 | -1.5454 | -0.90105 | -0.028525 | 1.5135 | 0.86581 | 0.40102 | 6.1345 | -3817.1259
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.0036322 | 0.57253 | -1.298 | -0.6113 | -0.25829 | -0.26428 | -0.14779 | 0.38068 | 0.8153 | 0.59308 | 0.10289 | -0.4192 | -0.13796 | 0.33872 | 0.77051 | 0.4967 | -0.027166 | -0.37279 | -0.031978 | 0.070881 | -0.42513
db2 | 0.15832 | -2.2917 | -0.61851 | -0.030437 | -0.78173 | -0.21129 | 0.18447 | -0.187 | -0.41323 | 0.24123 | 0.14814 | -0.023945 | -0.38757 | -0.032471 | -0.071806 | 0.35071 | 0.5187 | 0.58088 | 0.14167 | 0.59217 | 0.41568
db3 | -2.9197 | 42.8096 | -0.47414 | 0.21819 | -0.19467 | 0.75302 | -0.33054 | 0.66906 | -0.15491 | 0.93901 | -0.070332 | -0.044658 | -0.45036 | 0.077427 | -1.3006 | -0.9326 | -0.48777 | 0.51061 | 0.18866 | 1.0846 | -2.2261
db4 | 0.034382 | -1.0714 | -0.11251 | -0.059194 | -0.17842 | -0.11546 | 0.20908 | 0.037223 | 0.11506 | 0.040381 | -0.12766 | 0.075187 | -0.12112 | 0.005268 | -0.059323 | 0.079554 | 0.15238 | -0.08128 | 0.00091137 | 0.13992 | -0.070103
db5 | -0.55951 | -5.0911 | -0.23723 | -0.40086 | -0.027831 | 0.015002 | 0.091288 | 0.19665 | 0.24172 | 0.18962 | 0.2114 | 0.20996 | 0.03883 | -0.074939 | -0.081118 | -0.027596 | -0.0097483 | -0.16271 | -0.062675 | -0.10976 | -4.0119
db6 | 0.027339 | 0.072293 | 0.085942 | 0.032251 | 0.094505 | 0.061604 | -0.013352 | -0.063308 | -0.0017617 | -0.076962 | -0.006601 | -0.053017 | 0.052462 | -0.016731 | -0.025236 | -0.049866 | -0.10316 | -0.031347 | 0.040729 | 0.073847 | -0.0030074
db7 | 1.1149 | 2.8928 | -0.14255 | -0.058815 | 0.064469 | 0.036569 | 0.099341 | 0.019271 | 0.035012 | -0.027247 | 0.0061995 | 0.0037695 | -0.022761 | -0.063295 | 0.017803 | 0.0044085 | -0.011629 | -0.021932 | 0.03223 | 0.02916 | 0.88657
db8 | -0.0019433 | 0.058757 | 0.0633 | 0.018741 | -0.0046795 | 0.0090804 | -0.0078537 | -0.027075 | -0.00011567 | -0.02106 | -0.016714 | -0.032156 | -0.0046842 | 0.029701 | -0.025662 | -0.01467 | -0.018666 | 0.009992 | 0.034551 | 0.0079713 | -0.016947
db9 | 0.45994 | 0.32723 | 0.01551 | 0.005614 | -0.00084561 | 0.010007 | -0.0057118 | 0.005123 | -0.014694 | 0.014118 | -0.0089992 | -0.0053817 | -0.014187 | -0.0028275 | 0.00076489 | -0.01616 | -0.0050917 | -0.0073446 | 0.018803 | 0.011303 | 0.41336
db10 | -0.0031297 | 0.080641 | 0.087681 | 0.028456 | 0.017496 | 0.01205 | -0.020424 | -0.065051 | -0.0061064 | -0.028199 | -0.038 | -0.028813 | 0.020352 | 0.0023489 | -0.042541 | -0.038211 | -0.0016438 | 0.022328 | 0.034091 | 0.044187 | 0.01513
db11 | 0.73024 | 0.60331 | -0.0024127 | -0.011017 | -0.00017538 | -0.0051475 | -0.00085228 | -0.00043408 | 0.0046928 | 0.00067492 | 0.0011083 | 0.0081394 | 0.00040282 | 0.0030847 | 0.0028719 | 0.0065056 | 0.0040683 | -0.005143 | 0.0032282 | -0.0095954 | 0.63234
db12 | -0.0027656 | 0.015898 | 0.0088598 | 0.0030231 | 0.00028687 | 0.00069992 | -0.0013675 | -0.008641 | -0.0012444 | -0.0036263 | -0.0041514 | -0.005051 | 0.0037696 | 0.0010731 | -0.0056 | -0.0042579 | -0.00098077 | 0.0055769 | 0.0064292 | 0.0052017 | 0.0014948
db13 | 0.064607 | 0.085482 | 0.0029881 | 0.0044755 | -0.00040911 | -0.0016308 | -0.0012048 | -0.0028593 | -0.0026439 | -0.0033444 | -0.0033152 | 0.0010596 | -0.00090222 | 0.0012304 | 0.00026318 | 0.001734 | 0.00048334 | 0.00096258 | 0.0011902 | 0.0019227 | 0.050473
db14 | -0.0040781 | 0.0061579 | 0.006014 | 0.0021839 | -6.4743e-06 | 0.00087624 | -0.0014019 | -0.0054367 | -0.00071932 | -0.0013243 | -0.0026392 | -0.0017907 | 0.0033204 | -0.00014398 | -0.0032459 | -0.0028512 | -0.0009271 | 0.0017104 | 0.0021413 | 0.0042404 | 0.00060947
db15 | 0.037946 | -0.0016122 | 0.0078538 | 0.0011629 | -0.0023376 | -0.0034092 | -0.001044 | 0.00068641 | 0.00051262 | 5.846e-05 | 0.00078891 | -0.0022893 | -0.00054122 | 0.0010008 | 0.0010371 | -0.0017708 | -0.00019555 | -0.0033638 | 0.0047622 | -0.0029115 | 0.021605
da2 | 0.84385 | 10.5206 | 0.86182 | 0.092742 | -0.47336 | -0.56628 | -0.10582 | 0.11972 | -0.048052 | 0.35587 | -0.24755 | -0.21353 | -0.6924 | 0.080391 | 0.2913 | 0.0059807 | 0.63954 | 0.54897 | 0.22589 | -0.87525 | -4.3018
da3 | 0.34857 | 0.0076977 | 0.47298 | 0.014378 | 0.15492 | -0.095863 | 0.16188 | -0.11604 | -0.48659 | 0.15165 | -0.25067 | -0.15692 | 0.22937 | -0.20228 | -0.19365 | -0.26921 | 0.043841 | 0.24942 | 0.030764 | 0.26202 | -0.19691
da4 | -0.13119 | 2.6653 | 0.14234 | 0.28513 | 0.21276 | 0.049134 | 0.056535 | -0.32576 | 0.1525 | 0.09689 | 0.32061 | -0.15029 | -0.20943 | -0.19829 | -0.17239 | 0.053601 | -0.10635 | -0.18913 | -0.23472 | 0.21683 | -0.57401
da5 | -0.046571 | 2.9268 | -0.1113 | 0.088114 | 0.053692 | -0.033727 | -0.30213 | -0.10316 | -0.17444 | 0.0613 | 0.052648 | 0.064864 | 0.25923 | -0.1246 | 0.0068206 | -0.029076 | 0.00016606 | 0.051434 | 0.0012404 | 0.23894 | -0.048116
da6 | -0.010731 | 0.61022 | -0.13567 | -0.048171 | -0.012759 | 0.087012 | 0.065621 | 0.025875 | 1.2408e-05 | 0.02005 | 0.068616 | 0.066402 | 0.018372 | -0.028453 | -0.054057 | -0.048902 | -0.097109 | 0.023259 | -0.01551 | 0.065414 | -0.098417
da7 | -0.012271 | 1.5255 | -0.066464 | 0.021849 | -0.0028985 | -0.041745 | -0.072661 | -0.055877 | -0.01936 | 0.06222 | -0.0082774 | 0.008806 | 0.0064367 | -0.010822 | 0.038545 | -0.0013221 | -0.03433 | 0.038332 | 0.082101 | 0.055467 | 0.0048096
da8 | 0.007296 | 0.055266 | 0.012125 | -0.0042063 | -0.0035837 | 0.011774 | -0.019447 | -0.0037539 | -0.043342 | 0.0014702 | -1.212e-06 | 0.029621 | -0.0024372 | -0.0094357 | 0.0068352 | 0.023837 | -0.0020776 | 0.01623 | -0.020566 | 0.0069577 | -0.022263
da9 | -0.024177 | -0.28063 | 0.01835 | 0.011113 | 0.00039295 | -0.056873 | -0.033965 | -0.008712 | 0.012565 | 0.021828 | -0.032995 | 0.0062119 | 0.035937 | -0.0028356 | -0.015032 | -0.015573 | 0.015234 | 0.014595 | 0.018194 | 0.011565 | -0.07686
da10 | -0.0010312 | 0.063711 | -0.031722 | 0.0067694 | 0.015688 | 0.02884 | 0.0028854 | -0.0067094 | -0.015491 | 0.0052281 | 0.0056712 | 0.013155 | 0.0038254 | -0.011156 | -0.020384 | 0.0043337 | -0.0096605 | 0.021461 | -0.038362 | 0.025628 | -0.029928
da11 | -0.049509 | 0.13231 | -0.023665 | -0.0021547 | 7.0585e-05 | -0.010326 | -0.016407 | -0.0060403 | 0.0026228 | 0.0085072 | 0.0014919 | 0.0049857 | 0.0076792 | -0.00044499 | 0.0079398 | 0.0040522 | -0.0033135 | 0.0038926 | 0.010798 | 0.010312 | -0.033266
da12 | -0.0046726 | 0.0046056 | -0.0035964 | 0.0028351 | -0.00066389 | 0.0045841 | 0.0028725 | -0.00087364 | -0.0063782 | -0.0010859 | -0.00061941 | 0.0046376 | -0.00065525 | -0.0036228 | -0.0015509 | 0.00057623 | -0.00097083 | 0.0020121 | -0.0037948 | 0.0062945 | -0.006609
da13 | -0.0040046 | -0.031103 | 0.0018083 | 0.001571 | -0.0010185 | -0.0044806 | -0.0043162 | -0.00016559 | -0.0016119 | 0.0028497 | -0.00038658 | 0.00039151 | 0.0027619 | -0.0017803 | 0.00055228 | -0.00085428 | 3.0602e-05 | 0.0005271 | 0.0014777 | 0.0026439 | -0.0055844
da14 | -0.017828 | 0.013285 | -0.0016459 | 0.0010906 | 0.0015838 | 0.0029546 | -0.00018047 | -0.0028417 | -0.003325 | -0.0013105 | 0.00042085 | 0.0017206 | -6.3006e-05 | -0.0014962 | -0.0010541 | 0.0020711 | 0.00063374 | 0.0014282 | -0.002404 | 0.0024174 | -0.010144
da15 | -0.0071935 | 0.0074049 | 0.010061 | 0.0022201 | 0.0018283 | 0.0010377 | -0.0016717 | -0.0074364 | -0.00047258 | -0.0026492 | -0.0044517 | -0.0034148 | 0.0031458 | 0.00010818 | -0.0045052 | -0.0053285 | -0.0010957 | 0.0028747 | 0.0039554 | 0.0057946 | -0.0015182
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4053.5917
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.023 | -4250.819 | 3.5874 | -1.5919 | -0.58691 | 0.19662 | 0.13957 | 0.24527 | -0.70939 | 0.58251 | -0.1876 | 0.21507 | 0.98685 | -0.28826 | 0.85263 | 0.16474 | -0.024852 | 1.2687 | -1.2631 | 4.7501 | -3856.3644
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.064054 | 2.9746 | -1.456 | -0.52583 | -0.21478 | -0.19659 | -0.41776 | 0.29595 | 0.91874 | 0.70724 | 0.22674 | -0.29018 | 0.048259 | 0.19574 | 0.59722 | 0.50108 | -0.0050434 | -0.45463 | -0.10146 | 0.17125 | -0.78223
db2 | -0.075029 | -1.1213 | -0.78604 | -0.068711 | -0.3415 | 0.933 | -0.21925 | -0.050602 | 0.35721 | 0.089881 | 0.32635 | 0.099056 | 0.061908 | -0.63646 | 0.091087 | 0.21616 | 0.28044 | -0.010843 | 0.25211 | -0.59379 | -0.45903
db3 | -2.8532 | 43.1642 | -0.52608 | -0.05084 | -0.24007 | 0.41603 | -1.183 | 0.78787 | 0.6966 | 0.86368 | 0.086486 | -0.64672 | 0.23443 | 0.47717 | -2.65 | -0.87676 | 0.28975 | 1.1514 | 0.57435 | 0.59571 | -2.4028
db4 | -0.020483 | -0.97049 | 0.15936 | 0.026657 | -0.13184 | 0.10312 | 0.037248 | -0.043621 | -0.0044389 | 0.0058919 | 0.14751 | -0.0020952 | -0.13673 | -0.076183 | 0.11182 | -0.016509 | -0.13361 | 0.046026 | -0.059272 | -0.033344 | 0.028473
db5 | -0.81209 | -4.9043 | -0.2603 | -0.30818 | 0.065249 | 0.12603 | 0.13255 | 0.13851 | 0.14673 | 0.1627 | 0.20209 | 0.27098 | -0.0053812 | -0.027228 | 0.26639 | -0.075613 | -0.20934 | -0.26141 | -0.12984 | -0.23393 | -4.5695
db6 | -0.027688 | -0.077281 | 0.11742 | 0.07262 | 0.056911 | -0.19782 | -0.059632 | 0.042913 | 0.070586 | 0.095781 | -0.030831 | 0.0029071 | -0.060991 | -0.023478 | -0.11409 | -0.066985 | 0.033534 | 0.017495 | 0.035428 | 0.0082323 | -0.10951
db7 | 1.1436 | 2.9901 | -0.19045 | -0.13591 | -0.0099681 | -0.0022301 | 0.064594 | 0.059226 | 0.056189 | 0.015067 | 0.020747 | 0.11337 | 0.056812 | -0.01474 | 0.022152 | -0.033315 | -0.018325 | -0.016873 | 0.019192 | -0.0055391 | 0.84754
db8 | 0.0024881 | -0.041611 | 0.050671 | -0.015682 | -0.045487 | 0.0046218 | -0.043116 | -0.014202 | 0.023337 | -0.023139 | -0.0049753 | 0.011354 | 0.05347 | 0.02101 | -0.019115 | -0.031628 | -0.0051237 | -0.009712 | 0.0073502 | 0.040364 | -0.055548
db9 | 0.47553 | 0.37635 | 0.02976 | 0.0047947 | -0.023218 | -0.010168 | -0.028913 | 0.023593 | 0.009063 | 0.024562 | -0.0079277 | -0.018894 | -0.0048765 | -0.012829 | -0.010899 | -0.016297 | 0.026247 | 0.0061214 | 0.0068037 | 0.0030761 | 0.40526
db10 | -5.1133e-05 | 0.061494 | 0.10928 | -0.014979 | -0.032581 | -0.032065 | -0.053594 | -0.026635 | 0.037743 | 0.0019744 | 0.00014319 | -0.0065473 | 0.047297 | 0.021001 | -0.018123 | -0.034031 | -0.011681 | -0.014934 | -0.011242 | 0.038973 | -0.056143
db11 | 0.74354 | 0.59185 | -0.0021211 | -0.0053937 | 0.0013927 | 9.4178e-05 | 0.0019599 | -0.0056169 | -0.0017575 | 0.0066824 | -0.00081912 | 0.0072345 | 0.00046308 | 0.0018305 | 0.015188 | 0.006257 | -0.002186 | -0.0050255 | -0.0044053 | -0.013777 | 0.63944
db12 | 0.0015332 | 0.013211 | 0.0069454 | -0.0030429 | -0.0014232 | -0.0053811 | -0.0043279 | -0.0029813 | 0.0023598 | -0.00022118 | -0.00296 | -0.001156 | 0.0053309 | 0.002426 | -0.0034683 | -0.0019461 | 0.0046468 | -5.4101e-05 | 0.0010951 | 0.0041581 | -0.001132
db13 | 0.069485 | 0.088368 | 0.0018043 | 0.0012359 | -0.0011258 | -5.5236e-05 | 0.00040621 | -0.003633 | 0.00073826 | -2.2496e-05 | -0.0010654 | 0.0028351 | 0.00010356 | 0.0013771 | -0.0034571 | 0.00010009 | -0.0035869 | 0.00068824 | 0.0019749 | 0.0016823 | 0.054657
db14 | -0.0042852 | 0.0061101 | 0.0052654 | -0.0010274 | -0.00013618 | -0.0020766 | -0.0035816 | -0.00065733 | 0.0030547 | 0.00058379 | -0.00038378 | -0.00065029 | 0.001862 | 0.00041401 | -0.0020998 | -0.0015145 | 0.00024179 | -5.4625e-05 | -0.00092769 | 0.0016881 | -0.0043326
db15 | 0.026546 | -0.0065662 | 0.0045509 | 0.0014391 | -0.00097716 | -0.0024848 | -0.0022524 | -0.0031796 | -0.002285 | -0.0018449 | -0.0023411 | -0.00021878 | -0.002077 | 0.0017712 | 0.00024457 | -8.9107e-05 | 0.0014211 | 0.0015487 | 0.0044191 | 0.0023555 | 0.026174
da2 | -0.70142 | 5.9193 | 0.45675 | 0.83088 | 0.64773 | 1.1678 | 0.20055 | -1.2199 | -0.22221 | -0.65312 | 0.10601 | -0.45473 | -1.0063 | -0.03567 | -0.32535 | 0.41595 | 0.36082 | 0.028143 | 0.4312 | -0.72852 | 0.85576
da3 | 0.14951 | 0.82328 | 0.15745 | 0.10753 | -0.081519 | -0.16064 | -0.3717 | -0.15211 | -0.26963 | 0.14017 | -0.24175 | 0.48172 | 0.34159 | -0.16655 | -0.27843 | -0.17741 | -0.0087361 | 0.27112 | 0.011742 | 0.39714 | 0.18452
da4 | -0.017892 | 1.6229 | -0.04656 | -0.33398 | -0.42662 | -0.11168 | 0.0021393 | 0.37609 | 0.26719 | 0.37031 | -0.10417 | 0.23099 | -0.23413 | -0.36669 | -0.084488 | 0.020561 | 0.40044 | 0.0052977 | -0.063309 | 0.098612 | 0.47115
da5 | -0.021905 | 2.7125 | 0.029985 | -0.086418 | 0.075935 | 0.0075563 | -0.040403 | -0.1644 | 0.0084759 | 0.088275 | 0.082854 | 0.20039 | -0.0078299 | -0.16513 | -0.026191 | -0.18525 | -0.0040498 | -0.041753 | -0.013549 | 0.24149 | -0.15881
da6 | -0.040386 | 0.34408 | -0.015585 | -0.032853 | 0.062629 | 0.044273 | 0.10913 | 0.085559 | 0.10791 | 0.0068561 | 0.092655 | -0.029597 | 0.012814 | -0.11187 | -0.16577 | -0.11151 | -0.039307 | -0.032126 | -0.019671 | 0.03647 | -0.23156
da7 | 0.032304 | 1.8027 | -0.01607 | -0.020323 | -0.055637 | 0.019978 | -0.10797 | -0.0076996 | 0.07017 | 0.024415 | 0.032635 | -0.021793 | 0.0020601 | -0.056257 | 0.019626 | 0.021828 | 0.023643 | 0.035182 | -0.0017339 | 0.037949 | 0.11616
da8 | -0.0083551 | 0.0068092 | 0.011167 | 0.010202 | 0.024917 | 0.061652 | 0.047719 | 0.017488 | 0.025265 | -0.020286 | 0.013765 | -0.020369 | -0.028043 | -0.036245 | -0.030345 | -0.07203 | -0.019514 | -0.017709 | 0.013295 | 0.019072 | -0.080387
da9 | -0.0084965 | -0.23319 | 0.018501 | 0.012037 | -0.021103 | -0.070574 | -0.044225 | 0.0061849 | 0.034963 | 0.032673 | -0.018609 | 0.023143 | 0.020438 | 0.0001111 | -0.014724 | -0.036311 | 0.030453 | 0.017521 | 0.0040069 | 0.0055133 | -0.038955
da10 | 0.0054645 | -0.0018558 | -0.0026126 | -0.0037994 | -0.01195 | 0.018913 | -0.0020265 | 0.036163 | 0.04896 | 0.035428 | 0.019902 | -0.020412 | 0.01722 | -0.037376 | -0.020944 | -0.032576 | 0.0011766 | -0.02029 | -0.02728 | 0.0015042 | -0.15838
da11 | 0.00030324 | 0.18124 | -0.021153 | -0.010188 | 0.0014275 | -0.0058152 | -0.011719 | -0.001336 | 0.010955 | 0.015918 | -0.0016145 | 0.00096563 | 0.0048111 | -0.0049691 | 0.0035705 | -0.0076001 | 0.0019079 | 0.0067224 | 0.0037973 | 0.014319 | 0.020786
da12 | -0.0029525 | -8.7616e-05 | 0.0032774 | 0.0015267 | -0.0023763 | 0.0026595 | 0.001202 | 0.0024066 | 0.008168 | 0.0031739 | 0.0066697 | -0.0023177 | 0.0018479 | -0.009035 | -0.0089708 | -0.006091 | -0.0016155 | -0.0013004 | -0.0023789 | 0.0031539 | -0.016154
da13 | 0.00075938 | -0.025526 | -0.00016989 | 0.0012924 | -0.0032418 | -0.0015224 | -0.0056216 | -0.00045314 | 0.0033856 | 0.0015903 | -0.00099334 | 0.001157 | 0.00090038 | -0.0023356 | -0.00097785 | -0.00090448 | 0.0015465 | 0.0026067 | 0.0012139 | 0.0025272 | -0.00066543
da14 | -0.010486 | 0.012908 | -0.00016681 | -2.2725e-05 | -0.00015587 | 0.0016254 | 0.0017465 | 0.0013742 | 0.0026072 | 0.0013974 | 0.0018438 | 8.6055e-05 | -0.00018642 | -0.0030993 | -0.00090456 | -0.0016981 | -0.0010987 | -0.0013654 | -0.0019369 | -4.5795e-05 | -0.0056529
da15 | -0.0074156 | 0.00055144 | 0.0066449 | -0.0016419 | -0.0030867 | -0.0028299 | -0.0056222 | -0.0013604 | 0.0045914 | 0.0018188 | 0.00025146 | -0.00061431 | 0.0043956 | 0.0011954 | -0.0033194 | -0.0034753 | 0.0002954 | -0.00041748 | -0.00072562 | 0.0039002 | -0.017403
Field Colinearity | 19.6784
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 0.56889 | -1.3016 | -0.61493 | -0.26193 | -0.26792 | -0.15142 | 0.37705 | 0.81166 | 0.58944 | 0.099261 | -0.42283 | -0.14159 | 0.33508 | 0.76688 | 0.49306 | -0.030798 | -0.37643 | -0.03561 | 0.067249 | -0.42877
Twist (m2 mrad) | 4.2667 | -18.548 | -8.3015 | -3.3396 | -3.215 | -1.7035 | 3.959 | 7.9137 | 5.305 | 0.8189 | -3.1712 | -0.95575 | 2.0105 | 4.0261 | 2.2188 | -0.11549 | -1.1293 | -0.080122 | 0.10087 | -0.16079
Twist Integral (I0) (m2 rad) | -0.0075755
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.9106 | -1.52 | -0.58988 | -0.27883 | -0.26065 | -0.48182 | 0.2319 | 0.85468 | 0.64319 | 0.16269 | -0.35423 | -0.015795 | 0.13168 | 0.53317 | 0.43703 | -0.069097 | -0.51868 | -0.16551 | 0.1072 | -0.84629
Twist (m2 mrad) | 21.8292 | -21.6601 | -7.9634 | -3.5551 | -3.1278 | -5.4204 | 2.4349 | 8.3332 | 5.7887 | 1.3422 | -2.6567 | -0.10662 | 0.79009 | 2.7991 | 1.9666 | -0.25911 | -1.556 | -0.3724 | 0.1608 | -0.31736
Twist Integral (I0) (m2 rad) | -0.0011627

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.15832 | -2.9197 | 0.034382 | -0.55951 | 1.1149 | -0.036323 | 0.84385 | 0.34857 | -0.13119 | 15.1627 | 0.59601 | -0.075029 | -2.8532 | -0.020483 | -0.81209 | 1.1436 | -0.64054 | -0.70142 | 0.14951 | -0.017892 | 18.6916 | 0.59605 | 0.092379 | -1.5116 | -0.0024857 | -0.80801 | 1.1614 | 0 | 0.97139 | 0.32069 | -0.060697 | 14419.041 | -0.1187 | -1.4602 | -0.047373 | -1.055 | 1.1895 | 0 | -0.45325 | 0.17083 | 0.046925 | 14419.901
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8