## Deckblatt
| LHC - Projekt - CERN |  | Auftrag-Nr. / Order-No.: | N.911004
Magnetfeldmessung Collared Coil / Magnetic Field Measuring Collared Coil
Internes Ablage-Register/ Internal registration: | 'file:///tmp/in/input.xls'#$Deckblatt |  | Dok.-Nr./ Doc.ID-No.: | CF0053_0__N-3118C
Seite / Page:
BBP-SPA code: |  | --- | Datum / Date of report: | 2003-12-09
Beginn der Messung / start of measuring: |  | TT.12.JJ | Collared Coil-Nr. / collared coil-no.: | HCMB__A001-03000118_cc
Ende der Messung / end of measuring: |  | TT.12.JJ
ITP: |  | 11 | Status / status:
Spezifikations Nr. / Spezification No.: |  | LHC-MMS/98-198 Rev. 2.0, Annex B 18
Messgeräte / equipment: |  | Dipole Mole 15
| Prüfer / Tester |  | Prüfaufsicht / Supervisor | Freigabe / Approval
Name /Name | Kenklies |  | Zehner | Zehner
Firma / Company | BNM |  | BNN | BNN
Datum / Date | 2003-12-09 |  | 2003-12-09 | 2003-12-09
Unterschrift / Signature
CERN-Freigabe: | via E-Mail vom:
CERN release: | via e-mail dated:
Dieses Schriftstück unterliegt samt seinem Inhalt dem Schutz des Urheberrechts sowie der Pflicht zur vertraulichen Behandlung und darf nur mit ausdrücklicher Genehmigung von BNN GmbH genutzt, vervielfältigt, Dritten zugänglich gemacht oder in anderer Weise verwendet werden. Ist BNN GmbH nicht Urheber (Ersteller) des Schriftstücks, so kann die Zustimmung im Rahmen des mit BNN GmbH bestehenden Vertrages auch vom Urheber erteilt werden.

## Alarm sheet
File name | HCMB__A001-03000118_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000118
Date of test Ap 1 | 2003-12-09 |  |  |  | Date of test Ap 2 | 2003-12-09
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | yellow alarm | status ok | b9 | status ok | status ok | yellow alarm | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | yellow alarm
b15 | status ok | status ok | status ok | yellow alarm | b15 | status ok | status ok | yellow alarm | yellow alarm
a2 | status ok | status ok | status ok | red alarm | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | yellow alarm
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
258 | HCMB__A001 | 3000118 |  | 3 |  |  | 3
| COMMENTS: | Red alarm on a2 in coil head NCS, ap. 1 - asked for re-measurement.
| HOLDING POINT | HOLD | Data at CERN |  | 2003-12-09 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2003-12-09 | Answer to manufacturer |  |  | 2003-12-12
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-03000118_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000118 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3118_cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 3118_cct2 |  |  |  | a even | 1
Test Operator |  | Kenklies |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-12-09 |  |  | Time of test |  |  | 6:47 |  |  |  | Date of test |  | 2003-12-09 |  |  | Time of test |  |  | 6:48
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-02-05
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.61 |  |  | I. Gain |  |  | 50/500
Dx average |  | -0.00016679 |  |  | Dy average |  |  | 0.0010839 |  |  |  | Dx average |  | 0.0001639 |  |  | Dy average |  |  | 0.0010766
Magnetic length |  | 14.4206 |  |  | Tranfs. Function (mT/KA) |  |  | 595.9537 |  |  |  | Magnetic length |  | 14.4199 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0202
Phase |  | -13.803 |  |  | Integral field (T.m) |  |  | 85.8269 |  |  |  | Phase |  | -33.762 |  |  | Integral field (T.m) |  |  | 85.773
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.5101 | 5.9619 | 5.9586 | 5.9587 | 5.958 | 5.9595 | 5.9594 | 5.9597 | 5.959 | 5.9587 | 5.9589 | 5.959 | 5.9589 | 5.9594 | 5.9595 | 5.9591 | 5.9596 | 5.9597 | 5.964 | 3.67 | 85.8269 | C1 (T) | 3.4664 | 5.9629 | 5.9597 | 5.959 | 5.9596 | 5.9609 | 5.9602 | 5.9599 | 5.9596 | 5.9596 | 5.9596 | 5.9603 | 5.9597 | 5.9595 | 5.9596 | 5.9593 | 5.96 | 5.9602 | 5.9641 | 3.6293 | 85.773
Angle (mrad) | 2.1867 | 0.33994 | -0.25041 | -0.55251 | -0.14444 | -5.7e-05 | 0.36158 | 0.44185 | 0.060854 | -0.41061 | -0.46966 | -0.055777 | 0.50268 | 0.16167 | -0.21346 | -0.47312 | -0.48209 | -0.081662 | -0.24235 | 0.35648 | -13.803 | Angle (mrad) | 2.7283 | -0.4739 | -0.20135 | -0.6282 | -0.2964 | -0.063673 | 0.24605 | 0.73764 | 0.21517 | -0.46113 | -0.40638 | 0.11448 | 0.32998 | 0.21746 | -0.16246 | -0.16666 | -0.38128 | 0.060228 | 0.18369 | -0.73911 | -33.762
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -0.88982 | -0.60645 | -0.49848 | -1.2095 | -1.3474 | -0.48388 | -0.87483 | -0.94921 | -0.71443 | -1.2401 | -1.8852 | -1.7713 | -1.618 | -2.1984 | -1.7565 | -0.47418 | -1.0399 | -1.5225 | -0.92762 | 0.79543 | -1.1014 | b2 | 2.5836 | 2.2018 | 0.87891 | 1.3773 | 1.27 | 1.6981 | 2.1824 | 2.5068 | 1.8132 | 0.65963 | 0.92422 | 1.6428 | 1.5866 | 0.94932 | 0.81711 | 0.5209 | 1.958 | 1.7485 | 1.8935 | 2.2851 | 1.5385
b3 | 36.3962 | -1.9138 | -1.7863 | -3.8348 | -1.6965 | -4.4628 | -3.1195 | -3.5747 | -2.039 | -2.906 | -3.6847 | -2.8564 | -2.3534 | -4.4431 | -3.6059 | -3.2357 | -1.6853 | -2.3549 | -1.1811 | -0.96044 | -1.5563 | b3 | 35.1777 | -1.7858 | -1.6993 | -1.8215 | -1.7203 | -4.4014 | -2.4578 | -1.8597 | -1.5271 | -2.7151 | -3.2367 | -2.8244 | -1.9681 | -4.4347 | -3.403 | -2.4528 | -1.2843 | -1.7835 | -0.71293 | -0.96574 | -1.1577
b4 | -0.42137 | 0.05927 | 0.036442 | -0.087728 | 0.11185 | 0.24982 | 0.19754 | 0.11215 | 0.25325 | 0.097126 | -0.047117 | -0.038652 | 0.040115 | 0.056909 | 0.092941 | 0.27242 | -0.016537 | 0.001935 | 0.10207 | 0.22236 | 0.071992 | b4 | 0.052715 | 0.048906 | -0.094741 | -0.058613 | -0.10286 | -0.088754 | -0.020745 | 0.073975 | 0.012401 | -0.082304 | -0.29391 | 0.093596 | -0.11897 | -0.22679 | -0.32682 | -0.38139 | 0.092127 | 0.017841 | 0.089594 | 0.15728 | -0.064656
b5 | -5.4313 | -0.37235 | -0.51366 | -0.222 | -0.38463 | 0.22412 | 0.22092 | 0.31819 | -0.1506 | -0.0398 | 0.25003 | 0.086559 | 0.16392 | 0.18377 | -0.0028846 | -0.083263 | -0.2169 | -0.17315 | -0.22598 | -3.9022 | -0.34051 | b5 | -6.253 | -0.7604 | -0.93614 | -0.47975 | -0.43515 | 0.14011 | 0.09301 | -0.2116 | -0.30372 | -0.29701 | -0.062618 | -0.18746 | -0.20317 | -0.13863 | -0.16835 | -0.65271 | -0.52084 | -0.35844 | -0.50481 | -4.1706 | -0.63385
b6 | -0.36081 | -0.12958 | -0.071651 | -0.015327 | -0.078963 | -0.14924 | -0.044693 | -0.1478 | -0.12974 | -0.091991 | 0.0269 | 0.0072535 | 0.014056 | 0.041259 | -0.0054871 | -0.16528 | 0.069851 | 0.044077 | 0.0028 | -0.0021634 | -0.054016 | b6 | 0.1979 | -0.0061714 | 0.024088 | -0.067097 | -0.052607 | -0.061741 | 0.034713 | -0.067767 | -0.099918 | 0.033036 | -0.01614 | 0.044494 | 0.031835 | 0.01481 | 0.074146 | 0.14898 | 0.055256 | -0.0034513 | -0.039954 | -0.10571 | 0.0050645
b7 | 2.8582 | 0.9609 | 1.0641 | 1.2401 | 1.1794 | 1.2233 | 1.1839 | 1.143 | 1.1201 | 1.1453 | 1.1868 | 1.1593 | 1.13 | 1.1121 | 1.1504 | 1.2327 | 1.1328 | 1.1707 | 1.1218 | 0.93179 | 1.1931 | b7 | 2.6828 | 0.97597 | 1.0315 | 1.2195 | 1.2099 | 1.2473 | 1.1942 | 1.1616 | 1.1528 | 1.219 | 1.2325 | 1.2276 | 1.2284 | 1.1579 | 1.1562 | 1.2946 | 1.202 | 1.2038 | 1.1632 | 0.88713 | 1.2182
b8 | -0.042877 | 0.024875 | 0.034115 | 0.024765 | 0.031952 | -0.0065308 | 0.015943 | 0.030744 | 0.02112 | 0.018689 | 0.030658 | 0.025376 | 0.030851 | 0.042937 | 0.036457 | 0.074942 | 0.060302 | 0.047448 | 0.0075551 | -0.023701 | 0.026678 | b8 | 0.00062977 | 0.011913 | 0.029507 | -0.021473 | -0.010671 | 0.00072533 | 0.062404 | 0.047452 | -0.0079872 | -0.0053364 | 0.0060967 | -0.010537 | -0.0073326 | 0.032645 | -0.020203 | -0.029074 | 0.0020041 | 0.026347 | -0.012448 | -0.058832 | 0.0030524
b9 | 0.35889 | 0.5111 | 0.49978 | 0.49558 | 0.52633 | 0.48193 | 0.51052 | 0.5068 | 0.52486 | 0.48884 | 0.47973 | 0.49545 | 0.49778 | 0.48363 | 0.47327 | 0.50162 | 0.48516 | 0.51093 | 0.5065 | 0.44184 | 0.49276 | b9 | 0.33135 | 0.51687 | 0.5126 | 0.5313 | 0.5349 | 0.49476 | 0.50704 | 0.54029 | 0.53402 | 0.51896 | 0.51356 | 0.50925 | 0.5226 | 0.50794 | 0.49143 | 0.54244 | 0.51575 | 0.54222 | 0.52041 | 0.44726 | 0.51178
b10 | -0.036958 | -0.063961 | -0.0295 | -0.010429 | -0.013386 | -0.084003 | 0.015648 | -0.018507 | -0.028247 | 0.0059425 | 0.053075 | 0.021992 | 0.0099026 | 0.084641 | 0.027027 | 0.036761 | 0.061364 | 0.02971 | -0.048418 | -0.04512 | 0 | b10 | 0.067155 | -0.00090196 | -0.031521 | -0.038039 | -0.029922 | -0.077449 | 0.044675 | 0.03553 | -0.0098596 | 0.028164 | 0.052762 | 0.019709 | 0.0024553 | 0.05461 | -0.0047888 | -0.044425 | 0.023301 | 0.043977 | -0.048079 | -0.097251 | 0
b11 | 0.61314 | 0.73995 | 0.7358 | 0.74812 | 0.74812 | 0.7562 | 0.75099 | 0.75812 | 0.75222 | 0.7495 | 0.75917 | 0.7485 | 0.75053 | 0.76165 | 0.7566 | 0.74465 | 0.73927 | 0.7485 | 0.73651 | 0.63864 | 0.74142 | b11 | 0.5943 | 0.75323 | 0.75173 | 0.7561 | 0.76174 | 0.77685 | 0.76644 | 0.76381 | 0.76751 | 0.76664 | 0.77408 | 0.7607 | 0.76032 | 0.77615 | 0.76879 | 0.75068 | 0.75392 | 0.75291 | 0.74556 | 0.65913 | 0.75319
b12 | -0.0096978 | -0.0036694 | 0.0018511 | 0.00054768 | -0.0025883 | -0.010234 | 4.0911e-06 | -0.0026941 | -0.0071164 | -0.00017891 | 0.004156 | 0.0014234 | -0.0030667 | 0.0052295 | 0.0028448 | 0.0018248 | 0.0080635 | 0.0031929 | -0.0041067 | -0.0094632 | -0.00083621 | b12 | 0.015673 | -0.0013314 | -0.0036942 | -0.0018023 | -0.00018387 | -0.0024677 | 0.0093327 | 0.005791 | -0.0030429 | -1.9484e-05 | 0.0078914 | -0.00080159 | -0.0010387 | 0.0072915 | 0.00071879 | 0.003186 | 0.0018034 | 0.0018032 | -0.0060552 | -0.010648 | 0.0010425
b13 | 0.08906 | 0.066416 | 0.06957 | 0.065974 | 0.067146 | 0.062504 | 0.061473 | 0.058511 | 0.060464 | 0.061674 | 0.060573 | 0.061813 | 0.061851 | 0.060447 | 0.064439 | 0.062544 | 0.065126 | 0.068408 | 0.06622 | 0.050701 | 0.063986 | b13 | 0.094528 | 0.068824 | 0.069606 | 0.065785 | 0.064629 | 0.059606 | 0.060338 | 0.061113 | 0.064161 | 0.063238 | 0.063052 | 0.062327 | 0.061674 | 0.060428 | 0.059447 | 0.062715 | 0.063238 | 0.062439 | 0.066705 | 0.054207 | 0.063954
b14 | 0.0062204 | -0.0020234 | 0.0026105 | 0.0046605 | 0.0024698 | -0.0029696 | 0.0048532 | 0.0046605 | 0.0012752 | 0.0031977 | 0.0078156 | 0.0054856 | 0.0051364 | 0.011175 | 0.0075976 | 0.0063216 | 0.010971 | 0.0090549 | 0.0023489 | 0.0022297 | 0.0046694 | b14 | 0.0095971 | -0.0069836 | -0.0077409 | -0.0072695 | -0.007516 | -0.011321 | -0.0019394 | -0.0026517 | -0.0062269 | -0.0052303 | -0.001612 | -0.0056805 | -0.0078145 | -0.00083876 | -0.0061252 | -0.0072938 | -0.0050847 | -0.0028231 | -0.00988 | -0.012903 | -0.0055398
b15 | 0.011784 | 0.050736 | 0.048609 | 0.045914 | 0.047111 | 0.048856 | 0.046802 | 0.052828 | 0.052041 | 0.053026 | 0.053177 | 0.052579 | 0.054021 | 0.054596 | 0.053972 | 0.05082 | 0.048295 | 0.050735 | 0.048749 | 0.041698 | 0.049232 | b15 | 0.021227 | 0.043994 | 0.040858 | 0.038136 | 0.038574 | 0.044022 | 0.043423 | 0.044185 | 0.043409 | 0.046556 | 0.048377 | 0.047498 | 0.048651 | 0.053654 | 0.054519 | 0.044975 | 0.048121 | 0.045846 | 0.046762 | 0.045003 | 0.044882
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 21.8675 | 3.3994 | -2.5041 | -5.5251 | -1.4444 | -0.0005695 | 3.6158 | 4.4185 | 0.60854 | -4.1061 | -4.6966 | -0.55777 | 5.0268 | 1.6167 | -2.1346 | -4.7312 | -4.8209 | -0.81662 | -2.4235 | 3.5648 | 0 | a1 | 27.2826 | -4.739 | -2.0135 | -6.282 | -2.964 | -0.63673 | 2.4605 | 7.3764 | 2.1517 | -4.6113 | -4.0638 | 1.1448 | 3.2998 | 2.1746 | -1.6246 | -1.6666 | -3.8128 | 0.60228 | 1.8369 | -7.3911 | 0
a2 | 11.6554 | 0.37026 | 0.61628 | 0.98377 | 0.20402 | 0.35586 | 0.7975 | -0.27536 | -0.047081 | 1.2273 | 1.0927 | 1.2433 | 0.94919 | 1.788 | 0.84976 | 0.71476 | 0.35518 | -0.5615 | 1.2783 | 23.04 | 1.718 | a2 | 8.6461 | -1.1114 | -2.0139 | -2.1316 | -1.3134 | -2.5707 | -1.514 | -2.2981 | -1.2468 | -1.2381 | -0.92407 | -0.29956 | -0.43176 | -1.1659 | -1.7818 | -2.6697 | -1.9388 | -1.5342 | -0.92844 | 5.8207 | -0.96607
a3 | -0.63057 | 1.1665 | 0.97873 | 0.56709 | 0.64371 | 0.51523 | 0.69645 | 0.79672 | 0.5662 | 0.3273 | 0.55222 | 0.4609 | 0.56876 | 0.43611 | 0.33516 | 0.53841 | 0.76863 | 1.0662 | 1.223 | 0.17187 | 0.62185 | a3 | 0.49784 | 0.69483 | 0.74237 | 0.31794 | 0.24123 | 0.17545 | -0.14809 | 0.27531 | 0.64046 | 0.17455 | 0.15516 | -0.079982 | 0.088449 | 0.24941 | 0.17803 | 0.51343 | 0.72741 | 1.1547 | 1.3379 | 0.9429 | 0.43266
a4 | 1.8584 | 0.11513 | 0.18875 | -0.23845 | -0.078463 | 0.3894 | 0.30474 | 0.42144 | 0.07719 | -0.084352 | -0.018715 | -0.084636 | -0.23485 | -0.17244 | 0.20097 | -0.02385 | 0.34309 | 0.23666 | 0.18986 | 2.0608 | 0.20282 | a4 | 2.5362 | 0.24766 | 0.13665 | 0.34886 | 0.19377 | 0.010046 | -0.24801 | 0.20006 | -0.14531 | -0.43633 | 0.1872 | -0.48338 | -0.20779 | -0.091458 | -0.2453 | -0.088433 | -0.30459 | -0.43076 | -0.15304 | 1.0016 | 0.029946
a5 | 2.6802 | 0.28017 | 0.14574 | -0.0015221 | 0.21785 | -0.011434 | 0.017666 | 0.17299 | 0.20154 | 0.21029 | 0.1421 | 0.043401 | -0.040011 | 0.29872 | 0.20268 | 0.33134 | 0.067586 | 0.3587 | 0.58875 | -0.005734 | 0.25004 | a5 | 2.7915 | 0.15332 | -0.095916 | -0.10683 | 0.11914 | -0.0060381 | -0.064175 | -0.040437 | 0.17315 | 0.41318 | 0.13247 | -0.17191 | -0.16686 | 0.237 | 0.0087675 | -0.10697 | 0.10254 | 0.17357 | 0.37506 | -0.035274 | 0.14233
a6 | -0.1356 | 0.19074 | 0.10791 | 0.020566 | -0.00085962 | -0.017991 | -0.096111 | -0.095754 | -0.029721 | -0.082655 | -0.042595 | -0.12501 | -0.097572 | -0.12683 | -0.070673 | -0.16239 | -0.080376 | -0.070897 | 0.012361 | 0.11117 | -0.040571 | a6 | -0.33612 | 0.15058 | 0.099678 | 0.11374 | 0.0245 | -0.046041 | -0.11742 | -0.026773 | -0.041724 | -0.0014942 | -0.00057246 | 0.13349 | 0.053182 | 0.011385 | -0.088645 | -0.13487 | -0.00070445 | 0.097801 | 0.068615 | 0.012298 | 0.0055672
a7 | 1.4856 | -0.11288 | -0.055218 | -0.061075 | 0.0567 | 0.029371 | 0.015534 | 0.060163 | 0.018089 | -0.016879 | -0.0080649 | -0.0083233 | -0.046068 | 0.028263 | 0.057888 | -0.04792 | -0.013753 | 0.10058 | 0.12586 | 0.049812 | 0.053526 | a7 | 1.7031 | -0.063194 | -0.014134 | -0.030571 | -0.036077 | -0.028967 | 0.072772 | 0.14104 | 0.036236 | -0.027989 | 0.023597 | -0.048347 | -0.051843 | 0.082464 | 0.024156 | 0.012847 | 0.0017386 | 0.08181 | 0.096535 | 0.13211 | 0.069984
a8 | -0.081264 | 0.0082756 | 0.031268 | 0.059553 | 0.024176 | 0.0016517 | -0.00055808 | -0.024315 | -0.0025222 | 0.0063085 | -0.056901 | -0.058511 | -0.061901 | -0.03719 | -0.060976 | -0.034553 | -0.014822 | -0.0059821 | 0.0035826 | -0.020011 | -0.014767 | a8 | -0.15361 | 0.01511 | -0.0080508 | -0.0066461 | -0.021411 | -0.087487 | -0.035593 | -0.013277 | -0.02055 | 0.01834 | -0.023162 | 0.039391 | 0.0095868 | -0.0055268 | -0.063894 | -0.077092 | -0.030717 | -0.011108 | -0.012454 | -0.092381 | -0.025018
a9 | -0.25173 | -0.0074864 | -0.0033989 | 0.0014847 | -0.024538 | -0.047785 | -0.034794 | 0.0049853 | -0.01437 | -0.027454 | -0.031094 | -0.034511 | -0.01233 | -0.017609 | -0.044087 | -0.029035 | -0.020889 | 0.02436 | 0.025612 | -0.051088 | -0.024609 | a9 | -0.2555 | 0.061327 | 0.043039 | -0.0073909 | -0.0090511 | -0.021922 | -0.020674 | 0.015174 | 0.015451 | -0.012663 | -0.027781 | -0.0052284 | 0.020004 | -0.011452 | -0.034866 | -0.020504 | -0.0083515 | 0.024172 | 0.038036 | -0.010305 | -0.0061225
a10 | -0.033115 | 0.039878 | 0.036462 | 0.03109 | 0.018457 | 0.025969 | 0.04074 | 0.0016809 | 0.0050997 | -0.013068 | -0.014465 | -0.019153 | -0.025823 | -0.03988 | -0.03731 | -0.018127 | 0.0060422 | 0.0014047 | 0.027065 | -0.075666 | 0 | a10 | -0.12659 | 0.048015 | 0.06027 | 0.058703 | 0.049271 | -0.010853 | 0.030565 | 0.029682 | 0.020297 | -0.0023219 | -0.00035768 | 0.0046413 | -0.0076167 | -0.025931 | -0.064282 | -0.020286 | -0.017653 | 0.025426 | -0.01089 | -0.15284 | 0
a11 | 0.1366 | -0.064618 | -0.054665 | -0.05164 | -0.048674 | -0.042937 | -0.037404 | -0.026656 | -0.040334 | -0.051263 | -0.038053 | -0.039524 | -0.035738 | -0.040668 | -0.042205 | -0.042015 | -0.045175 | -0.028075 | -0.024982 | -0.017292 | -0.035659 | a11 | 0.18981 | -0.012155 | -0.0040235 | -0.016712 | -0.0072816 | -0.00052664 | 0.0051542 | 0.020149 | 0.011975 | -0.00093006 | 0.0020997 | 0.0038801 | 0.0074341 | 0.0050078 | -0.0016516 | -0.00035112 | 0.010568 | 0.016402 | 0.01291 | 0.016494 | 0.0089826
a12 | 0.0023268 | 0.0045927 | 0.0053574 | 0.00044927 | -0.0017174 | -0.0021592 | -0.0018541 | -0.0060054 | -0.0031146 | -0.0058507 | 4.6073e-05 | -0.0040185 | -0.0064513 | -0.011101 | -0.0094176 | -0.0066821 | -0.0017326 | -0.0056531 | 0.0019827 | -0.0077817 | -0.0029548 | a12 | -0.022269 | 0.0074455 | 0.0068676 | 0.0044131 | -0.00025201 | -0.0047019 | -0.0045897 | -0.003012 | -0.0055485 | -0.0062436 | -0.0031884 | -0.0032909 | -0.00088309 | -0.0036159 | -0.015648 | -0.011437 | -0.0083776 | 0.00020319 | -0.0022441 | -0.017579 | -0.0040516
a13 | -0.029564 | -0.00047472 | -0.0026933 | -0.0068428 | -0.003728 | -0.0042207 | -0.00281 | 0.00021539 | -0.0019801 | -0.0043624 | -0.0049175 | -0.0039812 | -0.0040942 | -0.0036175 | -0.0056027 | -0.0020303 | -0.00502 | 0.00019522 | 0.0037023 | -7.6587e-05 | -0.0036301 | a13 | -0.026176 | 0.0056589 | 0.0076274 | 0.00068463 | 0.0029973 | 0.0026072 | 5.7036e-06 | 0.0024021 | 0.0040071 | -3.9505e-05 | -0.0015399 | 0.0012798 | 0.0018436 | 0.0018159 | 0.0028379 | 0.0032331 | 0.0015302 | 0.0041567 | 0.0091047 | 0.0059631 | 0.0020128
a14 | 0.0084365 | -0.026356 | -0.02297 | -0.02042 | -0.021153 | -0.021544 | -0.020944 | -0.024939 | -0.025618 | -0.025692 | -0.027227 | -0.027519 | -0.028117 | -0.027581 | -0.026962 | -0.024254 | -0.02134 | -0.023293 | -0.021697 | -0.015767 | -0.023034 | a14 | 0.011983 | -0.020618 | -0.020694 | -0.019516 | -0.020548 | -0.022772 | -0.023826 | -0.021085 | -0.022371 | -0.022839 | -0.02416 | -0.023739 | -0.024133 | -0.026163 | -0.031128 | -0.024665 | -0.026555 | -0.027149 | -0.022883 | -0.016792 | -0.022308
a15 | 0.0025394 | -0.0018022 | 0.0027385 | 0.0050303 | 0.003609 | -0.0046651 | 0.0062339 | 0.0050313 | 0.001639 | 0.0047369 | 0.011376 | 0.0083537 | 0.0071364 | 0.015861 | 0.011201 | 0.010013 | 0.015142 | 0.012578 | 0.0022443 | 0.00017009 | 0.006147 | a15 | -0.0067136 | -0.01007 | -0.011175 | -0.012616 | -0.01259 | -0.019299 | -0.0030369 | -0.0036033 | -0.011222 | -0.0081272 | -0.0049296 | -0.010227 | -0.012329 | -0.0033532 | -0.011397 | -0.01313 | -0.0089961 | -0.0046439 | -0.015248 | -0.02424 | -0.010144
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 0.00011718 | 0.00015375 | 7.3988e-05 | 2.8443e-05 | 3.3008e-05 | 0.00019168 | -3.0761e-05 | 4.1582e-05 | 6.4284e-05 | -1.5432e-05 | -0.00012019 | -5.2095e-05 | -2.5155e-05 | -0.00019304 | -6.5162e-05 | -8.597e-05 | -0.00013979 | -6.7268e-05 | 0.00011375 | 0.00011432 | -0.00016679 | Dx (m) | -6.8027e-05 | 0 | 7.1957e-05 | 8.8375e-05 | 6.7837e-05 | 0.00016945 | -9.9592e-05 | -8.0788e-05 | 2.1135e-05 | -6.2456e-05 | -0.0001159 | -4.4097e-05 | 0 | -0.00011925 | 1.0274e-05 | 0.00010053 | -5.196e-05 | -0.0001005 | 0.00011004 | 0.000259 | 0.0001639
Dy (m) | 6.5751e-05 | -7.823e-05 | -7.873e-05 | -6.8684e-05 | -3.9796e-05 | -4.751e-05 | -9.3764e-05 | 0 | 0 | 2.8584e-05 | 2.6371e-05 | 4.0753e-05 | 5.7282e-05 | 7.8867e-05 | 8.0199e-05 | 3.6549e-05 | -2.2437e-05 | 0 | -5.8631e-05 | 0.00020463 | 0.0010839 | Dy (m) | 0.00038328 | -0.00010828 | -0.0001359 | -0.00013003 | -0.00010931 | 2.3915e-05 | -6.7158e-05 | -6.3949e-05 | -4.5293e-05 | 0 | 0 | -1.0144e-05 | 1.7082e-05 | 5.7611e-05 | 0.00014205 | 4.6006e-05 | 4.054e-05 | -5.522e-05 | 2.2944e-05 | 0.00038798 | 0.0010766
Magnetic length (m) |  | 14.4206 | [19 empty] | Magnetic length (m) |  | 14.4199
Transfer funct. (mT/KA) |  | 595.9537 | [19 empty] | Transfer funct. (mT/KA) |  | 596.0202
Coil positioning (mm) |  | 0.019342 | [19 empty] | Coil positioning (mm) |  | 0.02245
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-3000118 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-3000118 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 351.0104 | 596.1856 | 595.8594 | 595.8703 | 595.8033 | 595.9494 | 595.9414 | 595.9672 | 595.9042 | 595.8747 | 595.8917 | 595.9046 | 595.8865 | 595.9415 | 595.9508 | 595.9081 | 595.9567 | 595.9702 | 596.4007 | 366.9954 | 85.8269 | C1/i (mT/KA) | 346.6426 | 596.2939 | 595.9687 | 595.9034 | 595.9573 | 596.0899 | 596.0235 | 595.9923 | 595.955 | 595.9606 | 595.9619 | 596.0264 | 595.9721 | 595.9472 | 595.9643 | 595.9252 | 596.0011 | 596.0161 | 596.405 | 362.9273 | 85.773
Angle (mrad) | 2.1867 | 0.33994 | -0.25041 | -0.55251 | -0.14444 | -5.7e-05 | 0.36158 | 0.44185 | 0.060854 | -0.41061 | -0.46966 | -0.055777 | 0.50268 | 0.16167 | -0.21346 | -0.47312 | -0.48209 | -0.081662 | -0.24235 | 0.35648 | -13.803 | Angle (mrad) | 2.7283 | -0.4739 | -0.20135 | -0.6282 | -0.2964 | -0.063673 | 0.24605 | 0.73764 | 0.21517 | -0.46113 | -0.40638 | 0.11448 | 0.32998 | 0.21746 | -0.16246 | -0.16666 | -0.38128 | 0.060228 | 0.18369 | -0.73911 | -33.762
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -0.88982 | -0.60645 | -0.49848 | -1.2095 | -1.3474 | -0.48388 | -0.87483 | -0.94921 | -0.71443 | -1.2401 | -1.8852 | -1.7713 | -1.618 | -2.1984 | -1.7565 | -0.47418 | -1.0399 | -1.5225 | -0.92762 | 0.79543 | -1.1014 | b2 | 2.5836 | 2.2018 | 0.87891 | 1.3773 | 1.27 | 1.6981 | 2.1824 | 2.5068 | 1.8132 | 0.65963 | 0.92422 | 1.6428 | 1.5866 | 0.94932 | 0.81711 | 0.5209 | 1.958 | 1.7485 | 1.8935 | 2.2851 | 1.5385
b3 | 36.3962 | -1.9138 | -1.7863 | -3.8348 | -1.6965 | -4.4628 | -3.1195 | -3.5747 | -2.039 | -2.906 | -3.6847 | -2.8564 | -2.3534 | -4.4431 | -3.6059 | -3.2357 | -1.6853 | -2.3549 | -1.1811 | -0.96044 | -1.5563 | b3 | 35.1777 | -1.7858 | -1.6993 | -1.8215 | -1.7203 | -4.4014 | -2.4578 | -1.8597 | -1.5271 | -2.7151 | -3.2367 | -2.8244 | -1.9681 | -4.4347 | -3.403 | -2.4528 | -1.2843 | -1.7835 | -0.71293 | -0.96574 | -1.1577
b4 | -0.42137 | 0.05927 | 0.036442 | -0.087728 | 0.11185 | 0.24982 | 0.19754 | 0.11215 | 0.25325 | 0.097126 | -0.047117 | -0.038652 | 0.040115 | 0.056909 | 0.092941 | 0.27242 | -0.016537 | 0.001935 | 0.10207 | 0.22236 | 0.071992 | b4 | 0.052715 | 0.048906 | -0.094741 | -0.058613 | -0.10286 | -0.088754 | -0.020745 | 0.073975 | 0.012401 | -0.082304 | -0.29391 | 0.093596 | -0.11897 | -0.22679 | -0.32682 | -0.38139 | 0.092127 | 0.017841 | 0.089594 | 0.15728 | -0.064656
b5 | -5.4313 | -0.37235 | -0.51366 | -0.222 | -0.38463 | 0.22412 | 0.22092 | 0.31819 | -0.1506 | -0.0398 | 0.25003 | 0.086559 | 0.16392 | 0.18377 | -0.0028846 | -0.083263 | -0.2169 | -0.17315 | -0.22598 | -3.9022 | -0.34051 | b5 | -6.253 | -0.7604 | -0.93614 | -0.47975 | -0.43515 | 0.14011 | 0.09301 | -0.2116 | -0.30372 | -0.29701 | -0.062618 | -0.18746 | -0.20317 | -0.13863 | -0.16835 | -0.65271 | -0.52084 | -0.35844 | -0.50481 | -4.1706 | -0.63385
b6 | -0.36081 | -0.12958 | -0.071651 | -0.015327 | -0.078963 | -0.14924 | -0.044693 | -0.1478 | -0.12974 | -0.091991 | 0.0269 | 0.0072535 | 0.014056 | 0.041259 | -0.0054871 | -0.16528 | 0.069851 | 0.044077 | 0.0028 | -0.0021634 | -0.054016 | b6 | 0.1979 | -0.0061714 | 0.024088 | -0.067097 | -0.052607 | -0.061741 | 0.034713 | -0.067767 | -0.099918 | 0.033036 | -0.01614 | 0.044494 | 0.031835 | 0.01481 | 0.074146 | 0.14898 | 0.055256 | -0.0034513 | -0.039954 | -0.10571 | 0.0050645
b7 | 2.8582 | 0.9609 | 1.0641 | 1.2401 | 1.1794 | 1.2233 | 1.1839 | 1.143 | 1.1201 | 1.1453 | 1.1868 | 1.1593 | 1.13 | 1.1121 | 1.1504 | 1.2327 | 1.1328 | 1.1707 | 1.1218 | 0.93179 | 1.1931 | b7 | 2.6828 | 0.97597 | 1.0315 | 1.2195 | 1.2099 | 1.2473 | 1.1942 | 1.1616 | 1.1528 | 1.219 | 1.2325 | 1.2276 | 1.2284 | 1.1579 | 1.1562 | 1.2946 | 1.202 | 1.2038 | 1.1632 | 0.88713 | 1.2182
b8 | -0.042877 | 0.024875 | 0.034115 | 0.024765 | 0.031952 | -0.0065308 | 0.015943 | 0.030744 | 0.02112 | 0.018689 | 0.030658 | 0.025376 | 0.030851 | 0.042937 | 0.036457 | 0.074942 | 0.060302 | 0.047448 | 0.0075551 | -0.023701 | 0.026678 | b8 | 0.00062977 | 0.011913 | 0.029507 | -0.021473 | -0.010671 | 0.00072533 | 0.062404 | 0.047452 | -0.0079872 | -0.0053364 | 0.0060967 | -0.010537 | -0.0073326 | 0.032645 | -0.020203 | -0.029074 | 0.0020041 | 0.026347 | -0.012448 | -0.058832 | 0.0030524
b9 | 0.35889 | 0.5111 | 0.49978 | 0.49558 | 0.52633 | 0.48193 | 0.51052 | 0.5068 | 0.52486 | 0.48884 | 0.47973 | 0.49545 | 0.49778 | 0.48363 | 0.47327 | 0.50162 | 0.48516 | 0.51093 | 0.5065 | 0.44184 | 0.49276 | b9 | 0.33135 | 0.51687 | 0.5126 | 0.5313 | 0.5349 | 0.49476 | 0.50704 | 0.54029 | 0.53402 | 0.51896 | 0.51356 | 0.50925 | 0.5226 | 0.50794 | 0.49143 | 0.54244 | 0.51575 | 0.54222 | 0.52041 | 0.44726 | 0.51178
b10 | -0.036958 | -0.063961 | -0.0295 | -0.010429 | -0.013386 | -0.084003 | 0.015648 | -0.018507 | -0.028247 | 0.0059425 | 0.053075 | 0.021992 | 0.0099026 | 0.084641 | 0.027027 | 0.036761 | 0.061364 | 0.02971 | -0.048418 | -0.04512 | 0 | b10 | 0.067155 | -0.00090196 | -0.031521 | -0.038039 | -0.029922 | -0.077449 | 0.044675 | 0.03553 | -0.0098596 | 0.028164 | 0.052762 | 0.019709 | 0.0024553 | 0.05461 | -0.0047888 | -0.044425 | 0.023301 | 0.043977 | -0.048079 | -0.097251 | 0
b11 | 0.61314 | 0.73995 | 0.7358 | 0.74812 | 0.74812 | 0.7562 | 0.75099 | 0.75812 | 0.75222 | 0.7495 | 0.75917 | 0.7485 | 0.75053 | 0.76165 | 0.7566 | 0.74465 | 0.73927 | 0.7485 | 0.73651 | 0.63864 | 0.74142 | b11 | 0.5943 | 0.75323 | 0.75173 | 0.7561 | 0.76174 | 0.77685 | 0.76644 | 0.76381 | 0.76751 | 0.76664 | 0.77408 | 0.7607 | 0.76032 | 0.77615 | 0.76879 | 0.75068 | 0.75392 | 0.75291 | 0.74556 | 0.65913 | 0.75319
b12 | -0.0096978 | -0.0036694 | 0.0018511 | 0.00054768 | -0.0025883 | -0.010234 | 4.0911e-06 | -0.0026941 | -0.0071164 | -0.00017891 | 0.004156 | 0.0014234 | -0.0030667 | 0.0052295 | 0.0028448 | 0.0018248 | 0.0080635 | 0.0031929 | -0.0041067 | -0.0094632 | -0.00083621 | b12 | 0.015673 | -0.0013314 | -0.0036942 | -0.0018023 | -0.00018387 | -0.0024677 | 0.0093327 | 0.005791 | -0.0030429 | -1.9484e-05 | 0.0078914 | -0.00080159 | -0.0010387 | 0.0072915 | 0.00071879 | 0.003186 | 0.0018034 | 0.0018032 | -0.0060552 | -0.010648 | 0.0010425
b13 | 0.08906 | 0.066416 | 0.06957 | 0.065974 | 0.067146 | 0.062504 | 0.061473 | 0.058511 | 0.060464 | 0.061674 | 0.060573 | 0.061813 | 0.061851 | 0.060447 | 0.064439 | 0.062544 | 0.065126 | 0.068408 | 0.06622 | 0.050701 | 0.063986 | b13 | 0.094528 | 0.068824 | 0.069606 | 0.065785 | 0.064629 | 0.059606 | 0.060338 | 0.061113 | 0.064161 | 0.063238 | 0.063052 | 0.062327 | 0.061674 | 0.060428 | 0.059447 | 0.062715 | 0.063238 | 0.062439 | 0.066705 | 0.054207 | 0.063954
b14 | 0.0062204 | -0.0020234 | 0.0026105 | 0.0046605 | 0.0024698 | -0.0029696 | 0.0048532 | 0.0046605 | 0.0012752 | 0.0031977 | 0.0078156 | 0.0054856 | 0.0051364 | 0.011175 | 0.0075976 | 0.0063216 | 0.010971 | 0.0090549 | 0.0023489 | 0.0022297 | 0.0046694 | b14 | 0.0095971 | -0.0069836 | -0.0077409 | -0.0072695 | -0.007516 | -0.011321 | -0.0019394 | -0.0026517 | -0.0062269 | -0.0052303 | -0.001612 | -0.0056805 | -0.0078145 | -0.00083876 | -0.0061252 | -0.0072938 | -0.0050847 | -0.0028231 | -0.00988 | -0.012903 | -0.0055398
b15 | 0.011784 | 0.050736 | 0.048609 | 0.045914 | 0.047111 | 0.048856 | 0.046802 | 0.052828 | 0.052041 | 0.053026 | 0.053177 | 0.052579 | 0.054021 | 0.054596 | 0.053972 | 0.05082 | 0.048295 | 0.050735 | 0.048749 | 0.041698 | 0.049232 | b15 | 0.021227 | 0.043994 | 0.040858 | 0.038136 | 0.038574 | 0.044022 | 0.043423 | 0.044185 | 0.043409 | 0.046556 | 0.048377 | 0.047498 | 0.048651 | 0.053654 | 0.054519 | 0.044975 | 0.048121 | 0.045846 | 0.046762 | 0.045003 | 0.044882
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 21.8675 | 3.3994 | -2.5041 | -5.5251 | -1.4444 | -0.0005695 | 3.6158 | 4.4185 | 0.60854 | -4.1061 | -4.6966 | -0.55777 | 5.0268 | 1.6167 | -2.1346 | -4.7312 | -4.8209 | -0.81662 | -2.4235 | 3.5648 | 0 | a1 | 27.2826 | -4.739 | -2.0135 | -6.282 | -2.964 | -0.63673 | 2.4605 | 7.3764 | 2.1517 | -4.6113 | -4.0638 | 1.1448 | 3.2998 | 2.1746 | -1.6246 | -1.6666 | -3.8128 | 0.60228 | 1.8369 | -7.3911 | 0
a2 | 11.6554 | 0.37026 | 0.61628 | 0.98377 | 0.20402 | 0.35586 | 0.7975 | -0.27536 | -0.047081 | 1.2273 | 1.0927 | 1.2433 | 0.94919 | 1.788 | 0.84976 | 0.71476 | 0.35518 | -0.5615 | 1.2783 | 23.04 | 1.718 | a2 | 8.6461 | -1.1114 | -2.0139 | -2.1316 | -1.3134 | -2.5707 | -1.514 | -2.2981 | -1.2468 | -1.2381 | -0.92407 | -0.29956 | -0.43176 | -1.1659 | -1.7818 | -2.6697 | -1.9388 | -1.5342 | -0.92844 | 5.8207 | -0.96607
a3 | -0.63057 | 1.1665 | 0.97873 | 0.56709 | 0.64371 | 0.51523 | 0.69645 | 0.79672 | 0.5662 | 0.3273 | 0.55222 | 0.4609 | 0.56876 | 0.43611 | 0.33516 | 0.53841 | 0.76863 | 1.0662 | 1.223 | 0.17187 | 0.62185 | a3 | 0.49784 | 0.69483 | 0.74237 | 0.31794 | 0.24123 | 0.17545 | -0.14809 | 0.27531 | 0.64046 | 0.17455 | 0.15516 | -0.079982 | 0.088449 | 0.24941 | 0.17803 | 0.51343 | 0.72741 | 1.1547 | 1.3379 | 0.9429 | 0.43266
a4 | 1.8584 | 0.11513 | 0.18875 | -0.23845 | -0.078463 | 0.3894 | 0.30474 | 0.42144 | 0.07719 | -0.084352 | -0.018715 | -0.084636 | -0.23485 | -0.17244 | 0.20097 | -0.02385 | 0.34309 | 0.23666 | 0.18986 | 2.0608 | 0.20282 | a4 | 2.5362 | 0.24766 | 0.13665 | 0.34886 | 0.19377 | 0.010046 | -0.24801 | 0.20006 | -0.14531 | -0.43633 | 0.1872 | -0.48338 | -0.20779 | -0.091458 | -0.2453 | -0.088433 | -0.30459 | -0.43076 | -0.15304 | 1.0016 | 0.029946
a5 | 2.6802 | 0.28017 | 0.14574 | -0.0015221 | 0.21785 | -0.011434 | 0.017666 | 0.17299 | 0.20154 | 0.21029 | 0.1421 | 0.043401 | -0.040011 | 0.29872 | 0.20268 | 0.33134 | 0.067586 | 0.3587 | 0.58875 | -0.005734 | 0.25004 | a5 | 2.7915 | 0.15332 | -0.095916 | -0.10683 | 0.11914 | -0.0060381 | -0.064175 | -0.040437 | 0.17315 | 0.41318 | 0.13247 | -0.17191 | -0.16686 | 0.237 | 0.0087675 | -0.10697 | 0.10254 | 0.17357 | 0.37506 | -0.035274 | 0.14233
a6 | -0.1356 | 0.19074 | 0.10791 | 0.020566 | -0.00085962 | -0.017991 | -0.096111 | -0.095754 | -0.029721 | -0.082655 | -0.042595 | -0.12501 | -0.097572 | -0.12683 | -0.070673 | -0.16239 | -0.080376 | -0.070897 | 0.012361 | 0.11117 | -0.040571 | a6 | -0.33612 | 0.15058 | 0.099678 | 0.11374 | 0.0245 | -0.046041 | -0.11742 | -0.026773 | -0.041724 | -0.0014942 | -0.00057246 | 0.13349 | 0.053182 | 0.011385 | -0.088645 | -0.13487 | -0.00070445 | 0.097801 | 0.068615 | 0.012298 | 0.0055672
a7 | 1.4856 | -0.11288 | -0.055218 | -0.061075 | 0.0567 | 0.029371 | 0.015534 | 0.060163 | 0.018089 | -0.016879 | -0.0080649 | -0.0083233 | -0.046068 | 0.028263 | 0.057888 | -0.04792 | -0.013753 | 0.10058 | 0.12586 | 0.049812 | 0.053526 | a7 | 1.7031 | -0.063194 | -0.014134 | -0.030571 | -0.036077 | -0.028967 | 0.072772 | 0.14104 | 0.036236 | -0.027989 | 0.023597 | -0.048347 | -0.051843 | 0.082464 | 0.024156 | 0.012847 | 0.0017386 | 0.08181 | 0.096535 | 0.13211 | 0.069984
a8 | -0.081264 | 0.0082756 | 0.031268 | 0.059553 | 0.024176 | 0.0016517 | -0.00055808 | -0.024315 | -0.0025222 | 0.0063085 | -0.056901 | -0.058511 | -0.061901 | -0.03719 | -0.060976 | -0.034553 | -0.014822 | -0.0059821 | 0.0035826 | -0.020011 | -0.014767 | a8 | -0.15361 | 0.01511 | -0.0080508 | -0.0066461 | -0.021411 | -0.087487 | -0.035593 | -0.013277 | -0.02055 | 0.01834 | -0.023162 | 0.039391 | 0.0095868 | -0.0055268 | -0.063894 | -0.077092 | -0.030717 | -0.011108 | -0.012454 | -0.092381 | -0.025018
a9 | -0.25173 | -0.0074864 | -0.0033989 | 0.0014847 | -0.024538 | -0.047785 | -0.034794 | 0.0049853 | -0.01437 | -0.027454 | -0.031094 | -0.034511 | -0.01233 | -0.017609 | -0.044087 | -0.029035 | -0.020889 | 0.02436 | 0.025612 | -0.051088 | -0.024609 | a9 | -0.2555 | 0.061327 | 0.043039 | -0.0073909 | -0.0090511 | -0.021922 | -0.020674 | 0.015174 | 0.015451 | -0.012663 | -0.027781 | -0.0052284 | 0.020004 | -0.011452 | -0.034866 | -0.020504 | -0.0083515 | 0.024172 | 0.038036 | -0.010305 | -0.0061225
a10 | -0.033115 | 0.039878 | 0.036462 | 0.03109 | 0.018457 | 0.025969 | 0.04074 | 0.0016809 | 0.0050997 | -0.013068 | -0.014465 | -0.019153 | -0.025823 | -0.03988 | -0.03731 | -0.018127 | 0.0060422 | 0.0014047 | 0.027065 | -0.075666 | 0 | a10 | -0.12659 | 0.048015 | 0.06027 | 0.058703 | 0.049271 | -0.010853 | 0.030565 | 0.029682 | 0.020297 | -0.0023219 | -0.00035768 | 0.0046413 | -0.0076167 | -0.025931 | -0.064282 | -0.020286 | -0.017653 | 0.025426 | -0.01089 | -0.15284 | 0
a11 | 0.1366 | -0.064618 | -0.054665 | -0.05164 | -0.048674 | -0.042937 | -0.037404 | -0.026656 | -0.040334 | -0.051263 | -0.038053 | -0.039524 | -0.035738 | -0.040668 | -0.042205 | -0.042015 | -0.045175 | -0.028075 | -0.024982 | -0.017292 | -0.035659 | a11 | 0.18981 | -0.012155 | -0.0040235 | -0.016712 | -0.0072816 | -0.00052664 | 0.0051542 | 0.020149 | 0.011975 | -0.00093006 | 0.0020997 | 0.0038801 | 0.0074341 | 0.0050078 | -0.0016516 | -0.00035112 | 0.010568 | 0.016402 | 0.01291 | 0.016494 | 0.0089826
a12 | 0.0023268 | 0.0045927 | 0.0053574 | 0.00044927 | -0.0017174 | -0.0021592 | -0.0018541 | -0.0060054 | -0.0031146 | -0.0058507 | 4.6073e-05 | -0.0040185 | -0.0064513 | -0.011101 | -0.0094176 | -0.0066821 | -0.0017326 | -0.0056531 | 0.0019827 | -0.0077817 | -0.0029548 | a12 | -0.022269 | 0.0074455 | 0.0068676 | 0.0044131 | -0.00025201 | -0.0047019 | -0.0045897 | -0.003012 | -0.0055485 | -0.0062436 | -0.0031884 | -0.0032909 | -0.00088309 | -0.0036159 | -0.015648 | -0.011437 | -0.0083776 | 0.00020319 | -0.0022441 | -0.017579 | -0.0040516
a13 | -0.029564 | -0.00047472 | -0.0026933 | -0.0068428 | -0.003728 | -0.0042207 | -0.00281 | 0.00021539 | -0.0019801 | -0.0043624 | -0.0049175 | -0.0039812 | -0.0040942 | -0.0036175 | -0.0056027 | -0.0020303 | -0.00502 | 0.00019522 | 0.0037023 | -7.6587e-05 | -0.0036301 | a13 | -0.026176 | 0.0056589 | 0.0076274 | 0.00068463 | 0.0029973 | 0.0026072 | 5.7036e-06 | 0.0024021 | 0.0040071 | -3.9505e-05 | -0.0015399 | 0.0012798 | 0.0018436 | 0.0018159 | 0.0028379 | 0.0032331 | 0.0015302 | 0.0041567 | 0.0091047 | 0.0059631 | 0.0020128
a14 | 0.0084365 | -0.026356 | -0.02297 | -0.02042 | -0.021153 | -0.021544 | -0.020944 | -0.024939 | -0.025618 | -0.025692 | -0.027227 | -0.027519 | -0.028117 | -0.027581 | -0.026962 | -0.024254 | -0.02134 | -0.023293 | -0.021697 | -0.015767 | -0.023034 | a14 | 0.011983 | -0.020618 | -0.020694 | -0.019516 | -0.020548 | -0.022772 | -0.023826 | -0.021085 | -0.022371 | -0.022839 | -0.02416 | -0.023739 | -0.024133 | -0.026163 | -0.031128 | -0.024665 | -0.026555 | -0.027149 | -0.022883 | -0.016792 | -0.022308
a15 | 0.0025394 | -0.0018022 | 0.0027385 | 0.0050303 | 0.003609 | -0.0046651 | 0.0062339 | 0.0050313 | 0.001639 | 0.0047369 | 0.011376 | 0.0083537 | 0.0071364 | 0.015861 | 0.011201 | 0.010013 | 0.015142 | 0.012578 | 0.0022443 | 0.00017009 | 0.006147 | a15 | -0.0067136 | -0.01007 | -0.011175 | -0.012616 | -0.01259 | -0.019299 | -0.0030369 | -0.0036033 | -0.011222 | -0.0081272 | -0.0049296 | -0.010227 | -0.012329 | -0.0033532 | -0.011397 | -0.01313 | -0.0089961 | -0.0046439 | -0.015248 | -0.02424 | -0.010144
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.11718 | -0.15375 | -0.073988 | -0.028443 | -0.033008 | -0.19168 | 0.030761 | -0.041582 | -0.064284 | 0.015432 | 0.12019 | 0.052095 | 0.025155 | 0.19304 | 0.065162 | 0.08597 | 0.13979 | 0.067268 | -0.11375 | -0.11432 |  | Dx (mm) | 0.068027 | 0 | -0.071957 | -0.088375 | -0.067837 | -0.16945 | 0.099592 | 0.080788 | -0.021135 | 0.062456 | 0.1159 | 0.044097 | 0 | 0.11925 | -0.010274 | -0.10053 | 0.05196 | 0.1005 | -0.11004 | -0.259
Dy (mm) | -0.065751 | 0.07823 | 0.07873 | 0.068684 | 0.039796 | 0.04751 | 0.093764 | 0 | 0 | -0.028584 | -0.026371 | -0.040753 | -0.057282 | -0.078867 | -0.080199 | -0.036549 | 0.022437 | 0 | 0.058631 | -0.20463 |  | Dy (mm) | -0.38328 | 0.10828 | 0.1359 | 0.13003 | 0.10931 | -0.023915 | 0.067158 | 0.063949 | 0.045293 | 0 | 0 | 0.010144 | -0.017082 | -0.057611 | -0.14205 | -0.046006 | -0.04054 | 0.05522 | -0.022944 | -0.38798
Mag. len. (m) | 14.4206 | [20 empty] | Mag. len. (m) | 14.4199
TF (mT/KA) | 595.9537 | [20 empty] | TF (mT/KA) | 596.0202
Coil pos. (mm) | 0.019342 | [20 empty] | Coil pos. (mm) | 0.02245
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0072913 | [20 empty] | Twist (m2 rad) | 0.0038222

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -1.1014 | 0.66816 | -1.167 | 0.50523 | 1.718 | 5.4866 | 0.66462 | 0.59086 | 1.5385 | 0.62667 | 1.4828 | 0.58383 | -0.96607 | 2.8026 | -1.5037 | 0.67633 |  | 0
3 | -1.5563 | 8.8068 | -2.8171 | 0.99249 | 0.62185 | 0.40065 | 0.67903 | 0.27019 | -1.1577 | 8.4329 | -2.337 | 1.0127 | 0.43266 | 0.39772 | 0.41399 | 0.40146 |  | 0
4 | 0.071992 | 0.15524 | 0.081597 | 0.10737 | 0.20282 | 0.61152 | 0.086485 | 0.21206 | -0.064656 | 0.15156 | -0.07769 | 0.15002 | 0.029946 | 0.66776 | -0.081054 | 0.26198 |  | 0
5 | -0.34051 | 1.4607 | -0.052565 | 0.24697 | 0.25004 | 0.58258 | 0.18142 | 0.15913 | -0.63385 | 1.5621 | -0.3319 | 0.27939 | 0.14233 | 0.63361 | 0.063332 | 0.17445 |  | 0
6 | -0.054016 | 0.10168 | -0.048549 | 0.060651 | -0.040571 | 0.091854 | -0.045711 | 0.073517 | 0.0050645 | 0.077384 | 0.0017175 | 0.067032 | 0.0055672 | 0.11209 | 0.0077584 | 0.074985 |  | 0
7 | 1.1931 | 0.39284 | 1.1476 | 0.064247 | 0.053526 | 0.33522 | 0.0076204 | 0.060433 | 1.2182 | 0.35264 | 1.1818 | 0.075066 | 0.069984 | 0.38112 | 0.014382 | 0.059508 |  | 0
8 | 0.026678 | 0.02649 | 0.029803 | 0.017278 | -0.014767 | 0.035988 | -0.01418 | 0.024765 | 0.0030524 | 0.027718 | 0.0052239 | 0.021759 | -0.025018 | 0.045532 | -0.023783 | 0.030617 |  | 0
9 | 0.49276 | 0.036094 | 0.49736 | 0.016255 | -0.024609 | 0.056417 | -0.014579 | 0.02102 | 0.51178 | 0.046643 | 0.5194 | 0.015694 | -0.0061225 | 0.062883 | 0.0022835 | 0.027233 |  | 0
10 | 0 | 0.043957 | 2.2462e-05 | 5.2004e-05 | 0 | 0.03147 | 0.00039825 | 0.0010092 | 0 | 0.046107 | 3.3401e-05 | 0.00070541 | 0 | 0.055859 | -0.00015187 | 0.0005464 |  | 0
11 | 0.74142 | 0.038842 | 0.74901 | 0.0076385 | -0.035659 | 0.041119 | -0.041696 | 0.0099908 | 0.75319 | 0.043683 | 0.76152 | 0.0091118 | 0.0089826 | 0.042728 | 0.0029335 | 0.0095519 |  | 0
12 | -0.00083621 | 0.0051385 | -0.0004208 | 0.0024387 | -0.0029548 | 0.0045159 | -0.0032932 | 0.0026646 | 0.0010425 | 0.005892 | 0.00099783 | 0.0029602 | -0.0040516 | 0.0075863 | -0.003954 | 0.0034575 |  | 0
13 | 0.063986 | 0.007155 | 0.063268 | 0.002834 | -0.0036301 | 0.0064883 | -0.0026812 | 0.0025974 | 0.063954 | 0.0078778 | 0.062882 | 0.0023368 | 0.0020128 | 0.0070177 | 0.0029851 | 0.0024078 |  | 0
14 | 0.0046694 | 0.0037208 | 0.0042938 | 0.0016211 | -0.023034 | 0.0078876 | -0.024991 | 0.0017461 | -0.0055398 | 0.0047263 | -0.0061275 | 0.0012424 | -0.022308 | 0.0084671 | -0.024442 | 0.0014283 |  | 0
15 | 0.049232 | 0.0092013 | 0.050715 | 0.0027296 | 0.006147 | 0.0055085 | 0.0064698 | 0.0055658 | 0.044882 | 0.0068187 | 0.045642 | 0.0043268 | -0.010144 | 0.0053867 | -0.0097773 | 0.0044672 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.60446 | -0.46461
3 | [12 empty] | 3 | 9 | -0.65809 | -0.45
4 | [12 empty] | 4 | 16 | -1.8912 | -1.6182
5 | [12 empty] | 5 | 25 | -1.6183 | -1.5106
6 | [12 empty] | 6 | 36 | -2.7064 | -2.6465
7 | [12 empty] | 7 | 49 | -2.7756 | -2.7055
8 | [12 empty] | 8 | 64 | -3.8783 | -3.657
9 | [12 empty] | 9 | 81 | -3.9908 | -3.8789
10 | [12 empty] | sigma (mm) |  | 0.019342 | 0.02245
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.72285 | 0.18071 | 0.043371 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.42345 | 0.10586 | 0.025407 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.24505 | 0.061262 | 0.014703 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.14008 | 0.03502 | 0.0084049 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.079104 | 0.019776 | 0.0047462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.044126 | 0.011032 | 0.0026476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.024316 | 0.0060789 | 0.0014589 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.013236 | 0.003309 | 0.00079416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0071173 | 0.0017793 | 0.00042704 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0037806 | 0.00094515 | 0.00022684 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0019838 | 0.00049594 | 0.00011903 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0010283 | 0.00025706 | 6.1695e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.0005265 | 0.00013163 | 3.159e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -0.3832 | -0.63034 | -0.50497 | -1.2178 | -1.3509 | -0.58164 | -0.85586 | -0.9667 | -0.72985 | -1.2359 | -1.8348 | -1.7561 | -1.6149 | -2.1016 | -1.732 | -0.44377 | -1.0102 | -1.5039 | -0.93499 | 0.77838 | -1.1014
3 | 36.3659 | -1.9106 | -1.7832 | -3.8382 | -1.6964 | -4.451 | -3.1156 | -3.5738 | -2.0361 | -2.9059 | -3.6836 | -2.8555 | -2.3512 | -4.4426 | -3.6098 | -3.2397 | -1.6835 | -2.3549 | -1.1771 | -1.0304 | -1.5563
4 | -0.61258 | 0.050957 | 0.0302 | -0.089238 | 0.11091 | 0.2598 | 0.19633 | 0.11526 | 0.25098 | 0.095856 | -0.05507 | -0.040129 | 0.039684 | 0.043019 | 0.08916 | 0.27125 | -0.0090456 | 0.0046757 | 0.10414 | 0.11768 | 0.071992
5 | -5.4411 | -0.37382 | -0.51272 | -0.22172 | -0.38541 | 0.21545 | 0.21868 | 0.31638 | -0.15305 | -0.038688 | 0.24941 | 0.087946 | 0.16546 | 0.18437 | -0.0011124 | -0.077338 | -0.2203 | -0.17402 | -0.22568 | -3.909 | -0.34051
6 | -0.27708 | -0.08055 | -0.045397 | -0.0043591 | -0.064426 | -0.065985 | -0.057032 | -0.13102 | -0.10432 | -0.098059 | -0.023369 | -0.013943 | 0.0049547 | -0.035294 | -0.033584 | -0.20206 | 0.013851 | 0.016282 | 0.050441 | 0.031835 | -0.054016
7 | 2.8583 | 0.96275 | 1.0662 | 1.242 | 1.1803 | 1.2229 | 1.1836 | 1.1436 | 1.1206 | 1.1451 | 1.1859 | 1.1598 | 1.1311 | 1.1099 | 1.1515 | 1.2306 | 1.1292 | 1.1694 | 1.1222 | 0.93236 | 1.1931
8 | -0.015299 | 0.061579 | 0.05139 | 0.031446 | 0.039668 | 0.035872 | 0.0070173 | 0.040661 | 0.036998 | 0.015508 | 0.003911 | 0.013892 | 0.025291 | -0.00034281 | 0.023608 | 0.055148 | 0.028166 | 0.031274 | 0.035374 | 0.004988 | 0.026678
9 | 0.35775 | 0.50754 | 0.50015 | 0.49655 | 0.52648 | 0.47406 | 0.51228 | 0.50639 | 0.52389 | 0.48899 | 0.47655 | 0.49526 | 0.49844 | 0.47665 | 0.47392 | 0.5003 | 0.48069 | 0.50987 | 0.50442 | 0.4473 | 0.49276
10 | 2.0618e-05 | -1.1061e-05 | -8.4956e-06 | 1.1387e-06 | 9.8968e-07 | 5.9605e-05 | -4.2527e-06 | 3.703e-05 | 0.00019704 | 8.7111e-07 | -7.1412e-06 | 2.5895e-06 | 1.3646e-06 | 4.1956e-05 | 1.6959e-05 | 6.6403e-06 | -2.366e-05 | 9.1536e-05 | 1.1994e-06 | -9.1826e-05 | 0
11 | 0.61231 | 0.73982 | 0.73616 | 0.74815 | 0.74802 | 0.75486 | 0.75088 | 0.75805 | 0.75193 | 0.74961 | 0.75885 | 0.74856 | 0.75082 | 0.76157 | 0.75697 | 0.7447 | 0.73852 | 0.74836 | 0.73628 | 0.63897 | 0.74142
12 | -0.0095935 | 0.035126 | 0.053349 | 0.015405 | -0.011285 | -0.019189 | -0.015167 | -0.0097669 | -0.043727 | -0.0076271 | -0.0089141 | -0.0073512 | -0.039994 | -0.028057 | 0.0019805 | -0.019183 | 0.015576 | -0.00055391 | 0.013636 | -0.053608 | -0.0083621 | ten times!
13 | 0.89193 | 0.64601 | 0.68335 | 0.65003 | 0.66565 | 0.61286 | 0.59857 | 0.5866 | 0.60527 | 0.62197 | 0.60404 | 0.62452 | 0.62984 | 0.60461 | 0.65714 | 0.62806 | 0.63587 | 0.67942 | 0.65452 | 0.53363 | 0.63986 | ten times!
14 | 0.0722 | 0.042847 | 0.057498 | 0.060205 | 0.038687 | 0.0456 | 0.041489 | 0.064695 | 0.040302 | 0.024123 | 0.023051 | 0.029495 | 0.036807 | 0.014659 | 0.039616 | 0.024221 | 0.056913 | 0.062443 | 0.070239 | 0.061267 | 0.046694 | ten times!
15 | 0.11784 | 0.50736 | 0.48609 | 0.45914 | 0.47111 | 0.48856 | 0.46802 | 0.52828 | 0.52041 | 0.53026 | 0.53177 | 0.52579 | 0.54021 | 0.54596 | 0.53972 | 0.5082 | 0.48295 | 0.50735 | 0.48749 | 0.41698 | 0.49232 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 11.9283 | 0.40897 | 0.64135 | 1.0167 | 0.21447 | 0.39243 | 0.82939 | -0.27146 | -0.042799 | 1.217 | 1.0734 | 1.2268 | 0.93164 | 1.7369 | 0.81316 | 0.6954 | 0.34698 | -0.56994 | 1.3028 | 23.0192 | 1.718
3 | -0.59703 | 1.1688 | 0.98069 | 0.56696 | 0.64247 | 0.52631 | 0.69153 | 0.79981 | 0.56707 | 0.32802 | 0.5524 | 0.4614 | 0.57021 | 0.44278 | 0.33416 | 0.54053 | 0.76023 | 1.0634 | 1.2258 | 0.22147 | 0.62185
4 | 1.8483 | 0.13212 | 0.2008 | -0.23487 | -0.073169 | 0.38638 | 0.29974 | 0.42313 | 0.080239 | -0.085383 | -0.021182 | -0.084338 | -0.23241 | -0.1826 | 0.19781 | -0.031269 | 0.34201 | 0.23098 | 0.20873 | 1.8728 | 0.20282
5 | 2.6686 | 0.29178 | 0.14975 | -0.0010405 | 0.21876 | -0.010362 | 0.019768 | 0.17182 | 0.20098 | 0.20989 | 0.14382 | 0.045403 | -0.039052 | 0.30687 | 0.2039 | 0.33367 | 0.070429 | 0.3601 | 0.58912 | -0.0021261 | 0.25004
6 | -0.0078356 | 0.15808 | 0.076897 | -0.010108 | -0.016765 | -0.036517 | -0.13546 | -0.094871 | -0.029311 | -0.071009 | -0.031206 | -0.10818 | -0.074317 | -0.097798 | -0.03944 | -0.14503 | -0.088668 | -0.073285 | -0.0058003 | 0.18048 | -0.040571
7 | 1.4805 | -0.11316 | -0.055372 | -0.061078 | 0.056505 | 0.029629 | 0.014925 | 0.059746 | 0.018022 | -0.016699 | -0.004916 | -0.0066423 | -0.044699 | 0.032613 | 0.060728 | -0.045569 | -0.013457 | 0.10075 | 0.12584 | 0.046873 | 0.053526
8 | -0.084041 | -0.011082 | 0.012633 | 0.043555 | 0.013938 | -0.013433 | -0.02258 | -0.024217 | -0.0029569 | 0.013083 | -0.049189 | -0.048164 | -0.048336 | -0.017641 | -0.041763 | -0.024751 | -0.01857 | -0.0067533 | -0.0090212 | 0.019787 | -0.014767
9 | -0.25507 | -0.0015914 | -0.00074107 | 0.0023321 | -0.023933 | -0.043037 | -0.036234 | 0.0050223 | -0.014196 | -0.027257 | -0.029433 | -0.033508 | -0.011685 | -0.0099991 | -0.041653 | -0.027498 | -0.022065 | 0.02431 | 0.028745 | -0.060555 | -0.024609
10 | 1.4723e-05 | -1.7043e-05 | 7.22e-06 | 1.6061e-07 | -1.2705e-06 | -5.8214e-06 | -4.2071e-06 | 0.0010289 | 0.0035745 | -6.6172e-07 | 1.5346e-06 | 1.2443e-06 | -4.6784e-06 | 7.2832e-05 | 5.2241e-07 | 7.7202e-06 | -1.029e-07 | 0.0025156 | -7.9135e-06 | 4.2754e-05 | 0
11 | 0.13636 | -0.063975 | -0.054503 | -0.051656 | -0.048644 | -0.042891 | -0.037367 | -0.026818 | -0.040464 | -0.051208 | -0.037986 | -0.039351 | -0.035747 | -0.039014 | -0.04166 | -0.0416 | -0.045136 | -0.027829 | -0.024681 | -0.019121 | -0.035659 | ten times!
12 | 0.040151 | 0.008736 | 0.013504 | -0.028867 | -0.036906 | -0.048264 | -0.058617 | -0.059991 | -0.032044 | -0.045588 | 0.015908 | -0.02094 | -0.038778 | -0.072434 | -0.055119 | -0.049453 | -0.022687 | -0.056624 | -0.0046063 | -0.0046456 | -0.029548 | ten times!
13 | -0.28495 | -0.034524 | -0.0415 | -0.075317 | -0.043371 | -0.072706 | -0.026654 | -0.0057763 | -0.032395 | -0.039893 | -0.022575 | -0.02714 | -0.033283 | 0.011278 | -0.037932 | -0.0025914 | -0.029271 | 0.013934 | 0.017097 | -0.011061 | -0.036301 | ten times!
14 | 0.093197 | -0.29853 | -0.25955 | -0.22899 | -0.22599 | -0.24192 | -0.24716 | -0.24767 | -0.25532 | -0.24504 | -0.27198 | -0.26113 | -0.25716 | -0.26556 | -0.23998 | -0.23433 | -0.23975 | -0.2399 | -0.2384 | -0.087245 | -0.23034 | ten times!
15 | 0.025394 | -0.018022 | 0.027385 | 0.050303 | 0.03609 | -0.046651 | 0.062339 | 0.050313 | 0.01639 | 0.047369 | 0.11376 | 0.083537 | 0.071364 | 0.15861 | 0.11201 | 0.10013 | 0.15142 | 0.12578 | 0.022443 | 0.0017009 | 0.06147 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 2.2796 | 2.2107 | 0.87639 | 1.3632 | 1.2594 | 1.6099 | 2.21 | 2.5265 | 1.8128 | 0.67958 | 0.96836 | 1.6573 | 1.5865 | 1.0098 | 0.81003 | 0.48911 | 1.9624 | 1.777 | 1.8807 | 2.2127 | 1.5385
3 | 35.0055 | -1.781 | -1.6972 | -1.8145 | -1.7178 | -4.4041 | -2.4603 | -1.8585 | -1.5282 | -2.7142 | -3.2307 | -2.826 | -1.9675 | -4.429 | -3.3974 | -2.4589 | -1.2829 | -1.788 | -0.71057 | -1.0271 | -1.1577
4 | -0.098943 | 0.052812 | -0.11366 | -0.071857 | -0.10674 | -0.083133 | -0.023939 | 0.077389 | 0.012736 | -0.077939 | -0.2922 | 0.095131 | -0.1183 | -0.22612 | -0.32752 | -0.39567 | 0.097517 | 0.028573 | 0.074498 | -0.093667 | -0.064656
5 | -6.219 | -0.75561 | -0.93165 | -0.47714 | -0.43542 | 0.13736 | 0.089673 | -0.2105 | -0.3049 | -0.29762 | -0.062068 | -0.18764 | -0.20344 | -0.13934 | -0.16442 | -0.64648 | -0.52167 | -0.35675 | -0.50657 | -4.18 | -0.63385
6 | -0.096902 | -0.0085864 | 0.049608 | -0.030462 | -0.025031 | 0.013099 | -0.0055373 | -0.097704 | -0.090739 | 0.0061648 | -0.066557 | 0.025216 | 0.032148 | -0.035603 | 0.077127 | 0.1947 | 0.033188 | -0.044556 | 0.0044402 | -0.042708 | 0.0050645
7 | 2.7071 | 0.97664 | 1.032 | 1.2184 | 1.2086 | 1.2482 | 1.1906 | 1.1597 | 1.1524 | 1.2192 | 1.2322 | 1.2279 | 1.2283 | 1.1565 | 1.1599 | 1.2949 | 1.2025 | 1.2024 | 1.1627 | 0.89562 | 1.2182
8 | 0.036106 | 0.015038 | 0.049617 | 0.0001702 | 0.0059389 | 0.040425 | 0.037987 | 0.027368 | -0.0023465 | -0.020589 | -0.021914 | -0.02113 | -0.0074934 | 0.0044507 | -0.015496 | -0.0029698 | -0.010447 | 0.0013299 | 0.014091 | -0.0024366 | 0.0030524
9 | 0.35462 | 0.51962 | 0.51573 | 0.53356 | 0.53667 | 0.48795 | 0.50577 | 0.53978 | 0.5344 | 0.51803 | 0.51032 | 0.50882 | 0.52267 | 0.50529 | 0.49624 | 0.54057 | 0.51549 | 0.54063 | 0.51774 | 0.46532 | 0.51178
10 | 0.00057957 | -0.0016762 | -2.3542e-05 | -1.1305e-05 | 6.5489e-06 | -5.995e-06 | -2.2277e-05 | -9.9914e-06 | 1.4444e-06 | -1.3332e-06 | -1.4252e-05 | -7.7651e-08 | 0.0023806 | -4.0867e-06 | -4.637e-06 | -2.5879e-05 | 6.0431e-06 | -1.9902e-06 | 8.066e-06 | -0.00059411 | 0
11 | 0.59914 | 0.75375 | 0.75216 | 0.75637 | 0.76172 | 0.77665 | 0.76564 | 0.76338 | 0.7673 | 0.76664 | 0.77349 | 0.7607 | 0.76033 | 0.77572 | 0.77024 | 0.75122 | 0.75408 | 0.7528 | 0.74516 | 0.66176 | 0.75319
12 | 0.18215 | -0.0089888 | 0.0057302 | 0.023644 | 0.031421 | 0.04618 | 0.050912 | 0.024143 | -0.019576 | -0.028074 | 0.027328 | -0.027325 | -0.01061 | 0.02131 | 0.0086535 | 0.075311 | -0.0055979 | -0.024644 | -0.010211 | -0.023707 | 0.010425 | ten times!
13 | 0.90516 | 0.67117 | 0.6703 | 0.63354 | 0.62521 | 0.58555 | 0.59262 | 0.60245 | 0.63286 | 0.63487 | 0.63195 | 0.62334 | 0.61989 | 0.61657 | 0.6278 | 0.63022 | 0.64263 | 0.6151 | 0.66275 | 0.56633 | 0.63954 | ten times!
14 | 0.10527 | -0.078816 | -0.065704 | -0.058449 | -0.064944 | -0.047975 | -0.056688 | -0.057811 | -0.058899 | -0.076249 | -0.062296 | -0.074909 | -0.07641 | -0.059488 | -0.043306 | -0.03073 | -0.068435 | -0.068288 | -0.053541 | 0.044411 | -0.055398 | ten times!
15 | 0.21227 | 0.43994 | 0.40858 | 0.38136 | 0.38574 | 0.44022 | 0.43423 | 0.44185 | 0.43409 | 0.46556 | 0.48377 | 0.47498 | 0.48651 | 0.53654 | 0.54519 | 0.44975 | 0.48121 | 0.45846 | 0.46762 | 0.45003 | 0.44882 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.2283 | -1.0886 | -1.9804 | -2.1005 | -1.2894 | -2.5796 | -1.4928 | -2.2867 | -1.2371 | -1.2394 | -0.92618 | -0.29578 | -0.43571 | -1.1995 | -1.8385 | -2.6769 | -1.9494 | -1.5362 | -0.91305 | 5.8053 | -0.96607
3 | 0.47096 | 0.69389 | 0.74638 | 0.32473 | 0.24553 | 0.17537 | -0.14348 | 0.27162 | 0.63982 | 0.17936 | 0.15133 | -0.076388 | 0.08809 | 0.24903 | 0.16939 | 0.50877 | 0.73086 | 1.1621 | 1.3353 | 0.99944 | 0.43266
4 | 1.9276 | 0.26704 | 0.16496 | 0.36131 | 0.20687 | 0.010594 | -0.24797 | 0.20401 | -0.14121 | -0.4424 | 0.18358 | -0.48115 | -0.2086 | -0.099987 | -0.25091 | -0.098028 | -0.31081 | -0.4302 | -0.14605 | 0.61869 | 0.029946
5 | 2.8205 | 0.15352 | -0.094769 | -0.10131 | 0.12132 | -0.008767 | -0.061421 | -0.038526 | 0.17422 | 0.41321 | 0.13249 | -0.17378 | -0.1667 | 0.23685 | 0.011597 | -0.10894 | 0.10321 | 0.17074 | 0.37701 | -0.0464 | 0.14233
6 | -0.01409 | 0.11328 | 0.04984 | 0.056822 | -0.023043 | -0.037246 | -0.14829 | -0.057012 | -0.059882 | -0.00087728 | -0.0015378 | 0.12985 | 0.060588 | 0.03146 | -0.030591 | -0.1134 | 0.016462 | 0.071439 | 0.081783 | 0.14585 | 0.0055672
7 | 1.7075 | -0.063725 | -0.016024 | -0.029663 | -0.036195 | -0.035064 | 0.072506 | 0.14024 | 0.036206 | -0.028461 | 0.024703 | -0.049018 | -0.051895 | 0.08351 | 0.022704 | 0.0091057 | 0.0024292 | 0.08167 | 0.095853 | 0.11286 | 0.069984
8 | -0.085666 | -0.011227 | -0.039376 | -0.039463 | -0.049216 | -0.083667 | -0.050648 | -0.030113 | -0.031779 | 0.018712 | -0.021646 | 0.037068 | 0.013788 | 0.0088868 | -0.031212 | -0.066319 | -0.020673 | -0.026341 | -0.0048652 | -0.011978 | -0.025018
9 | -0.23732 | 0.061379 | 0.047603 | -0.0020258 | -0.00555 | -0.023876 | -0.023873 | 0.012701 | 0.015914 | -0.012587 | -0.027759 | -0.0054426 | 0.020026 | -0.0081498 | -0.035576 | -0.022665 | -0.0073657 | 0.021533 | 0.036817 | -0.051237 | -0.0061225
10 | -0.00019777 | 3.9501e-05 | 3.5997e-06 | 2.343e-06 | 1.1662e-08 | 2.305e-05 | -1.4466e-05 | -7.3372e-06 | -2.8772e-06 | -0.0022877 | -0.00050083 | 1.4544e-06 | 2.3206e-05 | 2.0471e-05 | -5.255e-05 | 8.0759e-06 | 2.3769e-06 | 1.3688e-07 | 7.8516e-06 | 0.00010432 | 0
11 | 0.19468 | -0.012062 | -0.0033789 | -0.016308 | -0.0072797 | -0.0010804 | 0.0050444 | 0.020067 | 0.011988 | -0.00067774 | 0.0023388 | 0.0039792 | 0.0074227 | 0.0055587 | -0.0016896 | -0.0010002 | 0.010897 | 0.016324 | 0.01266 | 0.010875 | 0.0089826
12 | 0.045622 | 0.021851 | 0.0057756 | -0.015823 | -0.050953 | -0.033838 | -0.074505 | -0.059076 | -0.075401 | -0.062418 | -0.030624 | -0.03777 | -0.0013943 | -0.013113 | -0.096662 | -0.091706 | -0.066241 | -0.025255 | -0.0045654 | -0.016431 | -0.040516 | ten times!
13 | -0.23986 | 0.062371 | 0.072931 | 0.00088568 | 0.025597 | -0.0055073 | 0.019199 | 0.038344 | 0.038612 | 0.010513 | 0.0060149 | 0.021244 | 0.017415 | 0.041648 | 0.01928 | 0.010805 | 0.024277 | 0.063624 | 0.070057 | -0.011908 | 0.020128 | ten times!
14 | 0.19059 | -0.24541 | -0.25929 | -0.24518 | -0.24724 | -0.24598 | -0.25979 | -0.23172 | -0.24185 | -0.22421 | -0.2369 | -0.23765 | -0.23448 | -0.23288 | -0.24847 | -0.24048 | -0.24563 | -0.28849 | -0.23381 | -0.075829 | -0.22308 | ten times!
15 | -0.067136 | -0.1007 | -0.11175 | -0.12616 | -0.1259 | -0.19299 | -0.030369 | -0.036033 | -0.11222 | -0.081272 | -0.049296 | -0.10227 | -0.12329 | -0.033532 | -0.11397 | -0.1313 | -0.089961 | -0.046439 | -0.15248 | -0.2424 | -0.10144 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00038915 | -0.00015821 | -0.00013992 | -0.00025234 | -7.1874e-06 | -2.0611e-05 | 2.2681e-05 | -8.3032e-05 | -0.00013253 | -0.00010401 | -8.2361e-05 | -0.00011273 | -2.0443e-05 | -4.8382e-06 | -7.6488e-05 | 5.0619e-06 | 2.7715e-05 | 0.00075009 |  | 595.9537
|  | 0.00045918 | -8.6434e-05 | -0.00019599 | -0.00010556 | 0.00011691 | 5.5088e-06 | -4.6838e-05 | -0.00010942 | -0.00010002 | -9.7843e-05 | 1.0374e-05 | -8.073e-05 | -0.00012251 | -9.3817e-05 | -0.00015942 | -3.2074e-05 | -6.9069e-06 | 0.00064559 |  | 596.0202
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3975.5643
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.9112 | -4109.6866 | 4.6039 | -0.8701 | -0.68718 | -1.8115 | 0.6402 | 0.50595 | 0.9389 | -0.11831 | -0.61335 | -0.32807 | -0.11159 | -0.41533 | 0.50763 | 0.66369 | -0.05286 | 0.7627 | 0.98924 | 8.2135 | -3841.442
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.083755 | 2.2705 | 0.42369 | -0.16665 | -0.46876 | -0.060689 | 0.083698 | 0.44533 | 0.5256 | 0.14461 | -0.32686 | -0.38591 | 0.027978 | 0.58643 | 0.24542 | -0.12971 | -0.38936 | -0.39833 | 0.0020927 | -0.1586 | 0.44024
db2 | -1.1732 | -0.88982 | 0.56677 | 0.67474 | -0.036311 | -0.17415 | 0.68934 | 0.2984 | 0.22401 | 0.45879 | -0.06689 | -0.71197 | -0.59812 | -0.44481 | -1.0252 | -0.58328 | 0.69905 | 0.13331 | -0.3493 | 0.24561 | 0.79543
db3 | -2.8185 | 36.3962 | 0.90473 | 1.0323 | -1.0163 | 1.1221 | -1.6442 | -0.30099 | -0.75612 | 0.77956 | -0.087482 | -0.86618 | -0.037871 | 0.46519 | -1.6245 | -0.78737 | -0.41714 | 1.1333 | 0.46364 | 1.6374 | -0.96044
db4 | 0.082989 | -0.42137 | -0.023719 | -0.046547 | -0.17072 | 0.028865 | 0.16683 | 0.11455 | 0.029162 | 0.17026 | 0.014137 | -0.13011 | -0.12164 | -0.042874 | -0.02608 | 0.0099515 | 0.18943 | -0.099526 | -0.081054 | 0.019079 | 0.22236
db5 | -0.052096 | -5.4313 | -0.32026 | -0.46156 | -0.16991 | -0.33254 | 0.27621 | 0.27302 | 0.37028 | -0.0985 | 0.012296 | 0.30212 | 0.13865 | 0.21601 | 0.23586 | 0.049212 | -0.031167 | -0.1648 | -0.12106 | -0.17389 | -3.9022
db6 | -0.045752 | -0.36081 | -0.083824 | -0.025898 | 0.030425 | -0.033211 | -0.10349 | 0.0010593 | -0.10205 | -0.083983 | -0.046239 | 0.072653 | 0.053006 | 0.059808 | 0.087012 | 0.040265 | -0.11953 | 0.1156 | 0.089829 | 0.048553 | -0.0021634
db7 | 1.1476 | 2.8582 | -0.1867 | -0.083454 | 0.092472 | 0.031832 | 0.075747 | 0.036252 | -0.0045753 | -0.027539 | -0.0022833 | 0.039214 | 0.011723 | -0.017605 | -0.035523 | 0.0028367 | 0.085133 | -0.014805 | 0.023088 | -0.025809 | 0.93179
db8 | 0.030678 | -0.042877 | -0.0058025 | 0.0034373 | -0.0059132 | 0.0012743 | -0.037209 | -0.014735 | 6.6477e-05 | -0.0095575 | -0.011989 | -1.9303e-05 | -0.0053017 | 0.00017295 | 0.012259 | 0.0057792 | 0.044264 | 0.029625 | 0.01677 | -0.023123 | -0.023701
db9 | 0.49888 | 0.35889 | 0.012219 | 0.00090642 | -0.0033 | 0.027451 | -0.01695 | 0.011639 | 0.0079223 | 0.025979 | -0.010041 | -0.019148 | -0.0034238 | -0.0010934 | -0.015244 | -0.025612 | 0.0027426 | -0.013722 | 0.012053 | 0.0076215 | 0.44184
db10 | 0.0027562 | -0.036958 | -0.066717 | -0.032257 | -0.013185 | -0.016142 | -0.086759 | 0.012892 | -0.021263 | -0.031004 | 0.0031863 | 0.050319 | 0.019236 | 0.0071464 | 0.081885 | 0.024271 | 0.034005 | 0.058607 | 0.026953 | -0.051174 | -0.04512
db11 | 0.74913 | 0.61314 | -0.0091812 | -0.013337 | -0.0010143 | -0.0010097 | 0.0070617 | 0.0018572 | 0.008986 | 0.0030906 | 0.00036317 | 0.010035 | -0.00063073 | 0.001398 | 0.012518 | 0.0074677 | -0.004486 | -0.0098621 | -0.00063053 | -0.012626 | 0.63864
db12 | -0.00025094 | -0.0096978 | -0.0034185 | 0.002102 | 0.00079861 | -0.0023373 | -0.0099833 | 0.00025503 | -0.0024432 | -0.0068655 | 7.2022e-05 | 0.004407 | 0.0016743 | -0.0028158 | 0.0054805 | 0.0030958 | 0.0020757 | 0.0083145 | 0.0034438 | -0.0038557 | -0.0094632
db13 | 0.06362 | 0.08906 | 0.0027962 | 0.0059505 | 0.0023544 | 0.0035266 | -0.0011159 | -0.0021465 | -0.0051083 | -0.0031555 | -0.0019461 | -0.0030467 | -0.0018063 | -0.0017689 | -0.0031722 | 0.00081909 | -0.001076 | 0.0015066 | 0.0047885 | 0.0026005 | 0.050701
db14 | 0.0047023 | 0.0062204 | -0.0067257 | -0.0020918 | -4.1835e-05 | -0.0022325 | -0.0076719 | 0.00015089 | -4.1817e-05 | -0.0034271 | -0.0015046 | 0.0031133 | 0.00078335 | 0.00043408 | 0.0064729 | 0.0028953 | 0.0016193 | 0.006269 | 0.0043526 | -0.0023534 | 0.0022297
db15 | 0.050715 | 0.011784 | 2.1448e-05 | -0.0021063 | -0.0048005 | -0.0036035 | -0.0018586 | -0.0039125 | 0.0021131 | 0.0013264 | 0.0023109 | 0.0024625 | 0.0018644 | 0.0033058 | 0.0038811 | 0.0032568 | 0.00010541 | -0.0024203 | 1.9888e-05 | -0.0019661 | 0.041698
da2 | 0.66346 | 11.6554 | -0.2932 | -0.047176 | 0.32031 | -0.45944 | -0.3076 | 0.13404 | -0.93882 | -0.71054 | 0.56387 | 0.42922 | 0.57983 | 0.28573 | 1.1246 | 0.1863 | 0.051301 | -0.30828 | -1.225 | 0.61487 | 23.04
da3 | 0.67819 | -0.63057 | 0.48829 | 0.30054 | -0.11109 | -0.034473 | -0.16295 | 0.018265 | 0.11854 | -0.11199 | -0.35088 | -0.12597 | -0.21729 | -0.10943 | -0.24208 | -0.34303 | -0.13978 | 0.090445 | 0.38805 | 0.54484 | 0.17187
da4 | 0.085081 | 1.8584 | 0.030049 | 0.10366 | -0.32353 | -0.16354 | 0.30432 | 0.21966 | 0.33636 | -0.0078912 | -0.16943 | -0.1038 | -0.16972 | -0.31993 | -0.25752 | 0.11589 | -0.10893 | 0.25801 | 0.15158 | 0.10478 | 2.0608
da5 | 0.17925 | 2.6802 | 0.10092 | -0.033509 | -0.18077 | 0.038594 | -0.19069 | -0.16159 | -0.0062576 | 0.022287 | 0.031038 | -0.037151 | -0.13585 | -0.21926 | 0.11946 | 0.023424 | 0.15208 | -0.11167 | 0.17944 | 0.4095 | -0.005734
da6 | -0.042659 | -0.1356 | 0.2334 | 0.15057 | 0.063225 | 0.041799 | 0.024668 | -0.053453 | -0.053095 | 0.012937 | -0.039996 | 6.4117e-05 | -0.082352 | -0.054913 | -0.084169 | -0.028014 | -0.11973 | -0.037718 | -0.028238 | 0.055019 | 0.11117
da7 | 0.0067923 | 1.4856 | -0.11968 | -0.06201 | -0.067868 | 0.049907 | 0.022579 | 0.0087416 | 0.05337 | 0.011296 | -0.023672 | -0.014857 | -0.015116 | -0.05286 | 0.02147 | 0.051096 | -0.054712 | -0.020546 | 0.093792 | 0.11906 | 0.049812
da8 | -0.012412 | -0.081264 | 0.020688 | 0.04368 | 0.071965 | 0.036588 | 0.014064 | 0.011854 | -0.011903 | 0.0098898 | 0.018721 | -0.044489 | -0.046099 | -0.049489 | -0.024778 | -0.048564 | -0.022141 | -0.0024099 | 0.0064299 | 0.015995 | -0.020011
da9 | -0.016274 | -0.25173 | 0.0087878 | 0.012875 | 0.017759 | -0.0082633 | -0.031511 | -0.01852 | 0.02126 | 0.0019043 | -0.011179 | -0.01482 | -0.018236 | 0.0039446 | -0.0013344 | -0.027813 | -0.012761 | -0.0046144 | 0.040635 | 0.041887 | -0.051088
da10 | 0.0036701 | -0.033115 | 0.036208 | 0.032792 | 0.02742 | 0.014787 | 0.022299 | 0.03707 | -0.0019892 | 0.0014295 | -0.016738 | -0.018136 | -0.022824 | -0.029493 | -0.04355 | -0.040981 | -0.021797 | 0.002372 | -0.0022654 | 0.023395 | -0.075666
da11 | -0.041924 | 0.1366 | -0.022694 | -0.012741 | -0.0097164 | -0.0067501 | -0.0010136 | 0.0045198 | 0.015268 | 0.0015897 | -0.0093395 | 0.0038704 | 0.0023994 | 0.0061857 | 0.0012561 | -0.00028096 | -9.1451e-05 | -0.0032516 | 0.013849 | 0.016941 | -0.017292
da12 | -0.0029628 | 0.0023268 | 0.0075555 | 0.0083202 | 0.003412 | 0.0012453 | 0.00080359 | 0.0011087 | -0.0030426 | -0.00015181 | -0.0028879 | 0.0030089 | -0.0010558 | -0.0034886 | -0.0081387 | -0.0064548 | -0.0037193 | 0.0012302 | -0.0026903 | 0.0049455 | -0.0077817
da13 | -0.0029035 | -0.029564 | 0.0024288 | 0.00021021 | -0.0039393 | -0.0008245 | -0.0013172 | 9.3433e-05 | 0.0031189 | 0.00092334 | -0.0014589 | -0.002014 | -0.0010777 | -0.0011907 | -0.00071402 | -0.0026992 | 0.00087315 | -0.0021166 | 0.0030987 | 0.0066058 | -7.6587e-05
da14 | -0.024313 | 0.0084365 | -0.0020432 | 0.0013427 | 0.0038923 | 0.0031599 | 0.0027687 | 0.0033687 | -0.00062656 | -0.0013058 | -0.00138 | -0.0029148 | -0.0032067 | -0.0038045 | -0.0032683 | -0.0026495 | 5.8566e-05 | 0.0029729 | 0.0010197 | 0.0026157 | -0.015767
da15 | 0.0064698 | 0.0025394 | -0.008272 | -0.0037313 | -0.0014395 | -0.0028608 | -0.011135 | -0.00023596 | -0.0014385 | -0.0048308 | -0.001733 | 0.0049059 | 0.0018838 | 0.0006666 | 0.0093914 | 0.0047309 | 0.0035431 | 0.008672 | 0.0061085 | -0.0042255 | 0.00017009
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -4047.0233
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.9791 | -4183.6447 | 5.2827 | -0.17387 | -1.2695 | -0.36516 | 1.8598 | 0.74562 | 0.22211 | -0.40375 | -0.30978 | -0.28797 | 0.79428 | -0.11682 | -0.53462 | -0.2477 | -0.90376 | 0.36977 | 0.62146 | 7.1469 | -3910.4018
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.063151 | 2.7914 | -0.41075 | -0.1382 | -0.56505 | -0.23324 | -0.00052183 | 0.3092 | 0.80079 | 0.27832 | -0.39797 | -0.34323 | 0.17763 | 0.39313 | 0.28062 | -0.099305 | -0.10351 | -0.31812 | 0.12338 | 0.24684 | -0.67596
db2 | 1.4794 | 2.5836 | 0.72244 | -0.60048 | -0.10213 | -0.20939 | 0.21871 | 0.703 | 1.0274 | 0.33379 | -0.81975 | -0.55517 | 0.16337 | 0.10725 | -0.53007 | -0.66227 | -0.95849 | 0.47861 | 0.26907 | 0.41416 | 2.2851
db3 | -2.3382 | 35.1777 | 0.55248 | 0.63896 | 0.51669 | 0.61791 | -2.0631 | -0.11954 | 0.47854 | 0.81115 | -0.37688 | -0.89844 | -0.48614 | 0.37012 | -2.0964 | -1.0648 | -0.11459 | 1.054 | 0.55478 | 1.6253 | -0.96574
db4 | -0.07597 | 0.052715 | 0.12488 | -0.018771 | 0.017357 | -0.026891 | -0.012783 | 0.055225 | 0.14995 | 0.088371 | -0.0063333 | -0.21794 | 0.16957 | -0.043001 | -0.15082 | -0.25085 | -0.30542 | 0.1681 | 0.093812 | 0.16556 | 0.15728
db5 | -0.33265 | -6.253 | -0.42775 | -0.60349 | -0.1471 | -0.1025 | 0.47276 | 0.42566 | 0.12105 | 0.028927 | 0.035638 | 0.27003 | 0.14519 | 0.12947 | 0.19402 | 0.1643 | -0.32006 | -0.18819 | -0.025792 | -0.17216 | -4.1706
db6 | 0.0025837 | 0.1979 | -0.0087551 | 0.021505 | -0.069681 | -0.055191 | -0.064324 | 0.032129 | -0.070351 | -0.1025 | 0.030452 | -0.018724 | 0.04191 | 0.029251 | 0.012226 | 0.071562 | 0.14639 | 0.052672 | -0.006035 | -0.042538 | -0.10571
db7 | 1.1821 | 2.6828 | -0.20614 | -0.15056 | 0.037408 | 0.027802 | 0.065156 | 0.012059 | -0.020519 | -0.029293 | 0.036915 | 0.050346 | 0.045457 | 0.046302 | -0.02416 | -0.025891 | 0.11253 | 0.019892 | 0.021642 | -0.018944 | 0.88713
db8 | 0.0052239 | 0.00062977 | 0.0066892 | 0.024283 | -0.026697 | -0.015895 | -0.0044985 | 0.05718 | 0.042228 | -0.013211 | -0.01056 | 0.00087286 | -0.015761 | -0.012556 | 0.027421 | -0.025427 | -0.034298 | -0.0032198 | 0.021123 | -0.017672 | -0.058832
db9 | 0.5198 | 0.33135 | -0.002925 | -0.0071983 | 0.011505 | 0.0151 | -0.025039 | -0.012762 | 0.020496 | 0.014228 | -0.00083883 | -0.0062397 | -0.010545 | 0.0028008 | -0.011853 | -0.028363 | 0.022645 | -0.0040477 | 0.022426 | 0.00061187 | 0.44726
db10 | 0.0011221 | 0.067155 | -0.002024 | -0.032643 | -0.039161 | -0.031044 | -0.078571 | 0.043553 | 0.034408 | -0.010982 | 0.027042 | 0.05164 | 0.018587 | 0.0013333 | 0.053488 | -0.0059108 | -0.045547 | 0.022179 | 0.042855 | -0.049202 | -0.097251
db11 | 0.76151 | 0.5943 | -0.008282 | -0.0097761 | -0.0054045 | 0.00023229 | 0.015343 | 0.0049285 | 0.0023001 | 0.0059974 | 0.0051287 | 0.012571 | -0.00080731 | -0.0011906 | 0.014641 | 0.0072843 | -0.010832 | -0.0075897 | -0.0085987 | -0.015946 | 0.65913
db12 | 0.00096559 | 0.015673 | -0.002297 | -0.0046598 | -0.0027678 | -0.0011495 | -0.0034333 | 0.0083671 | 0.0048254 | -0.0040085 | -0.00098507 | 0.0069258 | -0.0017672 | -0.0020043 | 0.0063259 | -0.00024679 | 0.0022204 | 0.00083783 | 0.00083763 | -0.0070207 | -0.010648
db13 | 0.063296 | 0.094528 | 0.0055286 | 0.0063106 | 0.0024897 | 0.0013329 | -0.0036902 | -0.0029576 | -0.0021832 | 0.0008652 | -5.8219e-05 | -0.00024374 | -0.00096887 | -0.0016219 | -0.002868 | -0.0038492 | -0.00058078 | -5.7609e-05 | -0.00085681 | 0.0034093 | 0.054207
db14 | -0.0057795 | 0.0095971 | -0.0012041 | -0.0019614 | -0.00149 | -0.0017365 | -0.0055413 | 0.0038402 | 0.0031279 | -0.00044737 | 0.0005492 | 0.0041676 | 9.9012e-05 | -0.0020349 | 0.0049408 | -0.00034563 | -0.0015143 | 0.00069487 | 0.0029564 | -0.0041004 | -0.012903
db15 | 0.045642 | 0.021227 | -0.001648 | -0.0047845 | -0.0075066 | -0.0070685 | -0.0016202 | -0.002219 | -0.0014572 | -0.0022333 | 0.00091359 | 0.002735 | 0.0018559 | 0.0030091 | 0.0080121 | 0.0088765 | -0.00066685 | 0.0024792 | 0.00020359 | 0.0011192 | 0.045003
da2 | -1.5062 | 8.6461 | 0.39487 | -0.50763 | -0.62541 | 0.19283 | -1.0645 | -0.0077364 | -0.79186 | 0.25942 | 0.26811 | 0.58217 | 1.2067 | 1.0745 | 0.34031 | -0.2756 | -1.1634 | -0.43261 | -0.027918 | 0.5778 | 5.8207
da3 | 0.41325 | 0.49784 | 0.28158 | 0.32912 | -0.095312 | -0.17202 | -0.23781 | -0.56134 | -0.13794 | 0.22721 | -0.2387 | -0.25809 | -0.49324 | -0.3248 | -0.16384 | -0.23523 | 0.10018 | 0.31415 | 0.74141 | 0.92468 | 0.9429
da4 | -0.083897 | 2.5362 | 0.33156 | 0.22055 | 0.43275 | 0.27767 | 0.093944 | -0.16411 | 0.28395 | -0.061409 | -0.35244 | 0.27109 | -0.39948 | -0.12389 | -0.007561 | -0.1614 | -0.0045356 | -0.22069 | -0.34686 | -0.069141 | 1.0016
da5 | 0.062727 | 2.7915 | 0.090595 | -0.15864 | -0.16956 | 0.056415 | -0.068765 | -0.1269 | -0.10316 | 0.11043 | 0.35046 | 0.069746 | -0.23464 | -0.22958 | 0.17428 | -0.053959 | -0.1697 | 0.039816 | 0.11084 | 0.31233 | -0.035274
da6 | 0.016373 | -0.33612 | 0.1342 | 0.083304 | 0.097368 | 0.0081264 | -0.062415 | -0.1338 | -0.043146 | -0.058097 | -0.017868 | -0.016946 | 0.11712 | 0.036808 | -0.0049878 | -0.10502 | -0.15125 | -0.017078 | 0.081428 | 0.052241 | 0.012298
da7 | 0.015116 | 1.7031 | -0.078309 | -0.02925 | -0.045686 | -0.051192 | -0.044082 | 0.057657 | 0.12593 | 0.02112 | -0.043104 | 0.0084818 | -0.063462 | -0.066959 | 0.067349 | 0.0090403 | -0.002268 | -0.013377 | 0.066694 | 0.081419 | 0.13211
da8 | -0.018586 | -0.15361 | 0.033696 | 0.010535 | 0.011939 | -0.0028257 | -0.068902 | -0.017007 | 0.0053088 | -0.0019648 | 0.036925 | -0.0045761 | 0.057976 | 0.028172 | 0.013059 | -0.045308 | -0.058507 | -0.012131 | 0.0074776 | 0.0061319 | -0.092381
da9 | 0.0020732 | -0.2555 | 0.059254 | 0.040965 | -0.0094641 | -0.011124 | -0.023995 | -0.022747 | 0.0131 | 0.013378 | -0.014737 | -0.029854 | -0.0073016 | 0.017931 | -0.013526 | -0.036939 | -0.022577 | -0.010425 | 0.022098 | 0.035962 | -0.010305
da10 | 0.00926 | -0.12659 | 0.038755 | 0.05101 | 0.049443 | 0.040011 | -0.020113 | 0.021306 | 0.020422 | 0.011037 | -0.011582 | -0.0096176 | -0.0046187 | -0.016877 | -0.035191 | -0.073542 | -0.029546 | -0.026913 | 0.016166 | -0.02015 | -0.15284
da11 | 0.002886 | 0.18981 | -0.015041 | -0.0069095 | -0.019598 | -0.010168 | -0.0034126 | 0.0022683 | 0.017263 | 0.0090889 | -0.003816 | -0.00078625 | 0.00099412 | 0.0045482 | 0.0021219 | -0.0045376 | -0.0032371 | 0.0076816 | 0.013516 | 0.010024 | 0.016494
da12 | -0.0030057 | -0.022269 | 0.010451 | 0.0098733 | 0.0074188 | 0.0027537 | -0.0016962 | -0.001584 | -6.307e-06 | -0.0025428 | -0.0032379 | -0.00018268 | -0.00028521 | 0.0021226 | -0.00061016 | -0.012642 | -0.008431 | -0.0053719 | 0.0032089 | 0.0007616 | -0.017579
da13 | 0.0027896 | -0.026176 | 0.0028693 | 0.0048378 | -0.002105 | 0.00020771 | -0.00018242 | -0.0027839 | -0.00038746 | 0.0012175 | -0.0028291 | -0.0043295 | -0.0015098 | -0.00094599 | -0.00097367 | 4.8318e-05 | 0.00044351 | -0.0012594 | 0.0013671 | 0.0063151 | 0.0059631
da14 | -0.023603 | 0.011983 | 0.0029844 | 0.0029084 | 0.0040868 | 0.0030546 | 0.00083029 | -0.00022397 | 0.0025173 | 0.0012317 | 0.00076307 | -0.0005578 | -0.00013692 | -0.00053025 | -0.0025606 | -0.0075258 | -0.0010625 | -0.0029524 | -0.0035462 | 0.00071984 | -0.016792
da15 | -0.0097773 | -0.0067136 | -0.00029276 | -0.0013977 | -0.0028383 | -0.0028123 | -0.0095213 | 0.0067404 | 0.006174 | -0.0014447 | 0.0016501 | 0.0048477 | -0.00044937 | -0.0025515 | 0.0064242 | -0.00162 | -0.0033527 | 0.00078118 | 0.0051334 | -0.0054705 | -0.02424
Field Colinearity | 19.959
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.1867 | 0.33994 | -0.25041 | -0.55251 | -0.14444 | -5.7e-05 | 0.36158 | 0.44185 | 0.060854 | -0.41061 | -0.46966 | -0.055777 | 0.50268 | 0.16167 | -0.21346 | -0.47312 | -0.48209 | -0.081662 | -0.24235 | 0.35648
Twist (m2 mrad) | 16.4006 | 4.8441 | -3.3805 | -7.0446 | -1.7333 | -0.00064125 | 3.7966 | 4.308 | 0.54769 | -3.3875 | -3.5225 | -0.37649 | 3.0161 | 0.84876 | -0.96058 | -1.7742 | -1.4463 | -0.18374 | -0.36353 | 0.13368
Twist Integral (I0) (m2 rad) | 0.0072913
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.7283 | -0.4739 | -0.20135 | -0.6282 | -0.2964 | -0.063673 | 0.24605 | 0.73764 | 0.21517 | -0.46113 | -0.40638 | 0.11448 | 0.32998 | 0.21746 | -0.16246 | -0.16666 | -0.38128 | 0.060228 | 0.18369 | -0.73911
Twist (m2 mrad) | 20.4619 | -6.7531 | -2.7182 | -8.0095 | -3.5567 | -0.71632 | 2.5835 | 7.1919 | 1.9365 | -3.8043 | -3.0479 | 0.77273 | 1.9799 | 1.1417 | -0.73105 | -0.62499 | -1.1438 | 0.13551 | 0.27554 | -0.27717
Twist Integral (I0) (m2 rad) | 0.0038222

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-1.1732 | -2.8185 | 0.082989 | -0.052096 | 1.1476 | -0.83755 | 0.66346 | 0.67819 | 0.085081 | 19.3424 | 0.59595 | 1.4794 | -2.3382 | -0.07597 | -0.33265 | 1.1821 | -0.63151 | -1.5062 | 0.41325 | -0.083897 | 22.4497 | 0.59602 | -1.1014 | -1.5563 | 0.071992 | -0.34051 | 1.1931 | 0 | 1.718 | 0.62185 | 0.20282 | 14420.581 | 1.5385 | -1.1577 | -0.064656 | -0.63385 | 1.2182 | 0 | -0.96607 | 0.43266 | 0.029946 | 14419.866
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8