## Alarm sheet
File name | HCMB__A001-03000169_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000169
Date of test Ap 1 | 2004-06-03 |  |  |  | Date of test Ap 2 | 2004-06-03
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
423 | HCMB__A001 | 3000169 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2004-06-03 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2004-06-03 | Answer to manufacturer |  |  | 2004-06-03
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000169_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000169 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 03-JUN-2004_3169cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 03-JUN-2004_3169cc2.txt |  |  |  | a even | 1
Test Operator |  | Mahler |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2004-06-03 |  |  | Time of test |  |  | 10:50 - 13:41 |  |  |  | Date of test |  | 2004-06-03 |  |  | Time of test |  |  | 10:50 - 13:42
Mole name |  | Dipole 7 | [9 empty] | Mole name |  | Dipole 1 | [5 empty] | 2004-03-26
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.63 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.63 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00013115 |  |  | Dy average |  |  | 6.4316e-05 |  |  |  | Dx average |  | 0.0002093 |  |  | Dy average |  |  | -0.00010806
Magnetic length |  | 14.4208 |  |  | Tranfs. Function (mT/KA) |  |  | 596.4884 |  |  |  | Magnetic length |  | 14.4211 |  |  | Tranfs. Function (mT/KA) |  |  | 596.4748
Phase |  | 48.5506 |  |  | Integral field (T.m) |  |  | 85.8761 |  |  |  | Phase |  | -27.2678 |  |  | Integral field (T.m) |  |  | 85.8459
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4552 | 5.9675 | 5.9639 | 5.9644 | 5.9639 | 5.9649 | 5.965 | 5.9645 | 5.9645 | 5.9645 | 5.9643 | 5.9649 | 5.9648 | 5.9642 | 5.9643 | 5.9645 | 5.9652 | 5.9644 | 5.9681 | 3.6935 | 85.8761 | C1 (T) | 3.3809 | 5.968 | 5.964 | 5.9654 | 5.964 | 5.9646 | 5.9646 | 5.9645 | 5.9643 | 5.9639 | 5.964 | 5.9645 | 5.9642 | 5.9638 | 5.9639 | 5.9644 | 5.9648 | 5.9644 | 5.9681 | 3.7301 | 85.8459
Angle (mrad) | 3.079 | -0.73659 | -0.16873 | -0.33907 | -0.23752 | -0.010278 | 0.18639 | 0.55898 | 0.55015 | -0.030081 | -0.30761 | 0.022119 | 0.39241 | 0.61335 | -0.15818 | -0.31489 | -0.50473 | -0.21851 | -0.13373 | -1.5288 | 48.5506 | Angle (mrad) | 4.3709 | -0.41554 | -0.055879 | -0.31868 | -0.2055 | -0.001745 | 0.18667 | 0.65607 | 0.58427 | 0.03403 | -0.41155 | -0.18344 | 0.13674 | 0.5 | -0.16406 | -0.58585 | -0.67477 | -0.41545 | -0.37773 | -1.2226 | -27.2678
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 4.1021 | -0.17046 | 0.2572 | 0.46098 | 0.62449 | 0.097974 | -0.28474 | -0.40184 | -1.1921 | -0.67015 | -1.0952 | -0.17379 | -0.50353 | 0.29905 | -0.42859 | -0.81506 | -0.56558 | -0.053186 | -0.062565 | 0.7764 | -0.094803 | b2 | 4.2407 | 0.15202 | 0.62743 | 0.44098 | -0.017658 | 0.17565 | -0.027841 | 0.20446 | 0.73434 | 0.55342 | 0.43607 | 0.53969 | -0.76435 | -0.55001 | -0.36012 | -0.35567 | -0.24594 | 0.11725 | 0.37101 | -2.1035 | 0.16252
b3 | 43.7899 | -2.0683 | -1.2981 | -2.2067 | -2.6251 | -2.7612 | -1.598 | -1.6678 | -1.3611 | -2.7581 | -2.9943 | -1.9157 | -2.0788 | -3.9833 | -3.2151 | -1.7991 | -1.7273 | -2.2584 | -1.7347 | -1.1771 | -0.80286 | b3 | 44.7014 | -1.6704 | -1.2668 | -2.7863 | -1.7962 | -3.79 | -1.5736 | -1.5128 | -0.84135 | -2.7921 | -2.6108 | -1.9909 | -1.85 | -4.5054 | -3.3743 | -1.7147 | -1.3361 | -1.566 | -1.601 | -0.69463 | -0.71253
b4 | -0.17053 | -0.17807 | 0.11622 | 0.25191 | 0.044349 | 0.074742 | 0.061935 | 0.18632 | -0.042112 | 0.0040847 | -0.22802 | 0.022167 | -0.051546 | 0.070692 | 0.075203 | -0.14072 | -0.080539 | 0.082175 | 0.083466 | 0.19401 | 0.019456 | b4 | -0.39645 | 0.029395 | 0.15245 | 0.13236 | 0.12934 | -0.21548 | -0.18256 | -0.030851 | 0.07187 | 0.2228 | 0.27071 | 0.13221 | 0.030811 | -0.21611 | -0.018297 | -0.0035977 | -0.056185 | 0.12524 | 0.07478 | -0.092656 | 0.019083
b5 | -4.4459 | -0.76158 | -0.7233 | -0.20204 | -0.0035324 | 0.26324 | 0.41404 | 0.2006 | -0.054896 | 0.13034 | 0.19345 | -0.04824 | 0.22456 | 0.02274 | -0.030959 | -0.10288 | -0.16122 | -0.37237 | -0.35414 | -3.6347 | -0.32256 | b5 | -5.3309 | -0.61035 | -0.91102 | -0.08497 | -0.17496 | 0.036016 | -0.21734 | -0.059709 | -0.22976 | 0.043364 | 0.25383 | -0.030037 | -0.12988 | -0.10413 | -0.19881 | -0.34305 | -0.48095 | -0.56183 | -0.55088 | -3.5826 | -0.50109
b6 | 0.18133 | -0.0023595 | -0.052477 | -0.070513 | -0.015912 | -0.017847 | -0.049867 | -0.035347 | -0.0043802 | -0.02198 | 0.011364 | -0.044344 | 0.057493 | -0.18177 | 0.032067 | 0.046521 | 0.0036953 | 0.055584 | 0.051497 | 0.14549 | -0.0022604 | b6 | 0.0099739 | -0.069858 | 0.0042527 | -0.022514 | -0.062856 | 0.018282 | 0.040391 | -0.014338 | 0.052263 | -0.0034647 | 0.04142 | -0.039674 | 0.046742 | -0.0039131 | 0.027599 | 0.077046 | 0.056407 | 0.063705 | 0.051567 | -0.12631 | 0.0098839
b7 | 3.1868 | 1.0641 | 1.127 | 1.1632 | 1.1882 | 1.1112 | 1.0914 | 1.0944 | 1.1037 | 1.0478 | 1.0886 | 1.1304 | 1.1773 | 1.1895 | 1.1636 | 1.1922 | 1.1732 | 1.1912 | 1.1717 | 1.006 | 1.1948 | b7 | 2.8539 | 0.97603 | 1.0379 | 1.201 | 1.1703 | 1.1535 | 1.1673 | 1.1232 | 1.149 | 1.1157 | 1.1678 | 1.1396 | 1.1153 | 1.1728 | 1.1346 | 1.1745 | 1.0983 | 1.1191 | 1.1096 | 0.93242 | 1.1737
b8 | 0.023022 | 0.040613 | -0.034655 | -0.041637 | 0.015791 | -0.010806 | 0.0049213 | 0.002374 | 0.0020826 | -0.04302 | -0.053986 | -0.015415 | 0.015782 | -0.010458 | -0.019179 | -0.0051638 | -0.028128 | 0.031678 | 0.0082716 | 0.0062503 | -0.0064429 | b8 | -0.064709 | 0.02464 | -0.022109 | -0.039045 | -0.015684 | 0.0072187 | -0.013119 | 0.029457 | 0.021483 | -0.036209 | -0.06776 | -0.063315 | 0.0018964 | -0.051836 | -0.011418 | 0.060668 | -0.015383 | 0.010264 | 0.0081599 | -0.02445 | -0.011672
b9 | 0.45376 | 0.50893 | 0.49139 | 0.50063 | 0.51684 | 0.49269 | 0.51079 | 0.52048 | 0.53105 | 0.50724 | 0.48758 | 0.48338 | 0.4807 | 0.51104 | 0.47534 | 0.50398 | 0.48669 | 0.50763 | 0.4938 | 0.43438 | 0.49702 | b9 | 0.38541 | 0.47765 | 0.47003 | 0.47556 | 0.50742 | 0.44869 | 0.48921 | 0.47697 | 0.52303 | 0.47203 | 0.46376 | 0.46267 | 0.47205 | 0.49497 | 0.46433 | 0.48127 | 0.47041 | 0.47484 | 0.47447 | 0.42725 | 0.47337
b10 | 0.098146 | 0.025658 | -0.067537 | -0.040466 | 0.014479 | -0.031283 | -0.028575 | 0.064543 | 0.087487 | 0.01844 | -0.048007 | -0.023167 | 0.036917 | -0.063775 | 0.016794 | -0.0013587 | -0.053747 | 0.026647 | -0.00072314 | 0.017451 | 0 | b10 | 0.080669 | 0.010005 | -0.054101 | -0.026546 | 0.035443 | -0.012534 | -0.010952 | 0.023077 | 0.057733 | 0.0097253 | -0.0086715 | -0.038117 | 0.013125 | -0.070089 | 0.014655 | 0.082318 | -0.054378 | 0.012758 | 0.020606 | -0.079628 | 0
b11 | 0.61724 | 0.74178 | 0.7404 | 0.74907 | 0.74982 | 0.75706 | 0.75331 | 0.75762 | 0.75713 | 0.76049 | 0.76567 | 0.75311 | 0.75501 | 0.76139 | 0.75791 | 0.74451 | 0.74216 | 0.74925 | 0.74264 | 0.64685 | 0.74466 | b11 | 0.59995 | 0.74215 | 0.7393 | 0.75585 | 0.74668 | 0.7591 | 0.75322 | 0.75272 | 0.74851 | 0.75264 | 0.75664 | 0.75088 | 0.75339 | 0.766 | 0.76074 | 0.75239 | 0.74592 | 0.74331 | 0.73996 | 0.64141 | 0.74304
b12 | 0.018194 | 0.0073049 | -0.00812 | -0.0094324 | -0.00099053 | -0.0066684 | -0.0048268 | -0.0024688 | -0.00061792 | -0.004036 | -0.007304 | -0.0077782 | 0.0041416 | -0.0080095 | -0.0031123 | -0.0015716 | -0.0059485 | 0.0033748 | -0.0022661 | -0.00324 | -0.0025935 | b12 | 0.0079502 | 0.0015148 | -0.0066834 | -0.0098478 | -0.0020719 | -0.0048758 | 0.0020563 | 0.00076334 | 0.0062689 | -0.0020085 | -0.00079789 | -0.0060428 | 0.0027849 | -0.0092644 | -0.0047702 | 0.0045836 | -0.0093149 | -0.00079765 | 0.0020317 | -0.006454 | -0.0018757
b13 | 0.091963 | 0.069945 | 0.072098 | 0.064671 | 0.063236 | 0.059854 | 0.057803 | 0.060159 | 0.061071 | 0.060469 | 0.06343 | 0.065474 | 0.064977 | 0.061449 | 0.065262 | 0.064121 | 0.064806 | 0.068782 | 0.068142 | 0.053233 | 0.064692 | b13 | 0.09425 | 0.071753 | 0.073854 | 0.065444 | 0.067449 | 0.066751 | 0.070581 | 0.064562 | 0.063539 | 0.065584 | 0.065895 | 0.067954 | 0.068113 | 0.065553 | 0.066994 | 0.070043 | 0.07028 | 0.072111 | 0.070684 | 0.055699 | 0.068538
b14 | -0.0021332 | -0.0010844 | -0.004737 | -0.0042495 | -0.0035445 | -0.0043285 | -0.0039805 | -0.00034702 | 0.0011236 | -0.0016252 | -0.0044196 | -0.0036134 | -0.0006303 | -0.0046623 | -0.0018409 | -0.003021 | -0.0035707 | -0.0012118 | -0.00215 | 0.00021738 | -0.0025519 | b14 | -0.00016937 | -0.0042418 | -0.0065605 | -0.0049788 | -0.0032174 | -0.0045228 | -0.0037733 | -0.0025527 | -0.0019041 | -0.0030635 | -0.0045264 | -0.0056384 | -0.0035738 | -0.0077945 | -0.0042738 | -0.0015251 | -0.0059494 | -0.0039556 | -0.0036312 | -0.0013992 | -0.003994
b15 | -0.016942 | 0.021918 | 0.019169 | 0.02026 | 0.022868 | 0.02249 | 0.022979 | 0.02341 | 0.023437 | 0.02494 | 0.022227 | 0.020716 | 0.022147 | 0.020474 | 0.021803 | 0.021794 | 0.020743 | 0.022333 | 0.022329 | 0.011736 | 0.020496 | b15 | -0.016569 | 0.022665 | 0.019956 | 0.019632 | 0.02224 | 0.020045 | 0.020341 | 0.021968 | 0.022342 | 0.022752 | 0.020357 | 0.020499 | 0.022334 | 0.019601 | 0.02089 | 0.022046 | 0.020471 | 0.022144 | 0.020823 | 0.0015392 | 0.019418
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 30.7906 | -7.3659 | -1.6873 | -3.3907 | -2.3752 | -0.10278 | 1.8639 | 5.5898 | 5.5015 | -0.30081 | -3.076 | 0.22119 | 3.9241 | 6.1335 | -1.5818 | -3.149 | -5.0473 | -2.1851 | -1.3373 | -15.2882 | 0 | a1 | 43.7096 | -4.1554 | -0.55879 | -3.1868 | -2.055 | -0.017448 | 1.8667 | 6.5608 | 5.8426 | 0.3403 | -4.1155 | -1.8344 | 1.3674 | 5 | -1.6406 | -5.8585 | -6.7477 | -4.1545 | -3.7773 | -12.2264 | 0
a2 | 10.6496 | -0.05481 | 0.32724 | 0.13276 | 0.43808 | -0.28881 | 0.34078 | 1.0979 | 0.37492 | 0.40356 | -0.60668 | 0.39448 | 0.2546 | 0.69309 | 0.72438 | 0.19867 | 0.90738 | 0.97012 | -0.28512 | -3.7382 | 0.51442 | a2 | 17.3359 | -0.20026 | -0.29887 | 0.71742 | -0.29802 | -0.87192 | -0.39974 | -0.0081708 | -0.88301 | -0.33421 | -0.74463 | -0.40275 | -0.82326 | -0.60623 | -0.40968 | -0.02223 | -0.10091 | 0.094201 | -1.5294 | 0.25069 | 0.14907
a3 | -2.283 | 0.72575 | 0.80875 | 0.71542 | 0.99752 | 0.66228 | 0.57478 | 0.90735 | 0.86956 | 0.94513 | 0.65777 | 0.783 | 0.52538 | 0.71634 | 0.5104 | 0.26506 | 0.38903 | 0.84836 | 1.1533 | 0.099661 | 0.61435 | a3 | 1.2678 | 0.9657 | 0.6707 | 1.0341 | 1.2438 | 0.90783 | 0.65711 | 1.156 | 1.0804 | 0.83355 | 1.2458 | 0.82076 | 0.4176 | 0.74042 | 0.48249 | 0.40478 | 0.15872 | 0.67488 | 0.76198 | 0.3423 | 0.79144
a4 | 3.0266 | -0.18524 | -0.007445 | 0.50864 | 0.2734 | 0.52958 | 0.55044 | -0.18253 | -0.21377 | 0.038346 | 0.065653 | -0.21943 | -0.18236 | -0.50198 | -0.10248 | 0.37185 | 0.24738 | -0.13165 | -0.20717 | 0.44589 | 0.13961 | a4 | 4.3737 | -0.10756 | -0.075668 | -0.033405 | 0.26487 | 0.22087 | -0.23983 | -0.2477 | 0.032135 | -0.46021 | 0.47512 | -0.088806 | -0.054693 | 0.084974 | 0.078235 | -0.024686 | 0.23674 | 0.1783 | 0.31521 | 0.70661 | 0.18106
a5 | 2.8158 | 0.06538 | 0.068351 | 0.068045 | 0.41076 | 0.19007 | 0.22625 | 0.24342 | 0.22562 | 0.33035 | 0.24851 | 0.12401 | 0.042993 | 0.2046 | 0.22756 | 0.13554 | 0.1098 | 0.27151 | 0.51744 | -0.092145 | 0.27525 | a5 | 2.8839 | 0.1181 | -0.024157 | 0.14672 | 0.28263 | 0.088626 | 0.01532 | 0.25581 | 0.27168 | 0.47676 | 0.50862 | 0.13774 | 0.1908 | 0.29205 | 0.10955 | 0.18536 | 0.069994 | 0.29805 | 0.42444 | -0.067398 | 0.28348
a6 | 0.56658 | 0.074936 | 0.10383 | 0.2076 | 0.21912 | -0.049329 | -0.052774 | -0.083304 | -0.11093 | -0.11809 | 0.012177 | -0.04902 | -0.085635 | -0.10526 | -0.009036 | 0.077872 | 0.11497 | 0.053486 | 0.14616 | -0.1002 | 0.031929 | a6 | 0.60415 | 0.13372 | 0.12634 | 0.073737 | 0.1674 | 0.064717 | -0.016493 | -0.0091423 | -0.053603 | -0.010369 | 0.022797 | 0.021999 | 0.00067353 | 0.090585 | 0.0030306 | -0.054708 | 0.0079889 | 0.030343 | 0.057887 | 0.027261 | 0.052962
a7 | 1.7084 | -0.12674 | -0.030068 | -0.0058989 | 0.071423 | 0.019008 | 0.068812 | 0.10208 | 0.19042 | 0.031693 | 0.026977 | 0.017906 | 0.092447 | 0.06208 | 0.0072112 | -0.011264 | 0.01682 | 0.059789 | 0.084797 | -0.015033 | 0.086348 | a7 | 1.8165 | -0.026021 | 0.017037 | -0.00016782 | 0.041287 | -0.016891 | -0.0056705 | 0.12646 | 0.092198 | 0.029237 | 0.091485 | 0.040113 | 0.055008 | 0.02227 | 0.050023 | -0.026869 | -0.038368 | 0.059819 | 0.068604 | 0.030394 | 0.084834
a8 | 0.034362 | 0.014098 | 0.014617 | 0.046537 | 0.023807 | 0.032133 | 0.040085 | 0.02468 | -0.028015 | 0.010227 | -0.0090819 | 0.0082441 | -0.02788 | 0.0052744 | 0.021328 | 0.029531 | 0.072198 | 0.010534 | 0.040914 | -0.075798 | 0.015742 | a8 | -0.019683 | -0.020504 | 0.027322 | 0.061587 | 0.076279 | 0.054934 | 0.0073342 | 0.0063882 | 0.0048768 | -0.010274 | 0.018445 | 0.023928 | 0.027674 | 0.054203 | 0.043317 | 0.028334 | 0.015177 | 0.014391 | 0.037174 | 0.017953 | 0.024522
a9 | -0.20933 | -0.019758 | -0.028738 | -0.020412 | -0.004879 | -0.030641 | -0.019875 | 0.015075 | 0.0020922 | 0.0043039 | -0.0083784 | -0.0065322 | -3.7229e-05 | -0.018475 | -0.033941 | -0.053496 | -0.024827 | 0.0061289 | 0.019246 | -0.06103 | -0.019907 | a9 | -0.24936 | 0.02896 | 0.0063714 | 0.032671 | 0.020192 | 0.0067075 | 0.013825 | 0.041395 | 0.02505 | 0.024098 | 0.043897 | -0.00011434 | 0.021146 | 0.0087241 | 0.0098341 | 0.02638 | 0.012154 | 0.014187 | 0.034059 | -0.022956 | 0.011143
a10 | 0.026551 | -0.0087519 | -2.7681e-05 | 0.027158 | 0.038424 | 0.00079419 | 0.0063927 | 0.013857 | -0.013083 | 7.5462e-05 | -0.0018377 | -0.0024279 | -0.031216 | -0.014924 | 0.046041 | -0.0037952 | 0.026039 | -0.020741 | 0.027106 | -0.16871 | 0 | a10 | -0.046832 | -0.0078389 | 0.028418 | 0.023135 | 0.034558 | 0.022907 | -0.010591 | 0.021466 | -0.013363 | -0.022786 | -0.0028741 | 0.012168 | -0.011072 | 0.039237 | 0.013747 | -0.01052 | -0.0075049 | -0.029468 | -0.012082 | -0.065517 | 0
a11 | 0.17964 | -0.064434 | -0.04779 | -0.054116 | -0.048071 | -0.043639 | -0.033762 | -0.024293 | -0.018343 | -0.035755 | -0.034999 | -0.046186 | -0.03333 | -0.036205 | -0.039966 | -0.044082 | -0.040192 | -0.036318 | -0.032287 | -0.045493 | -0.033226 | a11 | 0.19543 | -0.021927 | -0.010727 | -0.0094255 | -0.012046 | -0.0066715 | -0.0044922 | 0.0077183 | 0.0092731 | -0.0037538 | 0.0057997 | -0.0028883 | -0.0066268 | 0.0029302 | -0.0025175 | -0.0056248 | -0.0063784 | -0.0011832 | 0.0026926 | -0.0060676 | 0.0021426
a12 | -0.0082687 | -0.0047194 | -0.0025175 | -0.00030803 | 0.0026491 | -0.010431 | -0.0085873 | -0.0051082 | -0.0063806 | -0.0082175 | -0.00090533 | -0.0082781 | -0.011934 | -0.013122 | -0.0026748 | -0.0076894 | 0.0019647 | -0.00288 | 0.0011991 | -0.016069 | -0.0053482 | a12 | -0.02009 | -0.0017285 | -0.0003896 | -0.00085488 | -0.00044291 | -0.0067415 | -0.010547 | -0.0049602 | -0.0072792 | -0.0076465 | -0.0062574 | -0.0053537 | -0.0071318 | -0.0010604 | -0.0068837 | -0.0056645 | -0.0075284 | -0.0094212 | -0.0061014 | -0.013919 | -0.0060484
a13 | -0.017117 | -0.0037513 | -0.0033094 | -0.0032449 | -0.0022627 | -0.0045506 | -0.0028516 | 0.0020439 | 0.0017874 | -0.00021527 | -0.0028477 | -0.0022011 | -0.00045334 | -0.00034156 | -0.0047787 | -0.0062363 | -0.0040313 | -0.0013614 | 0.0021461 | -0.0027552 | -0.0025043 | a13 | -0.016671 | 0.0024558 | 0.00053246 | 0.0019198 | 0.0038708 | 0.001316 | 0.0024962 | 0.0053223 | 0.0041352 | 0.00452 | 0.0077559 | 0.0041864 | 0.0047764 | 0.0051001 | 0.0018785 | 0.0033497 | 0.0016 | 0.0039746 | 0.0065091 | -0.00086736 | 0.0029026
a14 | 0.0048304 | -0.003285 | -0.0019082 | -0.00063352 | -0.0021657 | 9.8697e-05 | 0.0010787 | -0.00061744 | -0.0024584 | 0.00019004 | -0.00090385 | -0.00019072 | -0.0037553 | -0.0031665 | -0.00027795 | -0.00075656 | 0.0017017 | -0.0028164 | -0.0020927 | 0.0018251 | -0.00093921 | a14 | 0.0030912 | -0.0013068 | 0.00068925 | 0.0016117 | 0.0034897 | 0.003456 | 0.0011393 | 0.0027419 | 0.0029271 | -0.00044233 | 0.00209 | 0.0027069 | 0.0031302 | 0.0033183 | 0.0041521 | 0.0031358 | 0.0025223 | 0.0012503 | 0.001775 | 0.0031733 | 0.0021947
a15 | 0.0077903 | -0.0023421 | -0.00072546 | -0.00063018 | -1.576e-05 | -0.00030521 | -0.00091196 | 0.0016907 | 0.001591 | -0.00081492 | -0.0015992 | -0.00042243 | 0.00018911 | -0.00032748 | -0.00066648 | -0.0010697 | -0.0011044 | -0.00074375 | -6.8545e-05 | 0.0017554 | -0.00013948 | a15 | 0.0041922 | 0.0013564 | 0.0033635 | 0.0038874 | 0.0033842 | 0.0036741 | 0.0025992 | 0.0036808 | 0.0020502 | 0.0026444 | 0.0038335 | 0.0049797 | 0.0034822 | 0.0060714 | 0.0034347 | 0.0013594 | 0.0027152 | 0.0038568 | 0.0025471 | -8.9882e-07 | 0.0031937
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00026917 | -6.0098e-05 | 0.00015447 | 9.5819e-05 | -2.7133e-05 | 7.0125e-05 | 6.5014e-05 | -0.00014367 | -0.00019695 | -4.1121e-05 | 0.00010621 | 5.1769e-05 | -8.6009e-05 | 0.00014047 | -3.2161e-05 | 0 | 0.00012599 | -6.2592e-05 | 0 | -7.6598e-05 | 0.00013115 | Dx (m) | -0.00016706 | -2.3428e-05 | 0.00012536 | 6.0354e-05 | -7.9415e-05 | 2.8525e-05 | 2.457e-05 | -5.2624e-05 | -0.00013073 | -2.2218e-05 | 1.9531e-05 | 8.6428e-05 | -2.9829e-05 | 0.00015529 | -3.2654e-05 | -0.00018617 | 0.00012381 | -2.9273e-05 | -4.7232e-05 | 0.00020897 | 0.0002093
Dy (m) | 0 | 1.4849e-05 | 1.0067e-05 | -5.4745e-05 | -8.8847e-05 | 0 | -1.15e-05 | -3.5634e-05 | 2.4689e-05 | 0 | 0 | 0 | 6.6519e-05 | 4.0127e-05 | -0.00010497 | 0 | -5.2853e-05 | 4.4032e-05 | -6.1859e-05 | 0.00043689 | 6.4316e-05 | Dy (m) | 0.00018723 | 1.7264e-05 | -6.3552e-05 | -5.1304e-05 | -7.9943e-05 | -5.1061e-05 | 2.4049e-05 | -4.7939e-05 | 3.2034e-05 | 5.1347e-05 | 0 | -2.721e-05 | 2.4723e-05 | -8.7735e-05 | -3.0823e-05 | 2.2464e-05 | 1.8224e-05 | 6.7334e-05 | 2.7933e-05 | 0.00017583 | -0.00010806
Magnetic length (m) |  | 14.4208 | [19 empty] | Magnetic length (m) |  | 14.4211
Transfer funct. (mT/KA) |  | 596.4884 | [19 empty] | Transfer funct. (mT/KA) |  | 596.4748
Coil positioning (mm) |  | 0.018731 | [19 empty] | Coil positioning (mm) |  | 0.017903
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.0.8 | [20 empty] | Soft version | DIMM V2.0.8
Level Average | 0.025727 | -0.10779 | -0.004092 | -0.057077 | 0.29645 | 0.12121 | 0.2276 | 0.18847 | 0.039694 | -0.27695 | 0.40857 | -0.42659 | -0.33961 | -0.16322 | -0.18337 | 0.005852 | 0.11756 | 0.52131 | -0.3352 | 0.50413 |  | Level Average | -0.040692 | -0.036747 | -0.031726 | -0.03661 | -0.053672 | -0.090526 | -0.070332 | -0.089204 | -0.397 | -0.38498 | -0.4454 | -0.44627 | -0.45027 | -0.38675 | -0.45363 | -0.41852 | -0.14448 | -0.13766 | -0.066352 | -0.40874
Level Sigma2 | 0.24589 | 0.23099 | 0.39039 | 0.23968 | 0.38956 | 0.2503 | 0.26138 | 0.18259 | 0.15047 | 0.008423 | 0.028786 | 0.022353 | 0.041143 | 0.045225 | 0.0933 | 0.21685 | 0.80536 | 7.117 | 1.6833 | 0.05539 |  | Level Sigma2 | 0.010166 | 0.002062 | 0.003476 | 0.003425 | 0.006776 | 0.116 | 0.012527 | 0.020976 | 0.14483 | 0.45851 | 0.30576 | 0.39504 | 0.68815 | 0.55965 | 0.22661 | 0.36977 | 0.63942 | 0.5539 | 0.047507 | 2.138
Position (mm) | 0 | 749 | 1501 | 2252 | 3006 | 3754 | 4506 | 5258 | 6010 | 6761 | 7512 | 8263 | 9014 | 9765 | 10517 | 11268 | 12018 | 12770 | 13521 | 14272 |  | Position (mm) | 0 | 751 | 1502 | 2254 | 3005 | 3757 | 4508 | 5260 | 6011 | 6763 | 7515 | 8266 | 9017 | 9769 | 10520 | 11271 | 12023 | 12774 | 13526 | 14277
Fist move (mm) | First move Ap1 | 941.2464 | First move Ap2 | 946.4588 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-3000169 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000169 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 345.5244 | 596.7513 | 596.3914 | 596.4447 | 596.3869 | 596.4915 | 596.4962 | 596.4456 | 596.4532 | 596.4525 | 596.4261 | 596.4947 | 596.4831 | 596.4228 | 596.4295 | 596.4466 | 596.5199 | 596.4442 | 596.8107 | 369.3524 | 8587.6061 | C1/i (mT/KA) | 338.0858 | 596.8003 | 596.4021 | 596.5393 | 596.3965 | 596.4581 | 596.4558 | 596.4463 | 596.4315 | 596.3949 | 596.3967 | 596.4516 | 596.4201 | 596.3788 | 596.3947 | 596.4432 | 596.4775 | 596.4449 | 596.8135 | 373.012 | 8584.5934
Angle (mrad) | 3.079 | -0.73659 | -0.16873 | -0.33907 | -0.23752 | -0.010278 | 0.18639 | 0.55898 | 0.55015 | -0.030081 | -0.30761 | 0.022119 | 0.39241 | 0.61335 | -0.15818 | -0.31489 | -0.50473 | -0.21851 | -0.13373 | -1.5288 | 48.5506 | Angle (mrad) | 4.3709 | -0.41554 | -0.055879 | -0.31868 | -0.2055 | -0.001745 | 0.18667 | 0.65607 | 0.58427 | 0.03403 | -0.41155 | -0.18344 | 0.13674 | 0.5 | -0.16406 | -0.58585 | -0.67477 | -0.41545 | -0.37773 | -1.2226 | -27.2678
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 4.1021 | -0.17046 | 0.2572 | 0.46098 | 0.62449 | 0.097974 | -0.28474 | -0.40184 | -1.1921 | -0.67015 | -1.0952 | -0.17379 | -0.50353 | 0.29905 | -0.42859 | -0.81506 | -0.56558 | -0.053186 | -0.062565 | 0.7764 |  | b2 | 4.2407 | 0.15202 | 0.62743 | 0.44098 | -0.017658 | 0.17565 | -0.027841 | 0.20446 | 0.73434 | 0.55342 | 0.43607 | 0.53969 | -0.76435 | -0.55001 | -0.36012 | -0.35567 | -0.24594 | 0.11725 | 0.37101 | -2.1035
b3 | 43.7899 | -2.0683 | -1.2981 | -2.2067 | -2.6251 | -2.7612 | -1.598 | -1.6678 | -1.3611 | -2.7581 | -2.9943 | -1.9157 | -2.0788 | -3.9833 | -3.2151 | -1.7991 | -1.7273 | -2.2584 | -1.7347 | -1.1771 |  | b3 | 44.7014 | -1.6704 | -1.2668 | -2.7863 | -1.7962 | -3.79 | -1.5736 | -1.5128 | -0.84135 | -2.7921 | -2.6108 | -1.9909 | -1.85 | -4.5054 | -3.3743 | -1.7147 | -1.3361 | -1.566 | -1.601 | -0.69463
b4 | -0.17053 | -0.17807 | 0.11622 | 0.25191 | 0.044349 | 0.074742 | 0.061935 | 0.18632 | -0.042112 | 0.0040847 | -0.22802 | 0.022167 | -0.051546 | 0.070692 | 0.075203 | -0.14072 | -0.080539 | 0.082175 | 0.083466 | 0.19401 |  | b4 | -0.39645 | 0.029395 | 0.15245 | 0.13236 | 0.12934 | -0.21548 | -0.18256 | -0.030851 | 0.07187 | 0.2228 | 0.27071 | 0.13221 | 0.030811 | -0.21611 | -0.018297 | -0.0035977 | -0.056185 | 0.12524 | 0.07478 | -0.092656
b5 | -4.4459 | -0.76158 | -0.7233 | -0.20204 | -0.0035324 | 0.26324 | 0.41404 | 0.2006 | -0.054896 | 0.13034 | 0.19345 | -0.04824 | 0.22456 | 0.02274 | -0.030959 | -0.10288 | -0.16122 | -0.37237 | -0.35414 | -3.6347 |  | b5 | -5.3309 | -0.61035 | -0.91102 | -0.08497 | -0.17496 | 0.036016 | -0.21734 | -0.059709 | -0.22976 | 0.043364 | 0.25383 | -0.030037 | -0.12988 | -0.10413 | -0.19881 | -0.34305 | -0.48095 | -0.56183 | -0.55088 | -3.5826
b6 | 0.18133 | -0.0023595 | -0.052477 | -0.070513 | -0.015912 | -0.017847 | -0.049867 | -0.035347 | -0.0043802 | -0.02198 | 0.011364 | -0.044344 | 0.057493 | -0.18177 | 0.032067 | 0.046521 | 0.0036953 | 0.055584 | 0.051497 | 0.14549 |  | b6 | 0.0099739 | -0.069858 | 0.0042527 | -0.022514 | -0.062856 | 0.018282 | 0.040391 | -0.014338 | 0.052263 | -0.0034647 | 0.04142 | -0.039674 | 0.046742 | -0.0039131 | 0.027599 | 0.077046 | 0.056407 | 0.063705 | 0.051567 | -0.12631
b7 | 3.1868 | 1.0641 | 1.127 | 1.1632 | 1.1882 | 1.1112 | 1.0914 | 1.0944 | 1.1037 | 1.0478 | 1.0886 | 1.1304 | 1.1773 | 1.1895 | 1.1636 | 1.1922 | 1.1732 | 1.1912 | 1.1717 | 1.006 |  | b7 | 2.8539 | 0.97603 | 1.0379 | 1.201 | 1.1703 | 1.1535 | 1.1673 | 1.1232 | 1.149 | 1.1157 | 1.1678 | 1.1396 | 1.1153 | 1.1728 | 1.1346 | 1.1745 | 1.0983 | 1.1191 | 1.1096 | 0.93242
b8 | 0.023022 | 0.040613 | -0.034655 | -0.041637 | 0.015791 | -0.010806 | 0.0049213 | 0.002374 | 0.0020826 | -0.04302 | -0.053986 | -0.015415 | 0.015782 | -0.010458 | -0.019179 | -0.0051638 | -0.028128 | 0.031678 | 0.0082716 | 0.0062503 |  | b8 | -0.064709 | 0.02464 | -0.022109 | -0.039045 | -0.015684 | 0.0072187 | -0.013119 | 0.029457 | 0.021483 | -0.036209 | -0.06776 | -0.063315 | 0.0018964 | -0.051836 | -0.011418 | 0.060668 | -0.015383 | 0.010264 | 0.0081599 | -0.02445
b9 | 0.45376 | 0.50893 | 0.49139 | 0.50063 | 0.51684 | 0.49269 | 0.51079 | 0.52048 | 0.53105 | 0.50724 | 0.48758 | 0.48338 | 0.4807 | 0.51104 | 0.47534 | 0.50398 | 0.48669 | 0.50763 | 0.4938 | 0.43438 |  | b9 | 0.38541 | 0.47765 | 0.47003 | 0.47556 | 0.50742 | 0.44869 | 0.48921 | 0.47697 | 0.52303 | 0.47203 | 0.46376 | 0.46267 | 0.47205 | 0.49497 | 0.46433 | 0.48127 | 0.47041 | 0.47484 | 0.47447 | 0.42725
b10 | 0.098146 | 0.025658 | -0.067537 | -0.040466 | 0.014479 | -0.031283 | -0.028575 | 0.064543 | 0.087487 | 0.01844 | -0.048007 | -0.023167 | 0.036917 | -0.063775 | 0.016794 | -0.0013587 | -0.053747 | 0.026647 | -0.00072314 | 0.017451 |  | b10 | 0.080669 | 0.010005 | -0.054101 | -0.026546 | 0.035443 | -0.012534 | -0.010952 | 0.023077 | 0.057733 | 0.0097253 | -0.0086715 | -0.038117 | 0.013125 | -0.070089 | 0.014655 | 0.082318 | -0.054378 | 0.012758 | 0.020606 | -0.079628
b11 | 0.61724 | 0.74178 | 0.7404 | 0.74907 | 0.74982 | 0.75706 | 0.75331 | 0.75762 | 0.75713 | 0.76049 | 0.76567 | 0.75311 | 0.75501 | 0.76139 | 0.75791 | 0.74451 | 0.74216 | 0.74925 | 0.74264 | 0.64685 |  | b11 | 0.59995 | 0.74215 | 0.7393 | 0.75585 | 0.74668 | 0.7591 | 0.75322 | 0.75272 | 0.74851 | 0.75264 | 0.75664 | 0.75088 | 0.75339 | 0.766 | 0.76074 | 0.75239 | 0.74592 | 0.74331 | 0.73996 | 0.64141
b12 | 0.018194 | 0.0073049 | -0.00812 | -0.0094324 | -0.00099053 | -0.0066684 | -0.0048268 | -0.0024688 | -0.00061792 | -0.004036 | -0.007304 | -0.0077782 | 0.0041416 | -0.0080095 | -0.0031123 | -0.0015716 | -0.0059485 | 0.0033748 | -0.0022661 | -0.00324 |  | b12 | 0.0079502 | 0.0015148 | -0.0066834 | -0.0098478 | -0.0020719 | -0.0048758 | 0.0020563 | 0.00076334 | 0.0062689 | -0.0020085 | -0.00079789 | -0.0060428 | 0.0027849 | -0.0092644 | -0.0047702 | 0.0045836 | -0.0093149 | -0.00079765 | 0.0020317 | -0.006454
b13 | 0.091963 | 0.069945 | 0.072098 | 0.064671 | 0.063236 | 0.059854 | 0.057803 | 0.060159 | 0.061071 | 0.060469 | 0.06343 | 0.065474 | 0.064977 | 0.061449 | 0.065262 | 0.064121 | 0.064806 | 0.068782 | 0.068142 | 0.053233 |  | b13 | 0.09425 | 0.071753 | 0.073854 | 0.065444 | 0.067449 | 0.066751 | 0.070581 | 0.064562 | 0.063539 | 0.065584 | 0.065895 | 0.067954 | 0.068113 | 0.065553 | 0.066994 | 0.070043 | 0.07028 | 0.072111 | 0.070684 | 0.055699
b14 | -0.0021332 | -0.0010844 | -0.004737 | -0.0042495 | -0.0035445 | -0.0043285 | -0.0039805 | -0.00034702 | 0.0011236 | -0.0016252 | -0.0044196 | -0.0036134 | -0.0006303 | -0.0046623 | -0.0018409 | -0.003021 | -0.0035707 | -0.0012118 | -0.00215 | 0.00021738 |  | b14 | -0.00016937 | -0.0042418 | -0.0065605 | -0.0049788 | -0.0032174 | -0.0045228 | -0.0037733 | -0.0025527 | -0.0019041 | -0.0030635 | -0.0045264 | -0.0056384 | -0.0035738 | -0.0077945 | -0.0042738 | -0.0015251 | -0.0059494 | -0.0039556 | -0.0036312 | -0.0013992
b15 | -0.016942 | 0.021918 | 0.019169 | 0.02026 | 0.022868 | 0.02249 | 0.022979 | 0.02341 | 0.023437 | 0.02494 | 0.022227 | 0.020716 | 0.022147 | 0.020474 | 0.021803 | 0.021794 | 0.020743 | 0.022333 | 0.022329 | 0.011736 |  | b15 | -0.016569 | 0.022665 | 0.019956 | 0.019632 | 0.02224 | 0.020045 | 0.020341 | 0.021968 | 0.022342 | 0.022752 | 0.020357 | 0.020499 | 0.022334 | 0.019601 | 0.02089 | 0.022046 | 0.020471 | 0.022144 | 0.020823 | 0.0015392
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 30.7906 | -7.3659 | -1.6873 | -3.3907 | -2.3752 | -0.10278 | 1.8639 | 5.5898 | 5.5015 | -0.30081 | -3.076 | 0.22119 | 3.9241 | 6.1335 | -1.5818 | -3.149 | -5.0473 | -2.1851 | -1.3373 | -15.2882 |  | a1 | 43.7096 | -4.1554 | -0.55879 | -3.1868 | -2.055 | -0.017448 | 1.8667 | 6.5608 | 5.8426 | 0.3403 | -4.1155 | -1.8344 | 1.3674 | 5 | -1.6406 | -5.8585 | -6.7477 | -4.1545 | -3.7773 | -12.2264
a2 | 10.6496 | -0.05481 | 0.32724 | 0.13276 | 0.43808 | -0.28881 | 0.34078 | 1.0979 | 0.37492 | 0.40356 | -0.60668 | 0.39448 | 0.2546 | 0.69309 | 0.72438 | 0.19867 | 0.90738 | 0.97012 | -0.28512 | -3.7382 |  | a2 | 17.3359 | -0.20026 | -0.29887 | 0.71742 | -0.29802 | -0.87192 | -0.39974 | -0.0081708 | -0.88301 | -0.33421 | -0.74463 | -0.40275 | -0.82326 | -0.60623 | -0.40968 | -0.02223 | -0.10091 | 0.094201 | -1.5294 | 0.25069
a3 | -2.283 | 0.72575 | 0.80875 | 0.71542 | 0.99752 | 0.66228 | 0.57478 | 0.90735 | 0.86956 | 0.94513 | 0.65777 | 0.783 | 0.52538 | 0.71634 | 0.5104 | 0.26506 | 0.38903 | 0.84836 | 1.1533 | 0.099661 |  | a3 | 1.2678 | 0.9657 | 0.6707 | 1.0341 | 1.2438 | 0.90783 | 0.65711 | 1.156 | 1.0804 | 0.83355 | 1.2458 | 0.82076 | 0.4176 | 0.74042 | 0.48249 | 0.40478 | 0.15872 | 0.67488 | 0.76198 | 0.3423
a4 | 3.0266 | -0.18524 | -0.007445 | 0.50864 | 0.2734 | 0.52958 | 0.55044 | -0.18253 | -0.21377 | 0.038346 | 0.065653 | -0.21943 | -0.18236 | -0.50198 | -0.10248 | 0.37185 | 0.24738 | -0.13165 | -0.20717 | 0.44589 |  | a4 | 4.3737 | -0.10756 | -0.075668 | -0.033405 | 0.26487 | 0.22087 | -0.23983 | -0.2477 | 0.032135 | -0.46021 | 0.47512 | -0.088806 | -0.054693 | 0.084974 | 0.078235 | -0.024686 | 0.23674 | 0.1783 | 0.31521 | 0.70661
a5 | 2.8158 | 0.06538 | 0.068351 | 0.068045 | 0.41076 | 0.19007 | 0.22625 | 0.24342 | 0.22562 | 0.33035 | 0.24851 | 0.12401 | 0.042993 | 0.2046 | 0.22756 | 0.13554 | 0.1098 | 0.27151 | 0.51744 | -0.092145 |  | a5 | 2.8839 | 0.1181 | -0.024157 | 0.14672 | 0.28263 | 0.088626 | 0.01532 | 0.25581 | 0.27168 | 0.47676 | 0.50862 | 0.13774 | 0.1908 | 0.29205 | 0.10955 | 0.18536 | 0.069994 | 0.29805 | 0.42444 | -0.067398
a6 | 0.56658 | 0.074936 | 0.10383 | 0.2076 | 0.21912 | -0.049329 | -0.052774 | -0.083304 | -0.11093 | -0.11809 | 0.012177 | -0.04902 | -0.085635 | -0.10526 | -0.009036 | 0.077872 | 0.11497 | 0.053486 | 0.14616 | -0.1002 |  | a6 | 0.60415 | 0.13372 | 0.12634 | 0.073737 | 0.1674 | 0.064717 | -0.016493 | -0.0091423 | -0.053603 | -0.010369 | 0.022797 | 0.021999 | 0.00067353 | 0.090585 | 0.0030306 | -0.054708 | 0.0079889 | 0.030343 | 0.057887 | 0.027261
a7 | 1.7084 | -0.12674 | -0.030068 | -0.0058989 | 0.071423 | 0.019008 | 0.068812 | 0.10208 | 0.19042 | 0.031693 | 0.026977 | 0.017906 | 0.092447 | 0.06208 | 0.0072112 | -0.011264 | 0.01682 | 0.059789 | 0.084797 | -0.015033 |  | a7 | 1.8165 | -0.026021 | 0.017037 | -0.00016782 | 0.041287 | -0.016891 | -0.0056705 | 0.12646 | 0.092198 | 0.029237 | 0.091485 | 0.040113 | 0.055008 | 0.02227 | 0.050023 | -0.026869 | -0.038368 | 0.059819 | 0.068604 | 0.030394
a8 | 0.034362 | 0.014098 | 0.014617 | 0.046537 | 0.023807 | 0.032133 | 0.040085 | 0.02468 | -0.028015 | 0.010227 | -0.0090819 | 0.0082441 | -0.02788 | 0.0052744 | 0.021328 | 0.029531 | 0.072198 | 0.010534 | 0.040914 | -0.075798 |  | a8 | -0.019683 | -0.020504 | 0.027322 | 0.061587 | 0.076279 | 0.054934 | 0.0073342 | 0.0063882 | 0.0048768 | -0.010274 | 0.018445 | 0.023928 | 0.027674 | 0.054203 | 0.043317 | 0.028334 | 0.015177 | 0.014391 | 0.037174 | 0.017953
a9 | -0.20933 | -0.019758 | -0.028738 | -0.020412 | -0.004879 | -0.030641 | -0.019875 | 0.015075 | 0.0020922 | 0.0043039 | -0.0083784 | -0.0065322 | -3.7229e-05 | -0.018475 | -0.033941 | -0.053496 | -0.024827 | 0.0061289 | 0.019246 | -0.06103 |  | a9 | -0.24936 | 0.02896 | 0.0063714 | 0.032671 | 0.020192 | 0.0067075 | 0.013825 | 0.041395 | 0.02505 | 0.024098 | 0.043897 | -0.00011434 | 0.021146 | 0.0087241 | 0.0098341 | 0.02638 | 0.012154 | 0.014187 | 0.034059 | -0.022956
a10 | 0.026551 | -0.0087519 | -2.7681e-05 | 0.027158 | 0.038424 | 0.00079419 | 0.0063927 | 0.013857 | -0.013083 | 7.5462e-05 | -0.0018377 | -0.0024279 | -0.031216 | -0.014924 | 0.046041 | -0.0037952 | 0.026039 | -0.020741 | 0.027106 | -0.16871 |  | a10 | -0.046832 | -0.0078389 | 0.028418 | 0.023135 | 0.034558 | 0.022907 | -0.010591 | 0.021466 | -0.013363 | -0.022786 | -0.0028741 | 0.012168 | -0.011072 | 0.039237 | 0.013747 | -0.01052 | -0.0075049 | -0.029468 | -0.012082 | -0.065517
a11 | 0.17964 | -0.064434 | -0.04779 | -0.054116 | -0.048071 | -0.043639 | -0.033762 | -0.024293 | -0.018343 | -0.035755 | -0.034999 | -0.046186 | -0.03333 | -0.036205 | -0.039966 | -0.044082 | -0.040192 | -0.036318 | -0.032287 | -0.045493 |  | a11 | 0.19543 | -0.021927 | -0.010727 | -0.0094255 | -0.012046 | -0.0066715 | -0.0044922 | 0.0077183 | 0.0092731 | -0.0037538 | 0.0057997 | -0.0028883 | -0.0066268 | 0.0029302 | -0.0025175 | -0.0056248 | -0.0063784 | -0.0011832 | 0.0026926 | -0.0060676
a12 | -0.0082687 | -0.0047194 | -0.0025175 | -0.00030803 | 0.0026491 | -0.010431 | -0.0085873 | -0.0051082 | -0.0063806 | -0.0082175 | -0.00090533 | -0.0082781 | -0.011934 | -0.013122 | -0.0026748 | -0.0076894 | 0.0019647 | -0.00288 | 0.0011991 | -0.016069 |  | a12 | -0.02009 | -0.0017285 | -0.0003896 | -0.00085488 | -0.00044291 | -0.0067415 | -0.010547 | -0.0049602 | -0.0072792 | -0.0076465 | -0.0062574 | -0.0053537 | -0.0071318 | -0.0010604 | -0.0068837 | -0.0056645 | -0.0075284 | -0.0094212 | -0.0061014 | -0.013919
a13 | -0.017117 | -0.0037513 | -0.0033094 | -0.0032449 | -0.0022627 | -0.0045506 | -0.0028516 | 0.0020439 | 0.0017874 | -0.00021527 | -0.0028477 | -0.0022011 | -0.00045334 | -0.00034156 | -0.0047787 | -0.0062363 | -0.0040313 | -0.0013614 | 0.0021461 | -0.0027552 |  | a13 | -0.016671 | 0.0024558 | 0.00053246 | 0.0019198 | 0.0038708 | 0.001316 | 0.0024962 | 0.0053223 | 0.0041352 | 0.00452 | 0.0077559 | 0.0041864 | 0.0047764 | 0.0051001 | 0.0018785 | 0.0033497 | 0.0016 | 0.0039746 | 0.0065091 | -0.00086736
a14 | 0.0048304 | -0.003285 | -0.0019082 | -0.00063352 | -0.0021657 | 9.8697e-05 | 0.0010787 | -0.00061744 | -0.0024584 | 0.00019004 | -0.00090385 | -0.00019072 | -0.0037553 | -0.0031665 | -0.00027795 | -0.00075656 | 0.0017017 | -0.0028164 | -0.0020927 | 0.0018251 |  | a14 | 0.0030912 | -0.0013068 | 0.00068925 | 0.0016117 | 0.0034897 | 0.003456 | 0.0011393 | 0.0027419 | 0.0029271 | -0.00044233 | 0.00209 | 0.0027069 | 0.0031302 | 0.0033183 | 0.0041521 | 0.0031358 | 0.0025223 | 0.0012503 | 0.001775 | 0.0031733
a15 | 0.0077903 | -0.0023421 | -0.00072546 | -0.00063018 | -1.576e-05 | -0.00030521 | -0.00091196 | 0.0016907 | 0.001591 | -0.00081492 | -0.0015992 | -0.00042243 | 0.00018911 | -0.00032748 | -0.00066648 | -0.0010697 | -0.0011044 | -0.00074375 | -6.8545e-05 | 0.0017554 |  | a15 | 0.0041922 | 0.0013564 | 0.0033635 | 0.0038874 | 0.0033842 | 0.0036741 | 0.0025992 | 0.0036808 | 0.0020502 | 0.0026444 | 0.0038335 | 0.0049797 | 0.0034822 | 0.0060714 | 0.0034347 | 0.0013594 | 0.0027152 | 0.0038568 | 0.0025471 | -8.9882e-07
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.26883 | 0.060091 | -0.15442 | -0.09579 | 0.027131 | -0.070116 | -0.065002 | 0.14368 | 0.19703 | 0.041123 | -0.10618 | -0.051765 | 0.086058 | -0.14049 | 0.032134 | -0.0025803 | -0.12597 | 0.062595 | -0.0043447 | 0.076666 |  | Dx (mm) | 0.16757 | 0.023428 | -0.12533 | -0.060346 | 0.079405 | -0.028519 | -0.024575 | 0.052611 | 0.13073 | 0.022223 | -0.019531 | -0.086403 | 0.029833 | -0.15522 | 0.032648 | 0.18616 | -0.12378 | 0.029285 | 0.047238 | -0.20938
Dy (mm) | -0.0051112 | -0.014838 | -0.010031 | 0.054714 | 0.088854 | -0.0022583 | 0.011513 | 0.035701 | -0.024602 | 0.0021022 | -0.0089338 | -0.0086551 | -0.066489 | -0.040003 | 0.10497 | -0.0088187 | 0.052824 | -0.044025 | 0.061859 | -0.43799 |  | Dy (mm) | -0.18729 | -0.017264 | 0.063528 | 0.051281 | 0.079961 | 0.051049 | -0.02405 | 0.04794 | -0.03197 | -0.051355 | -0.0063077 | 0.027217 | -0.02472 | 0.087672 | 0.030828 | -0.022379 | -0.018162 | -0.06735 | -0.027928 | -0.17563
Mag. len. (m) | 14.4208 | [20 empty] | Mag. len. (m) | 14.4211
TF (mT/KA) | 596.4884 | [20 empty] | TF (mT/KA) | 596.4748
Coil pos. (mm) | 0.018731 | [20 empty] | Coil pos. (mm) | 0.017903
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.010413 | [20 empty] | Twist (m2 rad) | 0.019241

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 596.4884 | 0.11258 | -0.0044155 |  | -0.46534 | 3.7752 |  |  | 596.4748 | 0.12675 | -0.004018 |  | -0.95107 | 3.8937 |  | 0 | 4
2 | -0.13955 |  | -0.26323 | 0.49985 | 0.51818 |  | 0.33754 | 0.45225 | 0.13286 |  | 0.11066 | 0.43205 | 0.18085 |  | -0.39178 | 0.48795 |  | 0 | 5
3 | -0.80296 |  | -2.2241 | 0.70761 | 0.61157 |  | 0.72664 | 0.21988 | -0.7174 |  | -2.1434 | 0.96904 | 0.78836 |  | 0.79243 | 0.30295 |  | 0 | 6
4 | 0.030502 |  | 0.019827 | 0.12044 | 0.12205 |  | 0.036185 | 0.3096 | 0.0156 |  | 0.035754 | 0.13899 | 0.16495 |  | 0.029837 | 0.22854 |  | 0 | 7
5 | -0.32268 |  | -0.07587 | 0.3195 | 0.27549 |  | 0.2071 | 0.12465 | -0.5021 |  | -0.24161 | 0.28641 | 0.28329 |  | 0.21467 | 0.15061 |  | 0 | 8
6 | -0.0084644 |  | -0.0091847 | 0.052186 | 0.027484 |  | 0.014192 | 0.10172 | 0.0032243 |  | 0.014282 | 0.048941 | 0.053713 |  | 0.032652 | 0.054528 |  | 0 | 9
7 | 1.1949 |  | 1.1369 | 0.047664 | 0.086943 |  | 0.038272 | 0.066074 | 1.1734 |  | 1.1288 | 0.053358 | 0.085055 |  | 0.032555 | 0.045814 |  | 0 | 10
8 | -0.0065836 |  | -0.0060418 | 0.02385 | 0.016881 |  | 0.015463 | 0.019647 | -0.010816 |  | -0.0099105 | 0.024205 | 0.025224 |  | 0.024054 | 0.019544 |  | 0 | 11
9 | 0.49625 |  | 0.49887 | 0.014064 | -0.019417 |  | -0.012136 | 0.018852 | 0.47203 |  | 0.47649 | 0.016601 | 0.011752 |  | 0.021306 | 0.012112 |  | 0 | 12
10 | 0 |  | 0 | 0 | 9.1754e-21 |  | 9.7879e-21 | 4.1527e-20 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.74457 |  | 0.75197 | 0.0076424 | -0.033135 |  | -0.0396 | 0.01075 | 0.74294 |  | 0.75089 | 0.0072061 | 0.002281 |  | -0.0035185 | 0.0077328 |  | 0 | 14
12 | -0.0002782 |  | -0.0028121 | 0.0033167 | -0.00052215 |  | -0.0054369 | 0.0036506 | -0.00018621 |  | -0.0020198 | 0.0027457 | -0.00057782 |  | -0.0057348 | 0.0020881 |  | 0 | 15
13 | 0.0064569 |  | 0.064085 | 0.0037999 | -0.00024798 |  | -0.0019659 | 0.002461 | 0.0068447 |  | 0.068113 | 0.0028711 | 0.00029388 |  | 0.0036906 | 0.0018655 |  | 0 | 16
14 | -0.00023965 |  | -0.0025673 | 0.00053903 | -0.00010763 |  | -0.0014502 | 0.0013867 | -0.00039562 |  | -0.0042251 | 0.00048296 | 0.00020415 |  | 0.0020602 | 0.0012722 |  | 0 | 17
15 | 0.0020496 |  | 0.022002 | 0.001364 | -1.3948e-05 |  | -0.00045982 | 0.00096489 | 0.0019418 |  | 0.021172 | 0.0011118 | 0.00031937 |  | 0.0032733 | 0.0011558 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.74349 | -0.77838
3 | [12 empty] | 3 | 9 | -0.93026 | -0.61282
4 | [12 empty] | 4 | 16 | -1.6445 | -1.7247
5 | [12 empty] | 5 | 25 | -1.6116 | -1.5717
6 | [12 empty] | 6 | 36 | -2.6193 | -2.9631
7 | [12 empty] | 7 | 49 | -2.8803 | -3.0069
8 | [12 empty] | 8 | 64 | -3.8329 | -3.8281
9 | [12 empty] | 9 | 81 | -4.1176 | -4.2559
10 | [12 empty] | sigma (mm) |  | 0.018731 | 0.017903
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 345.5244 | 596.7513 | 596.3914 | 596.4447 | 596.3869 | 596.4915 | 596.4962 | 596.4456 | 596.4532 | 596.4525 | 596.4261 | 596.4947 | 596.4831 | 596.4228 | 596.4295 | 596.4466 | 596.5199 | 596.4442 | 596.8107 | 369.3524
2 | 2.7186 | -0.1571 | 0.23266 | 0.44072 | 0.6433 | 0.07502 | -0.29619 | -0.36984 | -1.163 | -0.65658 | -1.1333 | -0.18625 | -0.4866 | 0.22984 | -0.41013 | -0.81588 | -0.58876 | -0.040949 | -0.055058 | 0.78188 | -0.13955
3 | 43.7953 | -2.0659 | -1.295 | -2.1975 | -2.6211 | -2.7605 | -1.5961 | -1.6737 | -1.3587 | -2.7581 | -2.9987 | -1.9152 | -2.0758 | -3.978 | -3.2174 | -1.7997 | -1.7268 | -2.2583 | -1.7369 | -1.2142 | -0.80296
4 | 0.10731 | -0.16753 | 0.089776 | 0.24823 | 0.052959 | 0.078984 | 0.06888 | 0.18159 | -0.040873 | 0.0029869 | -0.22371 | 0.021326 | -0.056766 | 0.069518 | 0.081057 | -0.14106 | -0.083953 | 0.084847 | 0.090635 | 0.26908 | 0.030502
5 | -4.4611 | -0.76187 | -0.72599 | -0.20069 | 0.0023209 | 0.2629 | 0.41291 | 0.20122 | -0.053839 | 0.13053 | 0.19378 | -0.04879 | 0.22478 | 0.016468 | -0.031541 | -0.10304 | -0.1593 | -0.37409 | -0.35141 | -3.6251 | -0.32268
6 | -0.12411 | -0.024264 | 0.00905 | -0.0313 | -0.02505 | 0.0096364 | -0.024549 | -0.089558 | -0.082784 | -0.037165 | 0.052075 | -0.023747 | 0.019566 | -0.12367 | 0.019137 | 0.047641 | 0.05617 | 0.028338 | 0.055145 | 0.1206 | -0.0084644
7 | 3.1841 | 1.063 | 1.1247 | 1.1627 | 1.1889 | 1.1108 | 1.0917 | 1.0946 | 1.1038 | 1.0486 | 1.0863 | 1.13 | 1.1775 | 1.1888 | 1.1648 | 1.1921 | 1.1733 | 1.1902 | 1.1727 | 1.0194 | 1.1949
8 | -0.033878 | 0.026359 | 0.0011887 | -0.019596 | 0.0089881 | 0.0054833 | 0.020438 | -0.032565 | -0.047181 | -0.052832 | -0.029588 | -0.0036128 | -0.0036841 | 0.023677 | -0.028043 | -0.0043299 | 0.00010637 | 0.016598 | 0.0098415 | 0.0031577 | -0.0065836
9 | 0.43972 | 0.50818 | 0.48587 | 0.49936 | 0.51844 | 0.49153 | 0.50984 | 0.51584 | 0.5221 | 0.50684 | 0.48489 | 0.48276 | 0.48012 | 0.50661 | 0.47761 | 0.504 | 0.48383 | 0.50723 | 0.49468 | 0.4728 | 0.49625
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.6141 | 0.74154 | 0.73961 | 0.74847 | 0.74999 | 0.75677 | 0.75305 | 0.75773 | 0.75731 | 0.76058 | 0.76517 | 0.75289 | 0.75529 | 0.761 | 0.75779 | 0.74456 | 0.74174 | 0.7492 | 0.74268 | 0.65156 | 0.74457
12 | 0.0080485 | 0.043773 | -0.0023782 | -0.051849 | -0.023435 | -0.036987 | -0.021977 | -0.085188 | -0.091428 | -0.057916 | -0.025319 | -0.053723 | 0.0021572 | -0.019059 | -0.049467 | -0.01416 | -0.0033612 | 0.00378 | -0.019634 | -0.05269 | -0.002782 |  | ten times!
13 | 0.92383 | 0.70032 | 0.71553 | 0.64334 | 0.63163 | 0.59622 | 0.57615 | 0.6018 | 0.60948 | 0.6052 | 0.63077 | 0.65333 | 0.65209 | 0.61045 | 0.65285 | 0.6412 | 0.6453 | 0.68934 | 0.68036 | 0.52609 | 0.064569 |  | ten times!
14 | 0.015847 | -0.021404 | -0.022933 | -0.026796 | -0.040566 | -0.030293 | -0.027591 | -0.030674 | -0.027116 | -0.024713 | -0.024643 | -0.027272 | -0.022102 | -0.022826 | -0.024755 | -0.029669 | -0.014668 | -0.023361 | -0.020736 | -0.011568 | -0.0023965 |  | ten times!
15 | -0.16942 | 0.21918 | 0.19169 | 0.2026 | 0.22868 | 0.2249 | 0.22979 | 0.2341 | 0.23437 | 0.2494 | 0.22227 | 0.20716 | 0.22147 | 0.20474 | 0.21803 | 0.21794 | 0.20743 | 0.22333 | 0.22329 | 0.11736 | 0.020496 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 30.6234 | -7.3658 | -1.6842 | -3.3914 | -2.3792 | -0.10396 | 1.8654 | 5.5814 | 5.4954 | -0.3017 | -3.0804 | 0.2223 | 3.9209 | 6.1399 | -1.5806 | -3.1493 | -5.0389 | -2.1888 | -1.3372 | -15.2513 | -0.0044155
2 | 10.7481 | -0.063551 | 0.3404 | 0.15502 | 0.46234 | -0.28408 | 0.34734 | 1.0896 | 0.35083 | 0.39967 | -0.60161 | 0.3973 | 0.23303 | 0.68619 | 0.76215 | 0.19688 | 0.92388 | 0.95217 | -0.27191 | -3.7997 | 0.51818
3 | -2.4267 | 0.72725 | 0.80875 | 0.72158 | 0.99551 | 0.66887 | 0.58096 | 0.91081 | 0.87681 | 0.94485 | 0.65864 | 0.78103 | 0.52754 | 0.70439 | 0.50959 | 0.26501 | 0.39528 | 0.85045 | 1.1522 | 0.10862 | 0.61157
4 | 2.8431 | -0.18882 | -0.0066687 | 0.51277 | 0.27085 | 0.53286 | 0.55278 | -0.19244 | -0.22455 | 0.035085 | 0.072268 | -0.21802 | -0.17971 | -0.495 | -0.10344 | 0.37171 | 0.25264 | -0.1395 | -0.20148 | 0.072972 | 0.12205
5 | 2.7712 | 0.064045 | 0.072912 | 0.075028 | 0.40942 | 0.18904 | 0.22541 | 0.24731 | 0.23202 | 0.3318 | 0.24892 | 0.12315 | 0.046285 | 0.19811 | 0.22665 | 0.13572 | 0.114 | 0.27124 | 0.51669 | -0.071143 | 0.27549
6 | 0.41023 | 0.083196 | 0.10618 | 0.18494 | 0.18117 | -0.047973 | -0.05563 | -0.10227 | -0.11459 | -0.11933 | 0.016621 | -0.04524 | -0.060817 | -0.085386 | -0.052226 | 0.081572 | 0.093845 | 0.070675 | 0.12071 | 0.055714 | 0.027484
7 | 1.7047 | -0.12684 | -0.029281 | -0.0031252 | 0.070579 | 0.019925 | 0.069861 | 0.10059 | 0.19271 | 0.031557 | 0.026381 | 0.018026 | 0.093867 | 0.062213 | 0.0077579 | -0.011251 | 0.021177 | 0.060092 | 0.084659 | -0.011513 | 0.086943
8 | 0.061935 | 0.01821 | 0.014848 | 0.032727 | 0.0022581 | 0.031646 | 0.03671 | 0.014916 | -0.022061 | 0.0096424 | -0.0074507 | 0.010054 | -0.012838 | 0.013673 | -0.0016381 | 0.031558 | 0.058627 | 0.02087 | 0.026579 | 0.015936 | 0.016881
9 | -0.21284 | -0.019278 | -0.029099 | -0.017862 | -0.006112 | -0.030649 | -0.019481 | 0.012801 | 0.0045964 | 0.0042817 | -0.0087088 | -0.0067049 | 0.0026844 | -0.020936 | -0.035658 | -0.053508 | -0.021587 | 0.0074373 | 0.019332 | -0.050137 | -0.019417
10 | 0 | 0 | 1.7618e-19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9.1754e-21
11 | 0.18114 | -0.06418 | -0.048094 | -0.053801 | -0.04806 | -0.044122 | -0.034087 | -0.023761 | -0.017539 | -0.035531 | -0.035103 | -0.046507 | -0.032488 | -0.037605 | -0.039699 | -0.044104 | -0.039828 | -0.036106 | -0.032193 | -0.045614 | -0.033135
12 | -0.046887 | -0.038277 | -0.023677 | -0.030252 | -0.012738 | -0.10561 | -0.091879 | -0.068315 | -0.055687 | -0.08301 | -0.0071876 | -0.079585 | -0.088569 | -0.1142 | -0.074018 | -0.073016 | -0.0081021 | -0.0068237 | -0.017697 | 0.0053823 | -0.0052215 |  | ten times!
13 | -0.18118 | -0.036127 | -0.035711 | -0.031135 | -0.019769 | -0.045528 | -0.02763 | 0.021212 | 0.021789 | -0.0021864 | -0.029512 | -0.022326 | -0.0023826 | -0.0082437 | -0.046241 | -0.062581 | -0.037232 | -0.012674 | 0.022409 | -0.027894 | -0.0024798 |  | ten times!
14 | 0.030345 | -0.029012 | -0.018421 | -0.015961 | -0.038386 | 0.001229 | 0.0081204 | -0.015058 | -0.022417 | 0.0017447 | -0.0088016 | -0.00061066 | -0.02556 | -0.025299 | -0.02145 | -0.0060055 | 0.0068475 | -0.019683 | -0.032304 | 0.059476 | -0.0010763 |  | ten times!
15 | 0.077903 | -0.023421 | -0.0072546 | -0.0063018 | -0.0001576 | -0.0030521 | -0.0091196 | 0.016907 | 0.01591 | -0.0081492 | -0.015992 | -0.0042243 | 0.0018911 | -0.0032748 | -0.0066648 | -0.010697 | -0.011044 | -0.0074375 | -0.00068545 | 0.017554 | -0.00013948 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 338.0858 | 596.8003 | 596.4021 | 596.5393 | 596.3965 | 596.4581 | 596.4558 | 596.4463 | 596.4315 | 596.3949 | 596.3967 | 596.4516 | 596.4201 | 596.3788 | 596.3947 | 596.4432 | 596.4775 | 596.4449 | 596.8135 | 373.012
2 | 3.3315 | 0.15466 | 0.61377 | 0.42744 | 0.010823 | 0.16839 | -0.03425 | 0.22034 | 0.74322 | 0.55568 | 0.42915 | 0.52208 | -0.75907 | -0.62465 | -0.34541 | -0.31918 | -0.26573 | 0.1173 | 0.3774 | -2.1277 | 0.13286
3 | 44.5686 | -1.6702 | -1.2642 | -2.7852 | -1.7943 | -3.7891 | -1.5733 | -1.5146 | -0.84319 | -2.7888 | -2.6104 | -1.9893 | -1.8499 | -4.51 | -3.3738 | -1.7145 | -1.3381 | -1.5688 | -1.6031 | -0.71995 | -0.7174
4 | -0.31335 | 0.032279 | 0.12522 | 0.13293 | 0.13793 | -0.21418 | -0.1839 | -0.027226 | 0.076893 | 0.21681 | 0.27112 | 0.13248 | 0.030613 | -0.21389 | -0.015975 | 0.010453 | -0.070491 | 0.12439 | 0.078114 | -0.26638 | 0.0156
5 | -5.3646 | -0.61055 | -0.90851 | -0.084257 | -0.16956 | 0.037141 | -0.21693 | -0.059616 | -0.23127 | 0.043543 | 0.25403 | -0.03087 | -0.1303 | -0.10197 | -0.19905 | -0.34691 | -0.47894 | -0.56298 | -0.55207 | -3.5918 | -0.5021
6 | -0.27889 | -0.07777 | 0.050546 | 0.003063 | -0.09449 | 0.029588 | 0.050564 | -0.033056 | -0.001791 | -0.012746 | 0.049266 | -0.0045355 | 0.034519 | 0.061025 | 0.01507 | 8.7792e-05 | 0.10463 | 0.050716 | 0.032391 | -0.05929 | 0.0032243
7 | 2.8599 | 0.97594 | 1.0375 | 1.2013 | 1.1734 | 1.1547 | 1.1671 | 1.1227 | 1.1478 | 1.1163 | 1.1672 | 1.1376 | 1.115 | 1.1715 | 1.1353 | 1.1696 | 1.0974 | 1.1186 | 1.109 | 0.92901 | 1.1734
8 | -0.073124 | 0.019139 | 0.0058022 | -0.024751 | -0.033885 | 0.013402 | -0.0076183 | 0.018582 | -0.01107 | -0.041727 | -0.063628 | -0.044504 | -0.0049768 | -0.015322 | -0.018409 | 0.018228 | 0.011913 | 0.0032702 | -0.0028351 | 0.019546 | -0.010816
9 | 0.38289 | 0.4776 | 0.4674 | 0.47534 | 0.50739 | 0.44912 | 0.48921 | 0.47687 | 0.51926 | 0.47254 | 0.46368 | 0.46111 | 0.47199 | 0.49103 | 0.4643 | 0.47328 | 0.46692 | 0.47569 | 0.47414 | 0.42451 | 0.47203
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.60152 | 0.74215 | 0.73874 | 0.75543 | 0.74676 | 0.75879 | 0.75342 | 0.75254 | 0.74813 | 0.75293 | 0.75665 | 0.75045 | 0.75345 | 0.76501 | 0.7607 | 0.75192 | 0.74527 | 0.74374 | 0.74001 | 0.64212 | 0.74294
12 | -0.0099447 | 0.0029823 | -0.0012599 | -0.069906 | -0.05634 | -0.034846 | 0.032383 | -0.014542 | 0.0031234 | -0.032012 | 0.00076062 | -0.018178 | 0.012671 | -0.017666 | -0.062733 | -0.046737 | -0.03195 | -0.024773 | -0.0045352 | 0.018859 | -0.0018621 |  | ten times!
13 | 0.93829 | 0.71847 | 0.73259 | 0.65277 | 0.67858 | 0.66788 | 0.70489 | 0.64765 | 0.63658 | 0.65653 | 0.65818 | 0.67638 | 0.68135 | 0.6485 | 0.67198 | 0.70207 | 0.69682 | 0.72135 | 0.70777 | 0.55049 | 0.068447 |  | ten times!
14 | 0.014706 | -0.046984 | -0.043249 | -0.03839 | -0.044489 | -0.038975 | -0.034132 | -0.033592 | -0.043633 | -0.035917 | -0.042188 | -0.040682 | -0.041934 | -0.048506 | -0.047482 | -0.0493 | -0.039033 | -0.047035 | -0.044998 | -0.011337 | -0.0039562 |  | ten times!
15 | -0.16569 | 0.22665 | 0.19956 | 0.19632 | 0.2224 | 0.20045 | 0.20341 | 0.21968 | 0.22342 | 0.22752 | 0.20357 | 0.20499 | 0.22334 | 0.19601 | 0.2089 | 0.22046 | 0.20471 | 0.22144 | 0.20823 | 0.015392 | 0.019418 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 43.5854 | -4.155 | -0.56333 | -3.1856 | -2.0536 | -0.019438 | 1.866 | 6.5602 | 5.8508 | 0.34241 | -4.1162 | -1.8373 | 1.3677 | 4.9973 | -1.6391 | -5.8588 | -6.7487 | -4.1542 | -3.7725 | -12.245 | -0.004018
2 | 18.2959 | -0.20631 | -0.27951 | 0.74157 | -0.29274 | -0.84611 | -0.40229 | -0.0067935 | -0.90279 | -0.35326 | -0.7437 | -0.38804 | -0.83011 | -0.54624 | -0.3993 | -0.03561 | -0.10146 | 0.079467 | -1.5389 | 0.24477 | 0.18085
3 | 1.1253 | 0.96623 | 0.66732 | 1.0325 | 1.2383 | 0.91089 | 0.65529 | 1.1586 | 1.0801 | 0.83738 | 1.2477 | 0.81877 | 0.41802 | 0.74609 | 0.48214 | 0.40557 | 0.16371 | 0.67545 | 0.75972 | 0.36554 | 0.78836
4 | 4.0251 | -0.11069 | -0.062762 | -0.030297 | 0.26288 | 0.22103 | -0.24097 | -0.25019 | 0.02205 | -0.46218 | 0.47783 | -0.085814 | -0.056788 | 0.097788 | 0.078835 | -0.034612 | 0.23672 | 0.16734 | 0.30687 | 0.55524 | 0.16495
5 | 2.8546 | 0.11683 | -0.01958 | 0.14837 | 0.2802 | 0.088894 | 0.015487 | 0.25615 | 0.27423 | 0.47677 | 0.50883 | 0.13862 | 0.19114 | 0.29629 | 0.10927 | 0.18886 | 0.070586 | 0.29905 | 0.42406 | -0.072243 | 0.28329
6 | 0.68536 | 0.13989 | 0.10382 | 0.051996 | 0.13322 | 0.043764 | -0.006633 | -0.030496 | -0.044892 | 0.0096251 | 0.026027 | 0.012276 | 0.0098251 | 0.055515 | -0.0098904 | -0.043666 | 0.013353 | 0.056326 | 0.067681 | 0.087303 | 0.053713
7 | 1.8129 | -0.025648 | 0.019026 | 0.002187 | 0.03931 | -0.016398 | -0.0057262 | 0.12574 | 0.092219 | 0.028565 | 0.091457 | 0.041674 | 0.054688 | 0.027606 | 0.049585 | -0.028482 | -0.03771 | 0.05993 | 0.067974 | 0.030174 | 0.085055
8 | 0.033949 | -0.016943 | 0.013646 | 0.051038 | 0.056431 | 0.044245 | 0.013031 | -0.0053971 | 0.011205 | 0.00088162 | 0.020225 | 0.017998 | 0.032869 | 0.034419 | 0.03643 | 0.031091 | 0.019905 | 0.029245 | 0.042653 | 0.051002 | 0.025224
9 | -0.2372 | 0.029149 | 0.010076 | 0.034131 | 0.017239 | 0.0073921 | 0.013548 | 0.040212 | 0.026952 | 0.02463 | 0.043839 | 0.0009915 | 0.021493 | 0.015201 | 0.0093573 | 0.028392 | 0.011139 | 0.015099 | 0.034666 | -0.037622 | 0.011752
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.19857 | -0.021884 | -0.010484 | -0.0091321 | -0.011916 | -0.0066349 | -0.0046279 | 0.0078634 | 0.010018 | -0.0037106 | 0.0057174 | -0.0030812 | -0.0064446 | 0.0033493 | -0.002277 | -0.0048761 | -0.0070908 | -0.0010394 | 0.0029158 | -0.0086867 | 0.002281
12 | -0.056575 | -0.0089475 | -0.036544 | -0.031421 | -0.044669 | -0.091203 | -0.09305 | -0.073427 | -0.062269 | -0.0534 | -0.058571 | -0.064039 | -0.060439 | -0.045584 | -0.083848 | -0.049983 | -0.064876 | -0.060751 | -0.04925 | -0.071417 | -0.0057782 |  | ten times!
13 | -0.17091 | 0.024232 | 0.0091723 | 0.021894 | 0.038556 | 0.015679 | 0.024483 | 0.053056 | 0.03796 | 0.044072 | 0.077653 | 0.044826 | 0.046374 | 0.060165 | 0.018756 | 0.028772 | 0.017561 | 0.037429 | 0.063675 | -0.0054718 | 0.0029388 |  | ten times!
14 | -0.00042954 | -0.010107 | -7.6264e-05 | 0.0097586 | 0.018039 | 0.026996 | 0.015947 | 0.017151 | 0.032946 | 0.0047151 | 0.022574 | 0.026017 | 0.034994 | 0.026791 | 0.035294 | 0.033337 | 0.031053 | 0.023855 | 0.021548 | 0.033958 | 0.0020415 |  | ten times!
15 | 0.041922 | 0.013564 | 0.033635 | 0.038874 | 0.033842 | 0.036741 | 0.025992 | 0.036808 | 0.020502 | 0.026444 | 0.038335 | 0.049797 | 0.034822 | 0.060714 | 0.034347 | 0.013594 | 0.027152 | 0.038568 | 0.025471 | -8.9882e-06 | 0.0031937 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00044077 | -0.00016259 | -7.3234e-05 | -0.00017013 | 5.225e-06 | 1.3104e-05 | -7.1725e-05 | -5.8984e-05 | -6.0158e-05 | -0.00010442 | 1.059e-05 | -8.8574e-06 | -0.00010995 | -9.8717e-05 | -7.0049e-05 | 5.2837e-05 | -7.4072e-05 | 0.00054036 |  | 596.4884
|  | 0.00054576 | -0.00012183 | 0.00010819 | -0.00013122 | -2.7942e-05 | -3.1798e-05 | -4.7725e-05 | -7.2537e-05 | -0.0001339 | -0.00013088 | -3.8839e-05 | -9.165e-05 | -0.00016089 | -0.00013423 | -5.2922e-05 | 4.5825e-06 | -5.0072e-05 | 0.00056789 |  | 596.4748
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -4007.2543
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 596.4518 | -4207.0022 | 5.0213 | -1.0128 | -0.11914 | -1.0882 | 0.6655 | 0.7443 | -0.10405 | 0.023367 | 0.011631 | -0.43099 | 0.71915 | 0.52467 | -0.48631 | -0.37398 | -0.087287 | 1.1416 | -0.12752 | 6.0171 | -3807.5064
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.046534 | 3.1089 | -0.69005 | -0.12188 | -0.29261 | -0.19138 | 0.036138 | 0.23307 | 0.60467 | 0.59607 | 0.016364 | -0.26151 | 0.068764 | 0.43862 | 0.66052 | -0.11152 | -0.2684 | -0.45735 | -0.17234 | -0.087182 | -1.4786 |  | 0
db2 | -0.26323 | 2.7186 | 0.10613 | 0.49589 | 0.70395 | 0.90653 | 0.33825 | -0.032958 | -0.10661 | -0.89981 | -0.39335 | -0.87006 | 0.076979 | -0.22337 | 0.49307 | -0.14691 | -0.55265 | -0.32554 | 0.22228 | 0.20817 | 0.78188 |  | 0
db3 | -2.2251 | 43.7953 | 0.15815 | 0.92913 | 0.026561 | -0.39698 | -0.53645 | 0.62796 | 0.55039 | 0.86537 | -0.53404 | -0.77463 | 0.30893 | 0.14825 | -1.7539 | -0.99336 | 0.42436 | 0.4973 | -0.034212 | 0.48722 | -1.2142 |  | 0
db4 | 0.019827 | 0.10731 | -0.18736 | 0.069949 | 0.22841 | 0.033132 | 0.059157 | 0.049053 | 0.16176 | -0.060701 | -0.01684 | -0.24354 | 0.0014992 | -0.076593 | 0.049691 | 0.06123 | -0.16089 | -0.10378 | 0.065019 | 0.070808 | 0.26908 |  | 0
db5 | -0.0759 | -4.4611 | -0.686 | -0.65012 | -0.12482 | 0.078191 | 0.33877 | 0.48878 | 0.27709 | 0.022031 | 0.2064 | 0.26965 | 0.027079 | 0.30065 | 0.092338 | 0.044328 | -0.027173 | -0.083431 | -0.29822 | -0.27554 | -3.6251 |  | 0
db6 | -0.0091847 | -0.12411 | -0.015079 | 0.018235 | -0.022116 | -0.015866 | 0.018821 | -0.015365 | -0.080373 | -0.073599 | -0.02798 | 0.06126 | -0.014562 | 0.028751 | -0.11448 | 0.028322 | 0.056826 | 0.065355 | 0.037522 | 0.064329 | 0.1206 |  | 0
db7 | 1.1371 | 3.1841 | -0.073906 | -0.012207 | 0.025744 | 0.052016 | -0.026078 | -0.045236 | -0.042329 | -0.033121 | -0.088361 | -0.050653 | -0.0068652 | 0.040567 | 0.051872 | 0.027878 | 0.055169 | 0.036387 | 0.053318 | 0.035806 | 1.0194 |  | 9.8686e-17
db8 | -0.0060418 | -0.033878 | 0.032401 | 0.0072306 | -0.013554 | 0.01503 | 0.011525 | 0.02648 | -0.026524 | -0.04114 | -0.046791 | -0.023546 | 0.0024291 | 0.0023577 | 0.029718 | -0.022001 | 0.0017119 | 0.0061482 | 0.02264 | 0.015883 | 0.0031577 |  | 0
db9 | 0.50057 | 0.43972 | 0.0093055 | -0.013002 | 0.00048703 | 0.019565 | -0.0073459 | 0.010969 | 0.016962 | 0.023224 | 0.0079637 | -0.013985 | -0.016114 | -0.018753 | 0.0077398 | -0.021265 | 0.0051263 | -0.015041 | 0.0083566 | -0.0041928 | 0.4728 |  | 5.2427e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.75213 | 0.6141 | -0.010425 | -0.012356 | -0.0034909 | -0.0019727 | 0.004806 | 0.0010807 | 0.0057651 | 0.0053466 | 0.008617 | 0.013204 | 0.00092835 | 0.003323 | 0.0090377 | 0.0058224 | -0.0074104 | -0.010227 | -0.002768 | -0.0092814 | 0.65156 |  | 6.1679e-18
db12 | -0.0028121 | 0.00080485 | 0.0071894 | 0.0025742 | -0.0023729 | 0.00046855 | -0.00088663 | 0.00061434 | -0.0057067 | -0.0063307 | -0.0029795 | 0.00028019 | -0.0025603 | 0.0030278 | 0.0009062 | -0.0021346 | 0.0013961 | 0.0024759 | 0.00319 | 0.00084869 | -0.005269 |  | 0
db13 | 0.064085 | 0.092383 | 0.0059466 | 0.0074676 | 0.00024841 | -0.00092257 | -0.0044633 | -0.0064705 | -0.0039052 | -0.0031374 | -0.003565 | -0.001008 | 0.0012474 | 0.0011239 | -0.0030405 | 0.0011994 | 3.4623e-05 | 0.00044523 | 0.0048491 | 0.0039503 | 0.052609 |  | 1.542e-18
db14 | -0.0025673 | 0.0015847 | 0.0004269 | 0.00027405 | -0.00011229 | -0.0014893 | -0.00046198 | -0.00019178 | -0.00050005 | -0.00014429 | 9.6055e-05 | 0.00010307 | -0.00015991 | 0.00035711 | 0.00028469 | 9.1778e-05 | -0.00039959 | 0.0011005 | 0.00023123 | 0.00049376 | -0.0011568 |  | 0
db15 | 0.022002 | -0.016942 | -8.3934e-05 | -0.002833 | -0.0017419 | 0.00086544 | 0.0004881 | 0.00097686 | 0.0014081 | 0.0014351 | 0.0029381 | 0.00022478 | -0.001286 | 0.0001445 | -0.0015281 | -0.00019886 | -0.00020786 | -0.0012592 | 0.00033103 | 0.00032691 | 0.011736 |  | 2.1202e-18
da2 | 0.33754 | 10.7481 | -0.40109 | 0.0028657 | -0.18251 | 0.1248 | -0.62162 | 0.0098085 | 0.75205 | 0.01329 | 0.062138 | -0.93915 | 0.059761 | -0.10451 | 0.34865 | 0.42462 | -0.14065 | 0.58635 | 0.61464 | -0.60944 | -3.7997 |  | 0
da3 | 0.72664 | -2.4267 | 0.00060419 | 0.082107 | -0.0050608 | 0.26887 | -0.057776 | -0.14568 | 0.18417 | 0.15017 | 0.21821 | -0.068 | 0.054384 | -0.1991 | -0.022251 | -0.21705 | -0.46164 | -0.33137 | 0.12381 | 0.4256 | 0.10862 |  | 0
da4 | 0.036185 | 2.8431 | -0.22501 | -0.042854 | 0.47659 | 0.23466 | 0.49667 | 0.5166 | -0.22862 | -0.26073 | -0.0011003 | 0.036083 | -0.2542 | -0.2159 | -0.53119 | -0.13962 | 0.33553 | 0.21646 | -0.17569 | -0.23767 | 0.072972 |  | 0
da5 | 0.2071 | 2.7712 | -0.14305 | -0.13419 | -0.13207 | 0.20233 | -0.018054 | 0.018309 | 0.040213 | 0.024921 | 0.1247 | 0.041825 | -0.083947 | -0.16081 | -0.0089873 | 0.019556 | -0.071377 | -0.0931 | 0.064144 | 0.30959 | -0.071143 |  | 0
da6 | 0.014192 | 0.41023 | 0.069005 | 0.091988 | 0.17074 | 0.16698 | -0.062164 | -0.069822 | -0.11646 | -0.12878 | -0.13352 | 0.0024291 | -0.059432 | -0.075008 | -0.099578 | -0.066418 | 0.067381 | 0.079653 | 0.056483 | 0.10652 | 0.055714 |  | 0
da7 | 0.038272 | 1.7047 | -0.16512 | -0.067553 | -0.041397 | 0.032307 | -0.018346 | 0.03159 | 0.062314 | 0.15444 | -0.0067148 | -0.01189 | -0.020245 | 0.055596 | 0.023941 | -0.030514 | -0.049523 | -0.017095 | 0.02182 | 0.046388 | -0.011513 |  | 0
da8 | 0.015463 | 0.061935 | 0.002747 | -0.00061451 | 0.017264 | -0.013205 | 0.016183 | 0.021247 | -0.00054656 | -0.037524 | -0.0058204 | -0.022914 | -0.005409 | -0.028301 | -0.0017897 | -0.017101 | 0.016095 | 0.043165 | 0.0054075 | 0.011116 | 0.015936 |  | 0
da9 | -0.012136 | -0.21284 | -0.0071421 | -0.016963 | -0.0057262 | 0.0060242 | -0.018513 | -0.0073449 | 0.024937 | 0.016733 | 0.016418 | 0.0034274 | 0.0054313 | 0.014821 | -0.0087999 | -0.023522 | -0.041372 | -0.0094511 | 0.019573 | 0.031469 | -0.050137 |  | 0
da10 | 9.7879e-21 | 0 | -9.7879e-21 | 1.6639e-19 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | -9.7879e-21 | 0 |  | 0
da11 | -0.0396 | 0.18114 | -0.02458 | -0.0084934 | -0.0142 | -0.0084598 | -0.0045218 | 0.0055131 | 0.01584 | 0.022061 | 0.0040692 | 0.0044973 | -0.0069065 | 0.0071129 | 0.0019957 | -9.8405e-05 | -0.0045036 | -0.0002277 | 0.003495 | 0.0074073 | -0.045614 |  | 3.8549e-18
da12 | -0.0054369 | -0.0046887 | 0.0016092 | 0.0030692 | 0.0024117 | 0.0041631 | -0.0051238 | -0.003751 | -0.0013946 | -0.00013174 | -0.0028641 | 0.0047181 | -0.0025216 | -0.00342 | -0.0059834 | -0.0019649 | -0.0018647 | 0.0046267 | 0.0047545 | 0.0036672 | 0.00053823 |  | 0
da13 | -0.0019659 | -0.018118 | -0.0016467 | -0.0016052 | -0.0011476 | -1.0994e-05 | -0.0025868 | -0.00079704 | 0.0040872 | 0.0041449 | 0.0017473 | -0.00098532 | -0.00026664 | 0.0017277 | 0.0011416 | -0.0026581 | -0.0042922 | -0.0017572 | 0.0006985 | 0.0042068 | -0.0027894 |  | 0
da14 | -0.0014502 | 0.0030345 | -0.0014511 | -0.00039193 | -0.00014598 | -0.0023884 | 0.0015731 | 0.0022622 | -5.5617e-05 | -0.00079156 | 0.0016246 | 0.00057 | 0.0013891 | -0.0011059 | -0.0010797 | -0.00069488 | 0.00084961 | 0.0021349 | -0.00051818 | -0.0017802 | 0.0059476 |  | 0
da15 | -0.00045982 | 0.0077903 | -0.0018823 | -0.00026565 | -0.00017037 | 0.00044406 | 0.0001546 | -0.00045214 | 0.0021505 | 0.0020509 | -0.0003551 | -0.0011393 | 3.7384e-05 | 0.00064892 | 0.00013234 | -0.00020666 | -0.00060984 | -0.00064461 | -0.00028393 | 0.00039127 | 0.0017554 |  | 0
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -4038.7478
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.4333 | -4331.54 | 6.1541 | -0.52227 | 1.7781 | -0.61616 | 0.41664 | 0.37808 | 0.2188 | -0.029341 | -0.64299 | -0.61281 | 0.30766 | -0.22048 | -0.91293 | -0.64634 | 0.16683 | 0.74191 | 0.19533 | 6.3754 | -3745.9556
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.095107 | 4.4537 | -0.32039 | 0.038773 | -0.22345 | -0.11025 | 0.093163 | 0.28171 | 0.75113 | 0.68019 | 0.12935 | -0.31651 | -0.088625 | 0.23188 | 0.59484 | -0.068808 | -0.49077 | -0.57977 | -0.32031 | -0.28214 | -1.1294 |  | 0
db2 | -0.11066 | 3.3315 | 0.043999 | 0.5031 | 0.31678 | -0.099841 | 0.057721 | -0.14491 | 0.10968 | 0.63255 | 0.44502 | 0.31849 | 0.41141 | -0.86974 | -0.73531 | -0.45607 | -0.42984 | -0.3764 | 0.0066314 | 0.26674 | -2.1277 |  | 0
db3 | -2.1433 | 44.5686 | 0.47319 | 0.87914 | -0.64183 | 0.34912 | -1.6457 | 0.57004 | 0.62874 | 1.3002 | -0.64542 | -0.46706 | 0.15411 | 0.29348 | -2.3667 | -1.2304 | 0.42888 | 0.80532 | 0.5746 | 0.54023 | -0.71995 |  | 0
db4 | -0.035754 | -0.31335 | -0.0034748 | 0.089466 | 0.097173 | 0.10217 | -0.24993 | -0.21966 | -0.06298 | 0.041139 | 0.18106 | 0.23537 | 0.096727 | -0.005141 | -0.24964 | -0.051729 | -0.025301 | -0.10625 | 0.088637 | 0.04236 | -0.26638 |  | 0
db5 | -0.24192 | -5.3646 | -0.36893 | -0.66689 | 0.15736 | 0.072057 | 0.27876 | 0.024682 | 0.182 | 0.010347 | 0.28516 | 0.49564 | 0.21075 | 0.11132 | 0.13964 | 0.042565 | -0.1053 | -0.23733 | -0.32137 | -0.31045 | -3.5918 |  | 0
db6 | -0.014282 | -0.27889 | -0.092052 | 0.036263 | -0.011219 | -0.10877 | 0.015306 | 0.036282 | -0.047338 | -0.016073 | -0.027028 | 0.034984 | -0.018818 | 0.020237 | 0.046743 | 0.00078817 | -0.014194 | 0.090348 | 0.036434 | 0.018109 | -0.05929 |  | 0
db7 | 1.1292 | 2.8599 | -0.15284 | -0.091262 | 0.072573 | 0.044598 | 0.025972 | 0.038368 | -0.0060412 | 0.019013 | -0.01248 | 0.038385 | 0.0088634 | -0.01378 | 0.042688 | 0.0064892 | 0.040831 | -0.031417 | -0.010184 | -0.019781 | 0.92901 |  | -4.9343e-17
db8 | 0.0099105 | -0.073124 | 0.029049 | 0.015713 | -0.014841 | -0.023974 | 0.023312 | 0.0022922 | 0.028493 | -0.0011596 | -0.031817 | -0.053717 | -0.034593 | 0.0049337 | -0.0054118 | -0.0084986 | 0.028139 | 0.021824 | 0.013181 | 0.0070755 | 0.019546 |  | 0
db9 | 0.47774 | 0.38289 | 0.0011054 | -0.0090971 | -0.0011491 | 0.030898 | -0.027375 | 0.012714 | 0.0003773 | 0.04277 | -0.0039547 | -0.01281 | -0.015386 | -0.0045047 | 0.014541 | -0.012192 | -0.0032098 | -0.0095711 | -0.00079941 | -0.0023565 | 0.42451 |  | 3.084e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.75108 | 0.60152 | -0.0087458 | -0.012156 | 0.0045408 | -0.0041284 | 0.007895 | 0.002523 | 0.0016479 | -0.0027647 | 0.0020334 | 0.0057578 | -0.00044692 | 0.0025536 | 0.014118 | 0.0098076 | 0.0010251 | -0.0056258 | -0.0071549 | -0.01088 | 0.64212 |  | 3.7007e-17
db12 | 0.0020198 | -0.00099447 | 0.002318 | 0.0018938 | -0.0049708 | -0.0036142 | -0.0014649 | 0.0052581 | 0.00056555 | 0.0023321 | -0.0011814 | 0.0020958 | 0.00020197 | 0.0032869 | 0.00025317 | -0.0042535 | -0.0026539 | -0.0011753 | -0.00045752 | 0.0015662 | 0.0018859 |  | 0
db13 | 0.068113 | 0.093829 | 0.0037336 | 0.0051457 | -0.0028355 | -0.00025504 | -0.0013252 | 0.0023762 | -0.0033477 | -0.0044552 | -0.00246 | -0.0022954 | -0.0004749 | 2.2247e-05 | -0.0032631 | -0.00091471 | 0.0020937 | 0.0015688 | 0.0040222 | 0.0026642 | 0.055049 |  | 1.542e-18
db14 | 0.0042251 | 0.0014706 | -0.00047331 | -9.9825e-05 | 0.00038615 | -0.00022375 | 0.00032758 | 0.00081195 | 0.00086594 | -0.00013823 | 0.00063341 | 6.2748e-06 | 0.00015696 | 3.1759e-05 | -0.00062551 | -0.00052313 | -0.00070494 | 0.00032177 | -0.0004784 | -0.0002747 | -0.0011337 |  | 7.228e-20
db15 | 0.021172 | -0.016569 | 0.0014922 | -0.0012169 | -0.0015406 | 0.0010678 | -0.0011278 | -0.00083135 | 0.00079556 | 0.0011693 | 0.0015794 | -0.0008154 | -0.00067352 | 0.0011613 | -0.0015712 | -0.00028249 | 0.00087359 | -0.00070165 | 0.00097104 | -0.00034938 | 0.0015392 |  | -7.7099e-19
da2 | -0.39178 | 18.2959 | 0.18547 | 0.11227 | 1.1334 | 0.09904 | -0.45433 | -0.010509 | 0.38499 | -0.51101 | 0.038523 | -0.35192 | 0.0037469 | -0.43832 | -0.15446 | -0.0075129 | 0.35617 | 0.29033 | 0.47125 | -1.1471 | 0.24477 |  | 0
da3 | 0.79243 | 1.1253 | 0.1738 | -0.12511 | 0.24007 | 0.44586 | 0.11845 | -0.13714 | 0.36614 | 0.28766 | 0.044943 | 0.45531 | 0.026336 | -0.37441 | -0.046344 | -0.31029 | -0.38686 | -0.62872 | -0.11698 | -0.032709 | 0.36554 |  | 4.3175e-17
da4 | 0.029837 | 4.0251 | -0.14053 | -0.092599 | -0.060133 | 0.23304 | 0.1912 | -0.27081 | -0.28003 | -0.0077867 | -0.49201 | 0.44799 | -0.11565 | -0.086624 | 0.067952 | 0.048999 | -0.064448 | 0.20689 | 0.13751 | 0.27703 | 0.55524 |  | 0
da5 | 0.21467 | 2.8546 | -0.097842 | -0.23425 | -0.066297 | 0.065532 | -0.12578 | -0.19918 | 0.041484 | 0.059559 | 0.2621 | 0.29416 | -0.076051 | -0.023534 | 0.081617 | -0.1054 | -0.025809 | -0.14408 | 0.084379 | 0.20939 | -0.072243 |  | 0
da6 | 0.032652 | 0.68536 | 0.10723 | 0.071169 | 0.019344 | 0.10056 | 0.011112 | -0.039285 | -0.063148 | -0.077543 | -0.023027 | -0.0066245 | -0.020376 | -0.022827 | 0.022863 | -0.042542 | -0.076317 | -0.019299 | 0.023674 | 0.035029 | 0.087303 |  | 0
da7 | 0.032555 | 1.8129 | -0.058203 | -0.013529 | -0.030368 | 0.0067547 | -0.048953 | -0.038281 | 0.093181 | 0.059664 | -0.0039898 | 0.058902 | 0.0091191 | 0.022133 | -0.0049496 | 0.01703 | -0.061037 | -0.070265 | 0.027375 | 0.035419 | 0.030174 |  | 0
da8 | 0.024054 | 0.033949 | -0.040997 | -0.010408 | 0.026985 | 0.032377 | 0.020191 | -0.011023 | -0.029451 | -0.012849 | -0.023172 | -0.0038288 | -0.0060562 | 0.0088147 | 0.010365 | 0.012376 | 0.0070368 | -0.0041488 | 0.0051913 | 0.018599 | 0.051002 |  | 0
da9 | 0.021306 | -0.2372 | 0.0078432 | -0.011229 | 0.012825 | -0.0040672 | -0.013914 | -0.0077583 | 0.018906 | 0.005646 | 0.0033245 | 0.022533 | -0.020314 | 0.00018684 | -0.0061044 | -0.011949 | 0.0070859 | -0.010167 | -0.006207 | 0.01336 | -0.037622 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.0035185 | 0.19857 | -0.018366 | -0.006965 | -0.0056135 | -0.0083974 | -0.0031163 | -0.0011093 | 0.011382 | 0.013537 | -0.00019207 | 0.0092359 | 0.00043734 | -0.002926 | 0.0068678 | 0.0012416 | -0.0013576 | -0.0035723 | 0.0024792 | 0.0064344 | -0.0086867 |  | 0
da12 | -0.0057348 | -0.0056575 | 0.0048401 | 0.0020805 | 0.0025928 | 0.0012679 | -0.0033855 | -0.0035702 | -0.0016078 | -0.00049205 | 0.00039485 | -0.00012225 | -0.00066903 | -0.00030904 | 0.0011764 | -0.00265 | 0.00073658 | -0.00075278 | -0.00034029 | 0.00080984 | -0.0071417 |  | -5.0596e-19
da13 | 0.0036906 | -0.017091 | -0.0012675 | -0.0027734 | -0.0015013 | 0.00016496 | -0.0021227 | -0.0012424 | 0.0016149 | 0.00010536 | 0.0007166 | 0.0040746 | 0.00079195 | 0.00094677 | 0.0023259 | -0.0018151 | -0.0008134 | -0.0019345 | 5.227e-05 | 0.0026768 | -0.00054718 |  | 0
da14 | 0.0020602 | -4.2954e-05 | -0.0030709 | -0.0020678 | -0.0010843 | -0.00025628 | 0.00063946 | -0.00046543 | -0.00034504 | 0.0012344 | -0.0015887 | 0.00019722 | 0.00054155 | 0.0014392 | 0.00061892 | 0.0014692 | 0.0012735 | 0.0010451 | 0.0003253 | 9.4612e-05 | 0.0033958 |  | -1.2047e-19
da15 | 0.0032733 | 0.0041922 | -0.0019169 | 9.0191e-05 | 0.00061405 | 0.00011082 | 0.00040078 | -0.00067415 | 0.00040747 | -0.0012231 | -0.00062892 | 0.00056013 | 0.0017063 | 0.00020884 | 0.002798 | 0.00016131 | -0.0019139 | -0.00055812 | 0.00058347 | -0.00072627 | -8.9882e-07 |  | -3.7345e-19
Field Colinearity | 75.8184
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.079 | -0.73659 | -0.16873 | -0.33907 | -0.23752 | -0.010278 | 0.18639 | 0.55898 | 0.55015 | -0.030081 | -0.30761 | 0.022119 | 0.39241 | 0.61335 | -0.15818 | -0.31489 | -0.50473 | -0.21851 | -0.13373 | -1.5288
Twist (m2 mrad) | 23.0929 | -10.4964 | -2.2779 | -4.3231 | -2.8502 | -0.11563 | 1.9571 | 5.4501 | 4.9513 | -0.24817 | -2.307 | 0.1493 | 2.3545 | 3.2201 | -0.71183 | -1.1809 | -1.5142 | -0.49164 | -0.2006 | -0.57331
Twist Integral (I0) (m2 rad) | 0.010413
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.3709 | -0.41554 | -0.055879 | -0.31868 | -0.2055 | -0.001745 | 0.18667 | 0.65607 | 0.58427 | 0.03403 | -0.41155 | -0.18344 | 0.13674 | 0.5 | -0.16406 | -0.58585 | -0.67477 | -0.41545 | -0.37773 | -1.2226
Twist (m2 mrad) | 32.782 | -5.9215 | -0.75437 | -4.0631 | -2.466 | -0.019631 | 1.96 | 6.3967 | 5.2584 | 0.28075 | -3.0866 | -1.2382 | 0.82042 | 2.625 | -0.73826 | -2.1969 | -2.0243 | -0.93476 | -0.5666 | -0.45849
Twist Integral (I0) (m2 rad) | 0.019241

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.26323 | -2.2241 | 0.019827 | -0.07587 | 1.1369 | -0.46534 | 0.33754 | 0.72664 | 0.036185 | 18.7313 | 0.59649 | 0.11066 | -2.1434 | 0.035754 | -0.24161 | 1.1288 | -0.95107 | -0.39178 | 0.79243 | 0.029837 | 17.9033 | 0.59647 | -0.13955 | -0.80296 | 0.030502 | -0.32268 | 1.1949 | -0.0044155 | 0.51818 | 0.61157 | 0.12205 | 14420.837 | 0.13286 | -0.7174 | 0.0156 | -0.5021 | 1.1734 | -0.004018 | 0.18085 | 0.78836 | 0.16495 | 14421.106
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8