## Alarm sheet
File name | HCMB__A001-03000172_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000172
Date of test Ap 1 | 2004-04-16 |  |  |  | Date of test Ap 2 | 2004-04-16
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | yellow alarm | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | yellow alarm | status ok | status ok | status ok | b15 | yellow alarm | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | yellow alarm | status ok | status ok | status ok | a14 | yellow alarm | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | yellow alarm | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
370 | HCMB__A001 | 3000172 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2004-04-19 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2004-04-19 | Answer to manufacturer |  |  | 2004-04-19
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000172_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000172 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 16-APR-2004_3172cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 16-APR-2004_3172cc2.txt |  |  |  | a even | 1
Test Operator |  | Mahler |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2004-04-16 |  |  | Time of test |  |  | 11:13 - 14:47 |  |  |  | Date of test |  | 2004-04-16 |  |  | Time of test |  |  | 11:14 - 14:47
Mole name |  | Dipole 2 | [9 empty] | Mole name |  | Dipole 3 | [5 empty] | 2003-05-02
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500
Dx average |  | -7.3681e-05 |  |  | Dy average |  |  | 0.0011602 |  |  |  | Dx average |  | 5.4815e-05 |  |  | Dy average |  |  | 0.0012021
Magnetic length |  | 14.4222 |  |  | Tranfs. Function (mT/KA) |  |  | 596.2092 |  |  |  | Magnetic length |  | 14.4203 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1396
Phase |  | -14.483 |  |  | Integral field (T.m) |  |  | 85.874 |  |  |  | Phase |  | -37.4218 |  |  | Integral field (T.m) |  |  | 85.7976
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4723 | 5.9655 | 5.9614 | 5.9625 | 5.9613 | 5.9621 | 5.9621 | 5.9618 | 5.9602 | 5.9616 | 5.962 | 5.9613 | 5.9611 | 5.9615 | 5.962 | 5.9619 | 5.9629 | 5.9603 | 5.966 | 3.7239 | 85.874 | C1 (T) | 3.3759 | 5.9641 | 5.9599 | 5.961 | 5.9593 | 5.9603 | 5.9607 | 5.961 | 5.9634 | 5.9603 | 5.9606 | 5.9622 | 5.9603 | 5.9611 | 5.9609 | 5.9611 | 5.9613 | 5.9634 | 5.9643 | 3.7326 | 85.7976
Angle (mrad) | 2.6368 | -0.33733 | -0.11853 | -0.29424 | -0.086655 | -0.32236 | 0.23387 | 0.50991 | 0.011552 | 0.21726 | -0.50016 | -0.091948 | 0.44585 | 0.34925 | 0.2295 | -0.088389 | -0.089113 | -0.45418 | -0.60166 | -0.8767 | -14.483 | Angle (mrad) | 4.4113 | -0.49409 | -0.1461 | -0.4688 | -0.29024 | -0.38383 | -0.074584 | 0.30467 | 0.063637 | -0.13103 | -0.63283 | 0.23532 | 0.36291 | 0.6397 | 0.28551 | -0.15939 | -0.25048 | -0.20032 | -0.52976 | -1.0028 | -37.4218
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.4983 | -0.81773 | -0.23139 | -0.6429 | -0.83155 | -1.2264 | -0.52522 | -0.23356 | 0.5617 | 0.1479 | 0.44904 | 0.30627 | 0.091103 | -0.82538 | -1.3329 | -0.39782 | -0.45393 | 0.23764 | -0.5481 | 0.28672 | -0.39309 | b2 | -2.4429 | 0.080227 | 0.26831 | 0.26262 | -0.33731 | -0.0072338 | 0.13396 | -0.25913 | -0.27543 | -0.30078 | 0.27337 | 0.22785 | 0.0099903 | 1.1526 | 0.71791 | 0.155 | -0.20856 | 0.1258 | -0.5628 | 0.15052 | 0.0087054
b3 | 38.0791 | -2.3734 | -1.9304 | -3.9856 | -1.8626 | -3.3401 | -2.3177 | -2.7195 | -2.7092 | -2.3024 | -4.2113 | -3.6988 | -3.3424 | -5.3929 | -4.3068 | -2.2695 | -1.7975 | -2.1334 | -1.7133 | 0.93603 | -1.5435 | b3 | 40.8481 | -3.1234 | -2.2457 | -4.1008 | -2.199 | -4.4595 | -3.1166 | -2.5096 | -2.62 | -2.3619 | -3.8865 | -3.753 | -3.9956 | -5.9933 | -4.2812 | -3.2933 | -2.4799 | -2.9445 | -2.291 | -0.24526 | -1.9111
b4 | -1.2576 | -0.2955 | -0.10651 | 0.0058815 | -0.031826 | -0.17769 | 0.10213 | 0.2551 | 0.28061 | 0.2387 | 0.47973 | 0.55118 | 0.25265 | -0.03976 | -0.11657 | 0.072359 | 0.082551 | 0.027884 | -0.12302 | 0.24245 | 0.045633 | b4 | -1.8384 | 0.035006 | 0.053938 | -0.06542 | -0.18255 | -0.11325 | -0.2208 | -0.074322 | -0.056623 | -0.17156 | 0.17889 | 0.030779 | -0.028675 | -0.0074764 | -0.016806 | 0.062316 | -0.063599 | -0.069136 | -0.19885 | -0.28719 | -0.11092
b5 | -5.5556 | -1.0101 | -0.94342 | -0.29169 | -0.48974 | -0.25481 | -0.16301 | -0.18006 | -0.35046 | -0.2386 | -0.050798 | -0.31392 | -0.22202 | -0.22539 | -0.4394 | -0.71774 | -0.56002 | -0.54907 | -0.50584 | -3.8306 | -0.68382 | b5 | -6.0115 | -0.98035 | -0.80958 | 0.0022496 | -0.32949 | -0.024243 | 0.022977 | 0.047935 | -0.23754 | -0.29331 | 0.089072 | -0.14295 | -0.27312 | 0.064792 | -0.10704 | -0.31847 | -0.64144 | -0.52474 | -0.46993 | -3.8121 | -0.55836
b6 | -0.14246 | -0.043881 | -0.035953 | 0.010342 | -0.0056563 | 0.15763 | 0.035439 | 0.0084221 | -0.060432 | -0.028488 | -0.018879 | -0.11324 | 0.0048804 | 0.1317 | 0.12858 | -0.045744 | 0.015514 | -0.064488 | -0.031011 | 0.015419 | -0.0014886 | b6 | -0.22428 | 0.051977 | -0.00071421 | 0.091298 | 0.059446 | 0.074209 | 0.14794 | 0.014593 | 0.10809 | 0.065044 | -0.10473 | -0.08049 | -0.037441 | -0.27114 | -0.24928 | -0.09585 | 0.012651 | -0.10487 | -0.098768 | -0.11197 | -0.032053
b7 | 2.9371 | 0.97815 | 1.0612 | 1.1915 | 1.1086 | 1.1239 | 1.1436 | 1.0631 | 1.0932 | 1.082 | 1.2216 | 1.1676 | 1.1438 | 1.1696 | 1.1231 | 1.1755 | 1.1395 | 1.1208 | 1.1004 | 0.87461 | 1.1696 | b7 | 2.8449 | 1.0215 | 1.0764 | 1.2411 | 1.1997 | 1.1642 | 1.1358 | 1.1375 | 1.1642 | 1.12 | 1.1248 | 1.1573 | 1.1701 | 1.1053 | 1.0802 | 1.2079 | 1.1285 | 1.1261 | 1.1071 | 0.94505 | 1.1812
b8 | -0.094317 | -0.027044 | -0.015665 | -0.017408 | -0.029919 | -0.039884 | -0.0073833 | -0.01048 | 0.039561 | 0.028024 | 0.005768 | -0.038002 | -0.014746 | -0.026246 | -0.049072 | -0.004351 | 0.0026365 | -0.04694 | -0.066551 | -0.0079366 | -0.019661 | b8 | -0.028523 | 0.030313 | 0.0029299 | -0.00098713 | -0.00536 | -0.00089206 | -0.0064508 | 0.0048733 | 0.010183 | -0.0016714 | -0.01325 | -0.0059094 | 0.020501 | 0.0076102 | -0.0078775 | -0.041886 | 0.039072 | 0.0087815 | -0.025666 | -0.011647 | -0.00047524
b9 | 0.34641 | 0.4714 | 0.45603 | 0.4459 | 0.49797 | 0.45814 | 0.48682 | 0.47006 | 0.46906 | 0.46902 | 0.46019 | 0.46785 | 0.4672 | 0.48462 | 0.47245 | 0.5059 | 0.4734 | 0.47053 | 0.47908 | 0.43504 | 0.46749 | b9 | 0.39 | 0.48351 | 0.47286 | 0.46854 | 0.51484 | 0.47281 | 0.50111 | 0.4947 | 0.50487 | 0.50326 | 0.47256 | 0.47101 | 0.45898 | 0.46027 | 0.46414 | 0.48107 | 0.48232 | 0.49176 | 0.48567 | 0.44263 | 0.47843
b10 | 0.039932 | 0.013759 | -0.02199 | -0.039104 | -0.028912 | 0.015606 | 0.009648 | 0.026001 | 0.023767 | 0.019036 | 0.0096329 | -0.026577 | 0.027432 | 0.042603 | -0.0031277 | -0.033734 | 0.010929 | -0.015306 | -0.048217 | -0.0074788 | 0 | b10 | 0.048832 | 0.063391 | 0.011878 | 0.013547 | -0.001043 | 0.036014 | 0.034802 | 0.051831 | 0.061518 | 0.034698 | 0.020077 | 0.0014298 | 0.010318 | -0.076018 | -0.07065 | -0.098833 | -0.013839 | -0.058168 | -0.040594 | -0.012789 | 0
b11 | 0.60491 | 0.72609 | 0.72402 | 0.73782 | 0.73103 | 0.73538 | 0.73136 | 0.73854 | 0.74018 | 0.7344 | 0.74243 | 0.73415 | 0.73941 | 0.74143 | 0.74743 | 0.73134 | 0.72457 | 0.72537 | 0.7276 | 0.63002 | 0.72673 | b11 | 0.59663 | 0.74781 | 0.74185 | 0.74889 | 0.73821 | 0.75451 | 0.75417 | 0.75287 | 0.75061 | 0.74589 | 0.75943 | 0.75176 | 0.74937 | 0.75911 | 0.75959 | 0.74892 | 0.74506 | 0.74824 | 0.74152 | 0.65379 | 0.74222
b12 | 0.0094977 | -0.00088779 | -0.0042158 | -0.0088971 | -0.0055111 | -0.00050464 | -0.0036427 | -0.0004182 | 0.0009418 | 0.0029167 | 0.0028743 | -0.0082541 | 0.0021588 | -0.0017592 | -0.0036722 | -0.0056414 | -0.0024751 | -0.0033436 | -0.0086585 | -0.004259 | -0.0024014 | b12 | 0.011639 | 0.0092048 | 0.0026495 | 0.0038202 | 0.0023897 | 0.0070107 | 0.0048244 | 0.0050791 | 0.0059817 | 0.0033302 | 0.0010486 | 0.00063999 | -0.00039614 | 0.00011562 | -0.0061575 | -0.010228 | 0.0021417 | -0.003607 | -0.0026055 | -0.00088089 | 0.0016299
b13 | 0.088679 | 0.064528 | 0.064116 | 0.057554 | 0.056189 | 0.054057 | 0.052685 | 0.048476 | 0.053239 | 0.056761 | 0.05807 | 0.056992 | 0.059207 | 0.055891 | 0.058575 | 0.059601 | 0.060645 | 0.061806 | 0.059596 | 0.046605 | 0.058247 | b13 | 0.09629 | 0.069703 | 0.067723 | 0.058714 | 0.059954 | 0.0554 | 0.05388 | 0.054275 | 0.055491 | 0.053432 | 0.055939 | 0.059192 | 0.062618 | 0.054047 | 0.059022 | 0.061763 | 0.062711 | 0.064316 | 0.061146 | 0.051888 | 0.06025
b14 | 0.0017375 | 0.00075784 | -0.0021744 | -0.0041764 | -0.003353 | 0.0035768 | 0.0015545 | 0.0024604 | 0.0032304 | 0.0036373 | 0.0024594 | -0.0017963 | 0.0042248 | 0.0055447 | 0.0014071 | -0.0022747 | 0.0027232 | -0.00074222 | -0.0039298 | 0.0010249 | 0.00076941 | b14 | 0.0022623 | 0.0052203 | -0.00028002 | 0.00014654 | -0.00075081 | 0.0038727 | 0.0026558 | 0.0056793 | 0.0074699 | 0.0024597 | 0.00083967 | -0.0025053 | -4.205e-05 | -0.0066973 | -0.0095904 | -0.011185 | -0.0023453 | -0.0063518 | -0.0056281 | -0.0015603 | -0.00087163
b15 | 0.0043102 | 0.05516 | 0.052915 | 0.056111 | 0.0563 | 0.058716 | 0.058468 | 0.061116 | 0.05905 | 0.05979 | 0.055741 | 0.055128 | 0.056725 | 0.05676 | 0.055629 | 0.053543 | 0.0578 | 0.055956 | 0.051842 | 0.038934 | 0.054333 | b15 | 0.0093028 | 0.056082 | 0.057005 | 0.061315 | 0.058378 | 0.061307 | 0.058863 | 0.062296 | 0.062033 | 0.064639 | 0.065221 | 0.057946 | 0.063287 | 0.06019 | 0.058296 | 0.05436 | 0.058568 | 0.056311 | 0.060172 | 0.048699 | 0.057941
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 26.3678 | -3.3733 | -1.1852 | -2.9424 | -0.86655 | -3.2235 | 2.3387 | 5.0991 | 0.11552 | 2.1726 | -5.0016 | -0.91948 | 4.4585 | 3.4925 | 2.295 | -0.88389 | -0.89113 | -4.5418 | -6.0166 | -8.7671 | 0 | a1 | 44.1134 | -4.9409 | -1.461 | -4.688 | -2.9024 | -3.8383 | -0.74584 | 3.0467 | 0.63637 | -1.3103 | -6.3283 | 2.3532 | 3.6291 | 6.397 | 2.8551 | -1.5939 | -2.5048 | -2.0032 | -5.2976 | -10.0285 | 0
a2 | 17.997 | -0.18991 | -0.22798 | 0.79184 | -0.67788 | -0.80793 | -0.87915 | -0.26535 | -0.55993 | 0.2512 | -0.80642 | 0.18374 | -1.0886 | -0.87978 | -0.90099 | -0.23754 | 0.59355 | -0.17578 | 0.80104 | 14.2409 | 0.7445 | a2 | 17.0869 | 0.67094 | -0.64703 | -1.3513 | -1.7599 | -1.3828 | -1.6126 | -1.1862 | -0.84414 | -0.34122 | 0.075593 | 0.93917 | 0.82702 | -0.95978 | -0.85203 | 0.60016 | 0.66382 | -0.62766 | 1.0319 | 7.0195 | 0.38119
a3 | -1.062 | 1.4165 | 0.81178 | 1.2625 | 1.0856 | 0.83237 | 1.2121 | 0.46147 | 1.0775 | 1.1566 | 0.87092 | 0.74554 | 1.0339 | 1.5315 | 1.1269 | 1.1102 | 1.1696 | 1.1005 | 1.0058 | 0.10427 | 0.96102 | a3 | -0.37134 | 1.0044 | 1.0718 | 1.198 | 1.0355 | 0.85235 | 0.84705 | 0.58896 | 0.95034 | 0.47996 | 0.90602 | 0.52362 | 0.70517 | 0.6118 | 0.80558 | 0.63703 | 0.59439 | 0.60009 | 1.4271 | 0.76367 | 0.78715
a4 | 2.7261 | -0.27979 | -0.12933 | -0.86892 | -0.52549 | -0.62618 | -0.7785 | -0.94344 | -0.37844 | -0.54687 | 0.052822 | -0.81049 | -0.60407 | -0.063195 | -0.45804 | -0.81589 | -0.75703 | -0.56547 | -0.83597 | 0.89307 | -0.40554 | a4 | 3.4849 | 0.10427 | 0.027826 | 0.0016777 | -0.32531 | -0.64375 | -0.62474 | -0.30262 | 0.026951 | -0.19492 | -0.53443 | -0.41191 | -0.44739 | 0.4311 | 0.27357 | -0.14372 | 0.15449 | 0.042195 | -0.22516 | 0.5061 | -0.026122
a5 | 2.5339 | 0.14342 | 0.060605 | 0.070492 | 0.20046 | 0.15532 | 0.2077 | 0.03397 | 0.34776 | 0.28173 | 0.18269 | -0.063573 | -0.097494 | 0.43238 | 0.30062 | 0.26482 | 0.087766 | -0.011455 | 0.31681 | 0.039283 | 0.22983 | a5 | 2.2881 | 0.22034 | 0.046541 | 0.09019 | 0.14523 | 0.13179 | 0.23158 | 0.087256 | 0.3116 | 0.33388 | 0.25957 | 0.037787 | -0.11167 | -0.26519 | -0.047564 | 0.20596 | -0.014016 | -0.022657 | 0.30567 | -0.26468 | 0.16029
a6 | 0.55177 | 0.1326 | 0.088128 | 0.092918 | 0.10426 | 0.071718 | -0.08693 | 0.010834 | -0.0057866 | 0.090884 | -0.055796 | -0.034605 | 0.015018 | -0.033271 | 0.041276 | -0.015026 | 0.049151 | 0.071006 | 0.045559 | 0.13313 | 0.05136 | a6 | 0.49806 | 0.001058 | 0.015183 | -0.017704 | -0.079835 | -0.12768 | -0.16802 | 0.042815 | 0.014592 | -0.07627 | 0.010197 | -0.066971 | -0.016693 | -0.059197 | -0.14207 | -0.095394 | 0.079187 | 0.048503 | 0.092003 | 0.12012 | -0.009843
a7 | 1.4617 | -0.021243 | -0.056814 | -0.10983 | 0.016395 | 0.03194 | -0.014791 | -0.011394 | -0.065782 | -0.035522 | 0.014725 | 0.011841 | 0.065672 | -0.024421 | 0.00081904 | -0.079073 | -0.047488 | -0.059277 | -0.004918 | 0.0047856 | 0.024213 | a7 | 1.7015 | 0.0045845 | 0.0062541 | -0.019001 | 0.0011937 | -0.013382 | 0.065586 | -0.0029348 | 0.018338 | 0.012929 | 0.061707 | 0.02966 | 0.11382 | 0.10227 | -0.032964 | -0.089672 | -0.06431 | 0.04116 | 0.083587 | 0.011793 | 0.067198
a8 | 0.016905 | -0.019087 | -0.025672 | -0.0049278 | 0.0064997 | -0.025062 | -0.044359 | -0.070973 | -0.031049 | 0.0025847 | -0.015263 | -0.0016417 | -0.026529 | 0.023963 | 0.00042866 | -0.0025413 | -0.0061054 | -0.011093 | -0.0087298 | 0.041874 | -0.011639 | a8 | -0.021808 | -0.02489 | -0.0042558 | -0.054733 | -0.049196 | -0.088758 | -0.091411 | -0.013483 | -0.03364 | -0.063712 | -0.010845 | 0.049693 | 0.010608 | 0.010007 | 0.00085692 | 0.047681 | 0.029944 | 0.014881 | 0.01177 | -0.0044679 | -0.014307
a9 | -0.25015 | -0.039397 | -0.020701 | -0.015776 | -0.063095 | -0.041088 | -0.027053 | -0.030123 | -0.018876 | -0.032292 | -0.021591 | -0.031984 | 0.0042373 | 0.0055831 | -0.013587 | -0.012464 | -0.019371 | -0.0095655 | -0.0028441 | -0.061562 | -0.029891 | a9 | -0.23753 | 0.021484 | 0.015398 | 0.0084096 | 0.0039477 | 0.0045838 | 0.0047198 | 0.029461 | 0.019215 | -0.016201 | 0.0089979 | 0.013596 | 0.042827 | -0.00051779 | -0.017195 | 0.0085284 | 0.02275 | 0.0073164 | 0.027662 | -0.03943 | 0.0023872
a10 | 0.07987 | 0.010937 | 0.024323 | -0.0041912 | 0.011409 | -0.023089 | -0.030248 | -0.037441 | -0.024281 | -0.016542 | 0.0049705 | -0.0035559 | -0.0058761 | 0.016025 | 0.015174 | -0.00011277 | -0.014006 | 0.014176 | 0.024161 | -0.013395 | 0 | a10 | 0.02793 | 0.017795 | 0.022042 | -0.01435 | -0.0029171 | -0.015206 | 0.0001332 | -0.017293 | -0.020386 | -0.028384 | -0.023872 | 0.018595 | -0.022897 | 0.020451 | 0.016942 | 0.02253 | 0.01951 | 0.032738 | 0.00086506 | -0.067294 | 0
a11 | 0.14522 | -0.060201 | -0.053174 | -0.058596 | -0.048936 | -0.047522 | -0.041093 | -0.037494 | -0.038178 | -0.041078 | -0.04416 | -0.045573 | -0.037971 | -0.031765 | -0.036726 | -0.041826 | -0.042996 | -0.045787 | -0.035601 | -0.027749 | -0.037562 | a11 | 0.1956 | -0.01044 | -0.0052628 | -0.0031232 | -0.0062861 | 0.0030742 | 0.016108 | 0.0083247 | 0.015879 | 0.0072037 | 0.0080637 | 0.0098469 | 0.014661 | 0.018307 | 0.0030996 | 0.0098693 | 0.00017226 | 0.011151 | 0.011837 | 0.010488 | 0.011974
a12 | 0.016676 | 0.00078386 | 0.0013538 | -0.00082069 | -0.0022598 | -0.008085 | -0.0082018 | -0.006722 | -0.0074432 | -0.0036683 | -0.0074512 | -0.0053141 | -0.003888 | -0.006265 | -0.0063074 | -0.0083183 | -0.0060923 | -0.0050784 | -0.0034206 | -0.00027188 | -0.004043 | a12 | 0.00031826 | -0.0020829 | -0.0023549 | -0.0072982 | -0.0089351 | -0.010094 | -0.0084474 | -0.0053354 | -0.0067206 | -0.0074798 | -0.0047731 | 5.2801e-05 | -0.0068214 | -0.0060565 | -0.0072528 | -0.0072909 | -0.0025045 | -0.0026769 | -0.0051218 | -0.0098735 | -0.005585
a13 | -0.019085 | -0.002172 | -0.0021642 | -0.0009208 | -0.0012621 | -0.0041748 | -0.003083 | -0.003332 | 8.7055e-05 | -0.00029776 | -0.00034186 | -0.0026993 | -0.00030963 | 0.0029146 | -0.0010647 | 0.00088045 | -0.0011435 | -0.001116 | 0.0018384 | -0.0015514 | -0.001585 | a13 | -0.016883 | 0.0013099 | 0.0023286 | 0.00074258 | 0.004084 | 0.00080206 | 0.0017953 | 0.003912 | 0.0033105 | 0.0029258 | 0.0017698 | 0.0040028 | 0.0041483 | 0.0048306 | 0.0037261 | 0.0051043 | 0.0067127 | 0.0070273 | 0.0091267 | 0.00053348 | 0.0030448
a14 | 0.012688 | -0.031903 | -0.029606 | -0.032588 | -0.032357 | -0.033128 | -0.032983 | -0.038819 | -0.033389 | -0.034012 | -0.028505 | -0.030248 | -0.031784 | -0.028982 | -0.030131 | -0.029302 | -0.032647 | -0.032598 | -0.030692 | -0.014736 | -0.029963 | a14 | 0.010795 | -0.028485 | -0.029557 | -0.032309 | -0.030537 | -0.035784 | -0.034362 | -0.034274 | -0.033643 | -0.036264 | -0.036717 | -0.028765 | -0.034152 | -0.031063 | -0.029115 | -0.025909 | -0.029635 | -0.029048 | -0.032013 | -0.019377 | -0.030098
a15 | 0.013961 | 0.00053542 | -0.003702 | -0.0056834 | -0.0061719 | 0.0010266 | 0.00042287 | 0.0031618 | 0.0048966 | 0.0046609 | 0.002053 | -0.0048797 | 0.0053373 | 0.0060355 | 0.00089014 | -0.0029739 | 0.0032745 | -0.0031023 | -0.0070258 | 0.00011768 | 0.00036205 | a15 | 0.0075096 | 0.0070923 | -0.00015008 | 0.00063077 | -0.0023784 | 0.0028741 | 0.0026893 | 0.0063235 | 0.0074856 | 0.0010849 | -0.0016349 | -0.0041501 | -0.0013211 | -0.010287 | -0.012964 | -0.016544 | -0.0020376 | -0.009085 | -0.0091796 | -0.0064875 | -0.0021551
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00015665 | -2.9885e-05 | 5.5527e-05 | 8.8797e-05 | 6.8716e-05 | -3.9349e-05 | -2.6288e-05 | -6.4035e-05 | -5.7293e-05 | -4.606e-05 | -2.1307e-05 | 6.0804e-05 | -6.3596e-05 | -9.5948e-05 | 0 | 7.8152e-05 | -2.749e-05 | 3.7824e-05 | 0.0001152 | 1.855e-05 | -7.3681e-05 | Dx (m) | -0.00014929 | -0.00014359 | -2.6861e-05 | -3.0882e-05 | 0 | -8.0993e-05 | -7.8427e-05 | -0.00011659 | -0.00013828 | -7.8423e-05 | -4.4363e-05 | 0 | -2.2377e-05 | 0.00016906 | 0.00015804 | 0.00022386 | 3.157e-05 | 0.00013106 | 9.3016e-05 | 3.5995e-05 | 5.4815e-05
Dy (m) | -0.00018695 | -2.8084e-05 | -5.3041e-05 | 1.6721e-05 | -2.1933e-05 | 5.0827e-05 | 6.8812e-05 | 8.2921e-05 | 5.2814e-05 | 3.5721e-05 | -1.2648e-05 | 1.2021e-05 | 1.0251e-05 | -4.0822e-05 | -3.4084e-05 | 0 | 3.1226e-05 | -3.0839e-05 | -5.0832e-05 | 3.6965e-05 | 0.0011602 | Dy (m) | -3.0681e-05 | -4.2471e-05 | -5.0706e-05 | 3.2449e-05 | 0 | 3.4628e-05 | 0 | 4.0376e-05 | 4.9165e-05 | 6.5465e-05 | 5.3908e-05 | -4.2e-05 | 5.2374e-05 | -4.9786e-05 | -3.8505e-05 | -5.4005e-05 | -4.4521e-05 | -7.633e-05 | 0 | 0.00017429 | 0.0012021
Magnetic length (m) |  | 14.4222 | [19 empty] | Magnetic length (m) |  | 14.4203
Transfer funct. (mT/KA) |  | 596.2092 | [19 empty] | Transfer funct. (mT/KA) |  | 596.1396
Coil positioning (mm) |  | 0.019501 | [19 empty] | Coil positioning (mm) |  | 0.02093
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.0.8 | [20 empty] | Soft version | DIMM V2.0.8
Level Average | 0.020028 | 0.22988 | 0.066837 | -0.055328 | -0.76829 | -0.043209 | 0.27506 | -0.064617 | -0.24698 | 0.021334 | 0.35946 | 0.32397 | -0.12356 | 0.29051 | -0.36744 | -0.42519 | -0.79943 | -0.030326 | -0.11937 | -0.01283 |  | Level Average | 1.444 | 0.14205 | 1.0755 | 1.8378 | 0.20009 | 1.9245 | 1.922 | 1.3902 | 0.61538 | 1.4347 | 0.98961 | -0.70178 | -1.366 | -0.92884 | -1.336 | -1.7612 | -1.1674 | -1.8033 | -1.0991 | -0.059357
Level Sigma2 | 0.000501 | 0.22971 | 0.007353 | 0.000122 | 0.010311 | 0.000149 | 0.002323 | 0.000221 | 0.000709 | 0.004514 | 0.40962 | 0.012508 | 0.012608 | 0.003895 | 0.003928 | 0.000191 | 34.9269 | 3.5e-05 | 0.004936 | 5e-05 |  | Level Sigma2 | 0.43257 | 0.003289 | 2.9993 | 0.72618 | 0.00283 | 0.032215 | 2.8224 | 0.25658 | 0.057285 | 0.17966 | 0.1711 | 61.9577 | 1.0094 | 23.4066 | 1.2585 | 4.0951 | 43.5688 | 1.1807 | 24.2597 | 0.000204
Position (mm) | 0 | 748 | 1499 | 2249 | 3003 | 3753 | 4503 | 5254 | 6006 | 6756 | 7508 | 8259 | 9010 | 9761 | 10512 | 11263 | 12015 | 12765 | 13516 | 14267 |  | Position (mm) | 0 | 751 | 1503 | 2254 | 3006 | 3757 | 4508 | 5259 | 6010 | 6762 | 7513 | 8265 | 9016 | 9767 | 10519 | 11270 | 12021 | 12773 | 13524 | 14276
Fist move (mm) | First move Ap1 | 941.446 | First move Ap2 | 946.8954 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-3000172 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000172 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 347.2305 | 596.5469 | 596.1428 | 596.2518 | 596.1345 | 596.2117 | 596.209 | 596.1803 | 596.0216 | 596.1642 | 596.1996 | 596.1322 | 596.1086 | 596.1495 | 596.1996 | 596.1904 | 596.2864 | 596.0347 | 596.6026 | 372.3875 | 8587.4024 | C1/i (mT/KA) | 337.5928 | 596.4091 | 595.9854 | 596.0952 | 595.9295 | 596.0315 | 596.0651 | 596.102 | 596.3428 | 596.0322 | 596.0593 | 596.2153 | 596.0338 | 596.1121 | 596.0861 | 596.1117 | 596.1287 | 596.3407 | 596.432 | 373.2635 | 8579.7562
Angle (mrad) | 2.6368 | -0.33733 | -0.11853 | -0.29424 | -0.086655 | -0.32236 | 0.23387 | 0.50991 | 0.011552 | 0.21726 | -0.50016 | -0.091948 | 0.44585 | 0.34925 | 0.2295 | -0.088389 | -0.089113 | -0.45418 | -0.60166 | -0.8767 | -14.483 | Angle (mrad) | 4.4113 | -0.49409 | -0.1461 | -0.4688 | -0.29024 | -0.38383 | -0.074584 | 0.30467 | 0.063637 | -0.13103 | -0.63283 | 0.23532 | 0.36291 | 0.6397 | 0.28551 | -0.15939 | -0.25048 | -0.20032 | -0.52976 | -1.0028 | -37.4218
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.4983 | -0.81773 | -0.23139 | -0.6429 | -0.83155 | -1.2264 | -0.52522 | -0.23356 | 0.5617 | 0.1479 | 0.44904 | 0.30627 | 0.091103 | -0.82538 | -1.3329 | -0.39782 | -0.45393 | 0.23764 | -0.5481 | 0.28672 |  | b2 | -2.4429 | 0.080227 | 0.26831 | 0.26262 | -0.33731 | -0.0072338 | 0.13396 | -0.25913 | -0.27543 | -0.30078 | 0.27337 | 0.22785 | 0.0099903 | 1.1526 | 0.71791 | 0.155 | -0.20856 | 0.1258 | -0.5628 | 0.15052
b3 | 38.0791 | -2.3734 | -1.9304 | -3.9856 | -1.8626 | -3.3401 | -2.3177 | -2.7195 | -2.7092 | -2.3024 | -4.2113 | -3.6988 | -3.3424 | -5.3929 | -4.3068 | -2.2695 | -1.7975 | -2.1334 | -1.7133 | 0.93603 |  | b3 | 40.8481 | -3.1234 | -2.2457 | -4.1008 | -2.199 | -4.4595 | -3.1166 | -2.5096 | -2.62 | -2.3619 | -3.8865 | -3.753 | -3.9956 | -5.9933 | -4.2812 | -3.2933 | -2.4799 | -2.9445 | -2.291 | -0.24526
b4 | -1.2576 | -0.2955 | -0.10651 | 0.0058815 | -0.031826 | -0.17769 | 0.10213 | 0.2551 | 0.28061 | 0.2387 | 0.47973 | 0.55118 | 0.25265 | -0.03976 | -0.11657 | 0.072359 | 0.082551 | 0.027884 | -0.12302 | 0.24245 |  | b4 | -1.8384 | 0.035006 | 0.053938 | -0.06542 | -0.18255 | -0.11325 | -0.2208 | -0.074322 | -0.056623 | -0.17156 | 0.17889 | 0.030779 | -0.028675 | -0.0074764 | -0.016806 | 0.062316 | -0.063599 | -0.069136 | -0.19885 | -0.28719
b5 | -5.5556 | -1.0101 | -0.94342 | -0.29169 | -0.48974 | -0.25481 | -0.16301 | -0.18006 | -0.35046 | -0.2386 | -0.050798 | -0.31392 | -0.22202 | -0.22539 | -0.4394 | -0.71774 | -0.56002 | -0.54907 | -0.50584 | -3.8306 |  | b5 | -6.0115 | -0.98035 | -0.80958 | 0.0022496 | -0.32949 | -0.024243 | 0.022977 | 0.047935 | -0.23754 | -0.29331 | 0.089072 | -0.14295 | -0.27312 | 0.064792 | -0.10704 | -0.31847 | -0.64144 | -0.52474 | -0.46993 | -3.8121
b6 | -0.14246 | -0.043881 | -0.035953 | 0.010342 | -0.0056563 | 0.15763 | 0.035439 | 0.0084221 | -0.060432 | -0.028488 | -0.018879 | -0.11324 | 0.0048804 | 0.1317 | 0.12858 | -0.045744 | 0.015514 | -0.064488 | -0.031011 | 0.015419 |  | b6 | -0.22428 | 0.051977 | -0.00071421 | 0.091298 | 0.059446 | 0.074209 | 0.14794 | 0.014593 | 0.10809 | 0.065044 | -0.10473 | -0.08049 | -0.037441 | -0.27114 | -0.24928 | -0.09585 | 0.012651 | -0.10487 | -0.098768 | -0.11197
b7 | 2.9371 | 0.97815 | 1.0612 | 1.1915 | 1.1086 | 1.1239 | 1.1436 | 1.0631 | 1.0932 | 1.082 | 1.2216 | 1.1676 | 1.1438 | 1.1696 | 1.1231 | 1.1755 | 1.1395 | 1.1208 | 1.1004 | 0.87461 |  | b7 | 2.8449 | 1.0215 | 1.0764 | 1.2411 | 1.1997 | 1.1642 | 1.1358 | 1.1375 | 1.1642 | 1.12 | 1.1248 | 1.1573 | 1.1701 | 1.1053 | 1.0802 | 1.2079 | 1.1285 | 1.1261 | 1.1071 | 0.94505
b8 | -0.094317 | -0.027044 | -0.015665 | -0.017408 | -0.029919 | -0.039884 | -0.0073833 | -0.01048 | 0.039561 | 0.028024 | 0.005768 | -0.038002 | -0.014746 | -0.026246 | -0.049072 | -0.004351 | 0.0026365 | -0.04694 | -0.066551 | -0.0079366 |  | b8 | -0.028523 | 0.030313 | 0.0029299 | -0.00098713 | -0.00536 | -0.00089206 | -0.0064508 | 0.0048733 | 0.010183 | -0.0016714 | -0.01325 | -0.0059094 | 0.020501 | 0.0076102 | -0.0078775 | -0.041886 | 0.039072 | 0.0087815 | -0.025666 | -0.011647
b9 | 0.34641 | 0.4714 | 0.45603 | 0.4459 | 0.49797 | 0.45814 | 0.48682 | 0.47006 | 0.46906 | 0.46902 | 0.46019 | 0.46785 | 0.4672 | 0.48462 | 0.47245 | 0.5059 | 0.4734 | 0.47053 | 0.47908 | 0.43504 |  | b9 | 0.39 | 0.48351 | 0.47286 | 0.46854 | 0.51484 | 0.47281 | 0.50111 | 0.4947 | 0.50487 | 0.50326 | 0.47256 | 0.47101 | 0.45898 | 0.46027 | 0.46414 | 0.48107 | 0.48232 | 0.49176 | 0.48567 | 0.44263
b10 | 0.039932 | 0.013759 | -0.02199 | -0.039104 | -0.028912 | 0.015606 | 0.009648 | 0.026001 | 0.023767 | 0.019036 | 0.0096329 | -0.026577 | 0.027432 | 0.042603 | -0.0031277 | -0.033734 | 0.010929 | -0.015306 | -0.048217 | -0.0074788 |  | b10 | 0.048832 | 0.063391 | 0.011878 | 0.013547 | -0.001043 | 0.036014 | 0.034802 | 0.051831 | 0.061518 | 0.034698 | 0.020077 | 0.0014298 | 0.010318 | -0.076018 | -0.07065 | -0.098833 | -0.013839 | -0.058168 | -0.040594 | -0.012789
b11 | 0.60491 | 0.72609 | 0.72402 | 0.73782 | 0.73103 | 0.73538 | 0.73136 | 0.73854 | 0.74018 | 0.7344 | 0.74243 | 0.73415 | 0.73941 | 0.74143 | 0.74743 | 0.73134 | 0.72457 | 0.72537 | 0.7276 | 0.63002 |  | b11 | 0.59663 | 0.74781 | 0.74185 | 0.74889 | 0.73821 | 0.75451 | 0.75417 | 0.75287 | 0.75061 | 0.74589 | 0.75943 | 0.75176 | 0.74937 | 0.75911 | 0.75959 | 0.74892 | 0.74506 | 0.74824 | 0.74152 | 0.65379
b12 | 0.0094977 | -0.00088779 | -0.0042158 | -0.0088971 | -0.0055111 | -0.00050464 | -0.0036427 | -0.0004182 | 0.0009418 | 0.0029167 | 0.0028743 | -0.0082541 | 0.0021588 | -0.0017592 | -0.0036722 | -0.0056414 | -0.0024751 | -0.0033436 | -0.0086585 | -0.004259 |  | b12 | 0.011639 | 0.0092048 | 0.0026495 | 0.0038202 | 0.0023897 | 0.0070107 | 0.0048244 | 0.0050791 | 0.0059817 | 0.0033302 | 0.0010486 | 0.00063999 | -0.00039614 | 0.00011562 | -0.0061575 | -0.010228 | 0.0021417 | -0.003607 | -0.0026055 | -0.00088089
b13 | 0.088679 | 0.064528 | 0.064116 | 0.057554 | 0.056189 | 0.054057 | 0.052685 | 0.048476 | 0.053239 | 0.056761 | 0.05807 | 0.056992 | 0.059207 | 0.055891 | 0.058575 | 0.059601 | 0.060645 | 0.061806 | 0.059596 | 0.046605 |  | b13 | 0.09629 | 0.069703 | 0.067723 | 0.058714 | 0.059954 | 0.0554 | 0.05388 | 0.054275 | 0.055491 | 0.053432 | 0.055939 | 0.059192 | 0.062618 | 0.054047 | 0.059022 | 0.061763 | 0.062711 | 0.064316 | 0.061146 | 0.051888
b14 | 0.0017375 | 0.00075784 | -0.0021744 | -0.0041764 | -0.003353 | 0.0035768 | 0.0015545 | 0.0024604 | 0.0032304 | 0.0036373 | 0.0024594 | -0.0017963 | 0.0042248 | 0.0055447 | 0.0014071 | -0.0022747 | 0.0027232 | -0.00074222 | -0.0039298 | 0.0010249 |  | b14 | 0.0022623 | 0.0052203 | -0.00028002 | 0.00014654 | -0.00075081 | 0.0038727 | 0.0026558 | 0.0056793 | 0.0074699 | 0.0024597 | 0.00083967 | -0.0025053 | -4.205e-05 | -0.0066973 | -0.0095904 | -0.011185 | -0.0023453 | -0.0063518 | -0.0056281 | -0.0015603
b15 | 0.0043102 | 0.05516 | 0.052915 | 0.056111 | 0.0563 | 0.058716 | 0.058468 | 0.061116 | 0.05905 | 0.05979 | 0.055741 | 0.055128 | 0.056725 | 0.05676 | 0.055629 | 0.053543 | 0.0578 | 0.055956 | 0.051842 | 0.038934 |  | b15 | 0.0093028 | 0.056082 | 0.057005 | 0.061315 | 0.058378 | 0.061307 | 0.058863 | 0.062296 | 0.062033 | 0.064639 | 0.065221 | 0.057946 | 0.063287 | 0.06019 | 0.058296 | 0.05436 | 0.058568 | 0.056311 | 0.060172 | 0.048699
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 26.3678 | -3.3733 | -1.1852 | -2.9424 | -0.86655 | -3.2235 | 2.3387 | 5.0991 | 0.11552 | 2.1726 | -5.0016 | -0.91948 | 4.4585 | 3.4925 | 2.295 | -0.88389 | -0.89113 | -4.5418 | -6.0166 | -8.7671 |  | a1 | 44.1134 | -4.9409 | -1.461 | -4.688 | -2.9024 | -3.8383 | -0.74584 | 3.0467 | 0.63637 | -1.3103 | -6.3283 | 2.3532 | 3.6291 | 6.397 | 2.8551 | -1.5939 | -2.5048 | -2.0032 | -5.2976 | -10.0285
a2 | 17.997 | -0.18991 | -0.22798 | 0.79184 | -0.67788 | -0.80793 | -0.87915 | -0.26535 | -0.55993 | 0.2512 | -0.80642 | 0.18374 | -1.0886 | -0.87978 | -0.90099 | -0.23754 | 0.59355 | -0.17578 | 0.80104 | 14.2409 |  | a2 | 17.0869 | 0.67094 | -0.64703 | -1.3513 | -1.7599 | -1.3828 | -1.6126 | -1.1862 | -0.84414 | -0.34122 | 0.075593 | 0.93917 | 0.82702 | -0.95978 | -0.85203 | 0.60016 | 0.66382 | -0.62766 | 1.0319 | 7.0195
a3 | -1.062 | 1.4165 | 0.81178 | 1.2625 | 1.0856 | 0.83237 | 1.2121 | 0.46147 | 1.0775 | 1.1566 | 0.87092 | 0.74554 | 1.0339 | 1.5315 | 1.1269 | 1.1102 | 1.1696 | 1.1005 | 1.0058 | 0.10427 |  | a3 | -0.37134 | 1.0044 | 1.0718 | 1.198 | 1.0355 | 0.85235 | 0.84705 | 0.58896 | 0.95034 | 0.47996 | 0.90602 | 0.52362 | 0.70517 | 0.6118 | 0.80558 | 0.63703 | 0.59439 | 0.60009 | 1.4271 | 0.76367
a4 | 2.7261 | -0.27979 | -0.12933 | -0.86892 | -0.52549 | -0.62618 | -0.7785 | -0.94344 | -0.37844 | -0.54687 | 0.052822 | -0.81049 | -0.60407 | -0.063195 | -0.45804 | -0.81589 | -0.75703 | -0.56547 | -0.83597 | 0.89307 |  | a4 | 3.4849 | 0.10427 | 0.027826 | 0.0016777 | -0.32531 | -0.64375 | -0.62474 | -0.30262 | 0.026951 | -0.19492 | -0.53443 | -0.41191 | -0.44739 | 0.4311 | 0.27357 | -0.14372 | 0.15449 | 0.042195 | -0.22516 | 0.5061
a5 | 2.5339 | 0.14342 | 0.060605 | 0.070492 | 0.20046 | 0.15532 | 0.2077 | 0.03397 | 0.34776 | 0.28173 | 0.18269 | -0.063573 | -0.097494 | 0.43238 | 0.30062 | 0.26482 | 0.087766 | -0.011455 | 0.31681 | 0.039283 |  | a5 | 2.2881 | 0.22034 | 0.046541 | 0.09019 | 0.14523 | 0.13179 | 0.23158 | 0.087256 | 0.3116 | 0.33388 | 0.25957 | 0.037787 | -0.11167 | -0.26519 | -0.047564 | 0.20596 | -0.014016 | -0.022657 | 0.30567 | -0.26468
a6 | 0.55177 | 0.1326 | 0.088128 | 0.092918 | 0.10426 | 0.071718 | -0.08693 | 0.010834 | -0.0057866 | 0.090884 | -0.055796 | -0.034605 | 0.015018 | -0.033271 | 0.041276 | -0.015026 | 0.049151 | 0.071006 | 0.045559 | 0.13313 |  | a6 | 0.49806 | 0.001058 | 0.015183 | -0.017704 | -0.079835 | -0.12768 | -0.16802 | 0.042815 | 0.014592 | -0.07627 | 0.010197 | -0.066971 | -0.016693 | -0.059197 | -0.14207 | -0.095394 | 0.079187 | 0.048503 | 0.092003 | 0.12012
a7 | 1.4617 | -0.021243 | -0.056814 | -0.10983 | 0.016395 | 0.03194 | -0.014791 | -0.011394 | -0.065782 | -0.035522 | 0.014725 | 0.011841 | 0.065672 | -0.024421 | 0.00081904 | -0.079073 | -0.047488 | -0.059277 | -0.004918 | 0.0047856 |  | a7 | 1.7015 | 0.0045845 | 0.0062541 | -0.019001 | 0.0011937 | -0.013382 | 0.065586 | -0.0029348 | 0.018338 | 0.012929 | 0.061707 | 0.02966 | 0.11382 | 0.10227 | -0.032964 | -0.089672 | -0.06431 | 0.04116 | 0.083587 | 0.011793
a8 | 0.016905 | -0.019087 | -0.025672 | -0.0049278 | 0.0064997 | -0.025062 | -0.044359 | -0.070973 | -0.031049 | 0.0025847 | -0.015263 | -0.0016417 | -0.026529 | 0.023963 | 0.00042866 | -0.0025413 | -0.0061054 | -0.011093 | -0.0087298 | 0.041874 |  | a8 | -0.021808 | -0.02489 | -0.0042558 | -0.054733 | -0.049196 | -0.088758 | -0.091411 | -0.013483 | -0.03364 | -0.063712 | -0.010845 | 0.049693 | 0.010608 | 0.010007 | 0.00085692 | 0.047681 | 0.029944 | 0.014881 | 0.01177 | -0.0044679
a9 | -0.25015 | -0.039397 | -0.020701 | -0.015776 | -0.063095 | -0.041088 | -0.027053 | -0.030123 | -0.018876 | -0.032292 | -0.021591 | -0.031984 | 0.0042373 | 0.0055831 | -0.013587 | -0.012464 | -0.019371 | -0.0095655 | -0.0028441 | -0.061562 |  | a9 | -0.23753 | 0.021484 | 0.015398 | 0.0084096 | 0.0039477 | 0.0045838 | 0.0047198 | 0.029461 | 0.019215 | -0.016201 | 0.0089979 | 0.013596 | 0.042827 | -0.00051779 | -0.017195 | 0.0085284 | 0.02275 | 0.0073164 | 0.027662 | -0.03943
a10 | 0.07987 | 0.010937 | 0.024323 | -0.0041912 | 0.011409 | -0.023089 | -0.030248 | -0.037441 | -0.024281 | -0.016542 | 0.0049705 | -0.0035559 | -0.0058761 | 0.016025 | 0.015174 | -0.00011277 | -0.014006 | 0.014176 | 0.024161 | -0.013395 |  | a10 | 0.02793 | 0.017795 | 0.022042 | -0.01435 | -0.0029171 | -0.015206 | 0.0001332 | -0.017293 | -0.020386 | -0.028384 | -0.023872 | 0.018595 | -0.022897 | 0.020451 | 0.016942 | 0.02253 | 0.01951 | 0.032738 | 0.00086506 | -0.067294
a11 | 0.14522 | -0.060201 | -0.053174 | -0.058596 | -0.048936 | -0.047522 | -0.041093 | -0.037494 | -0.038178 | -0.041078 | -0.04416 | -0.045573 | -0.037971 | -0.031765 | -0.036726 | -0.041826 | -0.042996 | -0.045787 | -0.035601 | -0.027749 |  | a11 | 0.1956 | -0.01044 | -0.0052628 | -0.0031232 | -0.0062861 | 0.0030742 | 0.016108 | 0.0083247 | 0.015879 | 0.0072037 | 0.0080637 | 0.0098469 | 0.014661 | 0.018307 | 0.0030996 | 0.0098693 | 0.00017226 | 0.011151 | 0.011837 | 0.010488
a12 | 0.016676 | 0.00078386 | 0.0013538 | -0.00082069 | -0.0022598 | -0.008085 | -0.0082018 | -0.006722 | -0.0074432 | -0.0036683 | -0.0074512 | -0.0053141 | -0.003888 | -0.006265 | -0.0063074 | -0.0083183 | -0.0060923 | -0.0050784 | -0.0034206 | -0.00027188 |  | a12 | 0.00031826 | -0.0020829 | -0.0023549 | -0.0072982 | -0.0089351 | -0.010094 | -0.0084474 | -0.0053354 | -0.0067206 | -0.0074798 | -0.0047731 | 5.2801e-05 | -0.0068214 | -0.0060565 | -0.0072528 | -0.0072909 | -0.0025045 | -0.0026769 | -0.0051218 | -0.0098735
a13 | -0.019085 | -0.002172 | -0.0021642 | -0.0009208 | -0.0012621 | -0.0041748 | -0.003083 | -0.003332 | 8.7055e-05 | -0.00029776 | -0.00034186 | -0.0026993 | -0.00030963 | 0.0029146 | -0.0010647 | 0.00088045 | -0.0011435 | -0.001116 | 0.0018384 | -0.0015514 |  | a13 | -0.016883 | 0.0013099 | 0.0023286 | 0.00074258 | 0.004084 | 0.00080206 | 0.0017953 | 0.003912 | 0.0033105 | 0.0029258 | 0.0017698 | 0.0040028 | 0.0041483 | 0.0048306 | 0.0037261 | 0.0051043 | 0.0067127 | 0.0070273 | 0.0091267 | 0.00053348
a14 | 0.012688 | -0.031903 | -0.029606 | -0.032588 | -0.032357 | -0.033128 | -0.032983 | -0.038819 | -0.033389 | -0.034012 | -0.028505 | -0.030248 | -0.031784 | -0.028982 | -0.030131 | -0.029302 | -0.032647 | -0.032598 | -0.030692 | -0.014736 |  | a14 | 0.010795 | -0.028485 | -0.029557 | -0.032309 | -0.030537 | -0.035784 | -0.034362 | -0.034274 | -0.033643 | -0.036264 | -0.036717 | -0.028765 | -0.034152 | -0.031063 | -0.029115 | -0.025909 | -0.029635 | -0.029048 | -0.032013 | -0.019377
a15 | 0.013961 | 0.00053542 | -0.003702 | -0.0056834 | -0.0061719 | 0.0010266 | 0.00042287 | 0.0031618 | 0.0048966 | 0.0046609 | 0.002053 | -0.0048797 | 0.0053373 | 0.0060355 | 0.00089014 | -0.0029739 | 0.0032745 | -0.0031023 | -0.0070258 | 0.00011768 |  | a15 | 0.0075096 | 0.0070923 | -0.00015008 | 0.00063077 | -0.0023784 | 0.0028741 | 0.0026893 | 0.0063235 | 0.0074856 | 0.0010849 | -0.0016349 | -0.0041501 | -0.0013211 | -0.010287 | -0.012964 | -0.016544 | -0.0020376 | -0.009085 | -0.0091796 | -0.0064875
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.15705 | 0.029885 | -0.055528 | -0.088773 | -0.068703 | 0.039362 | 0.026294 | 0.06406 | 0.057311 | 0.046063 | 0.021305 | -0.060797 | 0.063595 | 0.095933 | -0.0087884 | -0.078144 | 0.027494 | -0.037817 | -0.11514 | -0.018552 |  | Dx (mm) | 0.14919 | 0.14352 | 0.026859 | 0.030888 | -0.0023445 | 0.081003 | 0.078418 | 0.11659 | 0.13829 | 0.078452 | 0.044371 | 0.0037834 | 0.022382 | -0.16904 | -0.15796 | -0.22363 | -0.031566 | -0.13102 | -0.093009 | -0.036052
Dy (mm) | 0.18676 | 0.028085 | 0.053032 | -0.016707 | 0.021932 | -0.050831 | -0.068831 | -0.082932 | -0.052811 | -0.035716 | 0.012648 | -0.012008 | -0.010244 | 0.040854 | 0.03408 | -0.0047313 | -0.031229 | 0.030837 | 0.050817 | -0.036962 |  | Dy (mm) | 0.03067 | 0.042457 | 0.0507 | -0.032447 | -0.0067377 | -0.034592 | -0.0013746 | -0.040336 | -0.049096 | -0.06545 | -0.053908 | 0.042 | -0.052381 | 0.049875 | 0.038561 | 0.05409 | 0.044523 | 0.076334 | 0.0034679 | -0.1744
Mag. len. (m) | 14.4222 | [20 empty] | Mag. len. (m) | 14.4203
TF (mT/KA) | 596.2092 | [20 empty] | TF (mT/KA) | 596.1396
Coil pos. (mm) | 0.019501 | [20 empty] | Coil pos. (mm) | 0.02093
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0094441 | [20 empty] | Twist (m2 rad) | 0.0094987

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 596.2092 | 0.14857 | -0.0024774 |  | -0.54731 | 3.2988 |  |  | 596.1396 | 0.14709 | -0.0030659 |  | -1.0381 | 3.537 |  | 0 | 4
2 | -0.41374 |  | -0.34708 | 0.55755 | 0.71878 |  | -0.28325 | 0.60517 | -0.017473 |  | 0.076188 | 0.37787 | 0.37854 |  | -0.37374 | 0.95998 |  | 0 | 5
3 | -1.5387 |  | -2.9103 | 1.0506 | 0.95961 |  | 1.0556 | 0.24772 | -1.9082 |  | -3.3127 | 1.0109 | 0.78661 |  | 0.8266 | 0.25608 |  | 0 | 6
4 | 0.053085 |  | 0.079296 | 0.22586 | -0.40149 |  | -0.55132 | 0.29418 | -0.10546 |  | -0.051036 | 0.10572 | -0.032356 |  | -0.15508 | 0.31027 |  | 0 | 7
5 | -0.6831 |  | -0.41735 | 0.26401 | 0.22966 |  | 0.16225 | 0.14709 | -0.56084 |  | -0.27644 | 0.31055 | 0.16006 |  | 0.10865 | 0.16247 |  | 0 | 8
6 | -0.0022512 |  | 0.0036241 | 0.065584 | 0.045099 |  | 0.034137 | 0.058625 | -0.034665 |  | -0.022268 | 0.087338 | -0.012779 |  | -0.031722 | 0.07812 |  | 0 | 9
7 | 1.1697 |  | 1.1226 | 0.056454 | 0.024717 |  | -0.021294 | 0.043359 | 1.1813 |  | 1.1372 | 0.051969 | 0.06823 |  | 0.018785 | 0.053666 |  | 0 | 10
8 | -0.020284 |  | -0.016944 | 0.023775 | -0.010769 |  | -0.013348 | 0.016932 | -0.00078813 |  | 0.001057 | 0.02399 | -0.013708 |  | -0.015365 | 0.037298 |  | 0 | 11
9 | 0.46722 |  | 0.47209 | 0.014376 | -0.029735 |  | -0.021147 | 0.016739 | 0.47644 |  | 0.48023 | 0.016853 | 0.0035324 |  | 0.01279 | 0.014364 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.7267 |  | 0.73396 | 0.0069019 | -0.0376 |  | -0.04376 | 0.007787 | 0.74198 |  | 0.74962 | 0.0061145 | 0.012004 |  | 0.0062997 | 0.0084564 |  | 0 | 14
12 | -0.00025764 |  | -0.0025316 | 0.0021015 | -0.00041711 |  | -0.0047128 | 0.0025583 | 0.00015381 |  | 0.0015913 | 0.0019524 | -0.00055093 |  | -0.005754 | 0.0021865 |  | 0 | 15
13 | 0.0058354 |  | 0.057713 | 0.0032919 | -0.00016295 |  | -0.0010038 | 0.0022156 | 0.0059947 |  | 0.058996 | 0.0043458 | 0.00031734 |  | 0.0039622 | 0.0020688 |  | 0 | 16
14 | 8.0669e-05 |  | 0.00069709 | 0.00080971 | -0.00299 |  | -0.031766 | 0.0016708 | -9.6148e-05 |  | -0.0010948 | 0.00092766 | -0.0030538 |  | -0.032426 | 0.0013828 |  | 0 | 17
15 | 0.0054333 |  | 0.056486 | 0.0024012 | 3.6209e-05 |  | -6.9134e-05 | 0.0042944 | 0.0057941 |  | 0.059793 | 0.0030152 | -0.00021551 |  | -0.0023085 | 0.006917 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.54323 | -0.50703
3 | [12 empty] | 3 | 9 | -0.67305 | -0.6757
4 | [12 empty] | 4 | 16 | -1.3557 | -1.7086
5 | [12 empty] | 5 | 25 | -1.6243 | -1.4933
6 | [12 empty] | 6 | 36 | -2.7805 | -2.4937
7 | [12 empty] | 7 | 49 | -3.0063 | -2.941
8 | [12 empty] | 8 | 64 | -3.9089 | -3.5095
9 | [12 empty] | 9 | 81 | -4.1661 | -4.1631
10 | [12 empty] | sigma (mm) |  | 0.019501 | 0.02093
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 347.2305 | 596.5469 | 596.1428 | 596.2518 | 596.1345 | 596.2117 | 596.209 | 596.1803 | 596.0216 | 596.1642 | 596.1996 | 596.1322 | 596.1086 | 596.1495 | 596.1996 | 596.1904 | 596.2864 | 596.0347 | 596.6026 | 372.3875
2 | -3.2252 | -0.8047 | -0.23893 | -0.68701 | -0.8438 | -1.2159 | -0.52787 | -0.21757 | 0.57327 | 0.15551 | 0.4609 | 0.27876 | 0.11486 | -0.75715 | -1.3328 | -0.4193 | -0.45241 | 0.23214 | -0.56529 | 0.28831 | -0.41374
3 | 38.2038 | -2.3732 | -1.9326 | -3.9829 | -1.865 | -3.3333 | -2.3087 | -2.7086 | -2.7085 | -2.3009 | -4.213 | -3.6911 | -3.3442 | -5.3927 | -4.3097 | -2.2678 | -1.7938 | -2.1363 | -1.7233 | 0.931 | -1.5387
4 | -0.941 | -0.28745 | -0.11808 | -0.00048836 | -0.038708 | -0.17719 | 0.09977 | 0.25715 | 0.28102 | 0.23892 | 0.48053 | 0.54687 | 0.25621 | -0.030516 | -0.11507 | 0.058868 | 0.085529 | 0.022915 | -0.13294 | 0.22539 | 0.053085
5 | -5.5187 | -1.0086 | -0.94263 | -0.29187 | -0.48918 | -0.25771 | -0.16153 | -0.18049 | -0.34935 | -0.23917 | -0.050888 | -0.31583 | -0.22216 | -0.22951 | -0.43866 | -0.71877 | -0.5606 | -0.54914 | -0.50621 | -3.832 | -0.6831
6 | -0.20892 | -0.054408 | -0.016219 | 0.048321 | 0.021352 | 0.14144 | 0.025185 | -0.01528 | -0.081318 | -0.045631 | -0.028 | -0.088239 | -0.02103 | 0.091746 | 0.13208 | -0.013191 | 0.0049795 | -0.050173 | 0.01362 | 0.021083 | -0.0022512
7 | 2.9445 | 0.97826 | 1.0603 | 1.1909 | 1.1078 | 1.1251 | 1.145 | 1.0658 | 1.093 | 1.0815 | 1.2215 | 1.1667 | 1.1443 | 1.1711 | 1.1229 | 1.1754 | 1.1396 | 1.12 | 1.0971 | 0.87391 | 1.1697
8 | -0.1419 | -0.034195 | -0.0042653 | 0.0013442 | -0.014471 | -0.047388 | -0.012531 | -0.023474 | 0.027379 | 0.0184 | 0.0010257 | -0.024436 | -0.028748 | -0.048017 | -0.047336 | 0.014281 | -0.0032039 | -0.038705 | -0.040661 | -0.0030677 | -0.020284
9 | 0.35099 | 0.47134 | 0.45607 | 0.4441 | 0.49705 | 0.45844 | 0.48779 | 0.47082 | 0.46902 | 0.46887 | 0.46011 | 0.46702 | 0.46631 | 0.4828 | 0.47271 | 0.50451 | 0.47347 | 0.47045 | 0.47679 | 0.43522 | 0.46722
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.60596 | 0.72612 | 0.72392 | 0.73732 | 0.73075 | 0.73566 | 0.73179 | 0.73892 | 0.74039 | 0.7344 | 0.74233 | 0.73387 | 0.73935 | 0.74137 | 0.74727 | 0.73108 | 0.72474 | 0.72518 | 0.72684 | 0.62998 | 0.7267
12 | -0.028494 | -0.022921 | -0.017837 | -0.052797 | -0.028056 | -0.018568 | -0.044707 | -0.024152 | -0.012152 | 0.010786 | 0.019979 | -0.057854 | -0.004968 | -0.0546 | -0.033345 | -0.023568 | -0.036269 | -0.01718 | -0.037487 | -0.036082 | -0.0025764 |  | ten times!
13 | 0.90282 | 0.63825 | 0.62823 | 0.57687 | 0.55471 | 0.55237 | 0.5439 | 0.50818 | 0.54445 | 0.57562 | 0.57755 | 0.57186 | 0.59251 | 0.54579 | 0.57799 | 0.59571 | 0.61367 | 0.61015 | 0.58057 | 0.47036 | 0.058354 |  | ten times!
14 | 0.033272 | -0.0058731 | 0.00083664 | 3.9708e-05 | -0.0027906 | 0.016305 | 0.0026448 | -0.0097973 | 0.002305 | 0.012321 | 0.015028 | 0.010121 | 0.01209 | 0.012635 | 0.018347 | 0.011826 | 0.013303 | 0.0092166 | 0.0069197 | 0.016161 | 0.00080669 |  | ten times!
15 | 0.043102 | 0.5516 | 0.52915 | 0.56111 | 0.563 | 0.58716 | 0.58468 | 0.61116 | 0.5905 | 0.5979 | 0.55741 | 0.55128 | 0.56725 | 0.5676 | 0.55629 | 0.53543 | 0.578 | 0.55956 | 0.51842 | 0.38934 | 0.054333 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 26.229 | -3.3716 | -1.1853 | -2.9389 | -0.86821 | -3.2253 | 2.338 | 5.099 | 0.11915 | 2.1722 | -5.0009 | -0.91861 | 4.4626 | 3.4995 | 2.2972 | -0.88509 | -0.89292 | -4.5426 | -6.0096 | -8.7509 | -0.0024774
2 | 17.1799 | -0.18705 | -0.21063 | 0.79719 | -0.6643 | -0.83176 | -0.90167 | -0.29536 | -0.58402 | 0.23526 | -0.80233 | 0.18384 | -1.1003 | -0.87115 | -0.88256 | -0.2286 | 0.58316 | -0.16314 | 0.82491 | 14.2452 | 0.71878
3 | -1.0961 | 1.4194 | 0.81151 | 1.2489 | 1.0794 | 0.83513 | 1.2169 | 0.47587 | 1.0839 | 1.1626 | 0.86965 | 0.73802 | 1.0411 | 1.5329 | 1.1268 | 1.099 | 1.1737 | 1.0966 | 0.98989 | 0.10878 | 0.95961
4 | 2.8766 | -0.27412 | -0.11676 | -0.8686 | -0.51972 | -0.63067 | -0.78242 | -0.94746 | -0.38748 | -0.55193 | 0.052058 | -0.81228 | -0.60315 | -0.070788 | -0.4539 | -0.81182 | -0.76171 | -0.56159 | -0.82134 | 0.85993 | -0.40149
5 | 2.5162 | 0.14261 | 0.062605 | 0.072969 | 0.20261 | 0.15685 | 0.20909 | 0.033971 | 0.34692 | 0.2802 | 0.1831 | -0.064591 | -0.097761 | 0.43174 | 0.29944 | 0.26442 | 0.087511 | -0.01008 | 0.31882 | 0.040177 | 0.22966
6 | 0.27714 | 0.12312 | 0.067151 | 0.096502 | 0.096081 | 0.091437 | -0.05901 | 0.042208 | 0.015921 | 0.1051 | -0.061361 | -0.029403 | 0.01768 | -0.049309 | 0.027769 | -0.015244 | 0.062172 | 0.058015 | 0.025623 | 0.14457 | 0.045099
7 | 1.4679 | -0.020695 | -0.057059 | -0.11013 | 0.016849 | 0.031512 | -0.01452 | -0.0098802 | -0.064189 | -0.035159 | 0.014829 | 0.011612 | 0.066305 | -0.024926 | 0.0015092 | -0.079163 | -0.047385 | -0.058853 | -0.0039394 | 0.0049847 | 0.024717
8 | 0.0049485 | -0.024763 | -0.037594 | -0.0020811 | -0.00067981 | -0.013342 | -0.028255 | -0.05172 | -0.018883 | 0.011168 | -0.017786 | 8.696e-05 | -0.024404 | 0.014394 | -0.0072046 | -0.0018732 | 0.0011024 | -0.018091 | -0.02034 | 0.048903 | -0.010769
9 | -0.26074 | -0.039775 | -0.019368 | -0.016319 | -0.062344 | -0.040187 | -0.02628 | -0.027711 | -0.017475 | -0.031528 | -0.021712 | -0.032267 | 0.0045839 | 0.0038477 | -0.01346 | -0.012553 | -0.018986 | -0.0090318 | -7.4118e-05 | -0.06184 | -0.029735
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14237 | -0.0602 | -0.052981 | -0.058739 | -0.048959 | -0.047333 | -0.041116 | -0.037238 | -0.03787 | -0.040901 | -0.044081 | -0.045846 | -0.037796 | -0.03133 | -0.036681 | -0.042264 | -0.042938 | -0.045844 | -0.035571 | -0.027854 | -0.0376
12 | 0.071014 | -0.0044956 | -0.011312 | -0.0019964 | -0.031909 | -0.060294 | -0.055847 | -0.037335 | -0.054621 | -0.022276 | -0.079646 | -0.049469 | -0.03446 | -0.080742 | -0.077231 | -0.080707 | -0.047332 | -0.064536 | -0.054089 | 0.0092376 | -0.0041711 |  | ten times!
13 | -0.20857 | -0.014592 | -0.033332 | -0.031864 | -0.029058 | -0.030386 | -0.02338 | -0.012744 | 0.016808 | 0.0099967 | 0.00098745 | -0.041221 | 0.012692 | 0.048676 | -0.013039 | -0.008788 | -0.0039207 | -0.020412 | -0.007113 | -0.017315 | -0.0016295 |  | ten times!
14 | 0.1022 | -0.33192 | -0.32086 | -0.32231 | -0.33723 | -0.30703 | -0.29678 | -0.34812 | -0.31052 | -0.32431 | -0.29121 | -0.29947 | -0.31585 | -0.31369 | -0.31686 | -0.29285 | -0.31235 | -0.34115 | -0.33527 | -0.13549 | -0.0299 |  | ten times!
15 | 0.13961 | 0.0053542 | -0.03702 | -0.056834 | -0.061719 | 0.010266 | 0.0042287 | 0.031618 | 0.048966 | 0.046609 | 0.02053 | -0.048797 | 0.053373 | 0.060355 | 0.0089014 | -0.029739 | 0.032745 | -0.031023 | -0.070258 | 0.0011768 | 0.00036209 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 337.5928 | 596.4091 | 595.9854 | 596.0952 | 595.9295 | 596.0315 | 596.0651 | 596.102 | 596.3428 | 596.0322 | 596.0593 | 596.2153 | 596.0338 | 596.1121 | 596.0861 | 596.1117 | 596.1287 | 596.3407 | 596.432 | 373.2635
2 | -3.1612 | 0.13798 | 0.2818 | 0.27295 | -0.33874 | 0.031796 | 0.16257 | -0.2275 | -0.23829 | -0.28268 | 0.28791 | 0.2321 | 0.016166 | 1.037 | 0.64201 | 0.072408 | -0.21466 | 0.085801 | -0.58728 | 0.13381 | -0.017473
3 | 40.9153 | -3.1235 | -2.2457 | -4.1005 | -2.1987 | -4.454 | -3.1134 | -2.5059 | -2.6188 | -2.3573 | -3.8828 | -3.7561 | -3.9913 | -5.9897 | -4.2798 | -3.2922 | -2.4791 | -2.9455 | -2.2944 | -0.26266 | -1.9082
4 | -1.6108 | 0.070312 | 0.059609 | -0.066125 | -0.18297 | -0.11386 | -0.2213 | -0.076465 | -0.052493 | -0.17129 | 0.17467 | 0.03128 | -0.02586 | -0.0080114 | -0.021216 | 0.04818 | -0.06851 | -0.085719 | -0.20888 | -0.30867 | -0.10546
5 | -5.9971 | -0.98253 | -0.80934 | 0.0015891 | -0.32929 | -0.024712 | 0.019633 | 0.046927 | -0.24215 | -0.29334 | 0.090277 | -0.14369 | -0.27262 | 0.050443 | -0.12023 | -0.32629 | -0.64029 | -0.5277 | -0.47254 | -3.8194 | -0.56084
6 | -0.35566 | 0.00030249 | -0.010807 | 0.077986 | 0.060435 | 0.041089 | 0.11648 | -0.03217 | 0.050955 | 0.033735 | -0.12352 | -0.081596 | -0.048789 | -0.2034 | -0.18951 | -0.0022223 | 0.024213 | -0.051685 | -0.062322 | -0.10067 | -0.034665
7 | 2.8464 | 1.0193 | 1.0763 | 1.2418 | 1.1999 | 1.1655 | 1.1361 | 1.1374 | 1.1643 | 1.1217 | 1.1253 | 1.1581 | 1.1697 | 1.106 | 1.0797 | 1.2051 | 1.1295 | 1.1271 | 1.1062 | 0.94519 | 1.1813
8 | -0.059332 | -0.0019121 | -0.0026795 | -0.007926 | -0.0048045 | -0.01899 | -0.024946 | -0.022828 | -0.023116 | -0.019752 | -0.023346 | -0.0064793 | 0.014611 | 0.044211 | 0.026312 | 0.0089589 | 0.046713 | 0.039365 | -0.0043638 | -0.00090097 | -0.00078813
9 | 0.38659 | 0.47909 | 0.47329 | 0.46856 | 0.51485 | 0.47155 | 0.49967 | 0.49187 | 0.5009 | 0.50281 | 0.47277 | 0.47142 | 0.4595 | 0.45401 | 0.45858 | 0.47002 | 0.48255 | 0.48905 | 0.48367 | 0.4486 | 0.47644
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.59552 | 0.74689 | 0.74172 | 0.74897 | 0.73826 | 0.75437 | 0.75393 | 0.75263 | 0.75029 | 0.74603 | 0.75957 | 0.75176 | 0.74961 | 0.75893 | 0.75878 | 0.74718 | 0.74503 | 0.7478 | 0.74135 | 0.65489 | 0.74198
12 | 0.011331 | 0.021828 | 0.014488 | 0.02523 | 0.024695 | 0.038234 | 0.018402 | 0.0050096 | 0.0045019 | 0.0023608 | -0.0077083 | 0.0060058 | -0.015388 | 0.067347 | 0.0052493 | -0.002828 | 0.0375 | 0.027199 | 0.014313 | 0.0037389 | 0.0015381 |  | ten times!
13 | 0.96285 | 0.68206 | 0.66583 | 0.59512 | 0.6011 | 0.56107 | 0.53757 | 0.54826 | 0.55964 | 0.55099 | 0.57425 | 0.58275 | 0.63987 | 0.51997 | 0.57005 | 0.58779 | 0.61646 | 0.61984 | 0.6066 | 0.54429 | 0.059947 |  | ten times!
14 | 0.01309 | -0.0116 | -0.015472 | -0.0143 | -0.006249 | -0.0029883 | -0.011487 | -0.0051217 | 0.0010264 | -0.01775 | -0.01471 | -0.028294 | -0.011516 | 0.01259 | -0.024186 | -0.019108 | -0.0089752 | -0.0084702 | -0.010454 | 0.0081729 | -0.00096148 |  | ten times!
15 | 0.093028 | 0.56082 | 0.57005 | 0.61315 | 0.58378 | 0.61307 | 0.58863 | 0.62296 | 0.62033 | 0.64639 | 0.65221 | 0.57946 | 0.63287 | 0.6019 | 0.58296 | 0.5436 | 0.58568 | 0.56311 | 0.60172 | 0.48699 | 0.057941 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 43.9679 | -4.9467 | -1.4608 | -4.6851 | -2.9028 | -3.8317 | -0.73839 | 3.0542 | 0.64244 | -1.3099 | -6.3276 | 2.3524 | 3.628 | 6.384 | 2.8455 | -1.5865 | -2.503 | -2.0086 | -5.2918 | -10.012 | -0.0030659
2 | 16.946 | 0.66958 | -0.63702 | -1.3713 | -1.7614 | -1.4091 | -1.6209 | -1.2062 | -0.87474 | -0.36383 | 0.046215 | 0.95748 | 0.80054 | -0.91245 | -0.81764 | 0.63787 | 0.67902 | -0.59197 | 1.0485 | 7.0177 | 0.37854
3 | -0.45315 | 1.0015 | 1.0712 | 1.1976 | 1.0351 | 0.86086 | 0.85564 | 0.59465 | 0.94919 | 0.48068 | 0.91191 | 0.52367 | 0.70668 | 0.62472 | 0.81332 | 0.63077 | 0.59575 | 0.60199 | 1.4235 | 0.75806 | 0.78661
4 | 3.448 | 0.10663 | 0.037189 | 0.0010394 | -0.32576 | -0.64646 | -0.629 | -0.30456 | 0.014068 | -0.2056 | -0.53601 | -0.41054 | -0.45017 | 0.41979 | 0.27278 | -0.12883 | 0.16111 | 0.050922 | -0.21808 | 0.34743 | -0.032356
5 | 2.2683 | 0.21964 | 0.046432 | 0.091222 | 0.14529 | 0.13558 | 0.23551 | 0.085961 | 0.31257 | 0.33689 | 0.25777 | 0.038856 | -0.11214 | -0.26415 | -0.051337 | 0.20121 | -0.013447 | -0.018433 | 0.30829 | -0.26915 | 0.16006
6 | 0.37767 | -0.014482 | -0.0041382 | -0.0032839 | -0.076981 | -0.11308 | -0.16928 | 0.059129 | 0.03387 | -0.050757 | 0.030632 | -0.084166 | 0.0040405 | -0.072552 | -0.15861 | -0.12553 | 0.060738 | 0.020066 | 0.093392 | 0.17844 | -0.012779
7 | 1.7032 | 0.0055254 | 0.00624 | -0.018318 | 0.0011314 | -0.010434 | 0.068534 | -0.0022066 | 0.020459 | 0.014942 | 0.061611 | 0.029685 | 0.11416 | 0.10281 | -0.032783 | -0.084348 | -0.064637 | 0.041687 | 0.084075 | 0.01089 | 0.06823
8 | -0.01076 | -0.036001 | -0.015732 | -0.047702 | -0.047559 | -0.081236 | -0.091261 | -0.0057091 | -0.023226 | -0.047613 | 0.0009553 | 0.040359 | 0.02147 | -0.00083673 | -0.0088438 | 0.036334 | 0.020177 | -0.0023334 | 0.012188 | 0.03119 | -0.013708
9 | -0.24053 | 0.018708 | 0.014766 | 0.0088769 | 0.0039403 | 0.0058955 | 0.0047396 | 0.031635 | 0.022307 | -0.01382 | 0.010132 | 0.013526 | 0.043384 | 0.0033195 | -0.014336 | 0.014026 | 0.023402 | 0.011938 | 0.02778 | -0.041895 | 0.0035324
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.19534 | -0.0105 | -0.0053088 | -0.0028971 | -0.0062893 | 0.0037602 | 0.016541 | 0.0088598 | 0.01667 | 0.0077244 | 0.0082374 | 0.0098294 | 0.014746 | 0.01764 | 0.0025119 | 0.0091723 | 5.9406e-05 | 0.011103 | 0.011534 | 0.010159 | 0.012004
12 | 0.00011577 | -0.043045 | -0.048228 | -0.059697 | -0.086432 | -0.087868 | -0.084945 | -0.04112 | -0.051207 | -0.051733 | -0.026999 | -0.017127 | -0.045716 | -0.073828 | -0.084439 | -0.088434 | -0.043259 | -0.054925 | -0.046723 | -0.034722 | -0.0055093 |  | ten times!
13 | -0.18167 | 0.042665 | 0.029465 | 0.015094 | 0.040253 | 0.031211 | 0.038587 | 0.071429 | 0.071488 | 0.052244 | 0.030503 | 0.041665 | 0.047311 | 0.010706 | 0.0049195 | 0.011362 | 0.060772 | 0.044877 | 0.068647 | -0.0020884 | 0.0031734 |  | ten times!
14 | 0.096373 | -0.31284 | -0.31934 | -0.30687 | -0.30217 | -0.34229 | -0.34469 | -0.32812 | -0.31988 | -0.3285 | -0.33762 | -0.30756 | -0.31398 | -0.34967 | -0.32653 | -0.31377 | -0.31836 | -0.33568 | -0.32888 | -0.12576 | -0.030538 |  | ten times!
15 | 0.075096 | 0.070923 | -0.0015008 | 0.0063077 | -0.023784 | 0.028741 | 0.026893 | 0.063235 | 0.074856 | 0.010849 | -0.016349 | -0.041501 | -0.013211 | -0.10287 | -0.12964 | -0.16544 | -0.020376 | -0.09085 | -0.091796 | -0.064875 | -0.0021551 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00056634 | -0.00011144 | 7.1377e-05 | -0.00012537 | 4.1186e-06 | -4.1e-07 | -4.8547e-05 | -0.00031473 | -7.5551e-05 | -1.6176e-05 | -0.00012922 | -0.00016881 | -0.00010021 | -1.6176e-05 | -3.1607e-05 | 0.00012941 | -0.00029276 | 0.00065976 |  | 596.2092
|  | 0.0004521 | -0.00025864 | -7.4451e-05 | -0.00035241 | -0.00018131 | -0.00012494 | -6.3045e-05 | 0.00034089 | -0.00018013 | -0.00013467 | 0.00012701 | -0.00017745 | -4.6102e-05 | -8.9716e-05 | -4.6773e-05 | -1.8256e-05 | 0.00033737 | 0.00049052 |  | 596.1396
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3964.5925
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 596.1636 | -4175.5833 | 6.4302 | -0.34816 | 1.4802 | -0.48739 | 0.80756 | 0.76227 | 0.28086 | -2.3812 | 0.010798 | 0.60459 | -0.52597 | -0.92183 | -0.23578 | 0.60459 | 0.45027 | 2.0606 | -2.1614 | 7.3645 | -3753.6017
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.054731 | 2.6776 | -0.28243 | -0.063796 | -0.23916 | -0.03209 | -0.2678 | 0.28853 | 0.56463 | 0.066646 | 0.27195 | -0.44536 | -0.03713 | 0.50099 | 0.40468 | 0.28445 | -0.033778 | -0.034561 | -0.39953 | -0.54623 | -0.82036 |  | 0
db2 | -0.34708 | -3.2252 | -0.45763 | 0.10814 | -0.33993 | -0.49672 | -0.86885 | -0.18079 | 0.1295 | 0.92034 | 0.50259 | 0.80797 | 0.62583 | 0.46194 | -0.41007 | -0.98577 | -0.072225 | -0.10534 | 0.57921 | -0.21821 | 0.28831 |  | 0
db3 | -2.9115 | 38.2038 | 0.53709 | 0.97766 | -1.0726 | 1.0453 | -0.42294 | 0.60157 | 0.20174 | 0.20179 | 0.60939 | -1.3027 | -0.78083 | -0.43385 | -2.4824 | -1.3994 | 0.6425 | 1.1165 | 0.77403 | 1.187 | 0.931 |  | 0
db4 | 0.079296 | -0.941 | -0.36675 | -0.19737 | -0.079785 | -0.118 | -0.25648 | 0.020474 | 0.17785 | 0.20172 | 0.15962 | 0.40123 | 0.46757 | 0.17691 | -0.10981 | -0.19437 | -0.020429 | 0.0062325 | -0.056381 | -0.21224 | 0.22539 |  | 0
db5 | -0.41701 | -5.5187 | -0.59128 | -0.52528 | 0.12548 | -0.071834 | 0.15964 | 0.25583 | 0.23687 | 0.068003 | 0.17818 | 0.36646 | 0.10152 | 0.19519 | 0.18784 | -0.021304 | -0.30142 | -0.14325 | -0.13179 | -0.088856 | -3.832 |  | -3.8549e-17
db6 | 0.0036241 | -0.20892 | -0.058032 | -0.019843 | 0.044697 | 0.017728 | 0.13782 | 0.021561 | -0.018904 | -0.084943 | -0.049255 | -0.031624 | -0.091863 | -0.024654 | 0.088122 | 0.12845 | -0.016815 | 0.0013554 | -0.053797 | 0.009996 | 0.021083 |  | 0
db7 | 1.1226 | 2.9445 | -0.1443 | -0.062233 | 0.068335 | -0.01474 | 0.002511 | 0.022404 | -0.056771 | -0.029594 | -0.041097 | 0.098945 | 0.044119 | 0.02174 | 0.048514 | 0.00037544 | 0.052835 | 0.017033 | -0.0025798 | -0.025497 | 0.87391 |  | -8.6351e-17
db8 | -0.016944 | -0.1419 | -0.01725 | 0.012679 | 0.018289 | 0.0024736 | -0.030443 | 0.0044137 | -0.00653 | 0.044324 | 0.035345 | 0.01797 | -0.0074917 | -0.011804 | -0.031073 | -0.030391 | 0.031225 | 0.01374 | -0.021761 | -0.023716 | -0.0030677 |  | 0
db9 | 0.47253 | 0.35099 | -0.00075069 | -0.016022 | -0.027994 | 0.024957 | -0.013654 | 0.015697 | -0.0012743 | -0.0030774 | -0.0032243 | -0.011979 | -0.0050708 | -0.0057864 | 0.010712 | 0.0006159 | 0.032417 | 0.0013814 | -0.0016424 | 0.0046939 | 0.43522 |  | -3.7007e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.73403 | 0.60596 | -0.0078402 | -0.010044 | 0.0033557 | -0.0032117 | 0.0017024 | -0.0021711 | 0.0049568 | 0.0064336 | 0.00043826 | 0.0083724 | -9.12e-05 | 0.0053868 | 0.0074104 | 0.013305 | -0.0028857 | -0.0092219 | -0.0087765 | -0.0071185 | 0.62998 |  | 5.5511e-17
db12 | -0.0025316 | -0.0028494 | 0.00023957 | 0.00074795 | -0.002748 | -0.000274 | 0.00067484 | -0.0019391 | 0.00011644 | 0.0013164 | 0.0036103 | 0.0045296 | -0.0032537 | 0.0020348 | -0.0029284 | -0.00080283 | 0.00017486 | -0.0010952 | 0.00081364 | -0.0012171 | -0.0036082 |  | 0
db13 | 0.057713 | 0.090282 | 0.0061118 | 0.0051096 | -2.6161e-05 | -0.0022425 | -0.0024763 | -0.0033234 | -0.0068957 | -0.0032678 | -0.00015152 | 4.1462e-05 | -0.00052744 | 0.0015378 | -0.003134 | 8.6302e-05 | 0.0018577 | 0.0036539 | 0.0033022 | 0.00034413 | 0.047036 |  | -2.6985e-18
db14 | 0.00069709 | 0.0033272 | -0.0012844 | -0.00061342 | -0.00069312 | -0.00097615 | 0.0009334 | -0.00043261 | -0.0016768 | -0.00046659 | 0.00053498 | 0.00080572 | 0.00031499 | 0.00051189 | 0.00056638 | 0.0011376 | 0.00048551 | 0.00063322 | 0.00022457 | -5.1146e-06 | 0.0016161 |  | -8.4327e-20
db15 | 0.056486 | 0.0043102 | -0.0013263 | -0.0035715 | -0.00037469 | -0.00018587 | 0.0022299 | 0.0019824 | 0.0046299 | 0.0025636 | 0.0033041 | -0.00074516 | -0.001358 | 0.00023904 | 0.00027406 | -0.00085756 | -0.002943 | 0.0013135 | -0.00053034 | -0.0046441 | 0.038934 |  | 4.6259e-18
da2 | -0.28325 | 17.1799 | 0.096204 | 0.072618 | 1.0804 | -0.38105 | -0.54851 | -0.61842 | -0.01211 | -0.30077 | 0.51851 | -0.51908 | 0.4671 | -0.81707 | -0.58789 | -0.5993 | 0.054653 | 0.86641 | 0.12011 | 1.1082 | 14.2452 |  | 0
da3 | 1.0556 | -1.0961 | 0.36379 | -0.24412 | 0.19325 | 0.023759 | -0.22051 | 0.1613 | -0.57976 | 0.028279 | 0.10693 | -0.18598 | -0.31762 | -0.014508 | 0.47723 | 0.071211 | 0.043407 | 0.1181 | 0.040973 | -0.06574 | 0.10878 |  | 4.9343e-17
da4 | -0.55132 | 2.8766 | 0.2772 | 0.43455 | -0.31728 | 0.031591 | -0.079351 | -0.23111 | -0.39615 | 0.16383 | -0.00061208 | 0.60337 | -0.26097 | -0.051833 | 0.48053 | 0.097418 | -0.2605 | -0.2104 | -0.010272 | -0.27002 | 0.85993 |  | 7.7099e-17
da5 | 0.16225 | 2.5162 | -0.019632 | -0.099641 | -0.089276 | 0.040362 | -0.0054001 | 0.04684 | -0.12827 | 0.18468 | 0.11795 | 0.020859 | -0.22684 | -0.26001 | 0.2695 | 0.1372 | 0.10217 | -0.074734 | -0.17233 | 0.15657 | 0.040177 |  | 0
da6 | 0.034137 | 0.27714 | 0.088987 | 0.033015 | 0.062366 | 0.061944 | 0.0573 | -0.093146 | 0.0080715 | -0.018216 | 0.070965 | -0.095497 | -0.063539 | -0.016457 | -0.083446 | -0.0063676 | -0.049381 | 0.028036 | 0.023879 | -0.0085136 | 0.14457 |  | -5.0114e-18
da7 | -0.021294 | 1.4679 | 0.00059832 | -0.035765 | -0.08884 | 0.038143 | 0.052805 | 0.0067735 | 0.011413 | -0.042895 | -0.013865 | 0.036123 | 0.032905 | 0.087599 | -0.0036327 | 0.022803 | -0.057869 | -0.026091 | -0.03756 | 0.017354 | 0.0049847 |  | 0
da8 | -0.013348 | 0.0049485 | -0.011415 | -0.024246 | 0.011267 | 0.012668 | 6.4394e-06 | -0.014907 | -0.038372 | -0.0055345 | 0.024516 | -0.0044381 | 0.013435 | -0.011056 | 0.027742 | 0.0061435 | 0.011475 | 0.01445 | -0.004743 | -0.0069924 | 0.048903 |  | 0
da9 | -0.021147 | -0.26074 | -0.018628 | 0.0017785 | 0.0048275 | -0.041198 | -0.01904 | -0.0051334 | -0.0065647 | 0.0036717 | -0.010382 | -0.0005649 | -0.011121 | 0.025731 | 0.024994 | 0.007687 | 0.0085936 | 0.0021602 | 0.012115 | 0.021073 | -0.06184 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.04376 | 0.14237 | -0.01644 | -0.0092207 | -0.014979 | -0.0051981 | -0.0035725 | 0.0026443 | 0.0065224 | 0.0058906 | 0.002859 | -0.00032021 | -0.0020856 | 0.005964 | 0.012431 | 0.0070798 | 0.0014961 | 0.00082294 | -0.0020838 | 0.0081897 | -0.027854 |  | 0
da12 | -0.0047128 | 0.0071014 | 0.0042632 | 0.0035816 | 0.0045131 | 0.0015218 | -0.0013166 | -0.00087196 | 0.00097926 | -0.00074929 | 0.0024852 | -0.0032518 | -0.00023413 | 0.0012668 | -0.0033614 | -0.0030104 | -0.0033579 | -2.0441e-05 | -0.0017408 | -0.00069615 | 0.00092376 |  | 0
da13 | -0.0010038 | -0.020857 | -0.00045534 | -0.0023293 | -0.0021826 | -0.001902 | -0.0020348 | -0.0013342 | -0.00027052 | 0.0026846 | 0.0020035 | 0.0011026 | -0.0031183 | 0.002273 | 0.0058714 | -0.00030005 | 0.00012503 | 0.00061176 | -0.0010374 | 0.00029253 | -0.0017315 |  | -1.4456e-19
da14 | -0.031766 | 0.01022 | -0.0014266 | -0.00032055 | -0.00046599 | -0.0019572 | 0.0010626 | 0.0020872 | -0.0030461 | 0.00071392 | -0.00066517 | 0.0026442 | 0.0018181 | 0.00018052 | 0.0003965 | 7.939e-05 | 0.0024804 | 0.00053054 | -0.0023498 | -0.0017619 | -0.013549 |  | 3.4694e-18
da15 | -6.9134e-05 | 0.013961 | 0.00060456 | -0.0036329 | -0.0056143 | -0.0061028 | 0.0010957 | 0.000492 | 0.003231 | 0.0049657 | 0.00473 | 0.0021221 | -0.0048106 | 0.0054065 | 0.0061047 | 0.00095928 | -0.0029048 | 0.0033436 | -0.0030331 | -0.0069567 | 0.00011768 |  | 0
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -4037.4855
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.1045 | -4336.6839 | 5.1105 | -1.9973 | -0.15538 | -2.9351 | -1.224 | -0.66033 | -0.04131 | 3.9983 | -1.2122 | -0.75763 | 1.8594 | -1.1854 | 0.12812 | -0.30804 | 0.12141 | 0.4066 | 3.963 | 5.4946 | -3738.2871
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.10381 | 4.5006 | -0.39086 | -0.042263 | -0.36469 | -0.18647 | -0.27936 | 0.029973 | 0.40923 | 0.16806 | -0.027177 | -0.52895 | 0.33906 | 0.46662 | 0.74222 | 0.38837 | -0.054837 | -0.14649 | -0.097052 | -0.42537 | -0.89739 |  | 0
db2 | -0.076188 | -3.1612 | 0.061791 | 0.20561 | 0.19676 | -0.41492 | -0.044393 | 0.086383 | -0.30369 | -0.31448 | -0.35887 | 0.21173 | 0.15592 | -0.060023 | 0.96085 | 0.56582 | -0.0037808 | -0.29084 | 0.009613 | -0.66347 | 0.13381 |  | 0
db3 | -3.3142 | 40.9153 | 0.1892 | 1.067 | -0.7878 | 1.114 | -1.1413 | 0.19932 | 0.80681 | 0.69388 | 0.95543 | -0.57009 | -0.4434 | -0.67861 | -2.677 | -0.96706 | 0.020491 | 0.83362 | 0.36721 | 1.0183 | -0.26266 |  | -2.7139e-16
db4 | 0.051036 | -1.6108 | 0.12135 | 0.11064 | -0.015089 | -0.13193 | -0.062824 | -0.17027 | -0.02543 | -0.0014576 | -0.12025 | 0.22571 | 0.082316 | 0.025175 | 0.043024 | 0.02982 | 0.099215 | -0.017475 | -0.034684 | -0.15785 | -0.30867 |  | 0
db5 | -0.27362 | -5.9971 | -0.70609 | -0.53291 | 0.27803 | -0.052852 | 0.25172 | 0.29607 | 0.32336 | 0.034284 | -0.016909 | 0.36671 | 0.13275 | 0.0038207 | 0.32688 | 0.15621 | -0.049852 | -0.36385 | -0.25126 | -0.19611 | -3.8194 |  | 4.4717e-17
db6 | 0.022268 | -0.35566 | 0.02257 | 0.011461 | 0.10025 | 0.082703 | 0.063357 | 0.13874 | -0.0099023 | 0.073223 | 0.056003 | -0.10125 | -0.059328 | -0.026521 | -0.18113 | -0.16724 | 0.020046 | 0.04648 | -0.029417 | -0.040055 | -0.10067 |  | 0
db7 | 1.1371 | 2.8464 | -0.11786 | -0.060862 | 0.10466 | 0.062698 | 0.028314 | -0.0010951 | 0.00027388 | 0.027115 | -0.015446 | -0.011867 | 0.020969 | 0.032553 | -0.031159 | -0.057501 | 0.06795 | -0.0076393 | -0.010088 | -0.031017 | 0.94519 |  | 0
db8 | -0.001057 | -0.059332 | -0.0029691 | -0.0037366 | -0.008983 | -0.0058615 | -0.020047 | -0.026003 | -0.023885 | -0.024173 | -0.020809 | -0.024403 | -0.0075363 | 0.013554 | 0.043154 | 0.025255 | 0.0079018 | 0.045656 | 0.038308 | -0.0054209 | -0.00090097 |  | 0
db9 | 0.48246 | 0.38659 | -0.0011361 | -0.0069433 | -0.011667 | 0.03462 | -0.0086845 | 0.019437 | 0.011643 | 0.020665 | 0.022575 | -0.0074575 | -0.0088096 | -0.020734 | -0.026224 | -0.021652 | -0.010213 | 0.0023154 | 0.0088202 | 0.0034442 | 0.4486 |  | -3.084e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.74988 | 0.59552 | -0.0027242 | -0.0078934 | -0.00064689 | -0.01136 | 0.0047514 | 0.004315 | 0.0030109 | 0.00067159 | -0.0035836 | 0.0099505 | 0.0021379 | -6.0295e-06 | 0.0093104 | 0.0091653 | -0.0024329 | -0.0045828 | -0.0018173 | -0.0082659 | 0.65489 |  | -1.8504e-17
db12 | -0.0015913 | 0.0011331 | 0.00059143 | -0.00014252 | 0.00093169 | 0.00087822 | 0.0022321 | 0.0002489 | -0.0010904 | -0.0011411 | -0.0013552 | -0.0023622 | -0.00099074 | -0.0031301 | 0.0051433 | -0.0010664 | -0.0018741 | 0.0021586 | 0.0011286 | -0.00016006 | 0.00037389 |  | -7.228e-20
db13 | 0.058996 | 0.096285 | 0.0092101 | 0.0075875 | 0.00051641 | 0.0011142 | -0.0028884 | -0.0052389 | -0.0041697 | -0.0030317 | -0.0038965 | -0.001571 | -0.00072048 | 0.0049912 | -0.0069985 | -0.0019908 | -0.00021665 | 0.0026503 | 0.0029881 | 0.0016648 | 0.054429 |  | -1.9275e-18
db14 | 0.0010948 | 0.001309 | -6.5204e-05 | -0.0004524 | -0.00033519 | 0.0004699 | 0.00079597 | -5.3858e-05 | 0.00058263 | 0.0011974 | -0.00068018 | -0.00037618 | -0.0017346 | -5.6773e-05 | 0.0023538 | -0.0013238 | -0.00081601 | 0.00019729 | 0.00024779 | 4.9385e-05 | 0.00081729 |  | 6.6257e-20
db15 | 0.059793 | 0.0093028 | -0.0037111 | -0.0027876 | 0.0015223 | -0.001415 | 0.0015143 | -0.00092949 | 0.0025037 | 0.0022402 | 0.0048466 | 0.0054285 | -0.0018471 | 0.003494 | 0.00039714 | -0.0014965 | -0.0054327 | -0.0012246 | -0.0034814 | 0.00037881 | 0.048699 |  | 3.084e-18
da2 | -0.37374 | 16.946 | 1.0433 | -0.26329 | -0.99756 | -1.3876 | -1.0353 | -1.2472 | -0.83243 | -0.501 | 0.0099051 | 0.41995 | 1.3312 | 1.1743 | -0.53871 | -0.4439 | 1.0116 | 1.0528 | -0.21823 | 1.4222 | 7.0177 |  | 0
da3 | 0.8266 | -0.45315 | 0.17485 | 0.24464 | 0.37099 | 0.20854 | 0.03426 | 0.029045 | -0.23194 | 0.1226 | -0.34592 | 0.085308 | -0.30292 | -0.11992 | -0.20187 | -0.013272 | -0.19583 | -0.23084 | -0.2246 | 0.5969 | 0.75806 |  | 0
da4 | -0.15508 | 3.448 | 0.26171 | 0.19227 | 0.15612 | -0.17067 | -0.49137 | -0.47392 | -0.14947 | 0.16915 | -0.050519 | -0.38093 | -0.25545 | -0.29509 | 0.57487 | 0.42786 | 0.026255 | 0.31619 | 0.206 | -0.063002 | 0.34743 |  | 0
da5 | 0.10865 | 2.2683 | 0.11099 | -0.062219 | -0.01743 | 0.036641 | 0.026931 | 0.12686 | -0.022691 | 0.20391 | 0.22824 | 0.14912 | -0.069795 | -0.22079 | -0.37281 | -0.15999 | 0.092559 | -0.1221 | -0.12708 | 0.19964 | -0.26915 |  | 0
da6 | -0.031722 | 0.37767 | 0.01724 | 0.027584 | 0.028438 | -0.045259 | -0.081359 | -0.13756 | 0.090851 | 0.065592 | -0.019035 | 0.062354 | -0.052444 | 0.035762 | -0.04083 | -0.12689 | -0.093809 | 0.09246 | 0.051788 | 0.12511 | 0.17844 |  | 0
da7 | 0.018785 | 1.7032 | -0.01326 | -0.012545 | -0.037103 | -0.017654 | -0.02922 | 0.049749 | -0.020992 | 0.0016736 | -0.0038437 | 0.042826 | 0.010899 | 0.095378 | 0.084027 | -0.051569 | -0.10313 | -0.083422 | 0.022902 | 0.065289 | 0.01089 |  | 0
da8 | -0.015365 | -0.01076 | -0.020636 | -0.00036741 | -0.032337 | -0.032194 | -0.065871 | -0.075896 | 0.0096559 | -0.0078609 | -0.032248 | 0.01632 | 0.055725 | 0.036836 | 0.014528 | 0.0065213 | 0.051699 | 0.035542 | 0.013032 | 0.027553 | 0.03119 |  | 0
da9 | 0.01279 | -0.24053 | 0.0059177 | 0.0019758 | -0.003913 | -0.0088496 | -0.0068945 | -0.0080503 | 0.018845 | 0.0095168 | -0.02661 | -0.0026583 | 0.00073657 | 0.030594 | -0.0094704 | -0.027126 | 0.0012361 | 0.010612 | -0.00085197 | 0.01499 | -0.041895 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | 0.0062997 | 0.19534 | -0.016799 | -0.011608 | -0.0091967 | -0.012589 | -0.0025395 | 0.010241 | 0.0025601 | 0.01037 | 0.0014247 | 0.0019377 | 0.0035298 | 0.0084467 | 0.011341 | -0.0037877 | 0.0028726 | -0.0062403 | 0.0048029 | 0.0052345 | 0.010159 |  | 0
da12 | -0.005754 | 1.1577e-05 | 0.0014495 | 0.00093123 | -0.00021563 | -0.0028891 | -0.0030328 | -0.0027404 | 0.001642 | 0.00063332 | 0.00058073 | 0.0030542 | 0.0040413 | 0.0011824 | -0.0016288 | -0.0026899 | -0.0030893 | 0.0014282 | 0.00026157 | 0.0010818 | -0.0034722 |  | 7.7099e-19
da13 | 0.0039622 | -0.018167 | 0.00030432 | -0.0010157 | -0.0024529 | 6.3116e-05 | -0.00084114 | -0.0001035 | 0.0031807 | 0.0031866 | 0.0012622 | -0.00091193 | 0.00020429 | 0.0007689 | -0.0028916 | -0.0034703 | -0.002826 | 0.002115 | 0.0005255 | 0.0029025 | -0.00020884 |  | -6.2643e-19
da14 | -0.032426 | 0.0096373 | 0.0011426 | 0.00049219 | 0.0017398 | 0.0022091 | -0.0018027 | -0.002043 | -0.00038555 | 0.00043862 | -0.00042324 | -0.0013356 | 0.0016701 | 0.0010289 | -0.002541 | -0.00022647 | 0.001049 | 0.0005909 | -0.0011419 | -0.00046183 | -0.012576 |  | -1.1565e-18
da15 | -0.0023085 | 0.0075096 | 0.0094008 | 0.0021584 | 0.0029392 | -6.9938e-05 | 0.0051826 | 0.0049978 | 0.0086319 | 0.0097941 | 0.0033934 | 0.00067352 | -0.0018417 | 0.00098735 | -0.0079789 | -0.010656 | -0.014236 | 0.00027081 | -0.0067766 | -0.0068711 | -0.0064875 |  | 0
Field Colinearity | 22.9388
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.6368 | -0.33733 | -0.11853 | -0.29424 | -0.086655 | -0.32236 | 0.23387 | 0.50991 | 0.011552 | 0.21726 | -0.50016 | -0.091948 | 0.44585 | 0.34925 | 0.2295 | -0.088389 | -0.089113 | -0.45418 | -0.60166 | -0.8767
Twist (m2 mrad) | 19.7758 | -4.8069 | -1.6001 | -3.7516 | -1.0399 | -3.6265 | 2.4557 | 4.9716 | 0.10397 | 1.7924 | -3.7512 | -0.62065 | 2.6751 | 1.8336 | 1.0327 | -0.33146 | -0.26734 | -1.0219 | -0.90249 | -0.32876
Twist Integral (I0) (m2 rad) | 0.0094441
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.4113 | -0.49409 | -0.1461 | -0.4688 | -0.29024 | -0.38383 | -0.074584 | 0.30467 | 0.063637 | -0.13103 | -0.63283 | 0.23532 | 0.36291 | 0.6397 | 0.28551 | -0.15939 | -0.25048 | -0.20032 | -0.52976 | -1.0028
Twist (m2 mrad) | 33.0848 | -7.0407 | -1.9723 | -5.9772 | -3.4829 | -4.3181 | -0.78313 | 2.9705 | 0.57273 | -1.081 | -4.7462 | 1.5884 | 2.1775 | 3.3584 | 1.2848 | -0.59771 | -0.75143 | -0.45073 | -0.79464 | -0.37607
Twist Integral (I0) (m2 rad) | 0.0094987
Note on B168: 3.9s
Note on B170: 4.9s
Note on B183: 3.8s
Note on B205: 5.6s
Note on B218: 3.9s
Note on R219: 4.1s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.34708 | -2.9103 | 0.079296 | -0.41735 | 1.1226 | -0.54731 | -0.28325 | 1.0556 | -0.55132 | 19.5007 | 0.59621 | 0.076188 | -3.3127 | -0.051036 | -0.27644 | 1.1372 | -1.0381 | -0.37374 | 0.8266 | -0.15508 | 20.9302 | 0.59614 | -0.41374 | -1.5387 | 0.053085 | -0.6831 | 1.1697 | -0.0024774 | 0.71878 | 0.95961 | -0.40149 | 14422.221 | -0.017473 | -1.9082 | -0.10546 | -0.56084 | 1.1813 | -0.0030659 | 0.37854 | 0.78661 | -0.032356 | 14420.305
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8