## Alarm sheet
File name | HCMB__A001-03000245_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000245
Date of test Ap 1 | 2004-08-24 |  |  |  | Date of test Ap 2 | 2004-08-24
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | yellow alarm | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
499 | HCMB__A001 | 3000245 |  | 3 |  |  | 3
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2004-08-24 | Analysis by |  | E. Todesco | Answer to MMS-LD |  |  | 2004-08-24 | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000245_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000245 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 24-AUG-2004_3245_07-22cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 24-AUG-2004_3245_07-22cc2.txt |  |  |  | a even | 1
Test Operator |  | Mehler |  |  | Test Controller |  |  | E. Todesco
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2004-08-24 |  |  | Time of test |  |  | 07:22 - 10:25 |  |  |  | Date of test |  | 2004-08-24 |  |  | Time of test |  |  | 07:22 - 10:25
Mole name |  | Dipole 7 | [9 empty] | Mole name |  | Dipole 1 | [5 empty] | 2004-03-26
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.64 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00017656 |  |  | Dy average |  |  | 0.00011932 |  |  |  | Dx average |  | 0.00032403 |  |  | Dy average |  |  | 0.0001987
Magnetic length |  | 14.4229 |  |  | Tranfs. Function (mT/KA) |  |  | 596.1079 |  |  |  | Magnetic length |  | 14.4253 |  |  | Tranfs. Function (mT/KA) |  |  | 596.0965
Phase |  | 48.3829 |  |  | Integral field (T.m) |  |  | 85.8575 |  |  |  | Phase |  | -27.117 |  |  | Integral field (T.m) |  |  | 85.8287
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4928 | 5.9643 | 5.9603 | 5.9606 | 5.9602 | 5.9611 | 5.9603 | 5.9604 | 5.9605 | 5.9609 | 5.9608 | 5.9608 | 5.9595 | 5.96 | 5.9603 | 5.9604 | 5.9615 | 5.9619 | 5.9657 | 3.6998 | 85.8575 | C1 (T) | 3.4148 | 5.9636 | 5.9598 | 5.9608 | 5.9602 | 5.9611 | 5.9607 | 5.9608 | 5.9616 | 5.9611 | 5.9606 | 5.9606 | 5.9599 | 5.9597 | 5.96 | 5.9603 | 5.9612 | 5.9607 | 5.9644 | 3.7412 | 85.8287
Angle (mrad) | 3.9601 | -0.93235 | -0.48178 | -0.86056 | -0.56012 | -0.28021 | -0.047436 | 0.70026 | 0.036749 | -0.16874 | -0.44219 | -0.077241 | 0.33032 | 0.3157 | 0.046161 | -0.081006 | -0.48976 | -0.49884 | 0.29924 | 1.4048 | 48.3829 | Angle (mrad) | 2.8695 | -0.84719 | -0.34699 | -0.65296 | -0.49722 | -0.41764 | -0.19681 | 0.5181 | 0.21972 | -0.038793 | -0.39761 | -0.32227 | 0.19677 | 0.40224 | 0.054792 | -0.13809 | -0.42112 | -0.47233 | 0.40933 | 2.0781 | -27.117
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -7.035 | -0.45485 | 0.13023 | -0.13643 | 0.2926 | 0.22026 | 0.21817 | 0.51078 | 0.61403 | 1.0747 | 1.0239 | 0.42359 | 0.42757 | 0.097477 | 0.50697 | 0.38621 | -0.35172 | 0.1009 | 0.49054 | 1.0942 | 0.11099 | b2 | 0.60071 | -0.6809 | -0.80582 | -1.2824 | -1.9395 | -0.42457 | -0.38207 | -1.2995 | -0.66153 | -1.7747 | -1.1274 | -0.68417 | -0.40324 | -0.68399 | -0.41743 | -0.20987 | -0.68887 | -0.54257 | -0.13418 | -0.017365 | -0.7192
b3 | 40.4707 | -3.4286 | -3.327 | -3.5484 | -2.4866 | -4.2265 | -2.4903 | -2.2374 | -3.0792 | -2.9936 | -2.6775 | -2.8991 | -3.0044 | -4.5086 | -3.9759 | -2.7966 | -2.5243 | -2.5448 | -2.3142 | -0.60215 | -1.6517 | b3 | 44.1716 | -3.9498 | -3.566 | -3.5792 | -2.1227 | -3.5183 | -2.3288 | -2.7084 | -1.9504 | -2.7655 | -3.2772 | -3.268 | -3.2615 | -4.7528 | -4.1371 | -2.9459 | -3.1808 | -3.2468 | -2.6613 | -0.73378 | -1.6862
b4 | -2.799 | 0.013859 | -0.10703 | -0.13858 | 0.026607 | -0.01212 | -0.19758 | 0.04778 | -0.013175 | 0.18991 | 0.1641 | 0.30991 | 0.10389 | 0.08284 | 0.08822 | 0.11963 | -0.028404 | -0.044146 | -0.068987 | 0.21012 | -0.050659 | b4 | -0.31059 | -0.028996 | -0.05313 | -0.12399 | -0.30803 | -0.17885 | -0.091005 | -0.24207 | 0.04728 | -0.21985 | -0.069353 | -0.16801 | 0.018675 | -0.15583 | -0.11489 | -0.059768 | 0.051924 | -0.014018 | -0.024919 | -0.0068216 | -0.099837
b5 | -5.2369 | -0.8734 | -0.91341 | -0.57849 | -0.50894 | -0.40619 | -0.48759 | -0.46731 | -0.54147 | -0.38542 | -0.40045 | -0.38838 | -0.39972 | -0.43898 | -0.26247 | -0.46799 | -0.50896 | -0.51011 | -0.67552 | -3.8706 | -0.76462 | b5 | -5.0546 | -0.60704 | -0.62689 | -0.20199 | -0.26585 | -0.051092 | -0.10669 | -0.093433 | -0.36599 | -0.21993 | -0.075812 | -0.17937 | -0.34532 | -0.26976 | -0.3733 | -0.34925 | -0.39109 | -0.48302 | -0.58064 | -3.9296 | -0.57023
b6 | -0.3514 | 0.1865 | 0.014585 | 0.093245 | 0.028875 | -0.024437 | -0.033236 | -0.023761 | 0.05946 | 0.021801 | -0.0010799 | 0.14038 | -0.03661 | 0.046873 | -0.057899 | -0.0088297 | 0.035875 | 0.011485 | -0.015032 | -0.052471 | 0.010404 | b6 | -0.012318 | 0.081157 | 0.03928 | -0.062478 | -0.03314 | -0.12586 | 0.0079641 | 0.0063182 | 0.078089 | 0.011349 | 0.023944 | 0.0067193 | -0.11139 | -0.00034056 | -0.070183 | -0.048328 | 0.072066 | 0.010697 | -0.015708 | -0.1067 | -0.010615
b7 | 2.8986 | 0.92597 | 0.98987 | 1.0626 | 1.0585 | 1.009 | 1.0702 | 1.0546 | 1.0472 | 1.0588 | 1.0642 | 1.0745 | 1.0936 | 1.041 | 1.0602 | 1.0585 | 1.0259 | 1.0059 | 1.0425 | 0.93824 | 1.0946 | b7 | 2.8977 | 0.91395 | 0.9871 | 1.138 | 1.068 | 1.0696 | 1.0912 | 1.0555 | 1.0808 | 1.0885 | 1.0456 | 1.0506 | 1.0739 | 1.0636 | 1.0305 | 1.0548 | 1.0237 | 1.0373 | 1.0229 | 0.86193 | 1.0987
b8 | -0.13538 | 0.047352 | 0.024154 | -0.0067685 | 0.035773 | 0.018388 | -0.02949 | 0.026591 | 0.028855 | 0.051994 | 0.027157 | 0.058165 | -0.0012 | -0.035848 | -0.03481 | -0.01459 | 0.038962 | 0.038494 | 0.008652 | -0.018415 | 0.0099508 | b8 | -0.093127 | 0.020552 | 0.012737 | -0.0075789 | 0.0058613 | 0.0062391 | -0.0015255 | -0.0019094 | 0.01363 | 0.007782 | -0.0044143 | 0.0095951 | -0.023715 | -0.033287 | -0.057512 | -0.0058701 | 0.035062 | 0.04202 | 0.022295 | -0.081855 | -0.0033707
b9 | 0.39698 | 0.46278 | 0.45374 | 0.46553 | 0.47717 | 0.43862 | 0.46417 | 0.46519 | 0.45417 | 0.46233 | 0.46686 | 0.45479 | 0.46668 | 0.46796 | 0.47114 | 0.46988 | 0.44927 | 0.4591 | 0.45279 | 0.42664 | 0.45815 | b9 | 0.37813 | 0.46853 | 0.45771 | 0.46776 | 0.49047 | 0.45338 | 0.47616 | 0.44649 | 0.47993 | 0.46693 | 0.45381 | 0.43928 | 0.44975 | 0.44993 | 0.46392 | 0.45783 | 0.44749 | 0.45664 | 0.44971 | 0.42334 | 0.45614
b10 | 0.034337 | 0.10607 | 0.028574 | 0.0053191 | 0.0022532 | -0.021486 | -0.034464 | 0.022384 | 0.042552 | 0.031815 | -0.014264 | 0.08034 | -0.034655 | -0.088501 | -0.14047 | -0.043726 | 0.049222 | 0.021631 | 0.0064822 | -0.063327 | 0 | b10 | 0.090395 | 0.078818 | 0.030645 | 0.0075928 | -0.00853 | -0.05049 | -0.0066091 | 0.017536 | 0.04516 | 0.041811 | -0.013076 | 0.072518 | -0.052144 | -0.085043 | -0.13144 | -0.03896 | 0.10258 | 0.057217 | -0.019555 | -0.15915 | 0
b11 | 0.63728 | 0.75079 | 0.74794 | 0.7526 | 0.75241 | 0.75852 | 0.74693 | 0.74692 | 0.74632 | 0.75293 | 0.75219 | 0.7531 | 0.74957 | 0.76266 | 0.76039 | 0.75106 | 0.75033 | 0.75374 | 0.74694 | 0.65953 | 0.74548 | b11 | 0.60422 | 0.74838 | 0.74653 | 0.75648 | 0.75088 | 0.7605 | 0.74964 | 0.7521 | 0.75677 | 0.75835 | 0.76087 | 0.75589 | 0.74852 | 0.75371 | 0.76147 | 0.74733 | 0.75193 | 0.74817 | 0.74392 | 0.64916 | 0.74503
b12 | 0.0029081 | 0.0095896 | 0.0042842 | -0.00012808 | 0.0011681 | -0.0035041 | -0.0015461 | -0.00012941 | 0.0065052 | 0.0061819 | 0.0013195 | 0.011958 | -0.0012093 | -0.0068899 | -0.010588 | -0.0044045 | 0.0081229 | 0.0042931 | 0.0011463 | -0.0053342 | 0.0012794 | b12 | 0.014342 | 0.0067086 | 0.0029751 | 0.00052489 | 0.00093308 | -0.00021573 | 0.0051995 | 0.0024342 | 0.0025415 | 0.0021945 | -0.0033779 | 0.007979 | -0.0054333 | -0.008148 | -0.015127 | -0.0029221 | 0.011369 | 0.0071312 | 0.00024877 | -0.014259 | 0.00074423
b13 | 0.093477 | 0.077467 | 0.076833 | 0.073861 | 0.075171 | 0.072036 | 0.074585 | 0.072905 | 0.072279 | 0.074253 | 0.075888 | 0.075289 | 0.074317 | 0.073044 | 0.075301 | 0.076186 | 0.075378 | 0.077149 | 0.076683 | 0.064234 | 0.075144 | b13 | 0.091656 | 0.072881 | 0.073409 | 0.071574 | 0.071432 | 0.070546 | 0.068958 | 0.070697 | 0.072895 | 0.071719 | 0.069394 | 0.073175 | 0.073263 | 0.071118 | 0.073445 | 0.073911 | 0.073541 | 0.074079 | 0.073313 | 0.060072 | 0.072371
b14 | 0.0024977 | 0.0009902 | -0.002039 | -0.0023592 | -0.0028835 | -0.0038191 | -0.0038288 | -0.0021549 | -0.0008854 | -0.00094356 | -0.0027316 | 0.00098013 | -0.0036372 | -0.00649 | -0.0080293 | -0.0049638 | -0.00031138 | -0.0018658 | -0.0025283 | -0.0030361 | -0.0024949 | b14 | 0.0058498 | -0.0032904 | -0.0053862 | -0.006459 | -0.0056836 | -0.0075183 | -0.0055765 | -0.0055112 | -0.0047284 | -0.0050814 | -0.0064023 | -0.001645 | -0.0079764 | -0.0080444 | -0.0099634 | -0.0070616 | -0.0020812 | -0.0044283 | -0.0069988 | -0.0038025 | -0.0053577
b15 | -0.013194 | 0.024321 | 0.024152 | 0.023004 | 0.022097 | 0.021884 | 0.020596 | 0.02193 | 0.023844 | 0.02533 | 0.022523 | 0.023354 | 0.021543 | 0.020165 | 0.016757 | 0.020654 | 0.023626 | 0.024374 | 0.023218 | 0.0048803 | 0.020757 | b15 | -0.014357 | 0.024932 | 0.023361 | 0.021686 | 0.01997 | 0.018552 | 0.021092 | 0.021477 | 0.022323 | 0.023743 | 0.021112 | 0.023518 | 0.018231 | 0.016524 | 0.013332 | 0.018281 | 0.0235 | 0.024137 | 0.020038 | -0.0030462 | 0.019045
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 39.6009 | -9.3235 | -4.8178 | -8.6056 | -5.6012 | -2.8021 | -0.47436 | 7.0026 | 0.36749 | -1.6874 | -4.4219 | -0.77241 | 3.3032 | 3.157 | 0.46161 | -0.81006 | -4.8976 | -4.9885 | 2.9924 | 14.0475 | 0 | a1 | 28.6951 | -8.4719 | -3.4699 | -6.5296 | -4.9722 | -4.1764 | -1.9681 | 5.181 | 2.1972 | -0.38793 | -3.9761 | -3.2227 | 1.9677 | 4.0224 | 0.54792 | -1.3809 | -4.2112 | -4.7233 | 4.0933 | 20.7815 | 0
a2 | 11.5908 | -0.34982 | 1.1382 | -0.004809 | 0.81206 | 0.16492 | 0.4781 | 0.77617 | 1.5163 | 0.37158 | 0.72555 | -0.81616 | 0.89003 | -0.41795 | -0.23406 | -0.49033 | 0.17792 | -0.23275 | -0.64204 | 9.8747 | 0.87374 | a2 | 11.7037 | -0.027421 | 0.89426 | -0.73482 | -0.04278 | -0.57919 | -0.12845 | 0.094884 | -0.18457 | -0.8744 | -0.90906 | -1.5186 | -0.70194 | -0.0098752 | -1.2097 | -0.098863 | -0.011019 | -0.7258 | -0.14937 | -9.1462 | -0.31001
a3 | -0.2455 | 0.75075 | 0.7914 | 0.84409 | 0.50563 | 0.24669 | -0.19993 | 0.13071 | -0.019123 | 0.68188 | 0.48805 | 0.89182 | 0.67922 | 0.94077 | 0.9685 | 0.84233 | 1.0657 | 1.1586 | 1.2069 | 0.16538 | 0.62132 | a3 | -1.1307 | 0.71421 | 0.82944 | 0.53268 | 0.7464 | 1.0054 | 0.57687 | 0.638 | 0.58579 | 0.9141 | 0.89098 | 0.72305 | 0.17888 | 0.86389 | 0.52662 | 0.34738 | 0.67142 | 0.39383 | 0.85276 | 0.72217 | 0.61443
a4 | 1.7457 | 0.033707 | 0.012146 | -0.065456 | 0.064581 | 0.45274 | 0.27222 | 0.25219 | 0.24639 | 0.50839 | 0.15307 | 0.50337 | 0.3309 | 0.50052 | 0.44324 | 0.094063 | 0.40822 | -0.16893 | -0.044801 | 1.1122 | 0.29726 | a4 | 2.5376 | -0.4864 | -0.75453 | -0.6387 | -0.68203 | -0.36932 | -0.64361 | -0.42167 | -0.60413 | -0.30624 | -0.53193 | -0.29166 | -0.14124 | 0.13646 | 0.2082 | -0.30465 | -0.12843 | -0.58803 | -0.051022 | -0.33995 | -0.27909
a5 | 2.5266 | 0.048632 | 0.19039 | 0.33164 | 0.12305 | -0.068757 | -0.048085 | 0.067083 | -0.0088176 | 0.11332 | 0.1604 | 0.054615 | 0.12508 | 0.32177 | 0.18894 | 0.10453 | 0.28273 | 0.36514 | 0.46566 | -0.17826 | 0.21801 | a5 | 2.7368 | 0.10468 | 0.083338 | 0.088296 | 0.23741 | 0.14758 | 0.05697 | 0.4282 | 0.27309 | 0.46009 | 0.34523 | 0.13586 | -0.039818 | 0.30209 | 0.12497 | -0.046282 | 0.12492 | 0.10721 | 0.35045 | -0.20067 | 0.24615
a6 | -0.16038 | -0.0812 | -0.039952 | -0.050437 | -0.030697 | -0.0095359 | 0.013294 | 0.0011536 | -0.042076 | -0.013986 | 0.035518 | 0.11755 | 0.15789 | 0.23291 | 0.21666 | 0.06216 | 0.1874 | 0.025971 | 0.13595 | 0.021504 | 0.043626 | a6 | 0.24961 | 0.028426 | 0.0017788 | -0.17976 | 0.038992 | 0.033424 | 0.069351 | 0.05086 | -0.0042989 | -0.069162 | 0.029545 | 0.12267 | 0.17965 | 0.26928 | 0.12867 | 0.20946 | 0.13969 | -0.0066912 | 0.23309 | -0.064468 | 0.071744
a7 | 1.4301 | -0.092374 | -0.075373 | -0.094404 | -0.034336 | 0.002302 | -0.089687 | 0.039928 | -0.018873 | -0.027047 | -0.052706 | -0.086459 | -0.038693 | -0.060373 | -0.076855 | -0.08842 | 0.016256 | 0.045134 | 0.052754 | -0.028144 | 0.0073584 | a7 | 1.8208 | -0.050652 | -0.027458 | 0.012058 | -0.01042 | 0.016245 | -0.055937 | 0.086835 | 0.025435 | 0.083362 | 0.040625 | 0.031762 | -0.053906 | 0.074652 | -0.023439 | -0.080448 | -0.026738 | -0.032766 | 0.043017 | 0.0425 | 0.058435
a8 | -0.087734 | 0.020183 | -0.0024954 | -0.018408 | 0.0041376 | 0.018831 | 0.022298 | 0.0096449 | -0.0028988 | -0.0033335 | 0.0041781 | 0.023142 | 0.024009 | 0.015049 | -0.0054132 | 0.020034 | 0.042145 | 0.021546 | 0.064382 | 0.0068082 | 0.010929 | a8 | -0.12057 | 0.0016823 | -0.028077 | -0.13922 | -0.03784 | -0.065811 | -0.057358 | -0.019836 | -0.030633 | -0.061997 | -0.042485 | -0.022573 | 0.0020412 | 0.03154 | -0.013506 | 0.032418 | -0.032904 | -0.034325 | 0.031117 | -0.021482 | -0.029702
a9 | -0.303 | -0.046229 | -0.017702 | -0.009847 | -0.029155 | -0.050452 | -0.043016 | -0.021528 | -0.038423 | -0.053798 | -0.011906 | -0.01246 | 0.013815 | -0.0035184 | 0.00091269 | -0.019808 | 0.0047692 | 0.015971 | 0.030961 | -0.035738 | -0.025569 | a9 | -0.24073 | 0.0097597 | -0.0086778 | 0.0037321 | -0.001061 | 0.014606 | 0.0028985 | 0.052255 | 0.022298 | 0.023807 | 0.020783 | 0.0017346 | 0.0071898 | 0.0074201 | -0.029803 | -0.011047 | 0.0010856 | 0.0015101 | 0.041737 | 0.0023922 | 0.0012426
a10 | -0.058532 | -0.025933 | -0.029893 | -0.027892 | 0.0016686 | 0.015804 | -0.0080183 | 0.0068253 | -0.017297 | -0.035956 | -0.031023 | 0.0095543 | -0.012845 | -0.020498 | 0.025652 | 0.0049517 | 0.047593 | 0.05448 | 0.12291 | -0.073941 | 0 | a10 | -0.052447 | 0.048335 | -0.005093 | -0.037284 | 0.01588 | -0.015022 | -0.039838 | 0.043559 | -0.005786 | -0.039296 | -0.02062 | 0.0076164 | -0.0053984 | -0.018679 | 0.0049095 | 0.021631 | -0.0044043 | 0.025879 | 0.087168 | -0.053509 | 0
a11 | 0.15279 | -0.073229 | -0.058194 | -0.058367 | -0.055096 | -0.050915 | -0.047516 | -0.039031 | -0.048308 | -0.050553 | -0.053568 | -0.051584 | -0.044056 | -0.043665 | -0.057436 | -0.05908 | -0.045914 | -0.043512 | -0.033127 | -0.030725 | -0.043876 | a11 | 0.1953 | -0.030125 | -0.01617 | -0.011158 | -0.011536 | -0.013862 | -0.016851 | 0.009933 | -0.0014824 | 0.00099358 | -0.005569 | -0.010549 | -0.011691 | 0.0047966 | -0.010825 | -0.015831 | -0.013585 | -0.017215 | 0.0029577 | 0.012822 | -0.0024916
a12 | -0.021408 | -0.0063928 | -0.0061658 | -0.0075973 | -0.0017884 | -0.0026724 | -0.00062475 | 0.0012305 | -0.003647 | -0.0047313 | -0.00032619 | 0.0026873 | -0.0018916 | -0.0014221 | 0.0030289 | 0.00034084 | 0.0084539 | 0.0052235 | 0.016301 | -0.0081226 | -0.00091453 | a12 | -0.020596 | 0.0054632 | -0.00098872 | -0.016758 | -0.0018755 | -0.0084127 | -0.0068706 | -0.0017154 | -0.0050292 | -0.010438 | -0.0060515 | -0.0007614 | -0.0030547 | -0.0047374 | -0.003777 | 0.0051042 | -0.0011305 | 0.0013082 | 0.014261 | -0.0031253 | -0.0030838
a13 | -0.02414 | -0.0061509 | -0.0035679 | -0.0037567 | -0.0042804 | -0.0061575 | -0.0070221 | -0.0027995 | -0.0060911 | -0.0049515 | -0.0036402 | -0.0036716 | -0.00069026 | -0.0010558 | -0.0019747 | -0.0038174 | -0.0028549 | 0.00098334 | 0.0020666 | -0.0031454 | -0.0039323 | a13 | -0.020196 | 0.0017028 | 0.0014502 | 0.0015349 | 0.0035108 | 0.0068252 | 0.0030097 | 0.0064464 | 0.0030479 | 0.0058133 | 0.0045734 | 0.0024147 | 0.0025289 | 0.0055285 | -0.00063922 | 0.00020457 | 0.0013795 | 0.0015431 | 0.0068831 | 0.0021101 | 0.0024747
a14 | -0.00010631 | -0.0040366 | -0.0033692 | -0.0035235 | -0.00049154 | -0.00027065 | -0.0003986 | -0.00070342 | 2.9786e-05 | -0.0014077 | -0.0014971 | -0.0024812 | -0.0034793 | -0.0047477 | -0.0025167 | -0.0015488 | 0.0011886 | -0.001421 | 0.0030985 | 0.0036543 | -0.0013207 | a14 | -0.0011728 | -0.0037458 | -0.0072462 | -0.0094474 | -0.0062338 | -0.01059 | -0.0092533 | -0.0062791 | -0.0078812 | -0.01013 | -0.01054 | -0.0079173 | -0.0071612 | -0.0049702 | -0.0046071 | -0.0063143 | -0.0059538 | -0.0051861 | -0.0021378 | 0.0045296 | -0.0064279
a15 | 0.0051777 | -0.0015323 | -0.0024772 | -0.0022227 | -0.0022573 | -0.0031643 | -0.0044762 | -0.0013749 | -0.0019789 | -0.0021352 | -0.0038969 | -0.00049137 | -0.0029374 | -0.0054562 | -0.0034562 | -0.0033332 | 0.00018979 | 0.00098887 | 0.0020154 | -0.0040883 | -0.0019522 | a15 | 0.0036179 | -0.00097915 | -0.0019821 | -0.002228 | -8.7557e-06 | -0.0023259 | -0.0036299 | 0.0018564 | -0.00052541 | -0.0014923 | -0.0027339 | -7.9306e-05 | -0.005657 | -0.0057646 | -0.0063607 | -0.0038617 | -0.0014882 | -0.00080416 | 0.001146 | -0.012606 | -0.0022267
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -5.1086e-05 | -0.00024369 | -6.9793e-05 | -1.6794e-05 | 0 | 5.0313e-05 | 7.6967e-05 | -4.9997e-05 | -9.9074e-05 | -7.6923e-05 | 2.7111e-05 | -0.00017914 | 7.6616e-05 | 0.00019407 | 0.00031678 | 9.9252e-05 | -0.00010446 | -4.152e-05 | 0 | 0.00015384 | 0.00017656 | Dx (m) | -0.00018635 | -0.0001743 | -7.0008e-05 | -1.8273e-05 | 1.9862e-05 | 0.00011218 | 1.2964e-05 | -4.0935e-05 | -0.00010147 | -9.3549e-05 | 2.8869e-05 | -0.00016288 | 0.00011822 | 0.00019213 | 0.00029392 | 8.9615e-05 | -0.0002322 | -0.00012861 | 4.3834e-05 | 0.00042001 | 0.00032403
Dy (m) | 0.00016818 | 3.5125e-05 | 6.2527e-05 | 6.169e-05 | 0 | -3.2045e-05 | 2.3147e-05 | -1.8147e-05 | 3.3012e-05 | 7.6043e-05 | 7.2044e-05 | -3.3913e-05 | 3.3637e-05 | 5.6843e-05 | -3.3563e-05 | 0 | -0.00011428 | -0.00012527 | -0.00027945 | 0.0001978 | 0.00011932 | Dy (m) | 0.0002082 | -0.0001169 | 1.0084e-05 | 8.3473e-05 | -3.5647e-05 | 3.5656e-05 | 9.061e-05 | -9.7935e-05 | 1.2821e-05 | 8.8221e-05 | 4.6284e-05 | -1.9403e-05 | 1.4128e-05 | 4.0998e-05 | 0 | -4.7324e-05 | 0 | -6.1795e-05 | -0.00019918 | 0.00013238 | 0.0001987
Magnetic length (m) |  | 14.4229 | [19 empty] | Magnetic length (m) |  | 14.4253
Transfer funct. (mT/KA) |  | 596.1079 | [19 empty] | Transfer funct. (mT/KA) |  | 596.0965
Coil positioning (mm) |  | 0.01669 | [19 empty] | Coil positioning (mm) |  | 0.018103
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.0.8 | [20 empty] | Soft version | DIMM V2.0.8
Level Average | 0.55185 | 0.77129 | 0.51928 | 0.5641 | 0.55363 | 0.7454 | 0.66466 | 0.52283 | 0.64949 | 0.72369 | 0.34409 | 0.089972 | 0.14504 | 0.052275 | 0.5082 | -0.060681 | -0.024384 | -0.12322 | -0.09079 | -0.37293 |  | Level Average | 0.01007 | 0.031232 | -0.056077 | -0.086866 | -0.044379 | -0.044827 | -0.047977 | -0.04787 | -0.046232 | -0.038244 | -0.0689 | -0.12699 | -0.09828 | -0.10216 | -0.29041 | -0.26147 | -0.22556 | -0.12406 | -0.17818 | -0.22742
Level Sigma2 | 1.6782 | 2.5971 | 0.69026 | 1.168 | 2.1752 | 9.012 | 8.03 | 1.2528 | 5.5453 | 11.1781 | 1.95 | 0.83925 | 1.4576 | 0.42661 | 0.020438 | 0.82608 | 0.11245 | 0.20023 | 0.045655 | 0.26087 |  | Level Sigma2 | 0.000202 | 0.002387 | 0.003034 | 0.003403 | 0.003172 | 0.008335 | 0.006754 | 0.001971 | 0.003508 | 0.00257 | 0.000715 | 0.037182 | 0.19623 | 0.01632 | 0.14138 | 0.10608 | 0.061385 | 0.015128 | 0.007552 | 0.24512
Position (mm) | 0 | 750 | 1501 | 2251 | 3002 | 3752 | 4504 | 5255 | 6007 | 6757 | 7508 | 8259 | 9010 | 9762 | 10513 | 11264 | 12016 | 12766 | 13517 | 14268 |  | Position (mm) | 0 | 751 | 1502 | 2254 | 3005 | 3757 | 4508 | 5259 | 6011 | 6762 | 7513 | 8264 | 9015 | 9767 | 10518 | 11269 | 12021 | 12772 | 13523 | 14275
Fist move (mm) | First move Ap1 | 940.4482 | First move Ap2 | 945.4735 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-3000245 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000245 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 349.2776 | 596.4306 | 596.0305 | 596.0567 | 596.0186 | 596.1053 | 596.0273 | 596.0434 | 596.0473 | 596.0904 | 596.0751 | 596.0824 | 595.9475 | 596.0043 | 596.0275 | 596.0364 | 596.153 | 596.1929 | 596.5726 | 369.9789 | 8585.7542 | C1/i (mT/KA) | 341.485 | 596.3625 | 595.9815 | 596.0841 | 596.0229 | 596.1104 | 596.0713 | 596.0843 | 596.1622 | 596.1072 | 596.0608 | 596.0572 | 595.9949 | 595.9731 | 596.0022 | 596.0278 | 596.1236 | 596.073 | 596.4382 | 374.1237 | 8582.8651
Angle (mrad) | 3.9601 | -0.93235 | -0.48178 | -0.86056 | -0.56012 | -0.28021 | -0.047436 | 0.70026 | 0.036749 | -0.16874 | -0.44219 | -0.077241 | 0.33032 | 0.3157 | 0.046161 | -0.081006 | -0.48976 | -0.49884 | 0.29924 | 1.4048 | 48.3829 | Angle (mrad) | 2.8695 | -0.84719 | -0.34699 | -0.65296 | -0.49722 | -0.41764 | -0.19681 | 0.5181 | 0.21972 | -0.038793 | -0.39761 | -0.32227 | 0.19677 | 0.40224 | 0.054792 | -0.13809 | -0.42112 | -0.47233 | 0.40933 | 2.0781 | -27.117
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -7.035 | -0.45485 | 0.13023 | -0.13643 | 0.2926 | 0.22026 | 0.21817 | 0.51078 | 0.61403 | 1.0747 | 1.0239 | 0.42359 | 0.42757 | 0.097477 | 0.50697 | 0.38621 | -0.35172 | 0.1009 | 0.49054 | 1.0942 |  | b2 | 0.60071 | -0.6809 | -0.80582 | -1.2824 | -1.9395 | -0.42457 | -0.38207 | -1.2995 | -0.66153 | -1.7747 | -1.1274 | -0.68417 | -0.40324 | -0.68399 | -0.41743 | -0.20987 | -0.68887 | -0.54257 | -0.13418 | -0.017365
b3 | 40.4707 | -3.4286 | -3.327 | -3.5484 | -2.4866 | -4.2265 | -2.4903 | -2.2374 | -3.0792 | -2.9936 | -2.6775 | -2.8991 | -3.0044 | -4.5086 | -3.9759 | -2.7966 | -2.5243 | -2.5448 | -2.3142 | -0.60215 |  | b3 | 44.1716 | -3.9498 | -3.566 | -3.5792 | -2.1227 | -3.5183 | -2.3288 | -2.7084 | -1.9504 | -2.7655 | -3.2772 | -3.268 | -3.2615 | -4.7528 | -4.1371 | -2.9459 | -3.1808 | -3.2468 | -2.6613 | -0.73378
b4 | -2.799 | 0.013859 | -0.10703 | -0.13858 | 0.026607 | -0.01212 | -0.19758 | 0.04778 | -0.013175 | 0.18991 | 0.1641 | 0.30991 | 0.10389 | 0.08284 | 0.08822 | 0.11963 | -0.028404 | -0.044146 | -0.068987 | 0.21012 |  | b4 | -0.31059 | -0.028996 | -0.05313 | -0.12399 | -0.30803 | -0.17885 | -0.091005 | -0.24207 | 0.04728 | -0.21985 | -0.069353 | -0.16801 | 0.018675 | -0.15583 | -0.11489 | -0.059768 | 0.051924 | -0.014018 | -0.024919 | -0.0068216
b5 | -5.2369 | -0.8734 | -0.91341 | -0.57849 | -0.50894 | -0.40619 | -0.48759 | -0.46731 | -0.54147 | -0.38542 | -0.40045 | -0.38838 | -0.39972 | -0.43898 | -0.26247 | -0.46799 | -0.50896 | -0.51011 | -0.67552 | -3.8706 |  | b5 | -5.0546 | -0.60704 | -0.62689 | -0.20199 | -0.26585 | -0.051092 | -0.10669 | -0.093433 | -0.36599 | -0.21993 | -0.075812 | -0.17937 | -0.34532 | -0.26976 | -0.3733 | -0.34925 | -0.39109 | -0.48302 | -0.58064 | -3.9296
b6 | -0.3514 | 0.1865 | 0.014585 | 0.093245 | 0.028875 | -0.024437 | -0.033236 | -0.023761 | 0.05946 | 0.021801 | -0.0010799 | 0.14038 | -0.03661 | 0.046873 | -0.057899 | -0.0088297 | 0.035875 | 0.011485 | -0.015032 | -0.052471 |  | b6 | -0.012318 | 0.081157 | 0.03928 | -0.062478 | -0.03314 | -0.12586 | 0.0079641 | 0.0063182 | 0.078089 | 0.011349 | 0.023944 | 0.0067193 | -0.11139 | -0.00034056 | -0.070183 | -0.048328 | 0.072066 | 0.010697 | -0.015708 | -0.1067
b7 | 2.8986 | 0.92597 | 0.98987 | 1.0626 | 1.0585 | 1.009 | 1.0702 | 1.0546 | 1.0472 | 1.0588 | 1.0642 | 1.0745 | 1.0936 | 1.041 | 1.0602 | 1.0585 | 1.0259 | 1.0059 | 1.0425 | 0.93824 |  | b7 | 2.8977 | 0.91395 | 0.9871 | 1.138 | 1.068 | 1.0696 | 1.0912 | 1.0555 | 1.0808 | 1.0885 | 1.0456 | 1.0506 | 1.0739 | 1.0636 | 1.0305 | 1.0548 | 1.0237 | 1.0373 | 1.0229 | 0.86193
b8 | -0.13538 | 0.047352 | 0.024154 | -0.0067685 | 0.035773 | 0.018388 | -0.02949 | 0.026591 | 0.028855 | 0.051994 | 0.027157 | 0.058165 | -0.0012 | -0.035848 | -0.03481 | -0.01459 | 0.038962 | 0.038494 | 0.008652 | -0.018415 |  | b8 | -0.093127 | 0.020552 | 0.012737 | -0.0075789 | 0.0058613 | 0.0062391 | -0.0015255 | -0.0019094 | 0.01363 | 0.007782 | -0.0044143 | 0.0095951 | -0.023715 | -0.033287 | -0.057512 | -0.0058701 | 0.035062 | 0.04202 | 0.022295 | -0.081855
b9 | 0.39698 | 0.46278 | 0.45374 | 0.46553 | 0.47717 | 0.43862 | 0.46417 | 0.46519 | 0.45417 | 0.46233 | 0.46686 | 0.45479 | 0.46668 | 0.46796 | 0.47114 | 0.46988 | 0.44927 | 0.4591 | 0.45279 | 0.42664 |  | b9 | 0.37813 | 0.46853 | 0.45771 | 0.46776 | 0.49047 | 0.45338 | 0.47616 | 0.44649 | 0.47993 | 0.46693 | 0.45381 | 0.43928 | 0.44975 | 0.44993 | 0.46392 | 0.45783 | 0.44749 | 0.45664 | 0.44971 | 0.42334
b10 | 0.034337 | 0.10607 | 0.028574 | 0.0053191 | 0.0022532 | -0.021486 | -0.034464 | 0.022384 | 0.042552 | 0.031815 | -0.014264 | 0.08034 | -0.034655 | -0.088501 | -0.14047 | -0.043726 | 0.049222 | 0.021631 | 0.0064822 | -0.063327 |  | b10 | 0.090395 | 0.078818 | 0.030645 | 0.0075928 | -0.00853 | -0.05049 | -0.0066091 | 0.017536 | 0.04516 | 0.041811 | -0.013076 | 0.072518 | -0.052144 | -0.085043 | -0.13144 | -0.03896 | 0.10258 | 0.057217 | -0.019555 | -0.15915
b11 | 0.63728 | 0.75079 | 0.74794 | 0.7526 | 0.75241 | 0.75852 | 0.74693 | 0.74692 | 0.74632 | 0.75293 | 0.75219 | 0.7531 | 0.74957 | 0.76266 | 0.76039 | 0.75106 | 0.75033 | 0.75374 | 0.74694 | 0.65953 |  | b11 | 0.60422 | 0.74838 | 0.74653 | 0.75648 | 0.75088 | 0.7605 | 0.74964 | 0.7521 | 0.75677 | 0.75835 | 0.76087 | 0.75589 | 0.74852 | 0.75371 | 0.76147 | 0.74733 | 0.75193 | 0.74817 | 0.74392 | 0.64916
b12 | 0.0029081 | 0.0095896 | 0.0042842 | -0.00012808 | 0.0011681 | -0.0035041 | -0.0015461 | -0.00012941 | 0.0065052 | 0.0061819 | 0.0013195 | 0.011958 | -0.0012093 | -0.0068899 | -0.010588 | -0.0044045 | 0.0081229 | 0.0042931 | 0.0011463 | -0.0053342 |  | b12 | 0.014342 | 0.0067086 | 0.0029751 | 0.00052489 | 0.00093308 | -0.00021573 | 0.0051995 | 0.0024342 | 0.0025415 | 0.0021945 | -0.0033779 | 0.007979 | -0.0054333 | -0.008148 | -0.015127 | -0.0029221 | 0.011369 | 0.0071312 | 0.00024877 | -0.014259
b13 | 0.093477 | 0.077467 | 0.076833 | 0.073861 | 0.075171 | 0.072036 | 0.074585 | 0.072905 | 0.072279 | 0.074253 | 0.075888 | 0.075289 | 0.074317 | 0.073044 | 0.075301 | 0.076186 | 0.075378 | 0.077149 | 0.076683 | 0.064234 |  | b13 | 0.091656 | 0.072881 | 0.073409 | 0.071574 | 0.071432 | 0.070546 | 0.068958 | 0.070697 | 0.072895 | 0.071719 | 0.069394 | 0.073175 | 0.073263 | 0.071118 | 0.073445 | 0.073911 | 0.073541 | 0.074079 | 0.073313 | 0.060072
b14 | 0.0024977 | 0.0009902 | -0.002039 | -0.0023592 | -0.0028835 | -0.0038191 | -0.0038288 | -0.0021549 | -0.0008854 | -0.00094356 | -0.0027316 | 0.00098013 | -0.0036372 | -0.00649 | -0.0080293 | -0.0049638 | -0.00031138 | -0.0018658 | -0.0025283 | -0.0030361 |  | b14 | 0.0058498 | -0.0032904 | -0.0053862 | -0.006459 | -0.0056836 | -0.0075183 | -0.0055765 | -0.0055112 | -0.0047284 | -0.0050814 | -0.0064023 | -0.001645 | -0.0079764 | -0.0080444 | -0.0099634 | -0.0070616 | -0.0020812 | -0.0044283 | -0.0069988 | -0.0038025
b15 | -0.013194 | 0.024321 | 0.024152 | 0.023004 | 0.022097 | 0.021884 | 0.020596 | 0.02193 | 0.023844 | 0.02533 | 0.022523 | 0.023354 | 0.021543 | 0.020165 | 0.016757 | 0.020654 | 0.023626 | 0.024374 | 0.023218 | 0.0048803 |  | b15 | -0.014357 | 0.024932 | 0.023361 | 0.021686 | 0.01997 | 0.018552 | 0.021092 | 0.021477 | 0.022323 | 0.023743 | 0.021112 | 0.023518 | 0.018231 | 0.016524 | 0.013332 | 0.018281 | 0.0235 | 0.024137 | 0.020038 | -0.0030462
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 39.6009 | -9.3235 | -4.8178 | -8.6056 | -5.6012 | -2.8021 | -0.47436 | 7.0026 | 0.36749 | -1.6874 | -4.4219 | -0.77241 | 3.3032 | 3.157 | 0.46161 | -0.81006 | -4.8976 | -4.9885 | 2.9924 | 14.0475 |  | a1 | 28.6951 | -8.4719 | -3.4699 | -6.5296 | -4.9722 | -4.1764 | -1.9681 | 5.181 | 2.1972 | -0.38793 | -3.9761 | -3.2227 | 1.9677 | 4.0224 | 0.54792 | -1.3809 | -4.2112 | -4.7233 | 4.0933 | 20.7815
a2 | 11.5908 | -0.34982 | 1.1382 | -0.004809 | 0.81206 | 0.16492 | 0.4781 | 0.77617 | 1.5163 | 0.37158 | 0.72555 | -0.81616 | 0.89003 | -0.41795 | -0.23406 | -0.49033 | 0.17792 | -0.23275 | -0.64204 | 9.8747 |  | a2 | 11.7037 | -0.027421 | 0.89426 | -0.73482 | -0.04278 | -0.57919 | -0.12845 | 0.094884 | -0.18457 | -0.8744 | -0.90906 | -1.5186 | -0.70194 | -0.0098752 | -1.2097 | -0.098863 | -0.011019 | -0.7258 | -0.14937 | -9.1462
a3 | -0.2455 | 0.75075 | 0.7914 | 0.84409 | 0.50563 | 0.24669 | -0.19993 | 0.13071 | -0.019123 | 0.68188 | 0.48805 | 0.89182 | 0.67922 | 0.94077 | 0.9685 | 0.84233 | 1.0657 | 1.1586 | 1.2069 | 0.16538 |  | a3 | -1.1307 | 0.71421 | 0.82944 | 0.53268 | 0.7464 | 1.0054 | 0.57687 | 0.638 | 0.58579 | 0.9141 | 0.89098 | 0.72305 | 0.17888 | 0.86389 | 0.52662 | 0.34738 | 0.67142 | 0.39383 | 0.85276 | 0.72217
a4 | 1.7457 | 0.033707 | 0.012146 | -0.065456 | 0.064581 | 0.45274 | 0.27222 | 0.25219 | 0.24639 | 0.50839 | 0.15307 | 0.50337 | 0.3309 | 0.50052 | 0.44324 | 0.094063 | 0.40822 | -0.16893 | -0.044801 | 1.1122 |  | a4 | 2.5376 | -0.4864 | -0.75453 | -0.6387 | -0.68203 | -0.36932 | -0.64361 | -0.42167 | -0.60413 | -0.30624 | -0.53193 | -0.29166 | -0.14124 | 0.13646 | 0.2082 | -0.30465 | -0.12843 | -0.58803 | -0.051022 | -0.33995
a5 | 2.5266 | 0.048632 | 0.19039 | 0.33164 | 0.12305 | -0.068757 | -0.048085 | 0.067083 | -0.0088176 | 0.11332 | 0.1604 | 0.054615 | 0.12508 | 0.32177 | 0.18894 | 0.10453 | 0.28273 | 0.36514 | 0.46566 | -0.17826 |  | a5 | 2.7368 | 0.10468 | 0.083338 | 0.088296 | 0.23741 | 0.14758 | 0.05697 | 0.4282 | 0.27309 | 0.46009 | 0.34523 | 0.13586 | -0.039818 | 0.30209 | 0.12497 | -0.046282 | 0.12492 | 0.10721 | 0.35045 | -0.20067
a6 | -0.16038 | -0.0812 | -0.039952 | -0.050437 | -0.030697 | -0.0095359 | 0.013294 | 0.0011536 | -0.042076 | -0.013986 | 0.035518 | 0.11755 | 0.15789 | 0.23291 | 0.21666 | 0.06216 | 0.1874 | 0.025971 | 0.13595 | 0.021504 |  | a6 | 0.24961 | 0.028426 | 0.0017788 | -0.17976 | 0.038992 | 0.033424 | 0.069351 | 0.05086 | -0.0042989 | -0.069162 | 0.029545 | 0.12267 | 0.17965 | 0.26928 | 0.12867 | 0.20946 | 0.13969 | -0.0066912 | 0.23309 | -0.064468
a7 | 1.4301 | -0.092374 | -0.075373 | -0.094404 | -0.034336 | 0.002302 | -0.089687 | 0.039928 | -0.018873 | -0.027047 | -0.052706 | -0.086459 | -0.038693 | -0.060373 | -0.076855 | -0.08842 | 0.016256 | 0.045134 | 0.052754 | -0.028144 |  | a7 | 1.8208 | -0.050652 | -0.027458 | 0.012058 | -0.01042 | 0.016245 | -0.055937 | 0.086835 | 0.025435 | 0.083362 | 0.040625 | 0.031762 | -0.053906 | 0.074652 | -0.023439 | -0.080448 | -0.026738 | -0.032766 | 0.043017 | 0.0425
a8 | -0.087734 | 0.020183 | -0.0024954 | -0.018408 | 0.0041376 | 0.018831 | 0.022298 | 0.0096449 | -0.0028988 | -0.0033335 | 0.0041781 | 0.023142 | 0.024009 | 0.015049 | -0.0054132 | 0.020034 | 0.042145 | 0.021546 | 0.064382 | 0.0068082 |  | a8 | -0.12057 | 0.0016823 | -0.028077 | -0.13922 | -0.03784 | -0.065811 | -0.057358 | -0.019836 | -0.030633 | -0.061997 | -0.042485 | -0.022573 | 0.0020412 | 0.03154 | -0.013506 | 0.032418 | -0.032904 | -0.034325 | 0.031117 | -0.021482
a9 | -0.303 | -0.046229 | -0.017702 | -0.009847 | -0.029155 | -0.050452 | -0.043016 | -0.021528 | -0.038423 | -0.053798 | -0.011906 | -0.01246 | 0.013815 | -0.0035184 | 0.00091269 | -0.019808 | 0.0047692 | 0.015971 | 0.030961 | -0.035738 |  | a9 | -0.24073 | 0.0097597 | -0.0086778 | 0.0037321 | -0.001061 | 0.014606 | 0.0028985 | 0.052255 | 0.022298 | 0.023807 | 0.020783 | 0.0017346 | 0.0071898 | 0.0074201 | -0.029803 | -0.011047 | 0.0010856 | 0.0015101 | 0.041737 | 0.0023922
a10 | -0.058532 | -0.025933 | -0.029893 | -0.027892 | 0.0016686 | 0.015804 | -0.0080183 | 0.0068253 | -0.017297 | -0.035956 | -0.031023 | 0.0095543 | -0.012845 | -0.020498 | 0.025652 | 0.0049517 | 0.047593 | 0.05448 | 0.12291 | -0.073941 |  | a10 | -0.052447 | 0.048335 | -0.005093 | -0.037284 | 0.01588 | -0.015022 | -0.039838 | 0.043559 | -0.005786 | -0.039296 | -0.02062 | 0.0076164 | -0.0053984 | -0.018679 | 0.0049095 | 0.021631 | -0.0044043 | 0.025879 | 0.087168 | -0.053509
a11 | 0.15279 | -0.073229 | -0.058194 | -0.058367 | -0.055096 | -0.050915 | -0.047516 | -0.039031 | -0.048308 | -0.050553 | -0.053568 | -0.051584 | -0.044056 | -0.043665 | -0.057436 | -0.05908 | -0.045914 | -0.043512 | -0.033127 | -0.030725 |  | a11 | 0.1953 | -0.030125 | -0.01617 | -0.011158 | -0.011536 | -0.013862 | -0.016851 | 0.009933 | -0.0014824 | 0.00099358 | -0.005569 | -0.010549 | -0.011691 | 0.0047966 | -0.010825 | -0.015831 | -0.013585 | -0.017215 | 0.0029577 | 0.012822
a12 | -0.021408 | -0.0063928 | -0.0061658 | -0.0075973 | -0.0017884 | -0.0026724 | -0.00062475 | 0.0012305 | -0.003647 | -0.0047313 | -0.00032619 | 0.0026873 | -0.0018916 | -0.0014221 | 0.0030289 | 0.00034084 | 0.0084539 | 0.0052235 | 0.016301 | -0.0081226 |  | a12 | -0.020596 | 0.0054632 | -0.00098872 | -0.016758 | -0.0018755 | -0.0084127 | -0.0068706 | -0.0017154 | -0.0050292 | -0.010438 | -0.0060515 | -0.0007614 | -0.0030547 | -0.0047374 | -0.003777 | 0.0051042 | -0.0011305 | 0.0013082 | 0.014261 | -0.0031253
a13 | -0.02414 | -0.0061509 | -0.0035679 | -0.0037567 | -0.0042804 | -0.0061575 | -0.0070221 | -0.0027995 | -0.0060911 | -0.0049515 | -0.0036402 | -0.0036716 | -0.00069026 | -0.0010558 | -0.0019747 | -0.0038174 | -0.0028549 | 0.00098334 | 0.0020666 | -0.0031454 |  | a13 | -0.020196 | 0.0017028 | 0.0014502 | 0.0015349 | 0.0035108 | 0.0068252 | 0.0030097 | 0.0064464 | 0.0030479 | 0.0058133 | 0.0045734 | 0.0024147 | 0.0025289 | 0.0055285 | -0.00063922 | 0.00020457 | 0.0013795 | 0.0015431 | 0.0068831 | 0.0021101
a14 | -0.00010631 | -0.0040366 | -0.0033692 | -0.0035235 | -0.00049154 | -0.00027065 | -0.0003986 | -0.00070342 | 2.9786e-05 | -0.0014077 | -0.0014971 | -0.0024812 | -0.0034793 | -0.0047477 | -0.0025167 | -0.0015488 | 0.0011886 | -0.001421 | 0.0030985 | 0.0036543 |  | a14 | -0.0011728 | -0.0037458 | -0.0072462 | -0.0094474 | -0.0062338 | -0.01059 | -0.0092533 | -0.0062791 | -0.0078812 | -0.01013 | -0.01054 | -0.0079173 | -0.0071612 | -0.0049702 | -0.0046071 | -0.0063143 | -0.0059538 | -0.0051861 | -0.0021378 | 0.0045296
a15 | 0.0051777 | -0.0015323 | -0.0024772 | -0.0022227 | -0.0022573 | -0.0031643 | -0.0044762 | -0.0013749 | -0.0019789 | -0.0021352 | -0.0038969 | -0.00049137 | -0.0029374 | -0.0054562 | -0.0034562 | -0.0033332 | 0.00018979 | 0.00098887 | 0.0020154 | -0.0040883 |  | a15 | 0.0036179 | -0.00097915 | -0.0019821 | -0.002228 | -8.7557e-06 | -0.0023259 | -0.0036299 | 0.0018564 | -0.00052541 | -0.0014923 | -0.0027339 | -7.9306e-05 | -0.005657 | -0.0057646 | -0.0063607 | -0.0038617 | -0.0014882 | -0.00080416 | 0.001146 | -0.012606
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.051218 | 0.24359 | 0.06981 | 0.0168 | 0.0047893 | -0.050307 | -0.076967 | 0.049998 | 0.099063 | 0.076938 | -0.027106 | 0.17904 | -0.07662 | -0.19402 | -0.31658 | -0.09924 | 0.10453 | 0.041554 | 0.0023425 | -0.15402 |  | Dx (mm) | 0.18709 | 0.17434 | 0.070005 | 0.018295 | -0.01986 | -0.11222 | -0.012951 | 0.04093 | 0.10147 | 0.093613 | -0.028876 | 0.16282 | -0.11821 | -0.19208 | -0.29354 | -0.089627 | 0.23203 | 0.12859 | -0.043894 | -0.41939
Dy (mm) | -0.16842 | -0.034961 | -0.062512 | -0.061699 | 0.0041208 | 0.032044 | -0.023146 | 0.018147 | -0.032988 | -0.076017 | -0.072045 | 0.033831 | -0.033635 | -0.0568 | 0.033436 | 0.0034015 | 0.11423 | 0.12528 | 0.27983 | -0.19776 |  | Dy (mm) | -0.20803 | 0.11681 | -0.010081 | -0.083516 | 0.035647 | -0.035625 | -0.090634 | 0.097917 | -0.012799 | -0.088213 | -0.046283 | 0.019402 | -0.014107 | -0.040907 | 0.0067875 | 0.047308 | -0.0057655 | 0.06176 | 0.19937 | -0.13184
Mag. len. (m) | 14.4229 | [20 empty] | Mag. len. (m) | 14.4253
TF (mT/KA) | 596.1079 | [20 empty] | TF (mT/KA) | 596.0965
Coil pos. (mm) | 0.01669 | [20 empty] | Coil pos. (mm) | 0.018103
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | -0.0056834 | [20 empty] | Twist (m2 rad) | -0.0094894

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 596.1079 | 0.15536 | -0.00023572 |  | -1.7736 | 4.3289 |  |  | 596.0965 | 0.12247 | -0.011644 |  | -1.6385 | 3.9195 |  | 0 | 4
2 | 0.10297 |  | 0.30945 | 0.38937 | 0.89913 |  | 0.21591 | 0.65259 | -0.74959 |  | -0.7864 | 0.49963 | -0.27438 |  | -0.38232 | 0.57371 |  | 0 | 5
3 | -1.654 |  | -3.0595 | 0.66037 | 0.62079 |  | 0.66658 | 0.40257 | -1.6888 |  | -3.1793 | 0.70905 | 0.61416 |  | 0.66983 | 0.21698 |  | 0 | 6
4 | -0.051546 |  | 0.034625 | 0.12352 | 0.28581 |  | 0.224 | 0.21969 | -0.10755 |  | -0.094122 | 0.10605 | -0.29308 |  | -0.36489 | 0.28628 |  | 0 | 7
5 | -0.76491 |  | -0.51254 | 0.16534 | 0.21955 |  | 0.15872 | 0.14886 | -0.57147 |  | -0.31085 | 0.1807 | 0.24666 |  | 0.18388 | 0.14909 |  | 0 | 8
6 | 0.00853 |  | 0.024996 | 0.045707 | 0.046357 |  | 0.046978 | 0.090554 | -0.01786 |  | -0.0089747 | 0.04267 | 0.073647 |  | 0.068303 | 0.10407 |  | 0 | 9
7 | 1.0941 |  | 1.0406 | 0.039317 | 0.0068162 |  | -0.038032 | 0.050213 | 1.0984 |  | 1.0493 | 0.048367 | 0.05855 |  | 0.003266 | 0.051013 |  | 0 | 10
8 | 0.011614 |  | 0.015783 | 0.015406 | 0.011658 |  | 0.012519 | 0.012924 | -0.001952 |  | 0.0011367 | 0.015869 | -0.029063 |  | -0.029199 | 0.033842 |  | 0 | 11
9 | 0.45648 |  | 0.45921 | 0.010087 | -0.025819 |  | -0.016171 | 0.023131 | 0.45235 |  | 0.45692 | 0.015376 | 0.00043852 |  | 0.0083634 | 0.01925 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.74541 |  | 0.75181 | 0.004252 | -0.043867 |  | -0.050704 | 0.0087206 | 0.74473 |  | 0.75263 | 0.0051403 | -0.0025125 |  | -0.0094116 | 0.0098694 |  | 0 | 14
12 | 0.0001402 |  | 0.0013461 | 0.0022939 | -7.6352e-05 |  | -0.00052528 | 0.0022564 | 7.5469e-05 |  | 0.00052335 | 0.0024548 | -0.00028494 |  | -0.0030052 | 0.0037069 |  | 0 | 15
13 | 0.0074987 |  | 0.074789 | 0.0017887 | -0.0003929 |  | -0.0033079 | 0.0024541 | 0.0072193 |  | 0.072076 | 0.0015676 | 0.00026113 |  | 0.0032822 | 0.0024057 |  | 0 | 16
14 | -0.00025964 |  | -0.0027861 | 0.00053517 | -0.00016606 |  | -0.0018372 | 0.0015734 | -0.00056338 |  | -0.0061248 | 0.00068478 | -0.00069704 |  | -0.0072966 | 0.0013932 |  | 0 | 17
15 | 0.0020757 |  | 0.022409 | 0.0020181 | -0.00019522 |  | -0.0021109 | 0.0018894 | 0.0019045 |  | 0.020878 | 0.0030202 | -0.00022267 |  | -0.002051 | 0.0023127 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.68501 | -0.62476
3 | [12 empty] | 3 | 9 | -0.66243 | -0.93589
4 | [12 empty] | 4 | 16 | -1.8035 | -1.7473
5 | [12 empty] | 5 | 25 | -1.8523 | -1.8071
6 | [12 empty] | 6 | 36 | -2.7437 | -2.7085
7 | [12 empty] | 7 | 49 | -3.1138 | -3.0023
8 | [12 empty] | 8 | 64 | -4.2609 | -3.7647
9 | [12 empty] | 9 | 81 | -4.1815 | -4.0626
10 | [12 empty] | sigma (mm) |  | 0.01669 | 0.018103
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 349.2776 | 596.4306 | 596.0305 | 596.0567 | 596.0186 | 596.1053 | 596.0273 | 596.0434 | 596.0473 | 596.0904 | 596.0751 | 596.0824 | 595.9475 | 596.0043 | 596.0275 | 596.0364 | 596.153 | 596.1929 | 596.5726 | 369.9789
2 | -7.274 | -0.35969 | 0.15173 | -0.13554 | 0.29425 | 0.19618 | 0.19616 | 0.52422 | 0.64999 | 1.0957 | 1.0113 | 0.4882 | 0.3978 | -0.011723 | 0.3627 | 0.35389 | -0.30636 | 0.13042 | 0.53091 | 1.0794 | 0.10297
3 | 40.4441 | -3.4294 | -3.3258 | -3.5473 | -2.4865 | -4.224 | -2.4941 | -2.2371 | -3.0804 | -3.003 | -2.6787 | -2.9059 | -3.005 | -4.5108 | -3.9684 | -2.7945 | -2.5155 | -2.5482 | -2.3164 | -0.63526 | -1.654
4 | -2.8361 | 0.063517 | -0.09483 | -0.14111 | 0.027299 | -0.017447 | -0.20615 | 0.053564 | -0.00048529 | 0.19486 | 0.15883 | 0.32671 | 0.095698 | 0.058499 | 0.070156 | 0.10878 | -0.0082872 | -0.028395 | -0.037954 | 0.07815 | -0.051546
5 | -5.2236 | -0.88593 | -0.91297 | -0.57803 | -0.50902 | -0.40664 | -0.48844 | -0.46696 | -0.5428 | -0.3856 | -0.40121 | -0.3946 | -0.4021 | -0.4402 | -0.26573 | -0.46818 | -0.50377 | -0.50929 | -0.66432 | -3.8742 | -0.76491
6 | -0.48881 | 0.10803 | -0.0081412 | 0.089 | 0.027036 | -0.0064957 | -0.0034309 | -0.042115 | 0.023066 | -0.0062258 | 0.010442 | 0.071445 | -0.0065785 | 0.11937 | 0.059659 | 0.028141 | -0.0013195 | -0.0012726 | -0.010684 | 0.00049599 | 0.00853
7 | 2.9075 | 0.92093 | 0.98924 | 1.0631 | 1.0585 | 1.0096 | 1.0691 | 1.0541 | 1.0461 | 1.0573 | 1.0644 | 1.0706 | 1.0932 | 1.0378 | 1.0556 | 1.058 | 1.0262 | 1.0063 | 1.0499 | 0.93652 | 1.0941
8 | -0.12093 | -0.0049353 | 0.0097691 | -0.010163 | 0.034641 | 0.028011 | -0.012209 | 0.015462 | 0.0082794 | 0.03718 | 0.033516 | 0.019649 | 0.015408 | 0.0069722 | 0.035393 | 0.0073221 | 0.017119 | 0.030458 | 0.01223 | 0.015834 | 0.011614
9 | 0.40127 | 0.44958 | 0.45367 | 0.4664 | 0.47716 | 0.43831 | 0.46286 | 0.46466 | 0.45224 | 0.46248 | 0.46783 | 0.44735 | 0.4655 | 0.45948 | 0.44805 | 0.46759 | 0.44942 | 0.46224 | 0.47099 | 0.42922 | 0.45648
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.63952 | 0.74942 | 0.748 | 0.75291 | 0.7524 | 0.75835 | 0.74686 | 0.74694 | 0.74598 | 0.75285 | 0.75223 | 0.75177 | 0.74955 | 0.76185 | 0.75828 | 0.75078 | 0.75041 | 0.75405 | 0.74989 | 0.66004 | 0.74541
12 | 0.023985 | -0.035786 | 0.0065548 | -0.0084037 | 0.0090156 | -0.010853 | 0.026208 | -0.027383 | 0.015928 | 0.02415 | 0.029566 | 0.023557 | 0.028265 | 0.031562 | 0.061927 | 0.0092334 | 0.023308 | 0.021171 | 0.014277 | 0.020884 | 0.001402 |  | ten times!
13 | 0.93393 | 0.77391 | 0.77103 | 0.74058 | 0.7518 | 0.71882 | 0.74367 | 0.72978 | 0.72346 | 0.74391 | 0.75914 | 0.75091 | 0.74194 | 0.72287 | 0.73292 | 0.75805 | 0.75506 | 0.77072 | 0.77351 | 0.63324 | 0.074987 |  | ten times!
14 | 0.02336 | -0.038445 | -0.033 | -0.025645 | -0.029783 | -0.02996 | -0.02438 | -0.030784 | -0.027769 | -0.024148 | -0.019977 | -0.024769 | -0.021965 | -0.030128 | -0.037557 | -0.032852 | -0.023274 | -0.025978 | -0.021087 | -0.017513 | -0.0025964 |  | ten times!
15 | -0.13194 | 0.24321 | 0.24152 | 0.23004 | 0.22097 | 0.21884 | 0.20596 | 0.2193 | 0.23844 | 0.2533 | 0.22523 | 0.23354 | 0.21543 | 0.20165 | 0.16757 | 0.20654 | 0.23626 | 0.24374 | 0.23218 | 0.048803 | 0.020757 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 39.4963 | -9.3194 | -4.822 | -8.6061 | -5.6015 | -2.802 | -0.4719 | 6.9998 | 0.35984 | -1.6843 | -4.4164 | -0.76465 | 3.3081 | 3.1525 | 0.45625 | -0.813 | -4.8963 | -4.9886 | 2.9844 | 14.1497 | -0.00023572
2 | 12.3942 | -0.38544 | 1.1072 | -0.032234 | 0.81298 | 0.18232 | 0.46951 | 0.78018 | 1.5046 | 0.33864 | 0.70441 | -0.82341 | 0.88426 | -0.4266 | -0.18235 | -0.47938 | 0.19874 | -0.20091 | -0.56619 | 9.8636 | 0.89913
3 | -0.34447 | 0.74938 | 0.79007 | 0.84277 | 0.50556 | 0.25077 | -0.19704 | 0.12834 | -0.023507 | 0.67752 | 0.49087 | 0.87406 | 0.68431 | 0.95873 | 0.99274 | 0.84391 | 1.0587 | 1.1608 | 1.2104 | 0.20294 | 0.62079
4 | 1.5077 | 0.023735 | -0.004416 | -0.075165 | 0.064936 | 0.45499 | 0.26869 | 0.2534 | 0.24239 | 0.49944 | 0.1473 | 0.50416 | 0.32999 | 0.50934 | 0.45938 | 0.096878 | 0.41494 | -0.15746 | -0.00057909 | 0.92561 | 0.28581
5 | 2.5116 | 0.056367 | 0.19148 | 0.33358 | 0.12306 | -0.068668 | -0.048011 | 0.067193 | -0.0070148 | 0.11413 | 0.16066 | 0.047029 | 0.12827 | 0.33584 | 0.20968 | 0.10635 | 0.27576 | 0.3644 | 0.4668 | -0.18034 | 0.21955
6 | -0.013936 | -0.061833 | -0.016256 | -0.026737 | -0.032178 | -0.020907 | 0.019601 | -0.0063053 | -0.029224 | 0.015157 | 0.062074 | 0.11018 | 0.16983 | 0.24964 | 0.19556 | 0.057793 | 0.14544 | -0.019166 | 0.032944 | 0.085462 | 0.046357
7 | 1.4226 | -0.093717 | -0.074679 | -0.094449 | -0.034405 | 0.0024494 | -0.089262 | 0.039531 | -0.018363 | -0.025314 | -0.051854 | -0.088976 | -0.037952 | -0.060009 | -0.077082 | -0.087581 | 0.012609 | 0.042779 | 0.051695 | -0.029211 | 0.0068162
8 | -0.048968 | 0.033096 | 0.011434 | -0.0048137 | 0.003278 | 0.011022 | 0.025796 | 0.0061788 | 0.0059429 | 0.015153 | 0.019854 | 0.016951 | 0.031894 | 0.027235 | -0.01269 | 0.018357 | 0.017761 | -0.0058313 | 0.0047221 | 0.043924 | 0.011658
9 | -0.29835 | -0.040922 | -0.015651 | -0.0094252 | -0.029164 | -0.049666 | -0.043765 | -0.021923 | -0.036773 | -0.051053 | -0.012895 | -0.014805 | 0.012677 | -0.0082852 | 0.0076984 | -0.019469 | -0.00084121 | 0.013338 | 0.029849 | -0.048397 | -0.025819
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.15381 | -0.072005 | -0.057742 | -0.05829 | -0.055093 | -0.050929 | -0.04757 | -0.039069 | -0.047935 | -0.050013 | -0.053513 | -0.052157 | -0.044176 | -0.044097 | -0.056587 | -0.059049 | -0.047086 | -0.044001 | -0.033359 | -0.032217 | -0.043867
12 | -0.094226 | -0.034235 | -0.025997 | -0.043359 | -0.019926 | -0.045205 | 0.0021233 | 0.0039537 | -0.01538 | -0.0047805 | 0.034635 | 0.013534 | -0.0016446 | 0.013619 | 0.0081043 | -0.001095 | 0.025868 | -0.016276 | 0.011508 | 0.0050243 | -0.00076352 |  | ten times!
13 | -0.23815 | -0.053726 | -0.034855 | -0.038228 | -0.042695 | -0.060743 | -0.071133 | -0.027427 | -0.061157 | -0.049235 | -0.038217 | -0.033573 | -0.0098766 | -0.020421 | -0.023786 | -0.03922 | -0.029227 | 0.012072 | 0.026021 | -0.031741 | -0.003929 |  | ten times!
14 | -0.021547 | -0.03029 | -0.019834 | -0.023238 | -0.0055762 | -0.0097924 | -0.0028974 | -0.0097454 | 0.00839 | 0.0031331 | -0.0024783 | -0.030594 | -0.030679 | -0.046763 | -0.038792 | -0.018791 | -0.010502 | -0.039694 | -0.02256 | 0.039306 | -0.0016606 |  | ten times!
15 | 0.051777 | -0.015323 | -0.024772 | -0.022227 | -0.022573 | -0.031643 | -0.044762 | -0.013749 | -0.019789 | -0.021352 | -0.038969 | -0.0049137 | -0.029374 | -0.054562 | -0.034562 | -0.033332 | 0.0018979 | 0.0098887 | 0.020154 | -0.040883 | -0.0019522 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 341.485 | 596.3625 | 595.9815 | 596.0841 | 596.0229 | 596.1104 | 596.0713 | 596.0843 | 596.1622 | 596.1072 | 596.0608 | 596.0572 | 595.9949 | 595.9731 | 596.0022 | 596.0278 | 596.1236 | 596.073 | 596.4382 | 374.1237
2 | -0.34384 | -0.59008 | -0.77743 | -1.2799 | -1.9413 | -0.47523 | -0.39177 | -1.2791 | -0.63913 | -1.7537 | -1.1434 | -0.61992 | -0.44889 | -0.79555 | -0.55989 | -0.239 | -0.6025 | -0.49059 | -0.12792 | -0.064772 | -0.74959
3 | 44.0887 | -3.9589 | -3.564 | -3.5694 | -2.1281 | -3.5195 | -2.3188 | -2.7139 | -1.9499 | -2.7571 | -3.2732 | -3.2642 | -3.2608 | -4.7591 | -4.1428 | -2.9494 | -3.1828 | -3.2529 | -2.6633 | -0.72637 | -1.6888
4 | -0.22205 | -0.0012169 | -0.043002 | -0.12485 | -0.30728 | -0.18144 | -0.092545 | -0.2313 | 0.055196 | -0.22455 | -0.073628 | -0.16052 | 0.0092025 | -0.17093 | -0.14048 | -0.067649 | 0.073107 | 0.0021539 | -0.014476 | -0.38837 | -0.10755
5 | -5.0692 | -0.61022 | -0.62771 | -0.19724 | -0.26563 | -0.055597 | -0.10851 | -0.092044 | -0.36831 | -0.21845 | -0.076011 | -0.17899 | -0.34993 | -0.27302 | -0.37911 | -0.34761 | -0.39625 | -0.48354 | -0.56717 | -3.9403 | -0.57147
6 | -0.33734 | 0.022831 | 0.014989 | -0.070181 | -0.025785 | -0.083705 | 0.014741 | -0.0059285 | 0.039267 | -0.027212 | 0.033937 | -0.053438 | -0.06632 | 0.070683 | 0.03652 | -0.016306 | -0.01171 | -0.037095 | 0.0031666 | 0.018908 | -0.01786
7 | 2.9152 | 0.91256 | 0.98685 | 1.1428 | 1.0675 | 1.0708 | 1.0934 | 1.0547 | 1.0804 | 1.0905 | 1.0464 | 1.0498 | 1.0727 | 1.0604 | 1.0235 | 1.0552 | 1.0204 | 1.0342 | 1.0259 | 0.84896 | 1.0984
8 | -0.10285 | -0.017351 | -0.0023 | -0.011753 | 0.010427 | 0.029936 | 0.0012528 | -0.0081015 | -0.0094213 | -0.013776 | 0.0012997 | -0.024048 | 0.0012554 | 0.0072387 | 0.0064774 | 0.013194 | -0.013802 | 0.014431 | 0.0355 | 0.0015479 | -0.001952
9 | 0.37495 | 0.46425 | 0.45661 | 0.46933 | 0.49068 | 0.45067 | 0.47803 | 0.44837 | 0.47754 | 0.46669 | 0.45411 | 0.43311 | 0.44653 | 0.44169 | 0.44351 | 0.45653 | 0.4349 | 0.45359 | 0.45846 | 0.39174 | 0.45235
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.60526 | 0.74803 | 0.7464 | 0.75738 | 0.75085 | 0.76067 | 0.75009 | 0.75193 | 0.75665 | 0.75881 | 0.76098 | 0.75504 | 0.74813 | 0.75283 | 0.75858 | 0.74731 | 0.75022 | 0.74763 | 0.74577 | 0.64555 | 0.74473
12 | 0.052039 | -0.021199 | -0.0066274 | -0.0048989 | 0.020228 | 0.052006 | 0.056374 | 0.0083719 | -0.027073 | -0.029067 | -0.021129 | -0.0039824 | 0.0065463 | 0.013347 | 0.00088502 | 0.017608 | -0.0068203 | 0.0047445 | 0.03489 | 0.033283 | 0.00075469 |  | ten times!
13 | 0.91005 | 0.72985 | 0.73753 | 0.72268 | 0.71176 | 0.70189 | 0.69544 | 0.70399 | 0.73339 | 0.72766 | 0.69625 | 0.73263 | 0.7262 | 0.70092 | 0.71184 | 0.73198 | 0.73937 | 0.7427 | 0.72752 | 0.58396 | 0.072193 |  | ten times!
14 | 0.074419 | -0.069642 | -0.067165 | -0.066325 | -0.053573 | -0.057356 | -0.050806 | -0.060854 | -0.065883 | -0.068035 | -0.05796 | -0.047998 | -0.061359 | -0.052364 | -0.067762 | -0.058628 | -0.065645 | -0.070251 | -0.060863 | -0.034859 | -0.0056338 |  | ten times!
15 | -0.14357 | 0.24932 | 0.23361 | 0.21686 | 0.1997 | 0.18552 | 0.21092 | 0.21477 | 0.22323 | 0.23743 | 0.21112 | 0.23518 | 0.18231 | 0.16524 | 0.13332 | 0.18281 | 0.235 | 0.24137 | 0.20038 | -0.030462 | 0.019045 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 28.5737 | -8.467 | -3.4741 | -6.5351 | -4.9682 | -4.1811 | -1.9702 | 5.1882 | 2.1978 | -0.39233 | -3.9807 | -3.2074 | 1.9625 | 4.0206 | 0.5272 | -1.3808 | -4.2113 | -4.7158 | 4.0945 | 20.5557 | -0.011644
2 | 12.8096 | 0.012211 | 0.8832 | -0.77113 | -0.032134 | -0.58066 | -0.1524 | 0.12301 | -0.1945 | -0.91317 | -0.92388 | -1.525 | -0.70486 | -0.013227 | -1.1882 | -0.078804 | -0.031505 | -0.70817 | -0.08254 | -9.1219 | -0.27438
3 | -1.2258 | 0.72977 | 0.83867 | 0.53292 | 0.74595 | 0.99694 | 0.57394 | 0.64523 | 0.59672 | 0.91573 | 0.8877 | 0.73201 | 0.17598 | 0.86739 | 0.53754 | 0.34306 | 0.67673 | 0.40733 | 0.85324 | 0.69685 | 0.61416
4 | 2.1697 | -0.47401 | -0.75739 | -0.64305 | -0.67869 | -0.36585 | -0.64571 | -0.42364 | -0.61175 | -0.32094 | -0.53041 | -0.29605 | -0.14349 | 0.14752 | 0.21742 | -0.30174 | -0.13579 | -0.58426 | -0.020164 | -0.48166 | -0.29308
5 | 2.7223 | 0.10043 | 0.083418 | 0.087729 | 0.23799 | 0.14736 | 0.057447 | 0.42741 | 0.27352 | 0.46228 | 0.34581 | 0.12995 | -0.034034 | 0.3173 | 0.13622 | -0.040088 | 0.11551 | 0.10727 | 0.35438 | -0.21276 | 0.24666
6 | 0.34214 | -0.0061387 | 0.0059694 | -0.1463 | 0.025482 | 0.047516 | 0.104 | 0.013129 | -0.00032771 | -0.038025 | 0.047039 | 0.11366 | 0.18275 | 0.28969 | 0.12377 | 0.1893 | 0.14396 | -0.027815 | 0.16177 | -0.018068 | 0.073647
7 | 1.8221 | -0.051761 | -0.026596 | 0.012846 | -0.010815 | 0.013296 | -0.0563 | 0.087247 | 0.026787 | 0.086035 | 0.040036 | 0.033199 | -0.053944 | 0.076586 | -0.024911 | -0.079137 | -0.023511 | -0.032018 | 0.041749 | 0.034346 | 0.05855
8 | -0.062357 | -0.024875 | -0.025619 | -0.12087 | -0.046078 | -0.057439 | -0.037031 | -0.041416 | -0.028807 | -0.043663 | -0.032318 | -0.026717 | 0.0054268 | 0.040872 | -0.019105 | 0.02176 | -0.031809 | -0.047687 | -0.010213 | 0.0052562 | -0.029063
9 | -0.22558 | 0.00042411 | -0.0083255 | 0.0044289 | -0.00073309 | 0.012761 | 0.0023083 | 0.050402 | 0.022915 | 0.027707 | 0.020148 | 0.00033317 | 0.0064625 | 0.0036789 | -0.028568 | -0.0090451 | 0.0019397 | -0.0021223 | 0.045827 | -0.020597 | 0.00043852
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.19973 | -0.031248 | -0.016106 | -0.010931 | -0.011582 | -0.014477 | -0.016603 | 0.0098242 | -0.0011312 | 0.0017511 | -0.0057832 | -0.010569 | -0.011975 | 0.0039921 | -0.011476 | -0.015445 | -0.013373 | -0.017609 | 0.0033307 | 0.010758 | -0.0025125
12 | -0.044693 | -0.0075594 | -0.0053798 | -0.12558 | -0.036237 | -0.06098 | -0.024313 | -0.067881 | -0.045889 | -0.063561 | -0.036911 | -0.020411 | -0.021141 | -0.019343 | -0.042613 | 0.02649 | -0.010572 | -0.020614 | 0.041569 | 0.030902 | -0.0028494 |  | ten times!
13 | -0.19098 | 0.024961 | 0.017965 | 0.012546 | 0.035711 | 0.057117 | 0.025316 | 0.070556 | 0.036132 | 0.061957 | 0.041141 | 0.034249 | 0.017955 | 0.045468 | -0.016217 | 0.00027263 | 0.024268 | 0.022622 | 0.078783 | 0.031794 | 0.0026113 |  | ten times!
14 | -0.041898 | -0.060037 | -0.06938 | -0.079224 | -0.068202 | -0.1026 | -0.077177 | -0.080735 | -0.07602 | -0.082896 | -0.097999 | -0.082825 | -0.075001 | -0.053253 | -0.062193 | -0.073115 | -0.055578 | -0.063286 | -0.053864 | -0.0015492 | -0.0069704 |  | ten times!
15 | 0.036179 | -0.0097915 | -0.019821 | -0.02228 | -8.7557e-05 | -0.023259 | -0.036299 | 0.018564 | -0.0052541 | -0.014923 | -0.027339 | -0.00079306 | -0.05657 | -0.057646 | -0.063607 | -0.038617 | -0.014882 | -0.0080416 | 0.01146 | -0.12606 | -0.0022267 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00054138 | -0.0001298 | -8.5853e-05 | -0.00014977 | -4.3243e-06 | -0.00013517 | -0.00010816 | -0.00010162 | -2.932e-05 | -5.4986e-05 | -4.274e-05 | -0.00026904 | -0.00017376 | -0.00013484 | -0.00011991 | 7.5695e-05 | 0.00014263 | 0.00077959 |  | 596.1079
|  | 0.00044622 | -0.00019294 | -2.0821e-05 | -0.00012349 | 2.33e-05 | -4.2294e-05 | -2.0485e-05 | 0.0001102 | 1.7931e-05 | -5.9908e-05 | -6.5948e-05 | -0.00017046 | -0.00020703 | -0.00015821 | -0.00011527 | 4.5444e-05 | -3.9442e-05 | 0.00057321 |  | 596.0965
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3966.5628
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 596.0587 | -4140.2143 | 6.2399 | -0.47248 | -0.032925 | -0.67212 | 0.78243 | -0.52616 | -0.25606 | -0.19063 | 0.53246 | 0.27577 | 0.39824 | -1.865 | -0.91203 | -0.52281 | -0.3735 | 1.5827 | 2.2521 | 8.6223 | -3792.9113
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.17736 | 4.127 | -0.75458 | -0.30484 | -0.68325 | -0.38279 | -0.10284 | 0.13017 | 0.87734 | 0.21335 | 0.0089378 | -0.26427 | 0.1009 | 0.50817 | 0.49262 | 0.22299 | 0.096063 | -0.31227 | -0.3215 | 0.4758 | 1.5923 |  | 0
db2 | 0.30945 | -7.274 | -0.66914 | -0.15772 | -0.44499 | -0.015201 | -0.11327 | -0.11329 | 0.21477 | 0.34054 | 0.78622 | 0.70182 | 0.17875 | 0.088353 | -0.32117 | 0.053251 | 0.044444 | -0.61581 | -0.17903 | 0.22146 | 1.0794 |  | 0
db3 | -3.0591 | 40.4441 | -0.36988 | -0.26634 | -0.48776 | 0.57296 | -1.1645 | 0.5654 | 0.82244 | -0.02091 | 0.056468 | 0.38084 | 0.15357 | 0.054502 | -1.4513 | -0.90886 | 0.26501 | 0.54399 | 0.51133 | 0.74307 | -0.63526 |  | -1.727e-16
db4 | 0.034625 | -2.8361 | 0.028892 | -0.12945 | -0.17573 | -0.0073257 | -0.052072 | -0.24077 | 0.018938 | -0.03511 | 0.16023 | 0.1242 | 0.29208 | 0.061073 | 0.023874 | 0.035531 | 0.074159 | -0.042912 | -0.063021 | -0.072579 | 0.07815 |  | 0
db5 | -0.51193 | -5.2236 | -0.37339 | -0.40043 | -0.065491 | 0.003527 | 0.10591 | 0.024108 | 0.045588 | -0.030253 | 0.12694 | 0.11133 | 0.11794 | 0.11044 | 0.072347 | 0.24682 | 0.044359 | 0.0087753 | 0.0032532 | -0.15178 | -3.8742 |  | 5.2427e-17
db6 | 0.024996 | -0.48881 | 0.083035 | -0.033137 | 0.064004 | 0.00204 | -0.031491 | -0.028427 | -0.06711 | -0.0019297 | -0.031221 | -0.014554 | 0.046449 | -0.031574 | 0.094371 | 0.034663 | 0.0031452 | -0.026315 | -0.026268 | -0.03568 | 0.00049599 |  | 0
db7 | 1.0413 | 2.9075 | -0.11962 | -0.051316 | 0.022589 | 0.017917 | -0.030911 | 0.028519 | 0.013536 | 0.0054993 | 0.016749 | 0.023846 | 0.030021 | 0.052628 | -0.0027959 | 0.015079 | 0.017425 | -0.014337 | -0.034205 | 0.009379 | 0.93652 |  | -6.1679e-18
db8 | 0.015783 | -0.12093 | -0.020719 | -0.0060142 | -0.025946 | 0.018857 | 0.012227 | -0.027993 | -0.00032185 | -0.0075039 | 0.021396 | 0.017733 | 0.0038655 | -0.00037528 | -0.0088112 | 0.01961 | -0.0084613 | 0.0013354 | 0.014675 | -0.0035533 | 0.015834 |  | 1.7347e-18
db9 | 0.46123 | 0.40127 | -0.0096331 | -0.0055436 | 0.0071859 | 0.017951 | -0.020898 | 0.0036516 | 0.0054467 | -0.0069687 | 0.0032704 | 0.0086223 | -0.011867 | 0.0062911 | 0.00027036 | -0.011165 | 0.0083795 | -0.0097909 | 0.0030249 | 0.011773 | 0.42922 |  | -1.542e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.75196 | 0.63952 | -0.0023827 | -0.0038069 | 0.0011011 | 0.00059117 | 0.0065399 | -0.0049434 | -0.004865 | -0.005824 | 0.0010463 | 0.00042281 | -3.6582e-05 | -0.0022546 | 0.01004 | 0.0064776 | -0.0010248 | -0.0014004 | 0.002239 | -0.0019196 | 0.66004 |  | 6.7847e-17
db12 | 0.0013461 | 0.0023985 | -0.0049247 | -0.00069062 | -0.0021865 | -0.00044453 | -0.0024314 | 0.0012747 | -0.0040844 | 0.00024669 | 0.0010689 | 0.0016105 | 0.0010096 | 0.0014804 | 0.0018101 | 0.0048466 | -0.00042275 | 0.00098466 | 0.00077103 | 8.1625e-05 | 0.0020884 |  | -1.2047e-19
db13 | 0.074789 | 0.093393 | 0.0026015 | 0.002314 | -0.0007314 | 0.00039085 | -0.0029074 | -0.00042251 | -0.0018117 | -0.0024437 | -0.0003987 | 0.0011248 | 0.00030128 | -0.00059556 | -0.0025022 | -0.001497 | 0.001016 | 0.00071707 | 0.0022831 | 0.0025615 | 0.063324 |  | -1.0023e-17
db14 | -0.0027861 | 0.002336 | -0.0010583 | -0.00051389 | 0.00022158 | -0.00019223 | -0.00020991 | 0.00034813 | -0.00029227 | 9.2438e-06 | 0.0003713 | 0.00078844 | 0.00030922 | 0.00058957 | -0.00022669 | -0.00096954 | -0.00049906 | 0.00045873 | 0.00018827 | 0.00067746 | -0.0017513 |  | -3.3731e-19
db15 | 0.022409 | -0.013194 | 0.0019113 | 0.0017423 | 0.00059477 | -0.00031262 | -0.00052578 | -0.0018137 | -0.00047912 | 0.0014347 | 0.0029205 | 0.00011313 | 0.00094411 | -0.00086634 | -0.0022448 | -0.0056521 | -0.0017554 | 0.0012167 | 0.001964 | 0.00080822 | 0.0048803 |  | -1.9275e-19
da2 | 0.21591 | 12.3942 | -0.60135 | 0.89134 | -0.24814 | 0.59707 | -0.033594 | 0.2536 | 0.56427 | 1.2887 | 0.12273 | 0.4885 | -1.0393 | 0.66835 | -0.64251 | -0.39826 | -0.69529 | -0.017169 | -0.41682 | -0.7821 | 9.8636 |  | 0
da3 | 0.66658 | -0.34447 | 0.082808 | 0.12349 | 0.1762 | -0.16102 | -0.4158 | -0.86361 | -0.53824 | -0.69008 | 0.010945 | -0.17571 | 0.20749 | 0.017729 | 0.29216 | 0.32616 | 0.17733 | 0.39213 | 0.49425 | 0.54379 | 0.20294 |  | 0
da4 | 0.224 | 1.5077 | -0.20026 | -0.22841 | -0.29916 | -0.15906 | 0.23099 | 0.044695 | 0.029402 | 0.01839 | 0.27545 | -0.076693 | 0.28016 | 0.10599 | 0.28534 | 0.23538 | -0.12712 | 0.19095 | -0.38146 | -0.22458 | 0.92561 |  | 0
da5 | 0.15872 | 2.5116 | -0.10235 | 0.032764 | 0.17487 | -0.035657 | -0.22739 | -0.20673 | -0.091525 | -0.16573 | -0.044592 | 0.0019456 | -0.11169 | -0.030446 | 0.17712 | 0.050961 | -0.052366 | 0.11705 | 0.20568 | 0.30809 | -0.18034 |  | 0
da6 | 0.046978 | -0.013936 | -0.10881 | -0.063234 | -0.073715 | -0.079156 | -0.067885 | -0.027377 | -0.053283 | -0.076202 | -0.031821 | 0.015096 | 0.0632 | 0.12285 | 0.20266 | 0.14859 | 0.010815 | 0.098458 | -0.066144 | -0.014034 | 0.085462 |  | -5.0114e-18
da7 | -0.038032 | 1.4226 | -0.055685 | -0.036647 | -0.056417 | 0.0036275 | 0.040482 | -0.05123 | 0.077563 | 0.019669 | 0.012718 | -0.013822 | -0.050943 | 8.0176e-05 | -0.021977 | -0.039049 | -0.049549 | 0.050641 | 0.080811 | 0.089727 | -0.029211 |  | 0
da8 | 0.012519 | -0.048968 | 0.020577 | -0.001085 | -0.017333 | -0.0092409 | -0.0014968 | 0.013277 | -0.00634 | -0.0065759 | 0.0026343 | 0.0073356 | 0.0044326 | 0.019375 | 0.014717 | -0.025209 | 0.0058377 | 0.0052417 | -0.01835 | -0.0077967 | 0.043924 |  | 0
da9 | -0.016171 | -0.29835 | -0.024751 | 0.00051985 | 0.0067458 | -0.012994 | -0.033495 | -0.027595 | -0.0057525 | -0.020602 | -0.034882 | 0.003276 | 0.0013661 | 0.028848 | 0.0078858 | 0.023869 | -0.0032982 | 0.01533 | 0.029509 | 0.04602 | -0.048397 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.050704 | 0.15381 | -0.021301 | -0.0070381 | -0.0075857 | -0.0043893 | -0.00022487 | 0.0031339 | 0.011635 | 0.002769 | 0.00069055 | -0.0028087 | -0.0014534 | 0.006528 | 0.0066069 | -0.0058829 | -0.0083449 | 0.003618 | 0.0067031 | 0.017345 | -0.032217 |  | 6.1679e-18
da12 | -0.00052528 | -0.0094226 | -0.0028982 | -0.0020744 | -0.0038107 | -0.0014673 | -0.0039952 | 0.00073761 | 0.00092066 | -0.0010127 | 4.7237e-05 | 0.0039888 | 0.0018787 | 0.00036082 | 0.0018872 | 0.0013357 | 0.00041579 | 0.0031121 | -0.0011023 | 0.0016761 | 0.00050243 |  | 0
da13 | -0.0033079 | -0.023815 | -0.0020646 | -0.00017761 | -0.00051484 | -0.00096161 | -0.0027664 | -0.0038054 | 0.00056521 | -0.0028078 | -0.0016156 | -0.00051377 | -4.9357e-05 | 0.0023203 | 0.0012659 | 0.00092931 | -0.00061405 | 0.00038523 | 0.0045152 | 0.00591 | -0.0031741 |  | 0
da14 | -0.0018372 | -0.0021547 | -0.0011917 | -0.0001462 | -0.00048659 | 0.0012796 | 0.000858 | 0.0015475 | 0.00086269 | 0.0026762 | 0.0021505 | 0.0015894 | -0.0012222 | -0.0012306 | -0.0028391 | -0.0020419 | -4.1831e-05 | 0.00078705 | -0.0021322 | -0.00041871 | 0.0039306 |  | -1.2047e-19
da15 | -0.0021109 | 0.0051777 | 0.0005786 | -0.00036628 | -0.00011176 | -0.00014637 | -0.0010534 | -0.0023653 | 0.00073599 | 0.00013197 | -2.4261e-05 | -0.001786 | 0.0016195 | -0.00082646 | -0.0033453 | -0.0013453 | -0.0012223 | 0.0023007 | 0.0030998 | 0.0041263 | -0.0040883 |  | 0
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -3997.1608
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.0585 | -4270.9485 | 5.0996 | -1.2923 | 0.42896 | -0.59778 | 0.8702 | 0.21422 | 0.43232 | 1.7392 | 0.81651 | 0.038063 | -0.022334 | -1.0675 | -1.4333 | -0.94506 | -0.51557 | 1.0917 | 0.24274 | 6.3697 | -3723.3731
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.16385 | 3.0212 | -0.68284 | -0.18356 | -0.48966 | -0.33297 | -0.25426 | -0.033171 | 0.68267 | 0.38364 | 0.12462 | -0.23422 | -0.15689 | 0.3601 | 0.56591 | 0.21657 | 0.025768 | -0.25728 | -0.30773 | 0.5733 | 2.2194 |  | 0
db2 | 0.7864 | -0.34384 | 0.19633 | 0.0089711 | -0.49351 | -1.1549 | 0.31118 | 0.39463 | -0.49267 | 0.14727 | -0.96728 | -0.357 | 0.16648 | 0.33751 | -0.0091427 | 0.22652 | 0.5474 | 0.1839 | 0.29581 | 0.65849 | -0.064772 |  | 0
db3 | -3.1789 | 44.0887 | -0.77956 | -0.38469 | -0.39002 | 1.0512 | -0.34016 | 0.86058 | 0.4654 | 1.2294 | 0.42224 | -0.093893 | -0.084828 | -0.081433 | -1.5798 | -0.96349 | 0.22992 | -0.003431 | -0.073533 | 0.51602 | -0.72637 |  | 1.3569e-16
db4 | 0.094122 | -0.22205 | 0.092906 | 0.051121 | -0.030729 | -0.21316 | -0.087315 | 0.0015779 | -0.13718 | 0.14932 | -0.13043 | 0.020495 | -0.066396 | 0.10332 | -0.076811 | -0.046353 | 0.026474 | 0.16723 | 0.096276 | 0.079647 | -0.38837 |  | 0
db5 | -0.31036 | -5.0692 | -0.29937 | -0.31685 | 0.11362 | 0.045219 | 0.25526 | 0.20235 | 0.21881 | -0.057454 | 0.092399 | 0.23484 | 0.13186 | -0.039082 | 0.037835 | -0.068254 | -0.036762 | -0.085393 | -0.17269 | -0.25632 | -3.9403 |  | 0
db6 | 0.0089747 | -0.33734 | 0.031806 | 0.023964 | -0.061206 | -0.01681 | -0.07473 | 0.023716 | 0.0030462 | 0.048242 | -0.018237 | 0.042911 | -0.044463 | -0.057346 | 0.079658 | 0.045495 | -0.0073312 | -0.0027358 | -0.02812 | 0.012141 | 0.018908 |  | 0
db7 | 1.0498 | 2.9152 | -0.13678 | -0.062485 | 0.093488 | 0.01817 | 0.021494 | 0.04404 | 0.0054002 | 0.031045 | 0.041161 | -0.0029761 | 0.00044013 | 0.023412 | 0.011088 | -0.025852 | 0.0058386 | -0.028939 | -0.015106 | -0.023442 | 0.84896 |  | -4.9343e-17
db8 | -0.0011367 | -0.10285 | -0.018488 | -0.0034368 | -0.012889 | 0.0092906 | 0.028799 | 0.00011607 | -0.0092382 | -0.010558 | -0.014913 | 0.00016303 | -0.025185 | 0.00011874 | 0.006102 | 0.0053407 | 0.012058 | -0.014939 | 0.013295 | 0.034364 | 0.0015479 |  | 0
db9 | 0.45976 | 0.37495 | 0.0073276 | -0.00031584 | 0.01241 | 0.033758 | -0.0062533 | 0.021107 | -0.008551 | 0.020622 | 0.0097684 | -0.0028107 | -0.023811 | -0.010393 | -0.015236 | -0.013414 | -0.00039313 | -0.022018 | -0.00333 | 0.0015347 | 0.39174 |  | -3.3923e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.75286 | 0.60526 | -0.0045961 | -0.0062307 | 0.0047555 | -0.001781 | 0.0080463 | -0.0025419 | -0.00070302 | 0.004019 | 0.0061851 | 0.0083547 | 0.0024091 | -0.0044939 | 0.00019717 | 0.0059549 | -0.0053153 | -0.0024048 | -0.0049963 | -0.0068589 | 0.64555 |  | 8.0183e-17
db12 | -0.00052335 | 0.0052039 | -0.0026433 | -0.0011861 | -0.0010132 | 0.0014994 | 0.0046773 | 0.005114 | 0.00031383 | -0.0032306 | -0.00343 | -0.0026362 | -0.00092159 | 0.00013127 | 0.00081136 | -0.00043485 | 0.0012375 | -0.0012054 | -4.8907e-05 | 0.0029656 | 0.0033283 |  | 0
db13 | 0.072076 | 0.091005 | 0.00090918 | 0.0016774 | 0.00019231 | -0.00089975 | -0.0018863 | -0.0025318 | -0.0016766 | 0.0012637 | 0.0006909 | -0.0024502 | 0.0011871 | 0.00054408 | -0.0019839 | -0.00089121 | 0.0011228 | 0.0018614 | 0.0021943 | 0.00067648 | 0.058396 |  | 6.1679e-18
db14 | 0.0061248 | 0.0074419 | -0.00083934 | -0.00059172 | -0.00050767 | 0.00076756 | 0.00038926 | 0.0010442 | 3.9391e-05 | -0.0004635 | -0.00067864 | 0.00032882 | 0.001325 | -1.1037e-05 | 0.00088844 | -0.00065135 | 0.00026207 | -0.00043972 | -0.0009003 | 3.8531e-05 | -0.0034859 |  | 1.1565e-18
db15 | 0.020878 | -0.014357 | 0.0040537 | 0.0024823 | 0.00080769 | -0.00090784 | -0.0023262 | 0.00021379 | 0.00059895 | 0.001445 | 0.002865 | 0.00023358 | 0.0026398 | -0.0026468 | -0.004354 | -0.0075468 | -0.0025976 | 0.0026213 | 0.0032582 | -0.0008401 | -0.0030462 |  | 0
da2 | -0.38232 | 12.8096 | 0.39453 | 1.2655 | -0.38882 | 0.35018 | -0.19834 | 0.22991 | 0.50533 | 0.18782 | -0.53085 | -0.54156 | -1.1426 | -0.32255 | 0.36909 | -0.80586 | 0.30351 | 0.35081 | -0.32585 | 0.29978 | -9.1219 |  | 0
da3 | 0.66983 | -1.2258 | 0.059946 | 0.16885 | -0.13691 | 0.076124 | 0.32712 | -0.095882 | -0.024594 | -0.073109 | 0.24591 | 0.21787 | 0.062183 | -0.49384 | 0.19756 | -0.13229 | -0.32676 | 0.0069046 | -0.26249 | 0.18341 | 0.69685 |  | 0
da4 | -0.36489 | 2.1697 | -0.10912 | -0.3925 | -0.27816 | -0.3138 | -0.00096342 | -0.28083 | -0.058752 | -0.24686 | 0.043945 | -0.16552 | 0.068841 | 0.2214 | 0.51241 | 0.58231 | 0.063152 | 0.2291 | -0.21937 | 0.34472 | -0.48166 |  | 0
da5 | 0.18388 | 2.7223 | -0.083449 | -0.10046 | -0.096154 | 0.054102 | -0.036518 | -0.12644 | 0.24352 | 0.089634 | 0.2784 | 0.16193 | -0.053934 | -0.21792 | 0.13341 | -0.047667 | -0.22397 | -0.068376 | -0.076614 | 0.1705 | -0.21276 |  | 0
da6 | 0.068303 | 0.34214 | -0.074441 | -0.062333 | -0.2146 | -0.04282 | -0.020787 | 0.0357 | -0.055174 | -0.06863 | -0.10633 | -0.021264 | 0.045353 | 0.11445 | 0.22139 | 0.055471 | 0.121 | 0.07566 | -0.096118 | 0.093471 | -0.018068 |  | 0
da7 | 0.003266 | 1.8221 | -0.055027 | -0.029862 | 0.0095804 | -0.014081 | 0.01003 | -0.059566 | 0.083981 | 0.023521 | 0.082769 | 0.03677 | 0.029933 | -0.05721 | 0.07332 | -0.028177 | -0.082403 | -0.026777 | -0.035284 | 0.038483 | 0.034346 |  | 0
da8 | -0.029199 | -0.062357 | 0.0043249 | 0.00358 | -0.091672 | -0.016878 | -0.028239 | -0.0078319 | -0.012217 | 0.00039249 | -0.014463 | -0.0031187 | 0.0024829 | 0.034626 | 0.070071 | 0.010095 | 0.050959 | -0.0026092 | -0.018488 | 0.018986 | 0.0052562 |  | 0
da9 | 0.0083634 | -0.22558 | -0.0079393 | -0.016689 | -0.0039345 | -0.0090965 | 0.0043974 | -0.0060552 | 0.042039 | 0.014551 | 0.019344 | 0.011784 | -0.0080302 | -0.0019009 | -0.0046845 | -0.036931 | -0.017408 | -0.0064237 | -0.010486 | 0.037464 | -0.020597 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.0094116 | 0.19973 | -0.021836 | -0.006694 | -0.0015192 | -0.0021701 | -0.0050658 | -0.0071918 | 0.019236 | 0.0082805 | 0.011163 | 0.0036284 | -0.0011573 | -0.0025629 | 0.013404 | -0.0020644 | -0.0060336 | -0.003961 | -0.0081971 | 0.012742 | 0.010758 |  | 0
da12 | -0.0030052 | -0.0044693 | 0.0022492 | 0.0024672 | -0.0095529 | -0.00061859 | -0.0030929 | 0.00057381 | -0.0037829 | -0.0015838 | -0.003351 | -0.00068596 | 0.00096406 | 0.00089104 | 0.0010709 | -0.0012562 | 0.0056542 | 0.001948 | 0.00094379 | 0.0071621 | 0.0030902 |  | 0
da13 | 0.0032822 | -0.019098 | -0.00078615 | -0.0014857 | -0.0020277 | 0.00028887 | 0.0024294 | -0.00075065 | 0.0037734 | 0.00033092 | 0.0029134 | 0.00083188 | 0.00014271 | -0.0014867 | 0.0012646 | -0.0049039 | -0.003255 | -0.00085546 | -0.00102 | 0.0045961 | 0.0031794 |  | 0
da14 | -0.0072966 | -0.0041898 | 0.0012929 | 0.00035861 | -0.00062574 | 0.00047637 | -0.0029638 | -0.00042104 | -0.00077693 | -0.0003054 | -0.00099301 | -0.0025033 | -0.00098585 | -0.00020351 | 0.0019713 | 0.0010773 | -1.4917e-05 | 0.0017388 | 0.00096803 | 0.0019102 | -0.00015492 |  | 5.7824e-19
da15 | -0.002051 | 0.0036179 | 0.0010719 | 6.892e-05 | -0.00017693 | 0.0020423 | -0.00027487 | -0.0015789 | 0.0039075 | 0.0015256 | 0.0005587 | -0.00068287 | 0.0019717 | -0.0036059 | -0.0037136 | -0.0043096 | -0.0018107 | 0.0005628 | 0.0012469 | 0.003197 | -0.012606 |  | 0
Field Colinearity | 75.4999
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.9601 | -0.93235 | -0.48178 | -0.86056 | -0.56012 | -0.28021 | -0.047436 | 0.70026 | 0.036749 | -0.16874 | -0.44219 | -0.077241 | 0.33032 | 0.3157 | 0.046161 | -0.081006 | -0.48976 | -0.49884 | 0.29924 | 1.4048
Twist (m2 mrad) | 29.7005 | -13.286 | -6.5041 | -10.9722 | -6.7214 | -3.1524 | -0.49808 | 6.8275 | 0.33074 | -1.3921 | -3.3164 | -0.52138 | 1.9819 | 1.6574 | 0.20772 | -0.30377 | -1.4693 | -1.1224 | 0.44886 | 0.52678
Twist Integral (I0) (m2 rad) | -0.0056834
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.8695 | -0.84719 | -0.34699 | -0.65296 | -0.49722 | -0.41764 | -0.19681 | 0.5181 | 0.21972 | -0.038793 | -0.39761 | -0.32227 | 0.19677 | 0.40224 | 0.054792 | -0.13809 | -0.42112 | -0.47233 | 0.40933 | 2.0781
Twist (m2 mrad) | 21.5213 | -12.0725 | -4.6844 | -8.3252 | -5.9666 | -4.6984 | -2.0665 | 5.0515 | 1.9775 | -0.32004 | -2.9821 | -2.1753 | 1.1806 | 2.1117 | 0.24656 | -0.51784 | -1.2634 | -1.0627 | 0.61399 | 0.7793
Twist Integral (I0) (m2 rad) | -0.0094894
Note on B168: 4.2s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.30945 | -3.0595 | 0.034625 | -0.51254 | 1.0406 | -1.7736 | 0.21591 | 0.66658 | 0.224 | 16.6902 | 0.59611 | -0.7864 | -3.1793 | -0.094122 | -0.31085 | 1.0493 | -1.6385 | -0.38232 | 0.66983 | -0.36489 | 18.1027 | 0.5961 | 0.10297 | -1.654 | -0.051546 | -0.76491 | 1.0941 | -0.00023572 | 0.89913 | 0.62079 | 0.28581 | 14422.92 | -0.74959 | -1.6888 | -0.10755 | -0.57147 | 1.0984 | -0.011644 | -0.27438 | 0.61416 | -0.29308 | 14425.348
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8