## Alarm sheet
File name | HCMB__A001-03000258_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000258
Date of test Ap 1 | 2004-09-10 |  |  |  | Date of test Ap 2 | 2004-09-10
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | yellow alarm | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
519 | HCMB__A001 | 3000258 |  | 3 |  |  | 3
| COMMENTS: | Rather low main field (yellow alarm at 4.1 sigma in ap. 1)
| HOLDING POINT | OK | Data at CERN |  | 2004-09-10 | Analysis by |  | E. Todesco | Answer to MMS-LD |  |  | 2004-09-13 | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000258_cc.xls | [41 empty] | T
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000258 | [11 empty] | Signs | [23 empty] | T
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1 | [23 empty] | T
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 10-SEP-2004_3258_08-03cc1.txt |  |  |  | a odd | 1 | [23 empty] | T
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 10-SEP-2004_3258_08-03cc2.txt |  |  |  | a even | 1 | [23 empty] | T
Test Operator |  | Mahler |  |  | Test Controller |  |  | E. Todesco | [35 empty] | T
T
Aperture 1 | [11 empty] | Aperture 2 | [31 empty] | T
Date of test |  | 2004-09-10 |  |  | Time of test |  |  | 08:03 - 11:11 |  |  |  | Date of test |  | 2004-09-10 |  |  | Time of test |  |  | 08:04 - 11:11 | [23 empty] | T
Mole name |  | Dipole 7 | [9 empty] | Mole name |  | Dipole 1 | [5 empty] | 2004-03-26 | [23 empty] | T
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 | [23 empty] | T
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw | [23 empty] | T
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) | [23 empty] | T
Rotation speed |  | 0.63 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 | [23 empty] | T
Dx average |  | 0.00012488 |  |  | Dy average |  |  | 0.0002334 |  |  |  | Dx average |  | 0.00038711 |  |  | Dy average |  |  | 0.00023764 | [23 empty] | T
Magnetic length |  | 14.4207 |  |  | Tranfs. Function (mT/KA) |  |  | 595.3783 |  |  |  | Magnetic length |  | 14.4212 |  |  | Tranfs. Function (mT/KA) |  |  | 595.456 | [23 empty] | T
Phase |  | 48.2697 |  |  | Integral field (T.m) |  |  | 85.7331 |  |  |  | Phase |  | -27.1961 |  |  | Integral field (T.m) |  |  | 85.7301 | [23 empty] | T
T
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals | T
C1 (T) | 3.4529 | 5.9571 | 5.9534 | 5.9532 | 5.9524 | 5.9539 | 5.9534 | 5.953 | 5.9536 | 5.9526 | 5.9527 | 5.9527 | 5.9532 | 5.9531 | 5.9542 | 5.9536 | 5.9542 | 5.9538 | 5.9581 | 3.7053 | 85.7331 | C1 (T) | 3.3847 | 5.9576 | 5.9537 | 5.9541 | 5.9529 | 5.9546 | 5.9539 | 5.954 | 5.9541 | 5.9544 | 5.9545 | 5.955 | 5.9545 | 5.9538 | 5.9539 | 5.9541 | 5.9542 | 5.9547 | 5.958 | 3.7549 | 85.7301 | T
Angle (mrad) | 5.3121 | 0.015013 | -0.13958 | -0.46455 | -0.43304 | -0.24341 | -0.019084 | 0.25405 | 0.041848 | 0.20556 | -0.012781 | -0.10559 | 0.31344 | 0.52334 | 0.27771 | -0.081856 | -0.43061 | -0.60149 | -0.70215 | -2.3733 | 48.2697 | Angle (mrad) | 5.359 | 0.17261 | -0.051582 | -0.31633 | -0.31865 | -0.091435 | 0.20643 | 0.36994 | 0.13769 | -0.066656 | -0.03443 | -0.10605 | 0.10638 | 0.42982 | 0.048158 | -0.31687 | -0.43106 | -0.76363 | -0.99844 | -1.6199 | -27.1961 | T
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | T
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | T
b2 | -2.6812 | -1.5447 | -1.5455 | -0.53922 | -1.0187 | -1.0455 | -1.7974 | -1.2664 | -1.7331 | -1.4827 | -1.6082 | -1.8765 | -1.6966 | -1.2181 | -0.9171 | -0.96424 | -1.3411 | -1.6178 | -1.1891 | 0.027378 | -1.3509 | b2 | -1.5635 | 0.37894 | 0.71735 | 0.57859 | 0.62833 | 0.99417 | 1.4391 | 1.4849 | 0.53786 | 0.72911 | 1.4747 | 1.1292 | 0.8742 | 0.90523 | 1.2627 | 2.1323 | 1.4997 | 1.2031 | 1.4751 | 0.84764 | 0.99434 | T
b3 | 45.1174 | -2.1186 | -1.6706 | -3.1692 | -2.6504 | -3.908 | -2.9715 | -2.2413 | -2.1518 | -3.1051 | -5.3874 | -3.7256 | -3.7938 | -4.083 | -5.257 | -3.7179 | -1.6104 | -2.0889 | -1.6381 | -0.45599 | -1.5314 | b3 | 41.4021 | -1.6258 | -1.2039 | -2.2708 | -2.0742 | -3.7176 | -2.4762 | -1.5068 | -2.0938 | -2.4496 | -4.0473 | -1.3404 | -1.2562 | -3.263 | -3.3113 | -2.7738 | -1.419 | -1.9713 | -1.3052 | -1.2793 | -0.90521 | T
b4 | -0.82953 | 0.066807 | -0.06608 | 0.021773 | -0.027321 | -0.14138 | -0.15092 | -0.036118 | -0.11569 | -0.12004 | -0.27537 | -0.13122 | -0.11733 | -0.11132 | -0.07173 | 0.021306 | 0.1243 | -0.01369 | -0.132 | -0.43779 | -0.10569 | b4 | -1.2526 | -0.037174 | 0.12294 | -0.12503 | -0.099244 | 0.040971 | 0.085408 | -0.03694 | -0.04085 | -0.048865 | -0.027336 | -0.023646 | -0.15468 | -0.085752 | -0.10185 | -0.19052 | 0.10944 | -0.11805 | -0.12006 | -0.20712 | -0.088226 | T
b5 | -4.7375 | -0.58992 | -0.60137 | -0.3 | -0.24296 | -0.0014298 | -0.06874 | -0.23465 | -0.23358 | -0.32973 | 0.17277 | -0.13428 | -0.096851 | -0.19044 | -0.065099 | -0.15175 | -0.3393 | -0.21232 | -0.3431 | -3.6673 | -0.46832 | b5 | -5.5108 | -0.53032 | -0.74409 | -0.2822 | -0.3727 | -0.025824 | -0.027007 | -0.12266 | -0.17611 | 0.033745 | 0.46287 | 0.17108 | 0.0066188 | 0.019576 | -0.036988 | -0.3218 | -0.20599 | -0.25062 | -0.36005 | -3.4397 | -0.42002 | T
b6 | -0.020626 | -0.090857 | -0.097988 | -0.097469 | -0.10657 | -0.02401 | 0.038277 | 0.015846 | -0.015949 | -0.013424 | -0.066485 | 0.038182 | 0.08075 | 0.03694 | -0.031703 | -0.086197 | 0.025108 | 0.007593 | 0.035968 | 0.063229 | -0.016904 | b6 | -0.057109 | 0.1441 | 0.068658 | -0.055435 | -0.045914 | -0.041401 | -0.035688 | 0.043325 | -0.02294 | 0.079937 | -0.11813 | -0.074186 | -0.023355 | -0.033895 | -0.03517 | 0.12078 | -0.031956 | -0.046252 | -0.049922 | -0.046912 | -0.01143 | T
b7 | 2.946 | 1.0099 | 1.1311 | 1.2077 | 1.2031 | 1.1499 | 1.1821 | 1.188 | 1.1913 | 1.1101 | 1.1919 | 1.1909 | 1.2402 | 1.2188 | 1.2372 | 1.2587 | 1.1445 | 1.2055 | 1.2234 | 0.97672 | 1.229 | b7 | 2.921 | 1.0406 | 1.0769 | 1.1777 | 1.1933 | 1.1695 | 1.1725 | 1.1578 | 1.1632 | 1.1484 | 1.2158 | 1.126 | 1.1947 | 1.1212 | 1.1505 | 1.2138 | 1.173 | 1.1659 | 1.1901 | 1.0407 | 1.2067 | T
b8 | -0.086416 | -0.0068483 | -0.043889 | -0.051625 | -0.02326 | -0.020462 | -0.021952 | -0.0017407 | -0.014173 | 0.011534 | -0.014125 | -0.0027543 | 0.037892 | 0.013416 | 0.0071769 | 0.0030058 | 0.021312 | -0.0060623 | 0.00059618 | -0.015746 | -0.0089501 | b8 | -0.15998 | 0.014879 | 0.019571 | -0.020943 | -0.011203 | -0.024897 | 0.00032213 | -0.026246 | -0.025188 | 0.0131 | -0.013842 | -0.0049063 | -0.0021298 | -0.0027077 | 0.011294 | 0.02224 | 0.018203 | -0.010346 | -0.0063191 | -0.031477 | -0.0083282 | T
b9 | 0.35136 | 0.46725 | 0.45959 | 0.48558 | 0.50813 | 0.46279 | 0.47186 | 0.47122 | 0.48009 | 0.45619 | 0.45485 | 0.44959 | 0.4498 | 0.46849 | 0.44478 | 0.45535 | 0.46244 | 0.49118 | 0.48087 | 0.41636 | 0.4626 | b9 | 0.3634 | 0.47775 | 0.46587 | 0.49324 | 0.50778 | 0.45586 | 0.47109 | 0.47911 | 0.48725 | 0.46994 | 0.46561 | 0.48766 | 0.48991 | 0.47764 | 0.47396 | 0.4772 | 0.4825 | 0.48486 | 0.4863 | 0.4427 | 0.47499 | T
b10 | 0.042534 | -0.013325 | -0.065892 | -0.09374 | 0.010075 | -0.031502 | -0.0014904 | 0.0065954 | 0.0087729 | 0.012211 | -0.051476 | 0.033318 | 0.072059 | 0.032493 | -0.0014443 | -0.0089215 | 0.020837 | 0.0045827 | 0.030111 | 0.019369 | 0 | b10 | 0.064176 | 0.051743 | 0.039149 | -0.022096 | 0.039123 | -0.0015644 | -0.019161 | -0.0045161 | 0.0028269 | 0.0064704 | -0.05091 | -0.028522 | 0.0081267 | 0.0025801 | -0.0026738 | 0.023653 | 0.0067436 | -0.014915 | -0.023486 | -0.077778 | 0 | T
b11 | 0.58426 | 0.72121 | 0.71954 | 0.72625 | 0.72625 | 0.73846 | 0.73169 | 0.72098 | 0.72794 | 0.72543 | 0.74099 | 0.72672 | 0.73086 | 0.73536 | 0.74317 | 0.73344 | 0.72394 | 0.72548 | 0.72486 | 0.63585 | 0.72164 | b11 | 0.5881 | 0.72502 | 0.72636 | 0.73606 | 0.73844 | 0.74864 | 0.74095 | 0.73369 | 0.73995 | 0.74496 | 0.75727 | 0.74066 | 0.73654 | 0.73972 | 0.73989 | 0.73982 | 0.72211 | 0.73324 | 0.72468 | 0.63106 | 0.72921 | T
b12 | -0.006326 | -0.0046675 | -0.010253 | -0.012128 | -0.0015905 | -0.0032444 | -0.0028361 | -9.0372e-05 | -0.00042133 | 0.0024474 | -0.0030627 | 0.0030282 | 0.0080965 | 0.0040857 | -0.0008775 | -0.0037142 | -0.00015648 | -0.0023644 | -0.0006513 | -0.00079441 | -0.0016959 | b12 | 0.0080491 | 0.0059152 | 0.006147 | -0.001048 | 0.0041496 | -0.0021644 | -0.0011511 | 0.00072473 | 0.001443 | -0.0011602 | -0.0076154 | -0.0015495 | 0.0026482 | -0.0015827 | -0.0014263 | 0.0039715 | 0.0016618 | 0.00035006 | 0.00010489 | -0.0047493 | 0.00057292 | T
b13 | 0.092958 | 0.071871 | 0.073623 | 0.06742 | 0.068528 | 0.065539 | 0.064857 | 0.06652 | 0.067308 | 0.067885 | 0.065858 | 0.067346 | 0.069601 | 0.066588 | 0.066032 | 0.067749 | 0.068547 | 0.069839 | 0.071918 | 0.055464 | 0.068506 | b13 | 0.099837 | 0.074035 | 0.072932 | 0.069449 | 0.070177 | 0.069036 | 0.070462 | 0.069265 | 0.068561 | 0.06926 | 0.064919 | 0.06629 | 0.068612 | 0.063906 | 0.066091 | 0.068793 | 0.068 | 0.071208 | 0.070722 | 0.056442 | 0.069486 | T
b14 | 0.004062 | -0.003954 | -0.0053859 | -0.0054082 | -0.001717 | -0.0028533 | -0.0022791 | -0.0026433 | -0.0018947 | -0.0023621 | -0.0049071 | -0.00056701 | 0.00091403 | -0.0016587 | -0.003092 | -0.0029548 | -0.0019836 | -0.0029597 | -0.0010773 | -0.0014613 | -0.002361 | b14 | 0.0038474 | -0.0025808 | -0.0058235 | -0.0073404 | -0.0037879 | -0.0063264 | -0.0064866 | -0.0064499 | -0.0062149 | -0.0075369 | -0.0097511 | -0.0082952 | -0.0069383 | -0.0074126 | -0.0070658 | -0.0047508 | -0.0061695 | -0.006844 | -0.0072692 | -0.0034196 | -0.0060947 | T
b15 | -0.0077616 | 0.02397 | 0.020076 | 0.017584 | 0.02056 | 0.022428 | 0.022608 | 0.02118 | 0.022325 | 0.021861 | 0.020762 | 0.020955 | 0.023099 | 0.02291 | 0.021038 | 0.021551 | 0.021825 | 0.02301 | 0.022123 | 0.0097151 | 0.020384 | b15 | -0.012345 | 0.020947 | 0.019514 | 0.016519 | 0.01749 | 0.018555 | 0.01622 | 0.017853 | 0.019684 | 0.020797 | 0.015934 | 0.018734 | 0.021655 | 0.020636 | 0.019803 | 0.01864 | 0.01928 | 0.019175 | 0.017323 | -0.0019915 | 0.017214 | T
b16 | [21 empty] | b16 | [21 empty] | T
b17 | [21 empty] | b17 | [21 empty] | T
a1 | 53.1216 | 0.15013 | -1.3958 | -4.6455 | -4.3304 | -2.4341 | -0.19084 | 2.5405 | 0.41848 | 2.0556 | -0.12781 | -1.0559 | 3.1344 | 5.2334 | 2.7771 | -0.81856 | -4.3061 | -6.0149 | -7.0215 | -23.7331 | 0 | a1 | 53.5904 | 1.7261 | -0.51582 | -3.1633 | -3.1865 | -0.91435 | 2.0643 | 3.6994 | 1.3769 | -0.66656 | -0.3443 | -1.0605 | 1.0638 | 4.2982 | 0.48158 | -3.1687 | -4.3106 | -7.6363 | -9.9844 | -16.1986 | 0 | T
a2 | 7.425 | 0.24092 | 0.04623 | -0.43709 | 1.1755 | 0.3414 | -0.74456 | -0.10859 | -0.43934 | 0.13411 | 1.3619 | 0.27181 | 0.9508 | 1.3649 | 0.32823 | -1.2336 | 0.24771 | 0.018751 | -0.33398 | 1.4717 | 0.43777 | a2 | 20.9088 | -0.30254 | -0.4763 | -0.18592 | -0.23864 | -0.71172 | -0.050238 | -0.7674 | -0.58149 | 0.12347 | -0.85755 | -0.2502 | -0.53028 | 1.1733 | -1.2716 | -0.5866 | -0.57539 | 0.24949 | -0.13313 | 12.4633 | 0.71727 | T
a3 | 1.3024 | 0.8174 | 0.58323 | 0.54237 | 0.71183 | 0.84464 | 0.88558 | 0.78276 | 0.51033 | 0.42001 | 0.64021 | 0.77647 | 0.77694 | 0.70041 | 0.65646 | 0.75141 | 0.91277 | 0.93548 | 1.2168 | 0.73359 | 0.76436 | a3 | 0.17941 | 0.57414 | 0.39988 | -0.05024 | 0.47054 | 0.66983 | 0.88037 | 0.58163 | 0.26914 | 0.76204 | 1.1298 | 0.034919 | 0.18115 | -0.087243 | 0.052249 | 0.90998 | 0.96253 | 0.72395 | 0.87455 | 1.3537 | 0.53624 | T
a4 | 3.1018 | 0.15983 | -0.14852 | 0.054608 | -0.63105 | -0.66099 | -0.55667 | -0.25541 | -0.50826 | -0.32574 | -0.80926 | -1.0663 | -0.79823 | -0.35981 | -0.19451 | 0.050826 | -0.065972 | -0.24446 | -0.15435 | -0.33651 | -0.25643 | a4 | 4.0029 | 0.025022 | -0.17668 | -0.10529 | -0.25037 | -0.25407 | -0.57091 | -0.41827 | -0.2326 | -0.55992 | -0.027844 | -0.45358 | -0.3938 | -0.10675 | 0.23308 | -0.25893 | 0.29148 | -0.24659 | -0.073714 | 0.6869 | -0.045378 | T
a5 | 3.185 | -0.098249 | 0.013776 | 0.15691 | 0.39922 | 0.24195 | 0.27821 | 0.10095 | 0.15405 | 0.037839 | 0.22999 | 0.054044 | 0.076838 | 0.15996 | 0.20677 | 0.12903 | 0.304 | 0.29562 | 0.40885 | -0.011853 | 0.25984 | a5 | 2.5371 | -0.10519 | -0.045932 | -0.14275 | 0.096113 | 0.031748 | 0.20988 | 0.099575 | -0.008254 | 0.063327 | 0.25271 | 0.0065066 | -0.056697 | 0.081486 | 0.11605 | 0.11582 | 0.17146 | 0.26466 | 0.31588 | -0.10724 | 0.14797 | T
a6 | -0.021488 | 0.00019991 | 0.088834 | 0.0010833 | -0.16334 | -0.19678 | -0.13641 | -0.090408 | -0.031194 | -0.026118 | 0.013829 | -0.10883 | -0.16407 | -0.058694 | -0.048452 | -0.11926 | -0.0028763 | -0.0042427 | 0.0031439 | -0.26246 | -0.063496 | a6 | 0.19055 | 0.057107 | 0.052933 | 0.071033 | 0.1344 | 0.049871 | 0.062423 | 0.08738 | 0.18022 | 0.11118 | -0.0057188 | 0.067422 | 0.033291 | 0.10752 | 0.08518 | 0.11104 | 0.14544 | 0.10491 | 0.17601 | 0.092789 | 0.093676 | T
a7 | 1.7852 | -0.080676 | -0.043742 | -0.070977 | 0.016618 | 0.0098424 | 0.017519 | 0.01165 | -0.02897 | -0.051425 | -0.068724 | -0.0092093 | 0.092969 | 0.027932 | -0.052777 | -0.0020952 | 0.016216 | 0.018146 | 0.049214 | -0.060616 | 0.044216 | a7 | 1.694 | -0.029564 | -0.068653 | -0.077157 | 0.0039734 | 0.12363 | 0.050528 | 0.093146 | 0.017435 | -0.010214 | 0.070628 | -0.081377 | 0.053161 | 0.031913 | -0.022857 | 0.029042 | 0.10451 | 0.061587 | 0.072891 | 0.0025579 | 0.07225 | T
a8 | -0.033607 | -0.00018285 | 0.019666 | -0.0053045 | -0.04347 | -0.057305 | -0.061014 | -0.022863 | 0.0059449 | 0.0077902 | -0.01147 | -0.016424 | -0.041164 | -0.045553 | -0.033134 | -0.049283 | -0.00010466 | -0.0017284 | 0.0073488 | -0.098085 | -0.022328 | a8 | -0.018212 | 0.018236 | -0.0039557 | -0.022123 | -0.013355 | -0.059127 | 0.00055301 | -0.020532 | 0.0050842 | 0.0035859 | -0.067519 | -0.026125 | -0.046585 | -0.017208 | -0.064655 | -0.042208 | -0.008864 | -0.022438 | 0.011508 | -0.011579 | -0.020488 | T
a9 | -0.1916 | -0.0068612 | -0.028182 | -0.0093752 | 0.017556 | 0.010028 | 0.0078542 | -0.016943 | -0.0005622 | -0.023691 | 0.0067162 | -0.020005 | -0.0084254 | -0.0086489 | -0.018214 | -0.0011954 | -0.0080052 | 0.00027857 | 0.020434 | -0.056391 | -0.012157 | a9 | -0.27255 | 0.01565 | -0.018917 | -0.01741 | 0.0013902 | 0.022646 | 0.022185 | 0.016192 | -0.021664 | -0.010535 | 0.031237 | -0.021509 | -0.010087 | 0.0049182 | 0.0058746 | 0.0034488 | 0.012244 | 0.028353 | 0.029818 | -0.031218 | -0.0042056 | T
a10 | -0.13222 | -0.011926 | 0.034108 | 0.028789 | 0.016978 | -0.02687 | -0.0031675 | 0.005244 | 0.01913 | 0.013106 | 0.015751 | 0.018926 | -0.041944 | -0.046039 | -0.0010046 | -0.0085376 | 0.036951 | 0.030883 | 0.056674 | -0.097059 | 0 | a10 | -0.096668 | 0.024901 | 0.04608 | 0.025097 | 0.0037917 | -0.029309 | -0.00016343 | 0.0021924 | 0.00017492 | -0.024404 | -0.02007 | 0.01161 | -0.019402 | -0.024295 | -0.0039521 | -0.0016823 | 0.023144 | 0.0080259 | 0.035191 | -0.0031901 | 0 | T
a11 | 0.18502 | -0.067868 | -0.053247 | -0.056389 | -0.039945 | -0.044237 | -0.035921 | -0.04324 | -0.042803 | -0.048304 | -0.044993 | -0.04718 | -0.030699 | -0.039603 | -0.043544 | -0.040124 | -0.039415 | -0.038857 | -0.035409 | -0.060056 | -0.037593 | a11 | 0.19573 | -0.027384 | -0.017778 | -0.020833 | -0.0036997 | 0.003287 | 0.001541 | -0.0007708 | -0.0031283 | -0.01372 | -0.0027372 | -0.017843 | -0.0031408 | -2.1327e-05 | -0.013651 | -0.011696 | -0.0040712 | -0.0029345 | -0.006703 | -0.011694 | -0.0021572 | T
a12 | -0.026389 | -7.8405e-05 | 0.0045688 | 0.00030862 | -0.004242 | -0.0097302 | -0.0071372 | -0.0040744 | -0.0013038 | -0.0014866 | 0.0012134 | -0.0022051 | -0.0079161 | -0.0048455 | -0.0048086 | -0.0095463 | 0.0027795 | 0.0014123 | 0.0056183 | -0.016855 | -0.0035026 | a12 | -0.024597 | -0.00011397 | 0.00343 | 0.002351 | -0.00086046 | -0.0050969 | 0.0011003 | 0.0011421 | 0.00080325 | 0.0017068 | -0.0065898 | -0.00048879 | -0.0029924 | 0.0011271 | -0.0021555 | 0.0031124 | 0.0038866 | 0.0025498 | 0.0087256 | 0.0012401 | -8.1143e-05 | T
a13 | -0.01706 | -0.0017588 | -0.0036595 | -0.0024743 | 0.0007065 | 6.4531e-05 | -2.9597e-05 | -0.0025889 | -0.0013402 | -0.0032095 | -0.00063575 | -0.0027424 | 2.0505e-05 | -0.0004187 | -0.003306 | -0.0024328 | 0.00088834 | 0.0011434 | 0.0031311 | -0.0021836 | -0.0015567 | a13 | -0.017356 | 0.0030163 | -0.00089307 | -0.0020859 | 0.00088849 | 0.0055373 | 0.0056252 | 0.0038254 | -0.0012035 | 0.0032249 | 0.0040097 | -0.00038994 | 0.0022405 | 0.0018786 | 0.0025462 | 0.0014156 | 0.0035836 | 0.0045988 | 0.0057456 | 0.00059618 | 0.001775 | T
a14 | 0.01003 | -0.0042715 | -0.0031066 | -0.0043057 | -0.0064428 | -0.0082033 | -0.0083789 | -0.0058077 | -0.0052877 | -0.0037483 | -0.005946 | -0.0062878 | -0.005478 | -0.0055867 | -0.0055692 | -0.0053907 | -0.003628 | -0.0042804 | -0.0016404 | -0.0040198 | -0.0046891 | a14 | 0.01071 | -0.0025384 | -0.0022406 | -0.0033179 | -0.0037397 | -0.0077024 | -0.0044278 | -0.0062874 | -0.0059263 | -0.005638 | -0.005835 | -0.0073695 | -0.0076484 | -0.00546 | -0.0051671 | -0.0051641 | -0.0040708 | -0.0047474 | -0.0025459 | 0.0008965 | -0.0043321 | T
a15 | 0.0031876 | -0.0044601 | -0.0025325 | -0.0030479 | -0.00063312 | -0.0027946 | -0.00027577 | -0.00025458 | -0.00090582 | -0.0022144 | -0.0034419 | 0.0012585 | 0.0022019 | 0.00010394 | -0.001231 | -0.0014726 | 0.00016352 | -0.00031973 | 0.0012892 | 0.00093128 | -0.00084048 | a15 | -0.0036293 | -0.0023728 | -0.0012926 | -0.004349 | -0.0019256 | -0.0031906 | -0.0028732 | -0.002858 | -0.003663 | -0.0038174 | -0.003241 | -0.0040901 | -0.003033 | -0.0046119 | -0.0041719 | -0.00049116 | -0.0011799 | -0.0025556 | -0.001737 | -0.0046277 | -0.0029394 | T
a16 | [21 empty] | a16 | [21 empty] | T
a17 | [21 empty] | a17 | [21 empty] | T
Dx (m) | 0 | 2.8511e-05 | 0.00016081 | 0.00022351 | -2.1338e-05 | 6.8535e-05 | 0 | -1.4757e-05 | -1.7801e-05 | -2.6454e-05 | 0.00011986 | -7.4762e-05 | -0.00017137 | -8.0588e-05 | 0 | 1.9534e-05 | -4.4073e-05 | 0 | -6.3959e-05 | -7.5369e-05 | 0.00012488 | Dx (m) | -8.2721e-05 | -0.00011898 | -8.891e-05 | 5.2626e-05 | -9.003e-05 | 0 | 4.3962e-05 | 1.0469e-05 | 0 | -1.579e-05 | 0.00011412 | 6.6071e-05 | -1.8944e-05 | 0 | 0 | -5.4404e-05 | -1.5564e-05 | 3.4652e-05 | 5.5825e-05 | 0.00020934 | 0.00038711 | T
Dy (m) | 0.00038403 | 3.0798e-05 | -6.879e-05 | -5.0109e-05 | -4.0916e-05 | 6.5972e-05 | 0 | -1.325e-05 | -4.5723e-05 | -3.2477e-05 | -2.8866e-05 | -4.9128e-05 | 9.0497e-05 | 0.00010211 | 0 | 2.0858e-05 | -8.9163e-05 | -7.2734e-05 | -0.00013601 | 0.00025195 | 0.0002334 | Dy (m) | 0.00030635 | -6.29e-05 | -0.00011006 | -5.6474e-05 | 0 | 6.6532e-05 | 0 | 0 | 0 | 5.54e-05 | 4.5521e-05 | -2.5055e-05 | 4.4701e-05 | 5.5836e-05 | 0 | 0 | -5.4572e-05 | -1.847e-05 | -8.202e-05 | 1.2446e-05 | 0.00023764 | T
Magnetic length (m) |  | 14.4207 | [19 empty] | Magnetic length (m) |  | 14.4212 | [19 empty] | T
Transfer funct. (mT/KA) |  | 595.3783 | [19 empty] | Transfer funct. (mT/KA) |  | 595.456 | [19 empty] | T
Coil positioning (mm) |  | 0.016444 | [19 empty] | Coil positioning (mm) |  | 0.017371 | [19 empty] | T
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10 | [19 empty] | T
T
T
Soft version | DIMM V2.0.8 | [20 empty] | Soft version | DIMM V2.0.8 | [20 empty] | T
Level Average | -0.13953 | 0.70964 | 0.40535 | 0.38498 | 0.52541 | 0.53862 | 0.54526 | 0.23858 | 0.38315 | 0.036476 | -0.1388 | -0.39555 | 0.45623 | 0.42159 | -0.003411 | -0.080038 | 0.058147 | -0.086551 | -0.15069 | -0.14885 |  | Level Average | -0.026246 | -0.038313 | -0.019999 | -0.028855 | 0.046322 | -0.067584 | 0.001183 | 0.0864 | -0.13911 | -0.20181 | -0.10696 | -0.37279 | -0.14447 | 0.17806 | -0.027248 | -0.031454 | 0.02705 | -0.065585 | 0.035904 | -0.022051 |  | T
Level Sigma2 | 0.2461 | 1.0108 | 2.1866 | 2.8065 | 8.4538 | 7.0514 | 12.7171 | 2.8918 | 4.9297 | 0.41139 | 0.026852 | 0.017064 | 0.098327 | 0.18724 | 0.29621 | 1.064 | 0.73913 | 0.13454 | 0.13791 | 0.89542 |  | Level Sigma2 | 0.006283 | 0.000401 | 0.00234 | 0.002674 | 0.088164 | 0.060901 | 0.008849 | 0.041473 | 0.10054 | 0.034953 | 0.020567 | 1.0954 | 7.5e-05 | 0.22215 | 0.00207 | 0.003204 | 0.001889 | 0.082513 | 0.014267 | 0.004117 |  | T
Position (mm) | 0 | 751 | 1503 | 2254 | 3004 | 3755 | 4506 | 5258 | 6009 | 6760 | 7510 | 8262 | 9013 | 9764 | 10515 | 11266 | 12017 | 12768 | 13519 | 14269 |  | Position (mm) | 0 | 751 | 1502 | 2252 | 3005 | 3755 | 4506 | 5256 | 6009 | 6761 | 7511 | 8263 | 9015 | 9766 | 10517 | 11268 | 12019 | 12771 | 13521 | 14272 |  | T
Fist move (mm) | First move Ap1 | 940.8598 | First move Ap2 | 947.7809 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2 | [18 empty] | T
T
T
T
T
T
T
T
T
T
T
T

## Summary Data
File | HCMB__A001-3000258 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000258 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 345.2856 | 595.712 | 595.3362 | 595.3186 | 595.2377 | 595.3883 | 595.3437 | 595.3029 | 595.3553 | 595.2579 | 595.2708 | 595.2717 | 595.3241 | 595.3094 | 595.4152 | 595.3561 | 595.4241 | 595.377 | 595.8085 | 370.5259 | 8573.3136 | C1/i (mT/KA) | 338.4721 | 595.7588 | 595.3743 | 595.4069 | 595.2919 | 595.4598 | 595.3872 | 595.4027 | 595.4127 | 595.4389 | 595.452 | 595.5015 | 595.4452 | 595.3787 | 595.3934 | 595.4117 | 595.4225 | 595.4655 | 595.8035 | 375.4916 | 8573.0056
Angle (mrad) | 5.3121 | 0.015013 | -0.13958 | -0.46455 | -0.43304 | -0.24341 | -0.019084 | 0.25405 | 0.041848 | 0.20556 | -0.012781 | -0.10559 | 0.31344 | 0.52334 | 0.27771 | -0.081856 | -0.43061 | -0.60149 | -0.70215 | -2.3733 | 48.2697 | Angle (mrad) | 5.359 | 0.17261 | -0.051582 | -0.31633 | -0.31865 | -0.091435 | 0.20643 | 0.36994 | 0.13769 | -0.066656 | -0.03443 | -0.10605 | 0.10638 | 0.42982 | 0.048158 | -0.31687 | -0.43106 | -0.76363 | -0.99844 | -1.6199 | -27.1961
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.6812 | -1.5447 | -1.5455 | -0.53922 | -1.0187 | -1.0455 | -1.7974 | -1.2664 | -1.7331 | -1.4827 | -1.6082 | -1.8765 | -1.6966 | -1.2181 | -0.9171 | -0.96424 | -1.3411 | -1.6178 | -1.1891 | 0.027378 |  | b2 | -1.5635 | 0.37894 | 0.71735 | 0.57859 | 0.62833 | 0.99417 | 1.4391 | 1.4849 | 0.53786 | 0.72911 | 1.4747 | 1.1292 | 0.8742 | 0.90523 | 1.2627 | 2.1323 | 1.4997 | 1.2031 | 1.4751 | 0.84764
b3 | 45.1174 | -2.1186 | -1.6706 | -3.1692 | -2.6504 | -3.908 | -2.9715 | -2.2413 | -2.1518 | -3.1051 | -5.3874 | -3.7256 | -3.7938 | -4.083 | -5.257 | -3.7179 | -1.6104 | -2.0889 | -1.6381 | -0.45599 |  | b3 | 41.4021 | -1.6258 | -1.2039 | -2.2708 | -2.0742 | -3.7176 | -2.4762 | -1.5068 | -2.0938 | -2.4496 | -4.0473 | -1.3404 | -1.2562 | -3.263 | -3.3113 | -2.7738 | -1.419 | -1.9713 | -1.3052 | -1.2793
b4 | -0.82953 | 0.066807 | -0.06608 | 0.021773 | -0.027321 | -0.14138 | -0.15092 | -0.036118 | -0.11569 | -0.12004 | -0.27537 | -0.13122 | -0.11733 | -0.11132 | -0.07173 | 0.021306 | 0.1243 | -0.01369 | -0.132 | -0.43779 |  | b4 | -1.2526 | -0.037174 | 0.12294 | -0.12503 | -0.099244 | 0.040971 | 0.085408 | -0.03694 | -0.04085 | -0.048865 | -0.027336 | -0.023646 | -0.15468 | -0.085752 | -0.10185 | -0.19052 | 0.10944 | -0.11805 | -0.12006 | -0.20712
b5 | -4.7375 | -0.58992 | -0.60137 | -0.3 | -0.24296 | -0.0014298 | -0.06874 | -0.23465 | -0.23358 | -0.32973 | 0.17277 | -0.13428 | -0.096851 | -0.19044 | -0.065099 | -0.15175 | -0.3393 | -0.21232 | -0.3431 | -3.6673 |  | b5 | -5.5108 | -0.53032 | -0.74409 | -0.2822 | -0.3727 | -0.025824 | -0.027007 | -0.12266 | -0.17611 | 0.033745 | 0.46287 | 0.17108 | 0.0066188 | 0.019576 | -0.036988 | -0.3218 | -0.20599 | -0.25062 | -0.36005 | -3.4397
b6 | -0.020626 | -0.090857 | -0.097988 | -0.097469 | -0.10657 | -0.02401 | 0.038277 | 0.015846 | -0.015949 | -0.013424 | -0.066485 | 0.038182 | 0.08075 | 0.03694 | -0.031703 | -0.086197 | 0.025108 | 0.007593 | 0.035968 | 0.063229 |  | b6 | -0.057109 | 0.1441 | 0.068658 | -0.055435 | -0.045914 | -0.041401 | -0.035688 | 0.043325 | -0.02294 | 0.079937 | -0.11813 | -0.074186 | -0.023355 | -0.033895 | -0.03517 | 0.12078 | -0.031956 | -0.046252 | -0.049922 | -0.046912
b7 | 2.946 | 1.0099 | 1.1311 | 1.2077 | 1.2031 | 1.1499 | 1.1821 | 1.188 | 1.1913 | 1.1101 | 1.1919 | 1.1909 | 1.2402 | 1.2188 | 1.2372 | 1.2587 | 1.1445 | 1.2055 | 1.2234 | 0.97672 |  | b7 | 2.921 | 1.0406 | 1.0769 | 1.1777 | 1.1933 | 1.1695 | 1.1725 | 1.1578 | 1.1632 | 1.1484 | 1.2158 | 1.126 | 1.1947 | 1.1212 | 1.1505 | 1.2138 | 1.173 | 1.1659 | 1.1901 | 1.0407
b8 | -0.086416 | -0.0068483 | -0.043889 | -0.051625 | -0.02326 | -0.020462 | -0.021952 | -0.0017407 | -0.014173 | 0.011534 | -0.014125 | -0.0027543 | 0.037892 | 0.013416 | 0.0071769 | 0.0030058 | 0.021312 | -0.0060623 | 0.00059618 | -0.015746 |  | b8 | -0.15998 | 0.014879 | 0.019571 | -0.020943 | -0.011203 | -0.024897 | 0.00032213 | -0.026246 | -0.025188 | 0.0131 | -0.013842 | -0.0049063 | -0.0021298 | -0.0027077 | 0.011294 | 0.02224 | 0.018203 | -0.010346 | -0.0063191 | -0.031477
b9 | 0.35136 | 0.46725 | 0.45959 | 0.48558 | 0.50813 | 0.46279 | 0.47186 | 0.47122 | 0.48009 | 0.45619 | 0.45485 | 0.44959 | 0.4498 | 0.46849 | 0.44478 | 0.45535 | 0.46244 | 0.49118 | 0.48087 | 0.41636 |  | b9 | 0.3634 | 0.47775 | 0.46587 | 0.49324 | 0.50778 | 0.45586 | 0.47109 | 0.47911 | 0.48725 | 0.46994 | 0.46561 | 0.48766 | 0.48991 | 0.47764 | 0.47396 | 0.4772 | 0.4825 | 0.48486 | 0.4863 | 0.4427
b10 | 0.042534 | -0.013325 | -0.065892 | -0.09374 | 0.010075 | -0.031502 | -0.0014904 | 0.0065954 | 0.0087729 | 0.012211 | -0.051476 | 0.033318 | 0.072059 | 0.032493 | -0.0014443 | -0.0089215 | 0.020837 | 0.0045827 | 0.030111 | 0.019369 |  | b10 | 0.064176 | 0.051743 | 0.039149 | -0.022096 | 0.039123 | -0.0015644 | -0.019161 | -0.0045161 | 0.0028269 | 0.0064704 | -0.05091 | -0.028522 | 0.0081267 | 0.0025801 | -0.0026738 | 0.023653 | 0.0067436 | -0.014915 | -0.023486 | -0.077778
b11 | 0.58426 | 0.72121 | 0.71954 | 0.72625 | 0.72625 | 0.73846 | 0.73169 | 0.72098 | 0.72794 | 0.72543 | 0.74099 | 0.72672 | 0.73086 | 0.73536 | 0.74317 | 0.73344 | 0.72394 | 0.72548 | 0.72486 | 0.63585 |  | b11 | 0.5881 | 0.72502 | 0.72636 | 0.73606 | 0.73844 | 0.74864 | 0.74095 | 0.73369 | 0.73995 | 0.74496 | 0.75727 | 0.74066 | 0.73654 | 0.73972 | 0.73989 | 0.73982 | 0.72211 | 0.73324 | 0.72468 | 0.63106
b12 | -0.006326 | -0.0046675 | -0.010253 | -0.012128 | -0.0015905 | -0.0032444 | -0.0028361 | -9.0372e-05 | -0.00042133 | 0.0024474 | -0.0030627 | 0.0030282 | 0.0080965 | 0.0040857 | -0.0008775 | -0.0037142 | -0.00015648 | -0.0023644 | -0.0006513 | -0.00079441 |  | b12 | 0.0080491 | 0.0059152 | 0.006147 | -0.001048 | 0.0041496 | -0.0021644 | -0.0011511 | 0.00072473 | 0.001443 | -0.0011602 | -0.0076154 | -0.0015495 | 0.0026482 | -0.0015827 | -0.0014263 | 0.0039715 | 0.0016618 | 0.00035006 | 0.00010489 | -0.0047493
b13 | 0.092958 | 0.071871 | 0.073623 | 0.06742 | 0.068528 | 0.065539 | 0.064857 | 0.06652 | 0.067308 | 0.067885 | 0.065858 | 0.067346 | 0.069601 | 0.066588 | 0.066032 | 0.067749 | 0.068547 | 0.069839 | 0.071918 | 0.055464 |  | b13 | 0.099837 | 0.074035 | 0.072932 | 0.069449 | 0.070177 | 0.069036 | 0.070462 | 0.069265 | 0.068561 | 0.06926 | 0.064919 | 0.06629 | 0.068612 | 0.063906 | 0.066091 | 0.068793 | 0.068 | 0.071208 | 0.070722 | 0.056442
b14 | 0.004062 | -0.003954 | -0.0053859 | -0.0054082 | -0.001717 | -0.0028533 | -0.0022791 | -0.0026433 | -0.0018947 | -0.0023621 | -0.0049071 | -0.00056701 | 0.00091403 | -0.0016587 | -0.003092 | -0.0029548 | -0.0019836 | -0.0029597 | -0.0010773 | -0.0014613 |  | b14 | 0.0038474 | -0.0025808 | -0.0058235 | -0.0073404 | -0.0037879 | -0.0063264 | -0.0064866 | -0.0064499 | -0.0062149 | -0.0075369 | -0.0097511 | -0.0082952 | -0.0069383 | -0.0074126 | -0.0070658 | -0.0047508 | -0.0061695 | -0.006844 | -0.0072692 | -0.0034196
b15 | -0.0077616 | 0.02397 | 0.020076 | 0.017584 | 0.02056 | 0.022428 | 0.022608 | 0.02118 | 0.022325 | 0.021861 | 0.020762 | 0.020955 | 0.023099 | 0.02291 | 0.021038 | 0.021551 | 0.021825 | 0.02301 | 0.022123 | 0.0097151 |  | b15 | -0.012345 | 0.020947 | 0.019514 | 0.016519 | 0.01749 | 0.018555 | 0.01622 | 0.017853 | 0.019684 | 0.020797 | 0.015934 | 0.018734 | 0.021655 | 0.020636 | 0.019803 | 0.01864 | 0.01928 | 0.019175 | 0.017323 | -0.0019915
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 53.1216 | 0.15013 | -1.3958 | -4.6455 | -4.3304 | -2.4341 | -0.19084 | 2.5405 | 0.41848 | 2.0556 | -0.12781 | -1.0559 | 3.1344 | 5.2334 | 2.7771 | -0.81856 | -4.3061 | -6.0149 | -7.0215 | -23.7331 |  | a1 | 53.5904 | 1.7261 | -0.51582 | -3.1633 | -3.1865 | -0.91435 | 2.0643 | 3.6994 | 1.3769 | -0.66656 | -0.3443 | -1.0605 | 1.0638 | 4.2982 | 0.48158 | -3.1687 | -4.3106 | -7.6363 | -9.9844 | -16.1986
a2 | 7.425 | 0.24092 | 0.04623 | -0.43709 | 1.1755 | 0.3414 | -0.74456 | -0.10859 | -0.43934 | 0.13411 | 1.3619 | 0.27181 | 0.9508 | 1.3649 | 0.32823 | -1.2336 | 0.24771 | 0.018751 | -0.33398 | 1.4717 |  | a2 | 20.9088 | -0.30254 | -0.4763 | -0.18592 | -0.23864 | -0.71172 | -0.050238 | -0.7674 | -0.58149 | 0.12347 | -0.85755 | -0.2502 | -0.53028 | 1.1733 | -1.2716 | -0.5866 | -0.57539 | 0.24949 | -0.13313 | 12.4633
a3 | 1.3024 | 0.8174 | 0.58323 | 0.54237 | 0.71183 | 0.84464 | 0.88558 | 0.78276 | 0.51033 | 0.42001 | 0.64021 | 0.77647 | 0.77694 | 0.70041 | 0.65646 | 0.75141 | 0.91277 | 0.93548 | 1.2168 | 0.73359 |  | a3 | 0.17941 | 0.57414 | 0.39988 | -0.05024 | 0.47054 | 0.66983 | 0.88037 | 0.58163 | 0.26914 | 0.76204 | 1.1298 | 0.034919 | 0.18115 | -0.087243 | 0.052249 | 0.90998 | 0.96253 | 0.72395 | 0.87455 | 1.3537
a4 | 3.1018 | 0.15983 | -0.14852 | 0.054608 | -0.63105 | -0.66099 | -0.55667 | -0.25541 | -0.50826 | -0.32574 | -0.80926 | -1.0663 | -0.79823 | -0.35981 | -0.19451 | 0.050826 | -0.065972 | -0.24446 | -0.15435 | -0.33651 |  | a4 | 4.0029 | 0.025022 | -0.17668 | -0.10529 | -0.25037 | -0.25407 | -0.57091 | -0.41827 | -0.2326 | -0.55992 | -0.027844 | -0.45358 | -0.3938 | -0.10675 | 0.23308 | -0.25893 | 0.29148 | -0.24659 | -0.073714 | 0.6869
a5 | 3.185 | -0.098249 | 0.013776 | 0.15691 | 0.39922 | 0.24195 | 0.27821 | 0.10095 | 0.15405 | 0.037839 | 0.22999 | 0.054044 | 0.076838 | 0.15996 | 0.20677 | 0.12903 | 0.304 | 0.29562 | 0.40885 | -0.011853 |  | a5 | 2.5371 | -0.10519 | -0.045932 | -0.14275 | 0.096113 | 0.031748 | 0.20988 | 0.099575 | -0.008254 | 0.063327 | 0.25271 | 0.0065066 | -0.056697 | 0.081486 | 0.11605 | 0.11582 | 0.17146 | 0.26466 | 0.31588 | -0.10724
a6 | -0.021488 | 0.00019991 | 0.088834 | 0.0010833 | -0.16334 | -0.19678 | -0.13641 | -0.090408 | -0.031194 | -0.026118 | 0.013829 | -0.10883 | -0.16407 | -0.058694 | -0.048452 | -0.11926 | -0.0028763 | -0.0042427 | 0.0031439 | -0.26246 |  | a6 | 0.19055 | 0.057107 | 0.052933 | 0.071033 | 0.1344 | 0.049871 | 0.062423 | 0.08738 | 0.18022 | 0.11118 | -0.0057188 | 0.067422 | 0.033291 | 0.10752 | 0.08518 | 0.11104 | 0.14544 | 0.10491 | 0.17601 | 0.092789
a7 | 1.7852 | -0.080676 | -0.043742 | -0.070977 | 0.016618 | 0.0098424 | 0.017519 | 0.01165 | -0.02897 | -0.051425 | -0.068724 | -0.0092093 | 0.092969 | 0.027932 | -0.052777 | -0.0020952 | 0.016216 | 0.018146 | 0.049214 | -0.060616 |  | a7 | 1.694 | -0.029564 | -0.068653 | -0.077157 | 0.0039734 | 0.12363 | 0.050528 | 0.093146 | 0.017435 | -0.010214 | 0.070628 | -0.081377 | 0.053161 | 0.031913 | -0.022857 | 0.029042 | 0.10451 | 0.061587 | 0.072891 | 0.0025579
a8 | -0.033607 | -0.00018285 | 0.019666 | -0.0053045 | -0.04347 | -0.057305 | -0.061014 | -0.022863 | 0.0059449 | 0.0077902 | -0.01147 | -0.016424 | -0.041164 | -0.045553 | -0.033134 | -0.049283 | -0.00010466 | -0.0017284 | 0.0073488 | -0.098085 |  | a8 | -0.018212 | 0.018236 | -0.0039557 | -0.022123 | -0.013355 | -0.059127 | 0.00055301 | -0.020532 | 0.0050842 | 0.0035859 | -0.067519 | -0.026125 | -0.046585 | -0.017208 | -0.064655 | -0.042208 | -0.008864 | -0.022438 | 0.011508 | -0.011579
a9 | -0.1916 | -0.0068612 | -0.028182 | -0.0093752 | 0.017556 | 0.010028 | 0.0078542 | -0.016943 | -0.0005622 | -0.023691 | 0.0067162 | -0.020005 | -0.0084254 | -0.0086489 | -0.018214 | -0.0011954 | -0.0080052 | 0.00027857 | 0.020434 | -0.056391 |  | a9 | -0.27255 | 0.01565 | -0.018917 | -0.01741 | 0.0013902 | 0.022646 | 0.022185 | 0.016192 | -0.021664 | -0.010535 | 0.031237 | -0.021509 | -0.010087 | 0.0049182 | 0.0058746 | 0.0034488 | 0.012244 | 0.028353 | 0.029818 | -0.031218
a10 | -0.13222 | -0.011926 | 0.034108 | 0.028789 | 0.016978 | -0.02687 | -0.0031675 | 0.005244 | 0.01913 | 0.013106 | 0.015751 | 0.018926 | -0.041944 | -0.046039 | -0.0010046 | -0.0085376 | 0.036951 | 0.030883 | 0.056674 | -0.097059 |  | a10 | -0.096668 | 0.024901 | 0.04608 | 0.025097 | 0.0037917 | -0.029309 | -0.00016343 | 0.0021924 | 0.00017492 | -0.024404 | -0.02007 | 0.01161 | -0.019402 | -0.024295 | -0.0039521 | -0.0016823 | 0.023144 | 0.0080259 | 0.035191 | -0.0031901
a11 | 0.18502 | -0.067868 | -0.053247 | -0.056389 | -0.039945 | -0.044237 | -0.035921 | -0.04324 | -0.042803 | -0.048304 | -0.044993 | -0.04718 | -0.030699 | -0.039603 | -0.043544 | -0.040124 | -0.039415 | -0.038857 | -0.035409 | -0.060056 |  | a11 | 0.19573 | -0.027384 | -0.017778 | -0.020833 | -0.0036997 | 0.003287 | 0.001541 | -0.0007708 | -0.0031283 | -0.01372 | -0.0027372 | -0.017843 | -0.0031408 | -2.1327e-05 | -0.013651 | -0.011696 | -0.0040712 | -0.0029345 | -0.006703 | -0.011694
a12 | -0.026389 | -7.8405e-05 | 0.0045688 | 0.00030862 | -0.004242 | -0.0097302 | -0.0071372 | -0.0040744 | -0.0013038 | -0.0014866 | 0.0012134 | -0.0022051 | -0.0079161 | -0.0048455 | -0.0048086 | -0.0095463 | 0.0027795 | 0.0014123 | 0.0056183 | -0.016855 |  | a12 | -0.024597 | -0.00011397 | 0.00343 | 0.002351 | -0.00086046 | -0.0050969 | 0.0011003 | 0.0011421 | 0.00080325 | 0.0017068 | -0.0065898 | -0.00048879 | -0.0029924 | 0.0011271 | -0.0021555 | 0.0031124 | 0.0038866 | 0.0025498 | 0.0087256 | 0.0012401
a13 | -0.01706 | -0.0017588 | -0.0036595 | -0.0024743 | 0.0007065 | 6.4531e-05 | -2.9597e-05 | -0.0025889 | -0.0013402 | -0.0032095 | -0.00063575 | -0.0027424 | 2.0505e-05 | -0.0004187 | -0.003306 | -0.0024328 | 0.00088834 | 0.0011434 | 0.0031311 | -0.0021836 |  | a13 | -0.017356 | 0.0030163 | -0.00089307 | -0.0020859 | 0.00088849 | 0.0055373 | 0.0056252 | 0.0038254 | -0.0012035 | 0.0032249 | 0.0040097 | -0.00038994 | 0.0022405 | 0.0018786 | 0.0025462 | 0.0014156 | 0.0035836 | 0.0045988 | 0.0057456 | 0.00059618
a14 | 0.01003 | -0.0042715 | -0.0031066 | -0.0043057 | -0.0064428 | -0.0082033 | -0.0083789 | -0.0058077 | -0.0052877 | -0.0037483 | -0.005946 | -0.0062878 | -0.005478 | -0.0055867 | -0.0055692 | -0.0053907 | -0.003628 | -0.0042804 | -0.0016404 | -0.0040198 |  | a14 | 0.01071 | -0.0025384 | -0.0022406 | -0.0033179 | -0.0037397 | -0.0077024 | -0.0044278 | -0.0062874 | -0.0059263 | -0.005638 | -0.005835 | -0.0073695 | -0.0076484 | -0.00546 | -0.0051671 | -0.0051641 | -0.0040708 | -0.0047474 | -0.0025459 | 0.0008965
a15 | 0.0031876 | -0.0044601 | -0.0025325 | -0.0030479 | -0.00063312 | -0.0027946 | -0.00027577 | -0.00025458 | -0.00090582 | -0.0022144 | -0.0034419 | 0.0012585 | 0.0022019 | 0.00010394 | -0.001231 | -0.0014726 | 0.00016352 | -0.00031973 | 0.0012892 | 0.00093128 |  | a15 | -0.0036293 | -0.0023728 | -0.0012926 | -0.004349 | -0.0019256 | -0.0031906 | -0.0028732 | -0.002858 | -0.003663 | -0.0038174 | -0.003241 | -0.0040901 | -0.003033 | -0.0046119 | -0.0041719 | -0.00049116 | -0.0011799 | -0.0025556 | -0.001737 | -0.0046277
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.0017618 | -0.028511 | -0.16076 | -0.22331 | 0.021333 | -0.06857 | -0.0030941 | 0.014757 | 0.017799 | 0.026452 | -0.11985 | 0.07475 | 0.17141 | 0.080616 | -0.0031583 | -0.019538 | 0.044075 | 0.0068429 | 0.063973 | 0.07562 |  | Dx (mm) | 0.083283 | 0.11895 | 0.088933 | -0.05263 | 0.09002 | -0.0038445 | -0.043962 | -0.01047 | 0.006493 | 0.015786 | -0.11412 | -0.066069 | 0.018948 | 0.005931 | -0.0059738 | 0.054399 | 0.015568 | -0.034654 | -0.055854 | -0.20929
Dy (mm) | -0.38526 | -0.030794 | 0.068688 | 0.050051 | 0.040914 | -0.065964 | -0.0075112 | 0.01325 | 0.045722 | 0.032475 | 0.028858 | 0.049126 | -0.090363 | -0.10209 | -0.0024831 | -0.020858 | 0.08917 | 0.072735 | 0.13604 | -0.25235 |  | Dy (mm) | -0.30715 | 0.062879 | 0.11002 | 0.056473 | 0.0091799 | -0.066539 | -0.00028353 | 0.0050689 | 0.00042933 | -0.055398 | -0.045468 | 0.025056 | -0.044701 | -0.055834 | -0.0091907 | -0.0030058 | 0.054575 | 0.018469 | 0.082037 | -0.012472
Mag. len. (m) | 14.4207 | [20 empty] | Mag. len. (m) | 14.4212
TF (mT/KA) | 595.3783 | [20 empty] | TF (mT/KA) | 595.456
Coil pos. (mm) | 0.016444 | [20 empty] | Coil pos. (mm) | 0.017371
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.021627 | [20 empty] | Twist (m2 rad) | 0.026525

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 595.3783 | 0.14913 | -0.0013998 |  | -0.89 | 3.3877 |  |  | 595.456 | 0.12684 | 0.00089696 |  | -1.1249 | 3.6756 |  | 0 | 4
2 | -1.3543 |  | -1.3565 | 0.34827 | 0.50038 |  | 0.17855 | 0.69773 | 0.98021 |  | 1.0794 | 0.45016 | 0.7615 |  | -0.33298 | 0.52502 |  | 0 | 5
3 | -1.5374 |  | -3.072 | 1.162 | 0.76352 |  | 0.74954 | 0.18494 | -0.91102 |  | -2.2278 | 0.88914 | 0.53305 |  | 0.51857 | 0.37956 |  | 0 | 6
4 | -0.11165 |  | -0.070297 | 0.092801 | -0.27488 |  | -0.36018 | 0.33999 | -0.093525 |  | -0.044637 | 0.091196 | -0.056003 |  | -0.19553 | 0.24088 |  | 0 | 7
5 | -0.46833 |  | -0.22086 | 0.1897 | 0.26105 |  | 0.17599 | 0.13356 | -0.42106 |  | -0.15395 | 0.2718 | 0.14753 |  | 0.081105 | 0.12966 |  | 0 | 8
6 | -0.023259 |  | -0.01774 | 0.038763 | -0.056172 |  | -0.066095 | 0.065229 | -0.017373 |  | -0.009268 | 0.05498 | 0.099184 |  | 0.087929 | 0.046806 |  | 0 | 9
7 | 1.2292 |  | 1.1821 | 0.058025 | 0.044212 |  | -0.0078636 | 0.046872 | 1.2069 |  | 1.1584 | 0.044947 | 0.07137 |  | 0.023198 | 0.061986 |  | 0 | 10
8 | -0.0068976 |  | -0.0048588 | 0.011666 | -0.022456 |  | -0.023319 | 0.018424 | -0.0065298 |  | -0.0031387 | 0.015368 | -0.020273 |  | -0.022623 | 0.021907 |  | 0 | 11
9 | 0.46297 |  | 0.46689 | 0.016959 | -0.011049 |  | -0.0041716 | 0.012899 | 0.47488 |  | 0.47943 | 0.011836 | -0.004088 |  | 0.004905 | 0.01841 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.7218 |  | 0.72889 | 0.0070175 | -0.037485 |  | -0.043847 | 0.0086491 | 0.7293 |  | 0.73708 | 0.0087948 | -0.0021371 |  | -0.0081471 | 0.0087253 |  | 0 | 14
12 | -0.0001377 |  | -0.0012943 | 0.0016132 | -0.00032638 |  | -0.0032098 | 0.0027835 | 6.5545e-05 |  | 0.00038579 | 0.0017035 | 2.1456e-05 |  | 0.00022586 | 0.0026634 |  | 0 | 15
13 | 0.0068316 |  | 0.068031 | 0.0023255 | -0.00014948 |  | -0.001006 | 0.0019705 | 0.0069353 |  | 0.06895 | 0.0026486 | 0.00017852 |  | 0.0024202 | 0.0023238 |  | 0 | 16
14 | -0.00023916 |  | -0.0025719 | 0.00051374 | -0.00050596 |  | -0.0055702 | 0.0013451 | -0.00061114 |  | -0.0065656 | 0.00088199 | -0.00045703 |  | -0.005125 | 0.0012314 |  | 0 | 17
15 | 0.0020384 |  | 0.021659 | 0.0014403 | -8.4048e-05 |  | -0.0010315 | 0.001787 | 0.0017214 |  | 0.01882 | 0.0016757 | -0.00029394 |  | -0.0028585 | 0.0011813 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.70735 | -0.72124
3 | [12 empty] | 3 | 9 | -0.76879 | -0.54312
4 | [12 empty] | 4 | 16 | -1.7281 | -1.9091
5 | [12 empty] | 5 | 25 | -1.8378 | -1.6728
6 | [12 empty] | 6 | 36 | -2.9901 | -2.9813
7 | [12 empty] | 7 | 49 | -2.9536 | -2.9416
8 | [12 empty] | 8 | 64 | -4.2226 | -3.9982
9 | [12 empty] | 9 | 81 | -4.2138 | -4.2157
10 | [12 empty] | sigma (mm) |  | 0.016444 | 0.017371
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 345.2856 | 595.712 | 595.3362 | 595.3186 | 595.2377 | 595.3883 | 595.3437 | 595.3029 | 595.3553 | 595.2579 | 595.2708 | 595.2717 | 595.3241 | 595.3094 | 595.4152 | 595.3561 | 595.4241 | 595.377 | 595.8085 | 370.5259
2 | -2.7495 | -1.5547 | -1.5724 | -0.61929 | -1.0087 | -1.0836 | -1.7993 | -1.2613 | -1.7259 | -1.4714 | -1.682 | -1.8392 | -1.6283 | -1.1878 | -0.91925 | -0.97463 | -1.3232 | -1.6082 | -1.1573 | 0.0096551 | -1.3543
3 | 44.9068 | -2.1191 | -1.6742 | -3.1679 | -2.6549 | -3.902 | -2.9709 | -2.2418 | -2.1555 | -3.1064 | -5.3974 | -3.7332 | -3.7775 | -4.075 | -5.257 | -3.718 | -1.6124 | -2.092 | -1.6403 | -0.43516 | -1.5374
4 | -1.1163 | 0.063562 | -0.088605 | 0.0078574 | -0.022258 | -0.14516 | -0.15146 | -0.034988 | -0.11305 | -0.1177 | -0.26894 | -0.12823 | -0.11506 | -0.11155 | -0.0719 | 0.019975 | 0.1342 | -0.0082893 | -0.11375 | -0.37183 | -0.11165
5 | -4.735 | -0.59068 | -0.60421 | -0.30638 | -0.24425 | 0.0019037 | -0.068404 | -0.23507 | -0.23392 | -0.32988 | 0.17054 | -0.1367 | -0.096561 | -0.18955 | -0.065093 | -0.15152 | -0.3397 | -0.21243 | -0.34365 | -3.6492 | -0.46833
6 | -0.26519 | -0.079819 | -0.03487 | -0.0035325 | -0.11539 | 0.0035902 | 0.039521 | 0.009713 | -0.0239 | -0.024378 | -0.016769 | 0.0066044 | 0.0027579 | 0.0012564 | -0.030278 | -0.077502 | 0.0078145 | 0.0051474 | 0.010709 | 0.04256 | -0.023259
7 | 2.9514 | 1.0098 | 1.1288 | 1.2029 | 1.2026 | 1.1509 | 1.1823 | 1.1879 | 1.1916 | 1.1101 | 1.191 | 1.1906 | 1.239 | 1.2202 | 1.2373 | 1.2592 | 1.1441 | 1.2054 | 1.2238 | 0.98741 | 1.2292
8 | -0.05197 | -0.00047978 | -0.010031 | -0.00081735 | -0.028023 | -0.0058398 | -0.021293 | -0.0051187 | -0.018207 | 0.0054937 | 0.01162 | -0.019032 | 0.001969 | -0.0039417 | 0.0078592 | 0.0072041 | 0.011385 | -0.0076344 | -0.012572 | -0.023866 | -0.0068976
9 | 0.37829 | 0.46724 | 0.45522 | 0.47526 | 0.50838 | 0.46258 | 0.47187 | 0.47121 | 0.48047 | 0.45624 | 0.45182 | 0.44877 | 0.44526 | 0.4696 | 0.44478 | 0.45535 | 0.4637 | 0.49235 | 0.48393 | 0.42855 | 0.46297
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.59085 | 0.72112 | 0.71868 | 0.72451 | 0.72616 | 0.73874 | 0.73172 | 0.72095 | 0.72791 | 0.72536 | 0.74077 | 0.7265 | 0.73043 | 0.73547 | 0.74318 | 0.73352 | 0.7241 | 0.72556 | 0.72539 | 0.63864 | 0.7218
12 | -0.018023 | -0.031828 | -0.020763 | -0.015882 | -0.02602 | -0.00075195 | -0.026942 | -0.0080749 | -0.013103 | 0.011062 | 0.024959 | -0.0062034 | -0.0032596 | 0.0032666 | -0.007245 | -0.02744 | -0.022332 | -0.02643 | -0.035982 | -0.03366 | -0.001377 |  | ten times!
13 | 0.89998 | 0.71886 | 0.72798 | 0.66331 | 0.68354 | 0.65804 | 0.649 | 0.66491 | 0.67149 | 0.6784 | 0.65277 | 0.67142 | 0.69859 | 0.67127 | 0.66035 | 0.67791 | 0.68367 | 0.69616 | 0.718 | 0.56324 | 0.068316 |  | ten times!
14 | 0.030619 | -0.032781 | -0.028712 | -0.022999 | -0.020995 | -0.01435 | -0.022198 | -0.029035 | -0.02256 | -0.028976 | -0.029397 | -0.018061 | -0.025105 | -0.031884 | -0.030347 | -0.025827 | -0.027638 | -0.031085 | -0.020983 | -0.022599 | -0.0023916 |  | ten times!
15 | -0.077616 | 0.2397 | 0.20076 | 0.17584 | 0.2056 | 0.22428 | 0.22608 | 0.2118 | 0.22325 | 0.21861 | 0.20762 | 0.20955 | 0.23099 | 0.2291 | 0.21038 | 0.21551 | 0.21825 | 0.2301 | 0.22123 | 0.097151 | 0.020384 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 53.0601 | 0.14773 | -1.3892 | -4.6496 | -4.3294 | -2.4368 | -0.19177 | 2.5416 | 0.4236 | 2.0582 | -0.11547 | -1.0517 | 3.1158 | 5.2196 | 2.777 | -0.82116 | -4.2997 | -6.008 | -7.0107 | -23.7392 | -0.0013998
2 | 9.4696 | 0.23598 | 0.07076 | -0.40418 | 1.1865 | 0.31788 | -0.74687 | -0.10645 | -0.42883 | 0.14466 | 1.3892 | 0.28651 | 0.8948 | 1.3092 | 0.32694 | -1.241 | 0.25987 | 0.035873 | -0.31692 | 1.4516 | 0.50038
3 | 1.245 | 0.81857 | 0.57982 | 0.54433 | 0.7144 | 0.83499 | 0.88507 | 0.78351 | 0.51286 | 0.42222 | 0.6245 | 0.79167 | 0.79921 | 0.70352 | 0.65632 | 0.75166 | 0.91133 | 0.93595 | 1.2217 | 0.71859 | 0.76352
4 | 2.671 | 0.1549 | -0.13828 | 0.066386 | -0.63072 | -0.65711 | -0.55658 | -0.25503 | -0.50639 | -0.32346 | -0.80395 | -1.0657 | -0.80339 | -0.36742 | -0.1944 | 0.050674 | -0.062006 | -0.24131 | -0.14952 | -0.55405 | -0.27488
5 | 3.1826 | -0.09907 | 0.019956 | 0.15842 | 0.40153 | 0.23752 | 0.27817 | 0.10128 | 0.15443 | 0.03817 | 0.23105 | 0.055885 | 0.087255 | 0.16246 | 0.2067 | 0.12781 | 0.30338 | 0.29547 | 0.40735 | -0.0013224 | 0.26105
6 | 0.37798 | 0.010364 | 0.058931 | -0.025846 | -0.18084 | -0.16977 | -0.13325 | -0.096024 | -0.050237 | -0.038361 | -0.0012171 | -0.12923 | -0.13014 | -0.015575 | -0.047426 | -0.11 | -0.039148 | -0.035233 | -0.056707 | -0.17385 | -0.056172
7 | 1.7715 | -0.080765 | -0.041198 | -0.070401 | 0.017392 | 0.0076686 | 0.017373 | 0.011799 | -0.028747 | -0.051664 | -0.069123 | -0.0086481 | 0.097284 | 0.030008 | -0.052813 | -0.0024658 | 0.015435 | 0.018332 | 0.048987 | -0.059198 | 0.044212
8 | 0.030253 | 0.0064962 | 0.0026782 | -0.017727 | -0.05343 | -0.042616 | -0.059335 | -0.025683 | -0.0043802 | 0.0011135 | -0.017268 | -0.026114 | -0.021358 | -0.022717 | -0.032642 | -0.044824 | -0.019344 | -0.018541 | -0.024051 | -0.046634 | -0.022456
9 | -0.1828 | -0.0072585 | -0.022883 | -0.0034877 | 0.017146 | 0.0079525 | 0.0078431 | -0.01703 | -0.00095482 | -0.024085 | 0.008502 | -0.02162 | -0.001172 | -0.0049279 | -0.018217 | -0.0013822 | -0.0098511 | -9.7807e-06 | 0.016346 | -0.049918 | -0.011049
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.18347 | -0.067963 | -0.052316 | -0.055952 | -0.039845 | -0.044807 | -0.035949 | -0.0432 | -0.042776 | -0.04833 | -0.044842 | -0.04717 | -0.029348 | -0.03908 | -0.043556 | -0.040295 | -0.039486 | -0.038752 | -0.035585 | -0.059361 | -0.037485
12 | -0.010876 | 0.014485 | 0.0058391 | -0.024633 | -0.062318 | -0.066753 | -0.067934 | -0.046696 | -0.034592 | -0.029828 | -0.001819 | -0.043957 | -0.03479 | -0.00023034 | -0.047002 | -0.085824 | -0.015627 | -0.021789 | -0.014294 | -0.068588 | -0.0032638 |  | ten times!
13 | -0.15877 | -0.01945 | -0.037585 | -0.030025 | 0.0086532 | -0.0050954 | -0.00062513 | -0.024966 | -0.01202 | -0.03075 | -0.010724 | -0.023617 | 0.008017 | -0.0020379 | -0.033253 | -0.025605 | 0.011459 | 0.013304 | 0.033234 | -0.022331 | -0.0014948 |  | ten times!
14 | 0.075624 | -0.037684 | -0.045775 | -0.055911 | -0.071244 | -0.071428 | -0.082397 | -0.060357 | -0.06115 | -0.042847 | -0.067791 | -0.072131 | -0.040699 | -0.036674 | -0.055294 | -0.050442 | -0.052367 | -0.056569 | -0.041868 | -0.020588 | -0.0050596 |  | ten times!
15 | 0.031876 | -0.044601 | -0.025325 | -0.030479 | -0.0063312 | -0.027946 | -0.0027577 | -0.0025458 | -0.0090582 | -0.022144 | -0.034419 | 0.012585 | 0.022019 | 0.0010394 | -0.01231 | -0.014726 | 0.0016352 | -0.0031973 | 0.012892 | 0.0093128 | -0.00084048 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 338.4721 | 595.7588 | 595.3743 | 595.4069 | 595.2919 | 595.4598 | 595.3872 | 595.4027 | 595.4127 | 595.4389 | 595.452 | 595.5015 | 595.4452 | 595.3787 | 595.3934 | 595.4117 | 595.4225 | 595.4655 | 595.8035 | 375.4916
2 | -1.9756 | 0.40594 | 0.73512 | 0.5642 | 0.6508 | 0.98724 | 1.4263 | 1.4834 | 0.53947 | 0.72869 | 1.4143 | 1.1189 | 0.87605 | 0.90808 | 1.2603 | 2.1497 | 1.5085 | 1.1966 | 1.475 | 0.81415 | 0.98021
3 | 41.2036 | -1.6248 | -1.2093 | -2.273 | -2.073 | -3.7146 | -2.4755 | -1.5072 | -2.0937 | -2.444 | -4.0476 | -1.3427 | -1.2526 | -3.2619 | -3.3117 | -2.7718 | -1.4164 | -1.9728 | -1.3074 | -1.2885 | -0.91102
4 | -1.328 | -0.023888 | 0.13732 | -0.13042 | -0.091142 | 0.04045 | 0.085115 | -0.037123 | -0.040582 | -0.049816 | -0.017611 | -0.020948 | -0.15411 | -0.08685 | -0.10215 | -0.18648 | 0.1124 | -0.11894 | -0.11869 | -0.3762 | -0.093525
5 | -5.5266 | -0.53431 | -0.74417 | -0.28188 | -0.37112 | -0.026846 | -0.027473 | -0.1224 | -0.17604 | 0.031562 | 0.45898 | 0.17013 | 0.0063112 | 0.01787 | -0.03728 | -0.32383 | -0.20351 | -0.25052 | -0.35662 | -3.4429 | -0.42106
6 | -0.32661 | 0.099761 | 0.032191 | -0.035098 | -0.083815 | -0.042718 | -0.0175 | 0.04777 | -0.025603 | 0.073738 | -0.070295 | -0.04865 | -0.032183 | -0.036871 | -0.032671 | 0.097448 | -0.036388 | -0.03159 | -0.02435 | 0.02995 | -0.017373
7 | 2.9288 | 1.0403 | 1.076 | 1.1767 | 1.1937 | 1.171 | 1.1726 | 1.1576 | 1.1633 | 1.1483 | 1.2164 | 1.1256 | 1.1956 | 1.1216 | 1.1508 | 1.2133 | 1.1727 | 1.1656 | 1.1904 | 1.038 | 1.2069
8 | -0.13483 | -0.011401 | -0.00090558 | -0.009189 | -0.032708 | -0.024781 | 0.010065 | -0.023847 | -0.026681 | 0.0098838 | 0.010496 | 0.010002 | -0.006286 | -0.00417 | 0.012601 | 0.010019 | 0.014983 | -0.0021923 | 0.0076139 | 0.012308 | -0.0065298
9 | 0.37629 | 0.47532 | 0.46671 | 0.49338 | 0.50593 | 0.45689 | 0.47065 | 0.47909 | 0.48724 | 0.4706 | 0.46302 | 0.48682 | 0.49029 | 0.47835 | 0.47397 | 0.47652 | 0.48311 | 0.48466 | 0.48713 | 0.4341 | 0.47488
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.59255 | 0.72456 | 0.72625 | 0.73611 | 0.7382 | 0.74885 | 0.74091 | 0.7337 | 0.73994 | 0.74491 | 0.7569 | 0.74059 | 0.73659 | 0.73969 | 0.7399 | 0.73967 | 0.72223 | 0.73327 | 0.72514 | 0.63041 | 0.7293
12 | 0.059429 | -0.0016721 | 0.014993 | 0.014489 | -0.0030389 | -0.022372 | 0.010344 | 0.012503 | 0.011284 | -0.020581 | -0.025144 | 0.015352 | 0.016598 | -0.019243 | -0.011641 | 0.013268 | 0.010526 | 0.021519 | 0.032259 | 0.03584 | 0.00065545 |  | ten times!
13 | 0.97076 | 0.74147 | 0.73139 | 0.6901 | 0.70411 | 0.69409 | 0.70245 | 0.69189 | 0.6859 | 0.69589 | 0.64271 | 0.6573 | 0.68974 | 0.64172 | 0.66095 | 0.69002 | 0.67904 | 0.7096 | 0.70252 | 0.55886 | 0.069353 |  | ten times!
14 | 0.056121 | -0.047556 | -0.073698 | -0.068267 | -0.050991 | -0.060928 | -0.058987 | -0.063079 | -0.063214 | -0.076331 | -0.081322 | -0.073603 | -0.071646 | -0.073013 | -0.069368 | -0.055846 | -0.064697 | -0.063356 | -0.065897 | -0.037154 | -0.0061114 |  | ten times!
15 | -0.12345 | 0.20947 | 0.19514 | 0.16519 | 0.1749 | 0.18555 | 0.1622 | 0.17853 | 0.19684 | 0.20797 | 0.15934 | 0.18734 | 0.21655 | 0.20636 | 0.19803 | 0.1864 | 0.1928 | 0.19175 | 0.17323 | -0.019915 | 0.017214 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 53.4597 | 1.7268 | -0.51797 | -3.1658 | -3.1856 | -0.91062 | 2.0642 | 3.6985 | 1.3771 | -0.6643 | -0.34612 | -1.0632 | 1.0667 | 4.3007 | 0.48182 | -3.1664 | -4.3149 | -7.6371 | -9.992 | -16.0446 | 0.00089696
2 | 22.4031 | -0.29854 | -0.4649 | -0.17115 | -0.24138 | -0.74051 | -0.045767 | -0.76579 | -0.58159 | 0.10609 | -0.86403 | -0.24598 | -0.53729 | 1.1519 | -1.2752 | -0.5934 | -0.56804 | 0.25672 | -0.11478 | 12.4948 | 0.7615
3 | 0.052682 | 0.57403 | 0.40027 | -0.049972 | 0.47468 | 0.67014 | 0.87594 | 0.58089 | 0.26941 | 0.76313 | 1.1291 | 0.029735 | 0.18124 | -0.087976 | 0.05233 | 0.91237 | 0.96068 | 0.72282 | 0.87556 | 1.3786 | 0.53305
4 | 3.5549 | 0.035812 | -0.15646 | -0.10331 | -0.2516 | -0.25445 | -0.56874 | -0.41788 | -0.23257 | -0.55972 | -0.016106 | -0.45449 | -0.39348 | -0.1066 | 0.23316 | -0.26064 | 0.29349 | -0.24334 | -0.062612 | 0.67152 | -0.056003
5 | 2.5273 | -0.10986 | -0.049539 | -0.14073 | 0.092679 | 0.030994 | 0.21069 | 0.09978 | -0.0085952 | 0.064114 | 0.25094 | 0.0083635 | -0.057189 | 0.080742 | 0.1161 | 0.11415 | 0.17131 | 0.26598 | 0.31997 | -0.1017 | 0.14753
6 | 0.45741 | 0.035256 | 0.013271 | 0.046127 | 0.13041 | 0.077502 | 0.063324 | 0.085653 | 0.18001 | 0.1337 | 0.016637 | 0.055567 | 0.051785 | 0.12955 | 0.088863 | 0.11177 | 0.12227 | 0.098059 | 0.14299 | 0.097559 | 0.099184
7 | 1.6744 | -0.030842 | -0.069395 | -0.077149 | 0.0045108 | 0.12285 | 0.050538 | 0.093112 | 0.017426 | -0.0099385 | 0.067196 | -0.082037 | 0.053485 | 0.031893 | -0.022974 | 0.030015 | 0.10416 | 0.061345 | 0.073369 | 0.0013984 | 0.07137
8 | 0.044996 | 0.0032233 | -0.027285 | -0.035663 | -0.015608 | -0.044812 | 0.0010748 | -0.021595 | 0.0050519 | 0.015915 | -0.055879 | -0.032544 | -0.036189 | -0.0046722 | -0.062589 | -0.041621 | -0.021345 | -0.026189 | -0.0064822 | -0.012055 | -0.020273
9 | -0.25785 | 0.012359 | -0.023367 | -0.01605 | 0.0010194 | 0.022531 | 0.022178 | 0.016216 | -0.021666 | -0.010142 | 0.028798 | -0.020725 | -0.0097001 | 0.0050707 | 0.005849 | 0.0035349 | 0.011858 | 0.028646 | 0.031878 | -0.032085 | -0.004088
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.19866 | -0.027616 | -0.018413 | -0.020715 | -0.0036743 | 0.0031811 | 0.0015721 | -0.00076544 | -0.0031321 | -0.013779 | -0.0034479 | -0.017839 | -0.0030276 | -8.2833e-05 | -0.013668 | -0.011798 | -0.0041691 | -0.0028815 | -0.0063932 | -0.011565 | -0.0021371
12 | -0.01931 | -0.036533 | -0.021781 | -0.0049494 | -0.013717 | -0.018394 | 0.01289 | 0.0092256 | 0.0078799 | 0.043793 | -0.041833 | -0.016794 | -0.0085742 | 0.036379 | -0.017159 | 0.03204 | 0.012276 | 0.017339 | 0.048568 | 0.018251 | 0.00021456 |  | ten times!
13 | -0.17134 | 0.033713 | -0.0025072 | -0.019025 | 0.011725 | 0.051927 | 0.054749 | 0.038001 | -0.01172 | 0.029737 | 0.031614 | -0.0060333 | 0.021142 | 0.015869 | 0.024729 | 0.016195 | 0.038896 | 0.045696 | 0.060929 | 0.0070705 | 0.0017852 |  | ten times!
14 | 0.078366 | -0.033906 | -0.03914 | -0.042746 | -0.037292 | -0.066957 | -0.04528 | -0.063866 | -0.059137 | -0.046395 | -0.055429 | -0.079786 | -0.068039 | -0.044886 | -0.050378 | -0.05096 | -0.049222 | -0.051119 | -0.037961 | 0.0007841 | -0.0045703 |  | ten times!
15 | -0.036293 | -0.023728 | -0.012926 | -0.04349 | -0.019256 | -0.031906 | -0.028732 | -0.02858 | -0.03663 | -0.038174 | -0.03241 | -0.040901 | -0.03033 | -0.046119 | -0.041719 | -0.0049116 | -0.011799 | -0.025556 | -0.01737 | -0.046277 | -0.0029394 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00056047 | -7.0721e-05 | -0.00010028 | -0.00023616 | 1.6787e-05 | -5.8124e-05 | -0.00012665 | -3.864e-05 | -0.00020223 | -0.00018057 | -0.00017906 | -9.1044e-05 | -0.00011573 | 6.1968e-05 | -3.7297e-05 | 7.6917e-05 | -2.1928e-06 | 0.00072256 |  | 595.3783
|  | 0.00050859 | -0.00013713 | -8.2383e-05 | -0.00027551 | 6.4563e-06 | -0.00011547 | -8.9437e-05 | -7.2643e-05 | -2.8643e-05 | -6.6429e-06 | 7.6487e-05 | -1.8063e-05 | -0.00012974 | -0.00010505 | -7.4322e-05 | -5.6185e-05 | 1.6029e-05 | 0.00058366 |  | 595.456
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3988.1173
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 595.3306 | -4200.1029 | 6.4072 | 0.094695 | -0.20094 | -1.5598 | 0.96984 | 0.22068 | -0.46466 | 0.41553 | -1.2205 | -1.0039 | -0.98874 | -0.10855 | -0.35547 | 1.4217 | 0.42896 | 1.5712 | 0.78003 | 8.0281 | -3776.1317
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.089 | 5.395 | 0.10377 | -0.049916 | -0.37596 | -0.34394 | -0.15468 | 0.069823 | 0.34316 | 0.13136 | 0.29482 | 0.077452 | -0.016169 | 0.40058 | 0.61096 | 0.3667 | 0.0068838 | -0.34097 | -0.5118 | -0.61207 | -2.2849 |  | 0
db2 | -1.3565 | -2.7495 | -0.19825 | -0.21594 | 0.73718 | 0.34782 | 0.27284 | -0.44279 | 0.095138 | -0.36939 | -0.11493 | -0.3255 | -0.48277 | -0.27187 | 0.16865 | 0.43722 | 0.38184 | 0.033302 | -0.25169 | 0.19913 | 0.0096551 |  | -6.1679e-17
db3 | -3.0716 | 44.9068 | 0.95285 | 1.3978 | -0.095916 | 0.41708 | -0.83003 | 0.10112 | 0.83016 | 0.91644 | -0.034426 | -2.3254 | -0.66117 | -0.70556 | -1.003 | -2.185 | -0.64606 | 1.4596 | 0.97994 | 1.4317 | -0.43516 |  | 0
db4 | -0.070297 | -1.1163 | 0.13386 | -0.018308 | 0.078154 | 0.048039 | -0.074862 | -0.081166 | 0.035309 | -0.042754 | -0.0474 | -0.19864 | -0.057933 | -0.044764 | -0.041251 | -0.0016028 | 0.090272 | 0.20449 | 0.062008 | -0.043453 | -0.37183 |  | 0
db5 | -0.22015 | -4.735 | -0.36982 | -0.38335 | -0.085518 | -0.023388 | 0.22277 | 0.15246 | -0.01421 | -0.013056 | -0.10901 | 0.39141 | 0.084168 | 0.1243 | 0.031315 | 0.15577 | 0.069347 | -0.11884 | 0.0084352 | -0.12279 | -3.6492 |  | 0
db6 | -0.01774 | -0.26519 | -0.062078 | -0.01713 | 0.014208 | -0.097649 | 0.02133 | 0.057261 | 0.027453 | -0.0061602 | -0.0066374 | 0.00097085 | 0.024345 | 0.020498 | 0.018997 | -0.012538 | -0.059762 | 0.025555 | 0.022888 | 0.028449 | 0.04256 |  | 0
db7 | 1.1825 | 2.9514 | -0.1723 | -0.053312 | 0.02082 | 0.020504 | -0.031187 | 0.00017994 | 0.005792 | 0.009484 | -0.072011 | 0.0089478 | 0.0085458 | 0.056969 | 0.038155 | 0.055196 | 0.077099 | -0.037961 | 0.023376 | 0.041702 | 0.98741 |  | 8.6351e-17
db8 | -0.0048588 | -0.05197 | 0.0043791 | -0.0051722 | 0.0040415 | -0.023164 | -0.00098093 | -0.016434 | -0.00025982 | -0.013348 | 0.010353 | 0.016478 | -0.014173 | 0.0068279 | 0.00091715 | 0.012718 | 0.012063 | 0.016244 | -0.0027756 | -0.0077133 | -0.023866 |  | 0
db9 | 0.46778 | 0.37829 | 0.00035391 | -0.011668 | 0.0083687 | 0.04149 | -0.0043094 | 0.0049761 | 0.0043149 | 0.013584 | -0.01065 | -0.01507 | -0.018123 | -0.021627 | 0.0027053 | -0.022114 | -0.01154 | -0.0031898 | 0.025457 | 0.017041 | 0.42855 |  | -3.084e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.72903 | 0.59085 | -0.0077696 | -0.010215 | -0.0043856 | -0.0027277 | 0.0098438 | 0.0028251 | -0.007944 | -0.00098178 | -0.0035344 | 0.011881 | -0.0023895 | 0.0015374 | 0.006573 | 0.014286 | 0.0046308 | -0.004788 | -0.0033354 | -0.0035065 | 0.63864 |  | -4.9343e-17
db12 | -0.0012943 | -0.0018023 | -0.0018886 | -0.00078199 | -0.00029395 | -0.0013077 | 0.0012191 | -0.0014 | 0.00048678 | -1.5984e-05 | 0.0024005 | 0.0037902 | 0.00067394 | 0.00096831 | 0.0016209 | 0.00056978 | -0.0014497 | -0.00093892 | -0.0013487 | -0.002304 | -0.003366 |  | 0
db13 | 0.068031 | 0.089998 | 0.0038541 | 0.0047663 | -0.0017 | 0.00032296 | -0.0022279 | -0.0031315 | -0.0015401 | -0.00088255 | -0.00019139 | -0.0027548 | -0.00088965 | 0.0018279 | -0.00090481 | -0.0019967 | -0.00024061 | 0.00033508 | 0.001585 | 0.0037687 | 0.056324 |  | -3.084e-18
db14 | -0.0025719 | 0.0030619 | -0.00070627 | -0.00029938 | 0.00027192 | 0.00047233 | 0.0011368 | 0.0003521 | -0.0003316 | 0.00031582 | -0.00032575 | -0.0003678 | 0.00076579 | 6.1397e-05 | -0.00061655 | -0.00046287 | -1.089e-05 | -0.00019193 | -0.00053667 | 0.00047353 | -0.0022599 |  | 4.2163e-19
db15 | 0.021659 | -0.0077616 | 0.0023109 | -0.0015829 | -0.0040748 | -0.0010997 | 0.00076895 | 0.00094829 | -0.00047882 | 0.00066552 | 0.00020225 | -0.000897 | -0.00070404 | 0.0014399 | 0.0012507 | -0.00062148 | -0.0001087 | 0.0001662 | 0.0013512 | 0.00046355 | 0.0097151 |  | 7.7099e-19
da2 | 0.17855 | 9.4696 | 0.057431 | -0.10779 | -0.58273 | 1.0079 | 0.13933 | -0.92542 | -0.28501 | -0.60739 | -0.033891 | 1.2107 | 0.10796 | 0.71624 | 1.1307 | 0.14839 | -1.4196 | 0.081317 | -0.14268 | -0.49547 | 1.4516 |  | 0
da3 | 0.74954 | 1.245 | 0.06903 | -0.16971 | -0.2052 | -0.035137 | 0.085457 | 0.13554 | 0.033977 | -0.23668 | -0.32731 | -0.12503 | 0.042136 | 0.049674 | -0.046016 | -0.093215 | 0.0021256 | 0.1618 | 0.18642 | 0.47215 | 0.71859 |  | 0
da4 | -0.36018 | 2.671 | 0.51508 | 0.2219 | 0.42657 | -0.27053 | -0.29693 | -0.1964 | 0.10516 | -0.14621 | 0.036721 | -0.44377 | -0.70549 | -0.44321 | -0.0072361 | 0.16579 | 0.41086 | 0.29818 | 0.11887 | 0.21066 | -0.55405 |  | 0
da5 | 0.17599 | 3.1826 | -0.27506 | -0.15603 | -0.01757 | 0.22554 | 0.061532 | 0.10218 | -0.074704 | -0.021559 | -0.13782 | 0.05506 | -0.1201 | -0.088731 | -0.013522 | 0.030714 | -0.048173 | 0.12739 | 0.11948 | 0.23136 | -0.0013224 |  | 0
da6 | -0.066095 | 0.37798 | 0.076459 | 0.12503 | 0.04025 | -0.11474 | -0.10368 | -0.067157 | -0.029929 | 0.015858 | 0.027734 | 0.064878 | -0.063136 | -0.064046 | 0.050521 | 0.018669 | -0.043909 | 0.026947 | 0.030863 | 0.0093883 | -0.17385 |  | -3.4694e-18
da7 | -0.0078636 | 1.7715 | -0.072901 | -0.033335 | -0.062537 | 0.025256 | 0.015532 | 0.025237 | 0.019662 | -0.020883 | -0.0438 | -0.061259 | -0.00078447 | 0.10515 | 0.037872 | -0.04495 | 0.0053978 | 0.023299 | 0.026196 | 0.056851 | -0.059198 |  | 0
da8 | -0.023319 | 0.030253 | 0.029815 | 0.025997 | 0.0055922 | -0.030111 | -0.019297 | -0.036016 | -0.002364 | 0.018939 | 0.024433 | 0.0060508 | -0.0027946 | 0.0019611 | 0.0006016 | -0.0093226 | -0.021505 | 0.0039752 | 0.0047777 | -0.00073228 | -0.046634 |  | -1.9275e-19
da9 | -0.0041716 | -0.1828 | -0.0030868 | -0.018711 | 0.00068398 | 0.021317 | 0.012124 | 0.012015 | -0.012858 | 0.0032168 | -0.019913 | 0.012674 | -0.017449 | 0.0029997 | -0.00075626 | -0.014045 | 0.0027894 | -0.0056794 | 0.0041618 | 0.020518 | -0.049918 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.043847 | 0.18347 | -0.024116 | -0.0084689 | -0.012105 | 0.0040024 | -0.00095955 | 0.0078981 | 0.00064695 | 0.0010713 | -0.0044824 | -0.00099435 | -0.0033227 | 0.014499 | 0.0047672 | 0.00029153 | 0.0035521 | 0.0043616 | 0.0050948 | 0.0082626 | -0.059361 |  | -2.1202e-18
da12 | -0.0032098 | -0.0010876 | 0.0046583 | 0.0037937 | 0.00074646 | -0.0030219 | -0.0034655 | -0.0035836 | -0.0014598 | -0.00024945 | 0.00022696 | 0.0030279 | -0.0011859 | -0.00026925 | 0.0031868 | -0.0014904 | -0.0053726 | 0.0016471 | 0.0010309 | 0.0017805 | -0.0068588 |  | 0
da13 | -0.001006 | -0.015877 | -0.000939 | -0.0027525 | -0.0019965 | 0.0018714 | 0.00049649 | 0.00094353 | -0.0014905 | -0.00019595 | -0.002069 | -6.6367e-05 | -0.0013556 | 0.0018077 | 0.00080225 | -0.0023193 | -0.0015545 | 0.0021519 | 0.0023365 | 0.0043295 | -0.0022331 |  | 0
da14 | -0.0055702 | 0.0075624 | 0.0018018 | 0.00099262 | -2.0902e-05 | -0.0015543 | -0.0015726 | -0.0026696 | -0.00046557 | -0.00054487 | 0.0012854 | -0.0012089 | -0.0016429 | 0.0015003 | 0.0019027 | 4.0758e-05 | 0.00052597 | 0.00033349 | -8.6756e-05 | 0.0013834 | -0.0020588 |  | 0
da15 | -0.0010315 | 0.0031876 | -0.0034286 | -0.001501 | -0.0020164 | 0.00039839 | -0.0017631 | 0.00075573 | 0.00077692 | 0.00012569 | -0.0011829 | -0.0024104 | 0.00229 | 0.0032334 | 0.0011354 | -0.00019952 | -0.00044114 | 0.001195 | 0.00071178 | 0.0023207 | 0.00093128 |  | 0
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -4004.4899
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.4153 | -4315.3611 | 5.769 | -0.6887 | -0.14118 | -2.0726 | 0.74727 | -0.47204 | -0.21172 | -0.043772 | 0.39626 | 0.61627 | 1.4476 | 0.50207 | -0.6148 | -0.36792 | -0.060567 | 0.12082 | 0.843 | 6.5197 | -3693.6186
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.11249 | 5.4585 | 0.28517 | 0.060693 | -0.20409 | -0.20607 | 0.021428 | 0.31891 | 0.48234 | 0.2502 | 0.04606 | 0.077878 | 0.0061737 | 0.21916 | 0.54256 | 0.16067 | -0.20415 | -0.319 | -0.65122 | -0.88671 | -1.492 |  | 0
db2 | -1.0794 | -1.9756 | -0.67342 | -0.34425 | -0.51517 | -0.42856 | -0.092123 | 0.34689 | 0.40407 | -0.5399 | -0.35067 | 0.33492 | 0.039484 | -0.20332 | -0.17128 | 0.18091 | 1.0704 | 0.42914 | 0.11728 | 0.39563 | 0.81415 |  | 0
db3 | -2.2281 | 41.2036 | 0.60303 | 1.0185 | -0.045215 | 0.15477 | -1.4868 | -0.24771 | 0.72056 | 0.13404 | -0.2162 | -1.8198 | 0.88508 | 0.9752 | -1.0341 | -1.084 | -0.54403 | 0.81134 | 0.25498 | 0.92035 | -1.2885 |  | 0
db4 | 0.044637 | -1.328 | 0.020749 | 0.18196 | -0.085783 | -0.046505 | 0.085088 | 0.12975 | 0.0075142 | 0.0040554 | -0.005179 | 0.027027 | 0.023689 | -0.10947 | -0.042213 | -0.057511 | -0.14184 | 0.15703 | -0.074304 | -0.074056 | -0.3762 |  | 0
db5 | -0.15347 | -5.5266 | -0.38035 | -0.59022 | -0.12793 | -0.21717 | 0.12711 | 0.12648 | 0.031552 | -0.022086 | 0.18551 | 0.61294 | 0.32408 | 0.16026 | 0.17182 | 0.11667 | -0.16988 | -0.049555 | -0.096571 | -0.20267 | -3.4429 |  | 0
db6 | 0.009268 | -0.32661 | 0.10903 | 0.041459 | -0.02583 | -0.074547 | -0.03345 | -0.0082323 | 0.057038 | -0.016335 | 0.083006 | -0.061026 | -0.039382 | -0.022915 | -0.027603 | -0.023402 | 0.10672 | -0.02712 | -0.022322 | -0.015082 | 0.02995 |  | 0
db7 | 1.1584 | 2.9288 | -0.1181 | -0.082435 | 0.018276 | 0.035271 | 0.012634 | 0.014139 | -0.00079852 | 0.0049007 | -0.010147 | 0.058018 | -0.032833 | 0.037188 | -0.036832 | -0.0076611 | 0.054895 | 0.014313 | 0.0071992 | 0.031977 | 1.038 |  | 1.2336e-17
db8 | 0.0031387 | -0.13483 | -0.008262 | 0.0022331 | -0.0060503 | -0.029569 | -0.021643 | 0.013204 | -0.020708 | -0.023542 | 0.013022 | 0.013634 | 0.013141 | -0.0031473 | -0.0010313 | 0.01574 | 0.013158 | 0.018122 | 0.0009464 | 0.010753 | 0.012308 |  | 0
db9 | 0.47964 | 0.37629 | -0.0041088 | -0.012713 | 0.01395 | 0.026507 | -0.02254 | -0.0087796 | -0.00033779 | 0.0078125 | -0.0088295 | -0.016407 | 0.0073885 | 0.010862 | -0.0010731 | -0.0054519 | -0.0029065 | 0.0036868 | 0.0052349 | 0.0077055 | 0.4341 |  | -5.2427e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.73711 | 0.59255 | -0.01252 | -0.010827 | -0.00096813 | 0.001117 | 0.011775 | 0.0038335 | -0.0033833 | 0.002861 | 0.0078293 | 0.019826 | 0.0035084 | -0.00048866 | 0.0026102 | 0.0028214 | 0.0025931 | -0.014849 | -0.0038038 | -0.011935 | 0.63041 |  | -2.4672e-17
db12 | -0.00038579 | 0.0059429 | -0.00055301 | 0.0011135 | 0.0010632 | -0.00068969 | -0.002623 | 0.00064856 | 0.00086451 | 0.0007426 | -0.0024439 | -0.0029002 | 0.0011494 | 0.001274 | -0.0023101 | -0.0015499 | 0.00094106 | 0.00066676 | 0.0017661 | 0.0028401 | 0.003584 |  | 0
db13 | 0.06895 | 0.097076 | 0.0051979 | 0.0041896 | 6.0881e-05 | 0.0014618 | 0.00045952 | 0.0012955 | 0.00023973 | -0.00035935 | 0.00063993 | -0.0046789 | -0.0032193 | 2.4113e-05 | -0.0047772 | -0.0028542 | 5.2853e-05 | -0.0010457 | 0.0020101 | 0.0013027 | 0.055886 |  | -6.9389e-18
db14 | 0.0065656 | 0.0056121 | 0.00181 | -0.00080426 | -0.00026119 | 0.0014665 | 0.00047273 | 0.00066682 | 0.0002577 | 0.00024415 | -0.0010675 | -0.0015667 | -0.00079479 | -0.000599 | -0.0007358 | -0.00037122 | 0.00098093 | 9.5878e-05 | 0.00022994 | -2.4165e-05 | -0.0037154 |  | -3.3731e-19
db15 | 0.01882 | -0.012345 | 0.002127 | 0.00069401 | -0.002301 | -0.0013299 | -0.00026473 | -0.0025997 | -0.00096669 | 0.00086373 | 0.0019766 | -0.002886 | -8.6044e-05 | 0.0028355 | 0.0018162 | 0.00098296 | -0.00017967 | 0.00045998 | 0.00035488 | -0.001497 | -0.0019915 |  | -5.7824e-19
da2 | -0.33298 | 22.4031 | 0.034438 | -0.13192 | 0.16184 | 0.091597 | -0.40753 | 0.28721 | -0.4328 | -0.24861 | 0.43907 | -0.53105 | 0.087003 | -0.20431 | 1.4849 | -0.94221 | -0.26042 | -0.23506 | 0.5897 | 0.2182 | 12.4948 |  | 0
da3 | 0.51857 | 0.052682 | 0.055453 | -0.1183 | -0.56855 | -0.043893 | 0.15157 | 0.35737 | 0.062313 | -0.24916 | 0.24455 | 0.61049 | -0.48884 | -0.33733 | -0.60655 | -0.46625 | 0.39379 | 0.4421 | 0.20425 | 0.35699 | 1.3786 |  | 0
da4 | -0.19553 | 3.5549 | 0.23134 | 0.03907 | 0.092223 | -0.056068 | -0.058918 | -0.37321 | -0.22235 | -0.037042 | -0.36419 | 0.17942 | -0.25896 | -0.19795 | 0.088926 | 0.42869 | -0.065109 | 0.48902 | -0.047813 | 0.13292 | 0.67152 |  | 0
da5 | 0.081105 | 2.5273 | -0.19096 | -0.13064 | -0.22184 | 0.011573 | -0.050111 | 0.12958 | 0.018674 | -0.089701 | -0.016992 | 0.16983 | -0.072742 | -0.13829 | -0.00036361 | 0.034995 | 0.033043 | 0.090204 | 0.18487 | 0.23887 | -0.1017 |  | 0
da6 | 0.087929 | 0.45741 | -0.052674 | -0.074659 | -0.041802 | 0.042479 | -0.010427 | -0.024605 | -0.0022768 | 0.092077 | 0.045766 | -0.071293 | -0.032363 | -0.036145 | 0.041619 | 0.00093397 | 0.023837 | 0.034341 | 0.010129 | 0.05506 | 0.097559 |  | 0
da7 | 0.023198 | 1.6744 | -0.05404 | -0.092593 | -0.10035 | -0.018687 | 0.099653 | 0.02734 | 0.069914 | -0.0057724 | -0.033136 | 0.043998 | -0.10523 | 0.030287 | 0.0086946 | -0.046172 | 0.0068169 | 0.08096 | 0.038148 | 0.050171 | 0.0013984 |  | 0
da8 | -0.022623 | 0.044996 | 0.025846 | -0.0046623 | -0.01304 | 0.007015 | -0.022189 | 0.023698 | 0.0010279 | 0.027675 | 0.038538 | -0.033257 | -0.0099211 | -0.013566 | 0.017951 | -0.039966 | -0.018999 | 0.0012773 | -0.0035668 | 0.016141 | -0.012055 |  | 0
da9 | 0.004905 | -0.25785 | 0.0074544 | -0.028272 | -0.020955 | -0.0038856 | 0.017626 | 0.017273 | 0.011311 | -0.026571 | -0.015047 | 0.023893 | -0.02563 | -0.014605 | 0.00016569 | 0.00094403 | -0.0013701 | 0.0069531 | 0.023741 | 0.026973 | -0.032085 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.0081471 | 0.19866 | -0.019469 | -0.010266 | -0.012568 | 0.0044729 | 0.011328 | 0.0097192 | 0.0073817 | 0.005015 | -0.0056321 | 0.0046992 | -0.0096914 | 0.0051196 | 0.0080643 | -0.0055209 | -0.0036508 | 0.003978 | 0.0052657 | 0.0017539 | -0.011565 |  | -1.3492e-18
da12 | 0.00022586 | -0.001931 | -0.0038791 | -0.002404 | -0.00072081 | -0.0015975 | -0.0020652 | 0.0010631 | 0.0006967 | 0.00056212 | 0.0041534 | -0.0044091 | -0.0019053 | -0.0010833 | 0.0034121 | -0.0019418 | 0.0029782 | 0.0010017 | 0.0015081 | 0.0046309 | 0.0018251 |  | 0
da13 | 0.0024202 | -0.017134 | 0.0009511 | -0.0026709 | -0.0043227 | -0.0012477 | 0.0027725 | 0.0030547 | 0.0013799 | -0.0035922 | 0.00055345 | 0.0007412 | -0.0030235 | -0.00030601 | -0.0008333 | 5.272e-05 | -0.00080071 | 0.0014694 | 0.0021494 | 0.0036727 | 0.00070705 |  | 0
da14 | -0.005125 | 0.0078366 | 0.0017344 | 0.001211 | 0.00085036 | 0.0013958 | -0.0015707 | 0.00059696 | -0.0012616 | -0.00078866 | 0.00048546 | -0.00041793 | -0.0028536 | -0.0016789 | 0.00063635 | 8.7221e-05 | 2.9038e-05 | 0.00020281 | 1.307e-05 | 0.0013289 | 7.841e-05 |  | 0
da15 | -0.0028585 | -0.0036293 | 0.00048571 | 0.001566 | -0.0014905 | 0.00093292 | -0.00033202 | -1.4678e-05 | 5.5861e-07 | -0.0008045 | -0.00095882 | -0.0003825 | -0.0012315 | -0.00017444 | -0.0017533 | -0.0013134 | 0.0023674 | 0.0016786 | 0.00030296 | 0.0011215 | -0.0046277 |  | -2.8912e-19
Field Colinearity | 75.4657
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.3121 | 0.015013 | -0.13958 | -0.46455 | -0.43304 | -0.24341 | -0.019084 | 0.25405 | 0.041848 | 0.20556 | -0.012781 | -0.10559 | 0.31344 | 0.52334 | 0.27771 | -0.081856 | -0.43061 | -0.60149 | -0.70215 | -2.3733
Twist (m2 mrad) | 39.8408 | 0.21394 | -1.8844 | -5.923 | -5.1965 | -2.7384 | -0.20038 | 2.477 | 0.37663 | 1.6959 | -0.095857 | -0.71274 | 1.8807 | 2.7475 | 1.2497 | -0.30696 | -1.2918 | -1.3534 | -1.0532 | -0.88999
Twist Integral (I0) (m2 rad) | 0.021627
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.359 | 0.17261 | -0.051582 | -0.31633 | -0.31865 | -0.091435 | 0.20643 | 0.36994 | 0.13769 | -0.066656 | -0.03443 | -0.10605 | 0.10638 | 0.42982 | 0.048158 | -0.31687 | -0.43106 | -0.76363 | -0.99844 | -1.6199
Twist (m2 mrad) | 40.1924 | 2.4596 | -0.69636 | -4.0332 | -3.8238 | -1.0286 | 2.1675 | 3.6069 | 1.2392 | -0.54991 | -0.25822 | -0.71586 | 0.63827 | 2.2565 | 0.21671 | -1.1883 | -1.2932 | -1.7182 | -1.4977 | -0.60745
Twist Integral (I0) (m2 rad) | 0.026525
Note on B154: 4.1s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-1.3565 | -3.072 | -0.070297 | -0.22086 | 1.1821 | -0.89 | 0.17855 | 0.74954 | -0.36018 | 16.4437 | 0.59538 | 1.0794 | -2.2278 | -0.044637 | -0.15395 | 1.1584 | -1.1249 | -0.33298 | 0.51857 | -0.19553 | 17.3714 | 0.59546 | -1.3543 | -1.5374 | -0.11165 | -0.46833 | 1.2292 | -0.0013998 | 0.50038 | 0.76352 | -0.27488 | 14420.69 | 0.98021 | -0.91102 | -0.093525 | -0.42106 | 1.2069 | 0.00089696 | 0.7615 | 0.53305 | -0.056003 | 14421.245
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8