## Alarm sheet
File name | HCMB__A001-03000293_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000293
Date of test Ap 1 | 2004-11-15 |  |  |  | Date of test Ap 2 | 2004-11-15
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | yellow alarm | yellow alarm | status ok | status ok | Main field | yellow alarm | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
604 | HCMB__A001 | 3000293 |  | 3 |  |  | 3
| COMMENTS: | Low values on dB/B in both apertures.
| HOLDING POINT | OK | Data at CERN |  | 2004-11-16 | Analysis by |  | C. Vollinger | Answer to MMS-LD |  |  | 2004-11-16 | Answer to manufacturer |  |  | 2004-11-16
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000293_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000293 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 15-NOV-2004_3293_09-17cc1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 15-NOV-2004_3293_09-17cc2.txt |  |  |  | a even | 1
Test Operator |  | Mehler |  |  | Test Controller |  |  | C. Vollinger
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2004-11-15 |  |  | Time of test |  |  | 09:17 - 11:47 |  |  |  | Date of test |  | 2004-11-15 |  |  | Time of test |  |  | 09:18 - 11:47
Mole name |  | Dipole 7 | [9 empty] | Mole name |  | Dipole 1 | [5 empty] | 2004-03-26
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.63 |  |  | I. Gain |  |  | 50/500
Dx average |  | -5.0245e-05 |  |  | Dy average |  |  | 0.00052724 |  |  |  | Dx average |  | 0.00037151 |  |  | Dy average |  |  | 0.00043498
Magnetic length |  | 14.4192 |  |  | Tranfs. Function (mT/KA) |  |  | 595.2753 |  |  |  | Magnetic length |  | 14.4206 |  |  | Tranfs. Function (mT/KA) |  |  | 595.3163
Phase |  | 0.2755 |  |  | Integral field (T.m) |  |  | 85.7037 |  |  |  | Phase |  | 0.6379 |  |  | Integral field (T.m) |  |  | 85.6463
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.4654 | 5.9555 | 5.9523 | 5.953 | 5.953 | 5.9494 | 5.9521 | 5.9523 | 5.9525 | 5.9524 | 5.9524 | 5.9524 | 5.9532 | 5.9527 | 5.9525 | 5.9522 | 5.9534 | 5.9527 | 5.9557 | 3.6719 | 85.7037 | C1 (T) | 3.3656 | 5.9554 | 5.9525 | 5.9528 | 5.9515 | 5.9549 | 5.952 | 5.9526 | 5.9524 | 5.9521 | 5.9524 | 5.9529 | 5.9535 | 5.953 | 5.9536 | 5.953 | 5.9535 | 5.9528 | 5.9558 | 3.6877 | 85.6463
Angle (mrad) | 3.7003 | -0.31671 | -0.026917 | -0.075704 | 0.11753 | -0.14452 | 0.2947 | 0.40999 | -0.059969 | -0.38412 | -0.73617 | -0.096276 | 0.15679 | 0.27395 | 0.18006 | -0.059685 | -0.31665 | -0.34207 | -0.38127 | -1.0485 | 0.2755 | Angle (mrad) | 5.2253 | -0.71956 | -0.32887 | -0.40041 | -0.15854 | 0.25406 | 0.088068 | 0.36784 | -0.072298 | -0.24275 | -0.62192 | -0.10637 | 0.057442 | 0.27281 | 0.1242 | -0.21434 | -0.3387 | -0.43361 | -0.42528 | -0.089652 | 0.6379
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 3.0218 | -0.83009 | -0.72174 | -0.068625 | -0.12367 | -0.80256 | -0.70557 | -0.24975 | -0.39225 | -1.5723 | -1.2529 | -0.82461 | -1.0065 | -0.63208 | -0.14483 | -0.21486 | -1.1739 | -0.54663 | -0.42899 | -1.5599 | -0.56747 | b2 | -8.677 | 0.66207 | 0.8824 | 0.31045 | 0.82111 | 0.96669 | 0.71561 | 0.23672 | 0.34297 | 0.62012 | 0.28443 | 1.1849 | 1.1672 | 0.41335 | 0.3106 | 0.27754 | 0.63867 | 0.78735 | 0.23065 | 1.2472 | 0.35028
b3 | 45.8473 | -2.4364 | -2.1188 | -3.1636 | -3.0223 | -4.2266 | -2.7843 | -2.8219 | -2.4728 | -2.9942 | -4.0495 | -2.7575 | -2.0729 | -4.5313 | -4.1736 | -2.3966 | -1.9361 | -1.9285 | -1.8007 | 1.6305 | -1.2496 | b3 | 43.5448 | -2.7266 | -1.9309 | -2.9356 | -2.4588 | -4.0323 | -2.6502 | -2.1957 | -2.1334 | -3.5981 | -4.3346 | -2.4736 | -1.5912 | -4.4394 | -4.0401 | -2.274 | -1.0624 | -1.4072 | -1.1532 | 2.1841 | -1.119
b4 | -0.071552 | -0.058521 | -0.028808 | -0.025253 | 0.088084 | -0.12714 | -0.20389 | 0.092422 | 0.14323 | -0.33416 | -0.069979 | -0.20215 | -0.085558 | -0.050548 | -0.018714 | 0.044356 | -0.10718 | -0.079787 | -0.12163 | 0.23919 | -0.054135 | b4 | -3.0357 | 0.18348 | 0.11432 | 0.088 | 0.17085 | -0.024832 | 0.057261 | 0.0051337 | -0.059338 | -0.074383 | -0.0093602 | 0.095459 | 0.41476 | 0.0025421 | -0.021888 | -0.084936 | 0.25949 | 0.17645 | 0.051998 | 0.054361 | -0.017586
b5 | -4.3513 | -0.91195 | -0.96856 | -0.50645 | -0.3325 | -0.36248 | -0.3824 | -0.38827 | -0.56145 | -0.39714 | -0.11663 | -0.29628 | -0.1374 | -0.0438 | -0.19023 | -0.55175 | -0.30906 | -0.51619 | -0.66823 | -4.1088 | -0.66194 | b5 | -5.1921 | -0.74028 | -0.85264 | -0.41682 | -0.44269 | -0.094207 | -0.20273 | -0.38174 | -0.37976 | -0.22183 | 0.050589 | -0.15531 | -0.045381 | -0.12399 | -0.083923 | -0.39484 | -0.53714 | -0.50655 | -0.68464 | -4.1843 | -0.61201
b6 | -0.31861 | -0.099368 | -0.051688 | -0.084738 | -0.098816 | -0.015401 | -0.069484 | 0.010584 | -0.029232 | 0.098661 | -0.015747 | 6.5867e-05 | 0.047828 | -0.031543 | -0.12737 | -0.088441 | 0.11643 | 0.027312 | 0.05539 | 0.12991 | -0.024005 | b6 | -0.51814 | -0.031745 | -0.065929 | -0.056119 | -0.03561 | 0.014269 | -0.13293 | 0.085254 | 0.10517 | 0.020259 | 0.032018 | -0.17069 | 0.028337 | -0.10886 | -0.014532 | 0.12007 | 0.0090605 | -0.063627 | 0.012898 | 0.12773 | -0.024314
b7 | 3.0261 | 0.98096 | 1.0611 | 1.178 | 1.1797 | 1.1443 | 1.1619 | 1.1336 | 1.1316 | 1.1189 | 1.1732 | 1.1319 | 1.1815 | 1.1843 | 1.1088 | 1.1934 | 1.1422 | 1.1387 | 1.1502 | 1.008 | 1.1916 | b7 | 2.9766 | 0.99041 | 1.0617 | 1.1758 | 1.1764 | 1.1577 | 1.1301 | 1.1412 | 1.1559 | 1.0848 | 1.1331 | 1.1738 | 1.1918 | 1.0668 | 1.0993 | 1.1965 | 1.1372 | 1.1625 | 1.1656 | 1.0102 | 1.1837
b8 | -0.16925 | -0.065808 | -0.033141 | 0.0083243 | -0.0015473 | -0.00028786 | -0.049826 | 0.021691 | -0.00041538 | 0.0068573 | 0.020023 | -0.017202 | 0.016459 | -0.010789 | -0.011221 | -0.022918 | -0.0050147 | 0.026774 | 0.041504 | 0.0071811 | -0.0088876 | b8 | -0.15434 | -0.017551 | -0.027504 | 0.0026045 | 0.063263 | 0.030867 | 0.0036728 | 0.02082 | 0.033248 | 0.094154 | 0.045588 | 0.039603 | 0.019362 | -0.021842 | -0.058322 | -0.021792 | -0.014622 | -0.01885 | -0.014749 | 0.029836 | 0.004646
b9 | 0.39197 | 0.47009 | 0.45321 | 0.45957 | 0.46127 | 0.43392 | 0.46453 | 0.45888 | 0.46725 | 0.46784 | 0.43855 | 0.43626 | 0.46473 | 0.45384 | 0.4302 | 0.46035 | 0.46673 | 0.46671 | 0.47615 | 0.41486 | 0.45389 | b9 | 0.39567 | 0.46823 | 0.44885 | 0.45914 | 0.47862 | 0.43189 | 0.4633 | 0.45748 | 0.46916 | 0.45773 | 0.452 | 0.45379 | 0.47086 | 0.44442 | 0.43961 | 0.46849 | 0.47121 | 0.47568 | 0.48535 | 0.42303 | 0.45774
b10 | -0.041485 | -0.12212 | -0.057879 | -0.024879 | 0.00094859 | 0.031819 | -0.033242 | 0.020684 | -0.0029847 | 0.064738 | 0.064323 | 0.013097 | 0.035946 | 0.025935 | -0.088546 | -0.043495 | 0.041626 | 0.027995 | 0.056709 | 0.021898 | 0 | b10 | 0.035779 | -0.050643 | -0.063123 | 0.017142 | 0.072689 | 0.034247 | -0.023323 | 0.040951 | 0.056562 | 0.10651 | 0.045383 | -0.018372 | -0.030628 | -0.022271 | -0.07036 | -0.0078405 | -0.024915 | -0.049319 | -0.042528 | 0.015647 | 0
b11 | 0.59387 | 0.7192 | 0.71775 | 0.73026 | 0.73337 | 0.73786 | 0.73017 | 0.72947 | 0.72907 | 0.73262 | 0.73951 | 0.72802 | 0.73174 | 0.73858 | 0.74004 | 0.72617 | 0.72878 | 0.72603 | 0.71948 | 0.62507 | 0.7224 | b11 | 0.58566 | 0.71951 | 0.7154 | 0.72439 | 0.72595 | 0.736 | 0.72504 | 0.72049 | 0.72513 | 0.73377 | 0.73761 | 0.73104 | 0.7266 | 0.73624 | 0.74112 | 0.72187 | 0.72288 | 0.72188 | 0.71334 | 0.60839 | 0.7186
b12 | -0.0154 | -0.016459 | -0.0080845 | -0.0019568 | -0.0017368 | 0.0037686 | -0.0021582 | 0.00010856 | 0.00015383 | 0.0053906 | 0.0026315 | -0.0023909 | -0.00090236 | 0.0028107 | -0.0094913 | -0.0054047 | 0.0025482 | 0.0043286 | 0.0075534 | 0.0054198 | -0.0012978 | b12 | 0.0043515 | -0.0063733 | -0.0070259 | 0.0041543 | 0.0093802 | 0.0057288 | -0.0029823 | 0.0058179 | 0.0066506 | 0.012701 | 0.006249 | 0.00043656 | -0.0015166 | -0.0025414 | -0.0042327 | 0.0018989 | -0.0028299 | -0.0039512 | -0.0045741 | 0.0035292 | 0.0011273
b13 | 0.090558 | 0.072523 | 0.07187 | 0.067472 | 0.068544 | 0.067806 | 0.067347 | 0.067229 | 0.066643 | 0.066381 | 0.068345 | 0.06729 | 0.068479 | 0.065303 | 0.066181 | 0.070078 | 0.067868 | 0.069334 | 0.071346 | 0.056864 | 0.068641 | b13 | 0.09179 | 0.069177 | 0.071417 | 0.065677 | 0.066458 | 0.064736 | 0.065539 | 0.066423 | 0.06785 | 0.064183 | 0.062568 | 0.068625 | 0.067657 | 0.06447 | 0.066923 | 0.069377 | 0.069285 | 0.071024 | 0.071161 | 0.055371 | 0.067696
b14 | 0.0011534 | -0.0061406 | -0.0016883 | 0.0007201 | 0.0020673 | 0.0045026 | 0.0013338 | 0.0036235 | 0.0016841 | 0.0051253 | 0.0033024 | 0.0011564 | 0.0031409 | 0.0037634 | -0.0023067 | -0.00076555 | 0.0039572 | 0.003001 | 0.0040349 | 0.0041317 | 0.0017568 | b14 | 0.0036636 | -0.010472 | -0.009786 | -0.0067639 | -0.0041265 | -0.0055053 | -0.0087124 | -0.0065026 | -0.0049983 | -0.0027946 | -0.0059178 | -0.0079887 | -0.0088754 | -0.0081211 | -0.0094717 | -0.0065532 | -0.0081667 | -0.009272 | -0.0083461 | 0.0014195 | -0.0067463
b15 | -0.0067506 | 0.026979 | 0.026216 | 0.025579 | 0.026793 | 0.026681 | 0.025816 | 0.025858 | 0.027558 | 0.02647 | 0.024244 | 0.026328 | 0.025895 | 0.029526 | 0.026499 | 0.028158 | 0.026676 | 0.027747 | 0.024797 | 0.013208 | 0.025107 | b15 | -0.010975 | 0.021163 | 0.016926 | 0.021478 | 0.023382 | 0.024189 | 0.020045 | 0.021341 | 0.025254 | 0.026591 | 0.026934 | 0.021942 | 0.02043 | 0.022062 | 0.018509 | 0.020671 | 0.020827 | 0.019401 | 0.017827 | 0.00059117 | 0.019971
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 37.0033 | -3.1671 | -0.26917 | -0.75704 | 1.1753 | -1.4452 | 2.947 | 4.0999 | -0.59969 | -3.8412 | -7.3617 | -0.96276 | 1.5679 | 2.7395 | 1.8006 | -0.59685 | -3.1665 | -3.4207 | -3.8127 | -10.4849 | 0 | a1 | 52.2536 | -7.1956 | -3.2888 | -4.0041 | -1.5854 | 2.5406 | 0.88068 | 3.6784 | -0.72298 | -2.4275 | -6.2192 | -1.0637 | 0.57442 | 2.728 | 1.242 | -2.1434 | -3.387 | -4.3361 | -4.2528 | -0.89652 | 0
a2 | 13.6111 | -0.36086 | 0.33048 | -0.19634 | 0.81471 | -0.080608 | -0.70802 | 0.72269 | 1.2081 | -0.75461 | -0.49341 | -0.1554 | -0.52563 | 0.56495 | -0.70699 | 1.3209 | 0.64581 | 0.31164 | -0.32373 | 3.3982 | 0.6059 | a2 | 16.0894 | 0.39325 | 0.36331 | -0.33044 | 1.0122 | 0.39814 | -0.21481 | 0.84522 | 1.1153 | 0.50004 | -0.56983 | -0.085856 | 0.39798 | 0.83462 | 0.16192 | 0.67327 | 1.4355 | 0.32203 | 1.2287 | -4.9654 | 0.75584
a3 | -0.65632 | 0.36825 | 0.49567 | 0.75462 | 0.67117 | 0.59391 | 0.4908 | 0.46735 | 0.39023 | 0.67914 | 0.31853 | 0.20449 | 0.25602 | 0.50555 | 0.35316 | 0.50333 | 0.75088 | 0.41597 | 0.32424 | -0.63033 | 0.40483 | a3 | -0.42569 | 0.34028 | 0.1324 | 0.47146 | 0.59717 | 0.15071 | -0.025971 | 0.54049 | 0.1667 | 0.60005 | 0.55054 | 0.10339 | -0.35441 | 0.3641 | -0.096055 | 0.21684 | 0.50058 | 0.55571 | -0.030977 | 0.037177 | 0.23795
a4 | 3.2584 | -0.39801 | -0.17216 | -0.42592 | -0.63672 | -0.56341 | -0.27505 | -0.66689 | -0.44868 | -0.52513 | -0.33275 | -0.50874 | -0.29148 | -0.20465 | -0.080986 | -0.4342 | -0.22178 | -0.35582 | -0.38212 | 0.13463 | -0.25754 | a4 | 2.9149 | -0.19236 | -0.095007 | 0.12951 | -0.89646 | 0.071096 | 0.10236 | -0.23001 | -0.11907 | 0.0059358 | -0.27632 | 0.34939 | -0.4749 | 0.21412 | -0.22876 | -0.080817 | -0.29868 | -0.16549 | -0.031299 | -1.0784 | -0.064464
a5 | 3.0225 | -0.0034934 | 0.025439 | 0.05958 | 0.46939 | 0.068539 | -0.024757 | 0.051003 | 0.098638 | 0.25192 | 0.32635 | -0.033644 | -0.0020123 | 0.35397 | 0.27666 | 0.15759 | 0.066436 | 0.025817 | 0.13327 | -0.31247 | 0.20144 | a5 | 2.2879 | -0.058598 | -0.074714 | -0.008729 | 0.19359 | 0.0031974 | -0.045386 | 0.006054 | 0.06942 | 0.089736 | 0.13116 | -0.10574 | -0.39426 | 0.016136 | -0.05659 | 0.016961 | -0.11394 | -0.058075 | -0.1261 | -0.46999 | 0.025342
a6 | 0.55142 | 0.044007 | 0.031236 | 0.0081485 | -0.057398 | -0.096628 | -0.072814 | -0.010188 | 0.0023615 | -0.077903 | 0.028102 | -0.034453 | -0.096816 | 0.0091906 | -0.063031 | -0.17947 | -0.035324 | -0.047115 | 0.017975 | 0.026506 | -0.015244 | a6 | 0.085151 | 0.032424 | 0.05405 | 0.19474 | 0.1299 | 0.091749 | 0.097504 | 0.20696 | 0.11205 | 0.14625 | 0.15152 | 0.16982 | -0.030086 | 0.17472 | 0.023259 | 0.1185 | 0.077499 | 0.041399 | -0.0038487 | 0.16482 | 0.10105
a7 | 1.6826 | -0.088612 | -0.042212 | -0.09534 | -0.06995 | -0.007391 | 0.0092993 | -0.069113 | -0.086391 | -0.030123 | -0.032322 | 0.0029798 | -0.064892 | -0.047828 | 0.022448 | -0.020414 | 0.028701 | -0.0049121 | -0.038467 | -0.020148 | 0.017318 | a7 | 1.6195 | -0.063426 | -0.03839 | -0.094967 | 0.015867 | 0.0048997 | -0.010919 | 0.026625 | -0.0060459 | 0.031398 | -0.027858 | 0.053299 | -0.10965 | 0.01284 | 0.0082746 | -0.11498 | -0.00024474 | -0.029354 | -0.061512 | -0.004986 | 0.026492
a8 | 0.048617 | -0.034214 | -0.0031167 | -0.028662 | -0.027983 | -0.042758 | -0.061768 | -0.0066526 | -0.0069524 | -0.021183 | -0.032561 | -0.020497 | -0.050854 | -0.044218 | -0.049128 | -0.029714 | -0.027605 | -0.029783 | -0.0099763 | -0.010765 | -0.026353 | a8 | -0.058542 | -0.0018605 | 0.015869 | -0.036618 | -0.012696 | -0.033456 | -0.031662 | 0.033363 | -0.010337 | -0.024743 | -0.072238 | 0.00061304 | -0.065928 | 0.013346 | -0.0035018 | -0.029437 | 0.011093 | -0.031286 | -0.0076691 | 0.085196 | -0.013941
a9 | -0.20905 | -0.0069179 | -0.019692 | 0.0046727 | 0.016683 | 0.0063625 | 0.0064618 | 0.0071429 | -0.017207 | -0.014029 | -0.0096286 | -0.0093793 | -0.0031253 | -0.0040112 | -0.033557 | -0.017957 | 0.0034107 | -0.016746 | -0.0095156 | -0.039322 | -0.013698 | a9 | -0.24429 | 0.017982 | -0.020813 | 0.0057446 | 0.0001055 | -0.0012026 | -0.0078277 | 0.013622 | -0.017785 | 0.006525 | -0.0038491 | -0.013695 | -0.055882 | 0.0051158 | -0.016913 | -0.020954 | -0.029135 | -0.011619 | -0.02335 | -0.026345 | -0.017115
a10 | -0.010889 | -0.00047655 | 0.026293 | 0.028819 | -0.0085891 | -0.012087 | 0.0099317 | -0.00068572 | -0.0078234 | -0.017502 | 0.027405 | 0.0191 | 0.014685 | -0.024465 | 0.0049557 | -0.033311 | 0.0012551 | -0.012465 | 0.02062 | -0.047564 | 0 | a10 | -0.064516 | -0.0125 | 0.01217 | -0.000325 | -0.013289 | -0.01382 | 0.011437 | 0.035061 | -0.018209 | 0.0011001 | -0.037272 | 0.038289 | -0.011155 | 0.012408 | 0.010809 | -0.025795 | -0.020832 | -0.0031429 | -0.0057808 | 0.12483 | 0
a11 | 0.17922 | -0.0555 | -0.045257 | -0.046952 | -0.040622 | -0.044208 | -0.040045 | -0.036817 | -0.045221 | -0.051623 | -0.048747 | -0.040461 | -0.039358 | -0.037203 | -0.03535 | -0.043915 | -0.036498 | -0.038488 | -0.04252 | -0.055696 | -0.036399 | a11 | 0.20051 | -0.02853 | -0.017314 | -0.017754 | -0.0097139 | -0.0090772 | -0.0097417 | -0.0041523 | -0.0079675 | -0.017323 | -0.018494 | -0.0111 | -0.022448 | -0.0076941 | -0.0088015 | -0.01884 | -0.017058 | -0.019158 | -0.018219 | -0.014927 | -0.0083028
a12 | -0.0013196 | -0.0016292 | 0.0039893 | 0.0031919 | -0.0010333 | -0.00071881 | -0.0018353 | 0.0014403 | -0.0010534 | -0.0067633 | 0.0004473 | -0.0017881 | -0.0056603 | -0.003822 | -0.0069738 | -0.0066343 | -0.0012975 | -0.0033179 | 0.00086482 | -0.0026345 | -0.0018223 | a12 | -0.019217 | 0.0013323 | 0.0011087 | -0.00052964 | 0.0019172 | -0.00079323 | -0.00085576 | 0.0059721 | -0.0020405 | -0.0012981 | -0.004999 | 0.0018615 | -0.0088141 | 0.0025964 | 0.0016885 | -0.0034027 | -0.0028139 | -0.0035832 | -0.0049963 | 0.013393 | -0.0010539
a13 | -0.015889 | -0.00067132 | -0.0037264 | -0.0034336 | -0.0013318 | -0.0028957 | -0.0013116 | -0.0044861 | -0.0053662 | -0.0041188 | -0.0062772 | -0.0049984 | -0.0030907 | -0.00053833 | -0.0017514 | -0.0041645 | -0.002415 | -0.0035617 | -0.0040542 | -0.0072192 | -0.0037447 | a13 | -0.01717 | 0.0026596 | 0.00019918 | 0.00017348 | 0.00068453 | 5.8716e-05 | 4.3182e-07 | 0.0040234 | -2.9207e-05 | 0.0021119 | 0.0011928 | 0.00086757 | -0.0026365 | 0.0055299 | 0.00099695 | -0.0013324 | -0.0013095 | -1.1783e-05 | -0.0021812 | -0.0025895 | -1.6363e-05
a14 | 0.0054393 | -0.011407 | -0.0099904 | -0.010367 | -0.011695 | -0.011812 | -0.012344 | -0.009277 | -0.010553 | -0.011427 | -0.009983 | -0.011055 | -0.010544 | -0.013937 | -0.011742 | -0.011375 | -0.010128 | -0.010908 | -0.0087525 | -0.0023392 | -0.010187 | a14 | 0.0058151 | -0.0073328 | -0.0064958 | -0.0099543 | -0.0096868 | -0.0093735 | -0.0079442 | -0.0070278 | -0.010071 | -0.010997 | -0.013708 | -0.0080544 | -0.010722 | -0.0090907 | -0.0075312 | -0.010283 | -0.0077569 | -0.0074529 | -0.0062603 | -0.0026145 | -0.0082395
a15 | 0.0056186 | -0.0091252 | -0.0036069 | -0.00075981 | 0.0018578 | 0.0040362 | -0.00076377 | 0.002671 | -4.0557e-05 | 0.0051722 | 0.0032786 | 0.0011799 | 0.0030186 | 0.0043672 | -0.0035151 | -0.001784 | 0.0038541 | 0.0028372 | 0.0037375 | 0.0036079 | 0.0011411 | a15 | -0.004208 | -0.012065 | -0.01041 | -0.0070446 | -0.0031189 | -0.0063454 | -0.0093064 | -0.0031941 | -0.0052434 | -0.0017278 | -0.0069788 | -0.0068106 | -0.010514 | -0.0088958 | -0.010886 | -0.0091546 | -0.010571 | -0.0098197 | -0.011187 | -0.0026653 | -0.0076784
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 0.00011756 | 0.00028727 | 0.00014048 | 6.1967e-05 | 0 | -7.4715e-05 | 7.8431e-05 | -4.8161e-05 | 0 | -0.00015234 | -0.0001431 | -2.8022e-05 | -8.1453e-05 | -6.2365e-05 | 0.00020359 | 9.6718e-05 | -9.6716e-05 | -6.691e-05 | -0.00013069 | -7.0494e-05 | -5.0245e-05 | Dx (m) | -3.5469e-05 | 0.00011833 | 0.00015065 | -4.0226e-05 | -0.0001707 | -7.9493e-05 | 5.5039e-05 | -9.6115e-05 | -0.00013308 | -0.0002468 | -0.00010666 | 4.4067e-05 | 7.077e-05 | 5.1717e-05 | 0.00016168 | 1.6869e-05 | 5.7401e-05 | 0.00011587 | 0.00010094 | -3.4857e-05 | 0.00037151
Dy (m) | 0 | 2.3412e-05 | -5.3479e-05 | -6.3108e-05 | 1.9728e-05 | 2.3377e-05 | -1.8817e-05 | 0 | 1.8602e-05 | 2.9951e-05 | -7.2431e-05 | -4.6156e-05 | -3.8477e-05 | 5.3176e-05 | 0 | 8.3865e-05 | 0 | 2.5649e-05 | -5.648e-05 | 0.00012312 | 0.00052724 | Dy (m) | 0.00019914 | 3.4231e-05 | -2.5296e-05 | 0 | 2.8836e-05 | 3.0951e-05 | -2.6077e-05 | -8.3316e-05 | 4.1259e-05 | 0 | 8.3271e-05 | -8.8365e-05 | 2.8302e-05 | -2.8113e-05 | -2.2895e-05 | 6.1182e-05 | 5.035e-05 | 1.0499e-05 | 1.638e-05 | -0.00034914 | 0.00043498
Magnetic length (m) |  | 14.4192 | [19 empty] | Magnetic length (m) |  | 14.4206
Transfer funct. (mT/KA) |  | 595.2753 | [19 empty] | Transfer funct. (mT/KA) |  | 595.3163
Coil positioning (mm) |  | 0.015681 | [19 empty] | Coil positioning (mm) |  | 0.018592
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.1.0 | [20 empty] | Soft version | DIMM V2.1.0
Level Average | 0.32227 | 0.27448 | 0.43123 | 0.2871 | 0.27278 | 0.43911 | 0.39631 | 0.36026 | -0.15136 | 0.15327 | 0.17849 | -0.087153 | -0.073997 | -0.061512 | 0.024004 | 0.24615 | 0.19152 | -0.18924 | 0.37561 | 0.011747 |  | Level Average | -0.093541 | -0.15031 | -0.335 | -0.20685 | -0.34703 | -0.29568 | -0.45827 | -0.35208 | -0.46103 | -0.52577 | -0.5687 | -0.45384 | -0.52161 | -0.5783 | -0.53221 | -0.54331 | -0.45863 | -0.42818 | -0.38282 | -0.181
Level Sigma2 | 0.011179 | 0.051629 | 0.012812 | 0.058448 | 0.004229 | 0.002549 | 0.002805 | 0.002756 | 0.12909 | 0.005238 | 0.000575 | 0.21468 | 0.084631 | 0.29235 | 0.001572 | 0.027969 | 0.024653 | 0.020029 | 0.0032 | 0.001171 |  | Level Sigma2 | 0.042987 | 0.078044 | 0.008373 | 0.027824 | 0.035614 | 0.043207 | 0.084501 | 0.11809 | 0.093074 | 0.23633 | 0.174 | 0.18026 | 0.24568 | 0.18746 | 0.18307 | 0.46658 | 0.2756 | 0.30627 | 0.1991 | 0.000335
Position (mm) | 0 | 747 | 1501 | 2251 | 3000 | 3751 | 4505 | 5260 | 6012 | 6760 | 7511 | 8262 | 9011 | 9763 | 10515 | 11265 | 12016 | 12766 | 13518 | 14270 |  | Position (mm) | 0 | 759 | 1510 | 2262 | 3013 | 3765 | 4514 | 5266 | 6018 | 6767 | 7519 | 8270 | 9023 | 9773 | 10525 | 11277 | 12028 | 12779 | 13531 | 14282
Fist move (mm) | First move Ap1 | 942.2941 | First move Ap2 | 942.2306 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMB__A001-3000293 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000293 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 346.5394 | 595.5457 | 595.2306 | 595.3038 | 595.3003 | 594.9405 | 595.2083 | 595.2259 | 595.2457 | 595.237 | 595.2389 | 595.2374 | 595.321 | 595.2666 | 595.2462 | 595.2198 | 595.3407 | 595.2738 | 595.5738 | 367.1894 | 8570.3659 | C1/i (mT/KA) | 336.5642 | 595.542 | 595.247 | 595.2801 | 595.1544 | 595.4858 | 595.1995 | 595.2589 | 595.2446 | 595.2128 | 595.2449 | 595.2923 | 595.3488 | 595.3037 | 595.3579 | 595.3019 | 595.3523 | 595.2833 | 595.5838 | 368.7718 | 8564.632
Angle (mrad) | 3.7003 | -0.31671 | -0.026917 | -0.075704 | 0.11753 | -0.14452 | 0.2947 | 0.40999 | -0.059969 | -0.38412 | -0.73617 | -0.096276 | 0.15679 | 0.27395 | 0.18006 | -0.059685 | -0.31665 | -0.34207 | -0.38127 | -1.0485 | 0.2755 | Angle (mrad) | 5.2253 | -0.71956 | -0.32887 | -0.40041 | -0.15854 | 0.25406 | 0.088068 | 0.36784 | -0.072298 | -0.24275 | -0.62192 | -0.10637 | 0.057442 | 0.27281 | 0.1242 | -0.21434 | -0.3387 | -0.43361 | -0.42528 | -0.089652 | 0.6379
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 3.0218 | -0.83009 | -0.72174 | -0.068625 | -0.12367 | -0.80256 | -0.70557 | -0.24975 | -0.39225 | -1.5723 | -1.2529 | -0.82461 | -1.0065 | -0.63208 | -0.14483 | -0.21486 | -1.1739 | -0.54663 | -0.42899 | -1.5599 |  | b2 | -8.677 | 0.66207 | 0.8824 | 0.31045 | 0.82111 | 0.96669 | 0.71561 | 0.23672 | 0.34297 | 0.62012 | 0.28443 | 1.1849 | 1.1672 | 0.41335 | 0.3106 | 0.27754 | 0.63867 | 0.78735 | 0.23065 | 1.2472
b3 | 45.8473 | -2.4364 | -2.1188 | -3.1636 | -3.0223 | -4.2266 | -2.7843 | -2.8219 | -2.4728 | -2.9942 | -4.0495 | -2.7575 | -2.0729 | -4.5313 | -4.1736 | -2.3966 | -1.9361 | -1.9285 | -1.8007 | 1.6305 |  | b3 | 43.5448 | -2.7266 | -1.9309 | -2.9356 | -2.4588 | -4.0323 | -2.6502 | -2.1957 | -2.1334 | -3.5981 | -4.3346 | -2.4736 | -1.5912 | -4.4394 | -4.0401 | -2.274 | -1.0624 | -1.4072 | -1.1532 | 2.1841
b4 | -0.071552 | -0.058521 | -0.028808 | -0.025253 | 0.088084 | -0.12714 | -0.20389 | 0.092422 | 0.14323 | -0.33416 | -0.069979 | -0.20215 | -0.085558 | -0.050548 | -0.018714 | 0.044356 | -0.10718 | -0.079787 | -0.12163 | 0.23919 |  | b4 | -3.0357 | 0.18348 | 0.11432 | 0.088 | 0.17085 | -0.024832 | 0.057261 | 0.0051337 | -0.059338 | -0.074383 | -0.0093602 | 0.095459 | 0.41476 | 0.0025421 | -0.021888 | -0.084936 | 0.25949 | 0.17645 | 0.051998 | 0.054361
b5 | -4.3513 | -0.91195 | -0.96856 | -0.50645 | -0.3325 | -0.36248 | -0.3824 | -0.38827 | -0.56145 | -0.39714 | -0.11663 | -0.29628 | -0.1374 | -0.0438 | -0.19023 | -0.55175 | -0.30906 | -0.51619 | -0.66823 | -4.1088 |  | b5 | -5.1921 | -0.74028 | -0.85264 | -0.41682 | -0.44269 | -0.094207 | -0.20273 | -0.38174 | -0.37976 | -0.22183 | 0.050589 | -0.15531 | -0.045381 | -0.12399 | -0.083923 | -0.39484 | -0.53714 | -0.50655 | -0.68464 | -4.1843
b6 | -0.31861 | -0.099368 | -0.051688 | -0.084738 | -0.098816 | -0.015401 | -0.069484 | 0.010584 | -0.029232 | 0.098661 | -0.015747 | 6.5867e-05 | 0.047828 | -0.031543 | -0.12737 | -0.088441 | 0.11643 | 0.027312 | 0.05539 | 0.12991 |  | b6 | -0.51814 | -0.031745 | -0.065929 | -0.056119 | -0.03561 | 0.014269 | -0.13293 | 0.085254 | 0.10517 | 0.020259 | 0.032018 | -0.17069 | 0.028337 | -0.10886 | -0.014532 | 0.12007 | 0.0090605 | -0.063627 | 0.012898 | 0.12773
b7 | 3.0261 | 0.98096 | 1.0611 | 1.178 | 1.1797 | 1.1443 | 1.1619 | 1.1336 | 1.1316 | 1.1189 | 1.1732 | 1.1319 | 1.1815 | 1.1843 | 1.1088 | 1.1934 | 1.1422 | 1.1387 | 1.1502 | 1.008 |  | b7 | 2.9766 | 0.99041 | 1.0617 | 1.1758 | 1.1764 | 1.1577 | 1.1301 | 1.1412 | 1.1559 | 1.0848 | 1.1331 | 1.1738 | 1.1918 | 1.0668 | 1.0993 | 1.1965 | 1.1372 | 1.1625 | 1.1656 | 1.0102
b8 | -0.16925 | -0.065808 | -0.033141 | 0.0083243 | -0.0015473 | -0.00028786 | -0.049826 | 0.021691 | -0.00041538 | 0.0068573 | 0.020023 | -0.017202 | 0.016459 | -0.010789 | -0.011221 | -0.022918 | -0.0050147 | 0.026774 | 0.041504 | 0.0071811 |  | b8 | -0.15434 | -0.017551 | -0.027504 | 0.0026045 | 0.063263 | 0.030867 | 0.0036728 | 0.02082 | 0.033248 | 0.094154 | 0.045588 | 0.039603 | 0.019362 | -0.021842 | -0.058322 | -0.021792 | -0.014622 | -0.01885 | -0.014749 | 0.029836
b9 | 0.39197 | 0.47009 | 0.45321 | 0.45957 | 0.46127 | 0.43392 | 0.46453 | 0.45888 | 0.46725 | 0.46784 | 0.43855 | 0.43626 | 0.46473 | 0.45384 | 0.4302 | 0.46035 | 0.46673 | 0.46671 | 0.47615 | 0.41486 |  | b9 | 0.39567 | 0.46823 | 0.44885 | 0.45914 | 0.47862 | 0.43189 | 0.4633 | 0.45748 | 0.46916 | 0.45773 | 0.452 | 0.45379 | 0.47086 | 0.44442 | 0.43961 | 0.46849 | 0.47121 | 0.47568 | 0.48535 | 0.42303
b10 | -0.041485 | -0.12212 | -0.057879 | -0.024879 | 0.00094859 | 0.031819 | -0.033242 | 0.020684 | -0.0029847 | 0.064738 | 0.064323 | 0.013097 | 0.035946 | 0.025935 | -0.088546 | -0.043495 | 0.041626 | 0.027995 | 0.056709 | 0.021898 |  | b10 | 0.035779 | -0.050643 | -0.063123 | 0.017142 | 0.072689 | 0.034247 | -0.023323 | 0.040951 | 0.056562 | 0.10651 | 0.045383 | -0.018372 | -0.030628 | -0.022271 | -0.07036 | -0.0078405 | -0.024915 | -0.049319 | -0.042528 | 0.015647
b11 | 0.59387 | 0.7192 | 0.71775 | 0.73026 | 0.73337 | 0.73786 | 0.73017 | 0.72947 | 0.72907 | 0.73262 | 0.73951 | 0.72802 | 0.73174 | 0.73858 | 0.74004 | 0.72617 | 0.72878 | 0.72603 | 0.71948 | 0.62507 |  | b11 | 0.58566 | 0.71951 | 0.7154 | 0.72439 | 0.72595 | 0.736 | 0.72504 | 0.72049 | 0.72513 | 0.73377 | 0.73761 | 0.73104 | 0.7266 | 0.73624 | 0.74112 | 0.72187 | 0.72288 | 0.72188 | 0.71334 | 0.60839
b12 | -0.0154 | -0.016459 | -0.0080845 | -0.0019568 | -0.0017368 | 0.0037686 | -0.0021582 | 0.00010856 | 0.00015383 | 0.0053906 | 0.0026315 | -0.0023909 | -0.00090236 | 0.0028107 | -0.0094913 | -0.0054047 | 0.0025482 | 0.0043286 | 0.0075534 | 0.0054198 |  | b12 | 0.0043515 | -0.0063733 | -0.0070259 | 0.0041543 | 0.0093802 | 0.0057288 | -0.0029823 | 0.0058179 | 0.0066506 | 0.012701 | 0.006249 | 0.00043656 | -0.0015166 | -0.0025414 | -0.0042327 | 0.0018989 | -0.0028299 | -0.0039512 | -0.0045741 | 0.0035292
b13 | 0.090558 | 0.072523 | 0.07187 | 0.067472 | 0.068544 | 0.067806 | 0.067347 | 0.067229 | 0.066643 | 0.066381 | 0.068345 | 0.06729 | 0.068479 | 0.065303 | 0.066181 | 0.070078 | 0.067868 | 0.069334 | 0.071346 | 0.056864 |  | b13 | 0.09179 | 0.069177 | 0.071417 | 0.065677 | 0.066458 | 0.064736 | 0.065539 | 0.066423 | 0.06785 | 0.064183 | 0.062568 | 0.068625 | 0.067657 | 0.06447 | 0.066923 | 0.069377 | 0.069285 | 0.071024 | 0.071161 | 0.055371
b14 | 0.0011534 | -0.0061406 | -0.0016883 | 0.0007201 | 0.0020673 | 0.0045026 | 0.0013338 | 0.0036235 | 0.0016841 | 0.0051253 | 0.0033024 | 0.0011564 | 0.0031409 | 0.0037634 | -0.0023067 | -0.00076555 | 0.0039572 | 0.003001 | 0.0040349 | 0.0041317 |  | b14 | 0.0036636 | -0.010472 | -0.009786 | -0.0067639 | -0.0041265 | -0.0055053 | -0.0087124 | -0.0065026 | -0.0049983 | -0.0027946 | -0.0059178 | -0.0079887 | -0.0088754 | -0.0081211 | -0.0094717 | -0.0065532 | -0.0081667 | -0.009272 | -0.0083461 | 0.0014195
b15 | -0.0067506 | 0.026979 | 0.026216 | 0.025579 | 0.026793 | 0.026681 | 0.025816 | 0.025858 | 0.027558 | 0.02647 | 0.024244 | 0.026328 | 0.025895 | 0.029526 | 0.026499 | 0.028158 | 0.026676 | 0.027747 | 0.024797 | 0.013208 |  | b15 | -0.010975 | 0.021163 | 0.016926 | 0.021478 | 0.023382 | 0.024189 | 0.020045 | 0.021341 | 0.025254 | 0.026591 | 0.026934 | 0.021942 | 0.02043 | 0.022062 | 0.018509 | 0.020671 | 0.020827 | 0.019401 | 0.017827 | 0.00059117
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 37.0033 | -3.1671 | -0.26917 | -0.75704 | 1.1753 | -1.4452 | 2.947 | 4.0999 | -0.59969 | -3.8412 | -7.3617 | -0.96276 | 1.5679 | 2.7395 | 1.8006 | -0.59685 | -3.1665 | -3.4207 | -3.8127 | -10.4849 |  | a1 | 52.2536 | -7.1956 | -3.2888 | -4.0041 | -1.5854 | 2.5406 | 0.88068 | 3.6784 | -0.72298 | -2.4275 | -6.2192 | -1.0637 | 0.57442 | 2.728 | 1.242 | -2.1434 | -3.387 | -4.3361 | -4.2528 | -0.89652
a2 | 13.6111 | -0.36086 | 0.33048 | -0.19634 | 0.81471 | -0.080608 | -0.70802 | 0.72269 | 1.2081 | -0.75461 | -0.49341 | -0.1554 | -0.52563 | 0.56495 | -0.70699 | 1.3209 | 0.64581 | 0.31164 | -0.32373 | 3.3982 |  | a2 | 16.0894 | 0.39325 | 0.36331 | -0.33044 | 1.0122 | 0.39814 | -0.21481 | 0.84522 | 1.1153 | 0.50004 | -0.56983 | -0.085856 | 0.39798 | 0.83462 | 0.16192 | 0.67327 | 1.4355 | 0.32203 | 1.2287 | -4.9654
a3 | -0.65632 | 0.36825 | 0.49567 | 0.75462 | 0.67117 | 0.59391 | 0.4908 | 0.46735 | 0.39023 | 0.67914 | 0.31853 | 0.20449 | 0.25602 | 0.50555 | 0.35316 | 0.50333 | 0.75088 | 0.41597 | 0.32424 | -0.63033 |  | a3 | -0.42569 | 0.34028 | 0.1324 | 0.47146 | 0.59717 | 0.15071 | -0.025971 | 0.54049 | 0.1667 | 0.60005 | 0.55054 | 0.10339 | -0.35441 | 0.3641 | -0.096055 | 0.21684 | 0.50058 | 0.55571 | -0.030977 | 0.037177
a4 | 3.2584 | -0.39801 | -0.17216 | -0.42592 | -0.63672 | -0.56341 | -0.27505 | -0.66689 | -0.44868 | -0.52513 | -0.33275 | -0.50874 | -0.29148 | -0.20465 | -0.080986 | -0.4342 | -0.22178 | -0.35582 | -0.38212 | 0.13463 |  | a4 | 2.9149 | -0.19236 | -0.095007 | 0.12951 | -0.89646 | 0.071096 | 0.10236 | -0.23001 | -0.11907 | 0.0059358 | -0.27632 | 0.34939 | -0.4749 | 0.21412 | -0.22876 | -0.080817 | -0.29868 | -0.16549 | -0.031299 | -1.0784
a5 | 3.0225 | -0.0034934 | 0.025439 | 0.05958 | 0.46939 | 0.068539 | -0.024757 | 0.051003 | 0.098638 | 0.25192 | 0.32635 | -0.033644 | -0.0020123 | 0.35397 | 0.27666 | 0.15759 | 0.066436 | 0.025817 | 0.13327 | -0.31247 |  | a5 | 2.2879 | -0.058598 | -0.074714 | -0.008729 | 0.19359 | 0.0031974 | -0.045386 | 0.006054 | 0.06942 | 0.089736 | 0.13116 | -0.10574 | -0.39426 | 0.016136 | -0.05659 | 0.016961 | -0.11394 | -0.058075 | -0.1261 | -0.46999
a6 | 0.55142 | 0.044007 | 0.031236 | 0.0081485 | -0.057398 | -0.096628 | -0.072814 | -0.010188 | 0.0023615 | -0.077903 | 0.028102 | -0.034453 | -0.096816 | 0.0091906 | -0.063031 | -0.17947 | -0.035324 | -0.047115 | 0.017975 | 0.026506 |  | a6 | 0.085151 | 0.032424 | 0.05405 | 0.19474 | 0.1299 | 0.091749 | 0.097504 | 0.20696 | 0.11205 | 0.14625 | 0.15152 | 0.16982 | -0.030086 | 0.17472 | 0.023259 | 0.1185 | 0.077499 | 0.041399 | -0.0038487 | 0.16482
a7 | 1.6826 | -0.088612 | -0.042212 | -0.09534 | -0.06995 | -0.007391 | 0.0092993 | -0.069113 | -0.086391 | -0.030123 | -0.032322 | 0.0029798 | -0.064892 | -0.047828 | 0.022448 | -0.020414 | 0.028701 | -0.0049121 | -0.038467 | -0.020148 |  | a7 | 1.6195 | -0.063426 | -0.03839 | -0.094967 | 0.015867 | 0.0048997 | -0.010919 | 0.026625 | -0.0060459 | 0.031398 | -0.027858 | 0.053299 | -0.10965 | 0.01284 | 0.0082746 | -0.11498 | -0.00024474 | -0.029354 | -0.061512 | -0.004986
a8 | 0.048617 | -0.034214 | -0.0031167 | -0.028662 | -0.027983 | -0.042758 | -0.061768 | -0.0066526 | -0.0069524 | -0.021183 | -0.032561 | -0.020497 | -0.050854 | -0.044218 | -0.049128 | -0.029714 | -0.027605 | -0.029783 | -0.0099763 | -0.010765 |  | a8 | -0.058542 | -0.0018605 | 0.015869 | -0.036618 | -0.012696 | -0.033456 | -0.031662 | 0.033363 | -0.010337 | -0.024743 | -0.072238 | 0.00061304 | -0.065928 | 0.013346 | -0.0035018 | -0.029437 | 0.011093 | -0.031286 | -0.0076691 | 0.085196
a9 | -0.20905 | -0.0069179 | -0.019692 | 0.0046727 | 0.016683 | 0.0063625 | 0.0064618 | 0.0071429 | -0.017207 | -0.014029 | -0.0096286 | -0.0093793 | -0.0031253 | -0.0040112 | -0.033557 | -0.017957 | 0.0034107 | -0.016746 | -0.0095156 | -0.039322 |  | a9 | -0.24429 | 0.017982 | -0.020813 | 0.0057446 | 0.0001055 | -0.0012026 | -0.0078277 | 0.013622 | -0.017785 | 0.006525 | -0.0038491 | -0.013695 | -0.055882 | 0.0051158 | -0.016913 | -0.020954 | -0.029135 | -0.011619 | -0.02335 | -0.026345
a10 | -0.010889 | -0.00047655 | 0.026293 | 0.028819 | -0.0085891 | -0.012087 | 0.0099317 | -0.00068572 | -0.0078234 | -0.017502 | 0.027405 | 0.0191 | 0.014685 | -0.024465 | 0.0049557 | -0.033311 | 0.0012551 | -0.012465 | 0.02062 | -0.047564 |  | a10 | -0.064516 | -0.0125 | 0.01217 | -0.000325 | -0.013289 | -0.01382 | 0.011437 | 0.035061 | -0.018209 | 0.0011001 | -0.037272 | 0.038289 | -0.011155 | 0.012408 | 0.010809 | -0.025795 | -0.020832 | -0.0031429 | -0.0057808 | 0.12483
a11 | 0.17922 | -0.0555 | -0.045257 | -0.046952 | -0.040622 | -0.044208 | -0.040045 | -0.036817 | -0.045221 | -0.051623 | -0.048747 | -0.040461 | -0.039358 | -0.037203 | -0.03535 | -0.043915 | -0.036498 | -0.038488 | -0.04252 | -0.055696 |  | a11 | 0.20051 | -0.02853 | -0.017314 | -0.017754 | -0.0097139 | -0.0090772 | -0.0097417 | -0.0041523 | -0.0079675 | -0.017323 | -0.018494 | -0.0111 | -0.022448 | -0.0076941 | -0.0088015 | -0.01884 | -0.017058 | -0.019158 | -0.018219 | -0.014927
a12 | -0.0013196 | -0.0016292 | 0.0039893 | 0.0031919 | -0.0010333 | -0.00071881 | -0.0018353 | 0.0014403 | -0.0010534 | -0.0067633 | 0.0004473 | -0.0017881 | -0.0056603 | -0.003822 | -0.0069738 | -0.0066343 | -0.0012975 | -0.0033179 | 0.00086482 | -0.0026345 |  | a12 | -0.019217 | 0.0013323 | 0.0011087 | -0.00052964 | 0.0019172 | -0.00079323 | -0.00085576 | 0.0059721 | -0.0020405 | -0.0012981 | -0.004999 | 0.0018615 | -0.0088141 | 0.0025964 | 0.0016885 | -0.0034027 | -0.0028139 | -0.0035832 | -0.0049963 | 0.013393
a13 | -0.015889 | -0.00067132 | -0.0037264 | -0.0034336 | -0.0013318 | -0.0028957 | -0.0013116 | -0.0044861 | -0.0053662 | -0.0041188 | -0.0062772 | -0.0049984 | -0.0030907 | -0.00053833 | -0.0017514 | -0.0041645 | -0.002415 | -0.0035617 | -0.0040542 | -0.0072192 |  | a13 | -0.01717 | 0.0026596 | 0.00019918 | 0.00017348 | 0.00068453 | 5.8716e-05 | 4.3182e-07 | 0.0040234 | -2.9207e-05 | 0.0021119 | 0.0011928 | 0.00086757 | -0.0026365 | 0.0055299 | 0.00099695 | -0.0013324 | -0.0013095 | -1.1783e-05 | -0.0021812 | -0.0025895
a14 | 0.0054393 | -0.011407 | -0.0099904 | -0.010367 | -0.011695 | -0.011812 | -0.012344 | -0.009277 | -0.010553 | -0.011427 | -0.009983 | -0.011055 | -0.010544 | -0.013937 | -0.011742 | -0.011375 | -0.010128 | -0.010908 | -0.0087525 | -0.0023392 |  | a14 | 0.0058151 | -0.0073328 | -0.0064958 | -0.0099543 | -0.0096868 | -0.0093735 | -0.0079442 | -0.0070278 | -0.010071 | -0.010997 | -0.013708 | -0.0080544 | -0.010722 | -0.0090907 | -0.0075312 | -0.010283 | -0.0077569 | -0.0074529 | -0.0062603 | -0.0026145
a15 | 0.0056186 | -0.0091252 | -0.0036069 | -0.00075981 | 0.0018578 | 0.0040362 | -0.00076377 | 0.002671 | -4.0557e-05 | 0.0051722 | 0.0032786 | 0.0011799 | 0.0030186 | 0.0043672 | -0.0035151 | -0.001784 | 0.0038541 | 0.0028372 | 0.0037375 | 0.0036079 |  | a15 | -0.004208 | -0.012065 | -0.01041 | -0.0070446 | -0.0031189 | -0.0063454 | -0.0093064 | -0.0031941 | -0.0052434 | -0.0017278 | -0.0069788 | -0.0068106 | -0.010514 | -0.0088958 | -0.010886 | -0.0091546 | -0.010571 | -0.0098197 | -0.011187 | -0.0026653
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.11746 | -0.28686 | -0.14045 | -0.061974 | 0.0032916 | 0.07471 | -0.078426 | 0.048162 | -0.0058057 | 0.15233 | 0.14309 | 0.028017 | 0.08144 | 0.062373 | -0.20348 | -0.096754 | 0.096711 | 0.066908 | 0.13066 | 0.070523 |  | Dx (mm) | 0.03557 | -0.1183 | -0.15061 | 0.040223 | 0.17061 | 0.079483 | -0.055035 | 0.096143 | 0.13306 | 0.24656 | 0.10668 | -0.044066 | -0.070785 | -0.051718 | -0.16167 | -0.016867 | -0.057405 | -0.11586 | -0.10093 | 0.035142
Dy (mm) | 0.0042774 | -0.023263 | 0.053419 | 0.063104 | -0.019728 | -0.023373 | 0.018822 | 0.00083272 | -0.018602 | -0.029878 | 0.072432 | 0.046158 | 0.038496 | -0.053169 | 0.001664 | -0.083835 | 0.007771 | -0.025639 | 0.056442 | -0.12308 |  | Dy (mm) | -0.19945 | -0.034225 | 0.025275 | 0.0002231 | -0.028838 | -0.030941 | 0.026076 | 0.08328 | -0.041227 | 0.0083697 | -0.083227 | 0.088371 | -0.028286 | 0.028109 | 0.022874 | -0.061188 | -0.050345 | -0.010476 | -0.016354 | 0.34968
Mag. len. (m) | 14.4192 | [20 empty] | Mag. len. (m) | 14.4206
TF (mT/KA) | 595.2753 | [20 empty] | TF (mT/KA) | 595.3163
Coil pos. (mm) | 0.015681 | [20 empty] | Coil pos. (mm) | 0.018592
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.014391 | [20 empty] | Twist (m2 rad) | 0.011768

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 595.2753 | 0.13288 | 0.0023326 |  | -0.83695 | 2.9395 |  |  | 595.3163 | 0.1157 | -0.004456 |  | -1.6102 | 3.1013 |  | 0 | 4
2 | -0.54792 |  | -0.64903 | 0.40375 | 0.60658 |  | 0.090393 | 0.66075 | 0.34746 |  | 0.60561 | 0.31903 | 0.77994 |  | 0.46774 | 0.55378 |  | 0 | 5
3 | -1.2491 |  | -2.8709 | 0.85763 | 0.40813 |  | 0.47584 | 0.16408 | -1.1215 |  | -2.6333 | 1.0716 | 0.23537 |  | 0.26679 | 0.28331 |  | 0 | 6
4 | -0.058566 |  | -0.067205 | 0.11319 | -0.25884 |  | -0.38503 | 0.16378 | -0.022026 |  | 0.072146 | 0.12735 | -0.06351 |  | -0.12587 | 0.28158 |  | 0 | 7
5 | -0.66386 |  | -0.42607 | 0.25114 | 0.20223 |  | 0.12788 | 0.14656 | -0.61233 |  | -0.3461 | 0.25556 | 0.023208 |  | -0.029447 | 0.12314 |  | 0 | 8
6 | -0.022428 |  | -0.021381 | 0.042665 | -0.013738 |  | -0.037026 | 0.049915 | -0.027795 |  | -0.012545 | 0.074662 | 0.1042 |  | 0.10106 | 0.066372 |  | 0 | 9
7 | 1.1904 |  | 1.1376 | 0.052474 | 0.01722 |  | -0.035468 | 0.038773 | 1.1832 |  | 1.1322 | 0.054678 | 0.026223 |  | -0.022182 | 0.049583 |  | 0 | 10
8 | -0.0088245 |  | -0.0044789 | 0.017777 | -0.026967 |  | -0.030436 | 0.017701 | 0.0056773 |  | 0.0097502 | 0.019596 | -0.014072 |  | -0.015007 | 0.023546 |  | 0 | 11
9 | 0.45122 |  | 0.45438 | 0.014219 | -0.013754 |  | -0.0066524 | 0.012707 | 0.45627 |  | 0.45833 | 0.014565 | -0.01677 |  | -0.0092715 | 0.017613 |  | 0 | 12
10 | 0 |  | 0 | 0 | 9.3221e-20 |  | 9.9385e-20 | 4.2166e-19 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.72202 |  | 0.72952 | 0.0069871 | -0.03639 |  | -0.042718 | 0.0054921 | 0.71845 |  | 0.72623 | 0.0077861 | -0.0083103 |  | -0.014635 | 0.0063387 |  | 0 | 14
12 | -0.00011489 |  | -0.0010788 | 0.0019508 | -0.00018562 |  | -0.001977 | 0.0025635 | 0.00013859 |  | 0.0012848 | 0.001175 | -8.786e-05 |  | -0.000745 | 0.0026525 |  | 0 | 15
13 | 0.0068399 |  | 0.068074 | 0.0018526 | -0.00037399 |  | -0.003261 | 0.0015478 | 0.0067528 |  | 0.06724 | 0.0022428 | 3.8658e-06 |  | 0.00066781 | 0.0024547 |  | 0 | 16
14 | 0.00017132 |  | 0.0017078 | 0.00071057 | -0.0010512 |  | -0.011365 | 0.0010896 | -0.00068292 |  | -0.0074475 | 0.00062789 | -0.0008439 |  | -0.0090313 | 0.0012121 |  | 0 | 17
15 | 0.0025107 |  | 0.026546 | 0.0012155 | 0.00011411 |  | 0.00091196 | 0.0036689 | 0.0019971 |  | 0.02161 | 0.0027961 | -0.00076784 |  | -0.0079596 | 0.0030759 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.66067 | -0.86674
3 | [12 empty] | 3 | 9 | -0.98049 | -0.59605
4 | [12 empty] | 4 | 16 | -1.994 | -1.6641
5 | [12 empty] | 5 | 25 | -1.651 | -1.7294
6 | [12 empty] | 6 | 36 | -3.0759 | -2.6536
7 | [12 empty] | 7 | 49 | -3.0987 | -2.9552
8 | [12 empty] | 8 | 64 | -4.032 | -3.8406
9 | [12 empty] | 9 | 81 | -4.3094 | -4.1341
10 | [12 empty] | sigma (mm) |  | 0.015681 | 0.018592
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 346.5394 | 595.5457 | 595.2306 | 595.3038 | 595.3003 | 594.9405 | 595.2083 | 595.2259 | 595.2457 | 595.237 | 595.2389 | 595.2374 | 595.321 | 595.2666 | 595.2462 | 595.2198 | 595.3407 | 595.2738 | 595.5738 | 367.1894
2 | 3.6551 | -0.91332 | -0.75364 | -0.086089 | -0.12405 | -0.76704 | -0.73017 | -0.23371 | -0.39479 | -1.521 | -1.1821 | -0.81441 | -0.98544 | -0.60199 | -0.24468 | -0.24711 | -1.1511 | -0.5327 | -0.39916 | -1.5643 | -0.54792
3 | 45.8482 | -2.4377 | -2.1212 | -3.1686 | -3.0201 | -4.2226 | -2.7881 | -2.8227 | -2.4712 | -2.9824 | -4.052 | -2.7607 | -2.0736 | -4.5288 | -4.1743 | -2.3895 | -1.9345 | -1.9259 | -1.8017 | 1.6246 | -1.2491
4 | -0.18877 | -0.12006 | -0.060497 | -0.031754 | 0.086163 | -0.12115 | -0.21106 | 0.096832 | 0.14203 | -0.3217 | -0.060491 | -0.20056 | -0.082943 | -0.054333 | -0.027714 | 0.028686 | -0.10003 | -0.071817 | -0.099317 | 0.31642 | -0.058566
5 | -4.3616 | -0.92063 | -0.9702 | -0.50785 | -0.33207 | -0.36148 | -0.3844 | -0.38842 | -0.56151 | -0.40087 | -0.11536 | -0.29675 | -0.13964 | -0.043365 | -0.19788 | -0.54984 | -0.31246 | -0.51637 | -0.67006 | -4.1125 | -0.66386
6 | -0.19062 | 0.00067738 | 0.00011821 | -0.061096 | -0.0997 | -0.045512 | -0.037263 | -0.0087058 | -0.026346 | 0.038825 | -0.075822 | -0.011078 | 0.012986 | -0.056716 | -0.047721 | -0.047083 | 0.077522 | 0.00046678 | 0.0015845 | 0.10569 | -0.022428
7 | 3.018 | 0.97351 | 1.0591 | 1.1775 | 1.18 | 1.1447 | 1.1598 | 1.1332 | 1.1317 | 1.1187 | 1.171 | 1.1317 | 1.1801 | 1.1855 | 1.1079 | 1.1935 | 1.1424 | 1.1382 | 1.1477 | 1.0083 | 1.1904
8 | -0.14801 | -0.0022723 | -0.0036808 | 0.021866 | -0.0024167 | -0.015613 | -0.032625 | 0.011294 | 0.0010118 | -0.026482 | -0.009836 | -0.023157 | -0.0014082 | -0.02401 | 0.029947 | -0.0012491 | -0.026244 | 0.012281 | 0.011975 | -0.0043095 | -0.0088245
9 | 0.38936 | 0.45155 | 0.44965 | 0.45971 | 0.46136 | 0.43281 | 0.46325 | 0.45835 | 0.46732 | 0.4629 | 0.43473 | 0.43654 | 0.46348 | 0.45368 | 0.42066 | 0.4596 | 0.4646 | 0.46589 | 0.47284 | 0.41714 | 0.45122
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.5927 | 0.71617 | 0.71716 | 0.73031 | 0.73339 | 0.73769 | 0.73004 | 0.72947 | 0.72909 | 0.73222 | 0.73929 | 0.72801 | 0.73164 | 0.7386 | 0.73879 | 0.72619 | 0.72862 | 0.7259 | 0.71887 | 0.62503 | 0.72202
12 | -0.079395 | -0.017627 | -0.010996 | 0.0084186 | -0.018775 | 0.0024054 | 0.015528 | -0.021796 | 0.004974 | -0.016601 | -0.045926 | -0.038845 | -0.04923 | -0.00044267 | 0.00012671 | -0.0037212 | -0.020982 | 0.011184 | 0.0081166 | 0.032163 | -0.0011489 |  | ten times!
13 | 0.90679 | 0.71378 | 0.7128 | 0.67006 | 0.68715 | 0.6776 | 0.6725 | 0.67089 | 0.66801 | 0.66045 | 0.6743 | 0.66875 | 0.67973 | 0.6569 | 0.65808 | 0.70751 | 0.67515 | 0.69395 | 0.70566 | 0.56862 | 0.068399 |  | ten times!
14 | 0.0052019 | 0.0040772 | 0.011854 | 0.019861 | 0.019645 | 0.027834 | 0.029893 | 0.025997 | 0.018165 | 0.016774 | 0.0064098 | 0.005938 | 0.014998 | 0.020555 | 0.021291 | 0.016013 | 0.018572 | 0.014122 | 0.015404 | 0.02999 | 0.0017132 |  | ten times!
15 | -0.067506 | 0.26979 | 0.26216 | 0.25579 | 0.26793 | 0.26681 | 0.25816 | 0.25858 | 0.27558 | 0.2647 | 0.24244 | 0.26328 | 0.25895 | 0.29526 | 0.26499 | 0.28158 | 0.26676 | 0.27747 | 0.24797 | 0.13208 | 0.025107 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 37.0966 | -3.1743 | -0.26417 | -0.7575 | 1.175 | -1.4459 | 2.9445 | 4.0979 | -0.59971 | -3.8372 | -7.3522 | -0.96027 | 1.5727 | 2.7355 | 1.7921 | -0.59039 | -3.1696 | -3.4228 | -3.8088 | -10.5103 | 0.0023326
2 | 13.5789 | -0.3551 | 0.35198 | -0.16735 | 0.80743 | -0.09745 | -0.69733 | 0.72031 | 1.203 | -0.7773 | -0.46426 | -0.1411 | -0.51869 | 0.53289 | -0.69772 | 1.303 | 0.63904 | 0.30255 | -0.31676 | 3.427 | 0.60658
3 | -0.58872 | 0.34787 | 0.49167 | 0.75025 | 0.67185 | 0.60081 | 0.48767 | 0.473 | 0.39024 | 0.69149 | 0.32782 | 0.20865 | 0.26079 | 0.50733 | 0.35026 | 0.49658 | 0.75481 | 0.41981 | 0.33426 | -0.62681 | 0.40813
4 | 3.3463 | -0.40324 | -0.15914 | -0.41754 | -0.63862 | -0.56661 | -0.27382 | -0.66739 | -0.451 | -0.53695 | -0.34175 | -0.5053 | -0.2902 | -0.21039 | -0.067666 | -0.4415 | -0.22273 | -0.35934 | -0.37734 | 0.020823 | -0.25884
5 | 3.042 | -0.00046038 | 0.027541 | 0.061301 | 0.46887 | 0.070556 | -0.026052 | 0.051145 | 0.098482 | 0.25628 | 0.3255 | -0.033361 | -0.00023484 | 0.35331 | 0.27295 | 0.1503 | 0.067175 | 0.02695 | 0.13166 | -0.30832 | 0.20223
6 | 0.61662 | 0.043089 | 0.0091367 | -0.020173 | -0.049103 | -0.086994 | -0.080275 | -0.0093465 | 0.0096141 | -0.064485 | -0.00025641 | -0.052922 | -0.111 | 0.032466 | -0.06207 | -0.14486 | -0.039436 | -0.036695 | -0.0031633 | 0.070792 | -0.013738
7 | 1.6853 | -0.093284 | -0.041664 | -0.096288 | -0.069924 | -0.0060784 | 0.0076908 | -0.068988 | -0.086411 | -0.02871 | -0.031001 | 0.0035432 | -0.063448 | -0.046929 | 0.018339 | -0.022389 | 0.029816 | -0.0038089 | -0.038894 | -0.019471 | 0.01722
8 | 0.036273 | -0.030001 | -0.015811 | -0.042173 | -0.023727 | -0.038209 | -0.065644 | -0.0069943 | -0.0029091 | -0.013599 | -0.04686 | -0.029849 | -0.059153 | -0.032745 | -0.052678 | -0.01237 | -0.029467 | -0.023625 | -0.022038 | 0.014568 | -0.026967
9 | -0.20963 | -0.0084943 | -0.0161 | 0.0064494 | 0.016708 | 0.0072343 | 0.0072054 | 0.0071513 | -0.017261 | -0.011593 | -0.014171 | -0.0099826 | -0.004491 | -0.0024734 | -0.032945 | -0.021594 | 0.0031752 | -0.015925 | -0.012636 | -0.036119 | -0.013754
10 | 0 | 1.7889e-18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9.3221e-20
11 | 0.17916 | -0.05605 | -0.044615 | -0.046744 | -0.040642 | -0.044116 | -0.040112 | -0.036862 | -0.045223 | -0.050852 | -0.048912 | -0.040357 | -0.039037 | -0.036952 | -0.036258 | -0.044623 | -0.03643 | -0.038273 | -0.042869 | -0.055144 | -0.03639
12 | -0.029105 | -0.0057418 | 0.0090976 | 0.0003625 | -0.00075712 | 0.0055259 | -0.028027 | 0.015533 | -0.0020033 | -0.049204 | -0.02413 | -0.038816 | -0.073435 | -0.013474 | -0.073031 | -0.027717 | -0.015049 | -0.018948 | -0.016038 | 0.026651 | -0.0018562 |  | ten times!
13 | -0.15405 | -0.032829 | -0.047305 | -0.039596 | -0.012712 | -0.021404 | -0.020711 | -0.041468 | -0.053891 | -0.026706 | -0.053677 | -0.048023 | -0.025265 | 0.0027942 | -0.035756 | -0.050552 | -0.016895 | -0.029448 | -0.033538 | -0.067042 | -0.0037399 |  | ten times!
14 | 0.060066 | -0.13046 | -0.11561 | -0.11735 | -0.11265 | -0.11547 | -0.12793 | -0.094007 | -0.10131 | -0.11425 | -0.11816 | -0.12083 | -0.11567 | -0.12868 | -0.12367 | -0.095736 | -0.10606 | -0.10479 | -0.10307 | -0.0121 | -0.010512 |  | ten times!
15 | 0.056186 | -0.091252 | -0.036069 | -0.0075981 | 0.018578 | 0.040362 | -0.0076377 | 0.02671 | -0.00040557 | 0.051722 | 0.032786 | 0.011799 | 0.030186 | 0.043672 | -0.035151 | -0.01784 | 0.038541 | 0.028372 | 0.037375 | 0.036079 | 0.0011411 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 336.5642 | 595.542 | 595.247 | 595.2801 | 595.1544 | 595.4858 | 595.1995 | 595.2589 | 595.2446 | 595.2128 | 595.2449 | 595.2923 | 595.3488 | 595.3037 | 595.3579 | 595.3019 | 595.3523 | 595.2833 | 595.5838 | 368.7718
2 | -8.8493 | 0.62275 | 0.84858 | 0.32436 | 0.86843 | 1.0038 | 0.69837 | 0.26685 | 0.37556 | 0.72508 | 0.33344 | 1.1731 | 1.1552 | 0.38755 | 0.2335 | 0.27146 | 0.62853 | 0.76748 | 0.21701 | 1.2397 | 0.34746
3 | 43.4612 | -2.7216 | -1.9283 | -2.9362 | -2.4594 | -4.0323 | -2.6492 | -2.1992 | -2.1311 | -3.5949 | -4.3303 | -2.4674 | -1.5836 | -4.4383 | -4.0417 | -2.2734 | -1.0571 | -1.4033 | -1.1522 | 2.1172 | -1.1215
4 | -3.0996 | 0.16334 | 0.083658 | 0.091944 | 0.1873 | -0.023093 | 0.054357 | 0.013888 | -0.048122 | -0.061337 | -0.013199 | 0.09165 | 0.41663 | 0.0011399 | -0.025385 | -0.086747 | 0.25359 | 0.16278 | 0.036224 | 0.05029 | -0.022026
5 | -5.1916 | -0.74171 | -0.85516 | -0.41614 | -0.44201 | -0.095376 | -0.20414 | -0.37908 | -0.38524 | -0.22294 | 0.045875 | -0.15311 | -0.044541 | -0.1242 | -0.084458 | -0.39638 | -0.53813 | -0.50885 | -0.68424 | -4.1687 | -0.61233
6 | -0.66952 | 0.010374 | -0.009837 | -0.072818 | -0.10661 | -0.018261 | -0.11108 | 0.047314 | 0.05098 | -0.074047 | -0.0098297 | -0.15077 | 0.059206 | -0.089263 | 0.048257 | 0.12968 | 0.032106 | -0.01598 | 0.054774 | 0.11459 | -0.027795
7 | 2.9837 | 0.98958 | 1.0601 | 1.1757 | 1.1721 | 1.1571 | 1.1298 | 1.1415 | 1.1542 | 1.0751 | 1.1336 | 1.1745 | 1.1931 | 1.0665 | 1.0954 | 1.197 | 1.1366 | 1.1617 | 1.165 | 1.022 | 1.1832
8 | -0.13803 | 0.0082256 | 0.004061 | -0.0060857 | 0.024836 | 0.01473 | 0.015576 | 0.0006559 | 0.0042166 | 0.041071 | 0.023047 | 0.048443 | 0.035791 | -0.010958 | -0.02506 | -0.01747 | -0.0012022 | 0.0071438 | 0.0084838 | 0.018505 | 0.0056773
9 | 0.40181 | 0.46529 | 0.44398 | 0.45877 | 0.47226 | 0.43068 | 0.46278 | 0.45694 | 0.46558 | 0.44383 | 0.45108 | 0.45515 | 0.46988 | 0.44399 | 0.43371 | 0.46926 | 0.47101 | 0.47267 | 0.48313 | 0.44585 | 0.45627
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.58804 | 0.71899 | 0.71473 | 0.72428 | 0.72488 | 0.73573 | 0.72492 | 0.72045 | 0.72461 | 0.73174 | 0.73745 | 0.73116 | 0.72669 | 0.7362 | 0.7407 | 0.72202 | 0.72287 | 0.72161 | 0.71309 | 0.61134 | 0.71845
12 | 0.044641 | -0.0066106 | 0.0057025 | 0.022896 | 0.013629 | 0.020955 | -0.0043616 | 0.015465 | 0.0027877 | 0.01543 | 0.014673 | 0.026253 | 0.019166 | -0.00078116 | 0.034204 | 0.027825 | 0.00024135 | 0.018577 | 0.0052099 | 0.015165 | 0.0013859 |  | ten times!
13 | 0.90804 | 0.68421 | 0.70164 | 0.65883 | 0.6721 | 0.65292 | 0.65014 | 0.66453 | 0.68676 | 0.6464 | 0.63923 | 0.67811 | 0.67409 | 0.63953 | 0.6562 | 0.69773 | 0.69225 | 0.70262 | 0.70595 | 0.54634 | 0.067528 |  | ten times!
14 | 0.046762 | -0.0807 | -0.079033 | -0.074767 | -0.073376 | -0.069269 | -0.080038 | -0.084114 | -0.075875 | -0.082058 | -0.078058 | -0.07688 | -0.074395 | -0.073874 | -0.072125 | -0.058047 | -0.067438 | -0.07336 | -0.067137 | 0.0063482 | -0.0068292 |  | ten times!
15 | -0.10975 | 0.21163 | 0.16926 | 0.21478 | 0.23382 | 0.24189 | 0.20045 | 0.21341 | 0.25254 | 0.26591 | 0.26934 | 0.21942 | 0.2043 | 0.22062 | 0.18509 | 0.20671 | 0.20827 | 0.19401 | 0.17827 | 0.0059117 | 0.019971 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 52.1181 | -7.1915 | -3.2868 | -4.0033 | -1.5942 | 2.5405 | 0.87888 | 3.6725 | -0.73088 | -2.4351 | -6.2142 | -1.0701 | 0.57802 | 2.7299 | 1.2431 | -2.1417 | -3.3803 | -4.3334 | -4.2453 | -0.9119 | -0.004456
2 | 17.1129 | 0.387 | 0.3714 | -0.3326 | 0.99192 | 0.38205 | -0.20685 | 0.86062 | 1.1024 | 0.48618 | -0.61918 | -0.059603 | 0.38973 | 0.85152 | 0.17097 | 0.65733 | 1.4325 | 0.32788 | 1.2261 | -5.0554 | 0.77994
3 | -0.55083 | 0.33737 | 0.12936 | 0.47053 | 0.62503 | 0.14957 | -0.025241 | 0.54431 | 0.16907 | 0.5999 | 0.55561 | 0.10462 | -0.35827 | 0.36604 | -0.10249 | 0.21568 | 0.49986 | 0.55265 | -0.031384 | 0.04051 | 0.23537
4 | 2.6521 | -0.19995 | -0.092584 | 0.12962 | -0.90723 | 0.07035 | 0.10301 | -0.22267 | -0.12492 | 0.0011667 | -0.27862 | 0.35152 | -0.48177 | 0.21514 | -0.23046 | -0.086434 | -0.30658 | -0.16832 | -0.036929 | -0.73019 | -0.06351
5 | 2.2566 | -0.05779 | -0.071829 | -0.011029 | 0.18677 | 0.0011824 | -0.042789 | -0.0018866 | 0.06631 | 0.079081 | 0.12719 | -0.099102 | -0.39466 | 0.019694 | -0.055386 | 0.01971 | -0.1125 | -0.05686 | -0.12615 | -0.48483 | 0.023208
6 | 0.27436 | 0.041739 | 0.042538 | 0.196 | 0.14092 | 0.10425 | 0.086892 | 0.17252 | 0.12915 | 0.14032 | 0.18585 | 0.13404 | -0.020927 | 0.16437 | 0.014857 | 0.14366 | 0.097701 | 0.044497 | 0.00068769 | 0.040217 | 0.1042
7 | 1.6077 | -0.063764 | -0.037119 | -0.094361 | 0.01751 | 0.0063879 | -0.011676 | 0.02459 | -0.0049151 | 0.033585 | -0.023122 | 0.051869 | -0.11134 | 0.013377 | 0.0085908 | -0.11574 | -0.00028566 | -0.030928 | -0.06193 | -0.010515 | 0.026223
8 | -0.017315 | 0.0066818 | 0.009055 | -0.036775 | -0.0062088 | -0.027123 | -0.03755 | 0.014818 | -0.00012149 | -0.027303 | -0.054342 | -0.018542 | -0.061522 | 0.0075916 | -0.0095206 | -0.016113 | 0.021469 | -0.029575 | -0.0050429 | 0.01602 | -0.014072
9 | -0.23929 | 0.016282 | -0.018998 | 0.0057495 | 0.0024155 | -6.0092e-05 | -0.0071725 | 0.010032 | -0.015268 | 0.0059094 | 0.00025558 | -0.011942 | -0.056758 | 0.0057869 | -0.015136 | -0.021438 | -0.030432 | -0.012085 | -0.024027 | -0.031564 | -0.01677
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.20152 | -0.028569 | -0.017091 | -0.017741 | -0.0097505 | -0.0089217 | -0.0097218 | -0.0048374 | -0.0076144 | -0.017185 | -0.017813 | -0.011072 | -0.02288 | -0.007561 | -0.0085622 | -0.018802 | -0.017255 | -0.019453 | -0.018593 | -0.01603 | -0.0083103
12 | -0.058634 | 0.032255 | -0.0014431 | -0.005449 | 0.031875 | 0.0061735 | -0.020621 | 0.017944 | -0.00063239 | -0.020448 | -0.014131 | -0.023923 | -0.075949 | 0.015191 | 0.007217 | -0.0042204 | -0.0040481 | -0.030589 | -0.043302 | -0.0021008 | -0.0008786 |  | ten times!
13 | -0.16769 | 0.017222 | -0.0035981 | 0.0048082 | 0.018573 | 0.0049819 | -0.0016018 | 0.049542 | 0.0083789 | 0.042032 | 0.019345 | 0.01136 | -0.034089 | 0.053449 | 0.0023158 | -0.017717 | -0.019645 | -0.007464 | -0.027688 | -0.028988 | 3.8658e-05 |  | ten times!
14 | 0.041357 | -0.079116 | -0.081393 | -0.097249 | -0.086933 | -0.083418 | -0.087965 | -0.082385 | -0.086386 | -0.10829 | -0.11249 | -0.098984 | -0.10859 | -0.099803 | -0.093291 | -0.093682 | -0.073932 | -0.082224 | -0.0695 | -0.027076 | -0.008439 |  | ten times!
15 | -0.04208 | -0.12065 | -0.1041 | -0.070446 | -0.031189 | -0.063454 | -0.093064 | -0.031941 | -0.052434 | -0.017278 | -0.069788 | -0.068106 | -0.10514 | -0.088958 | -0.10886 | -0.091546 | -0.10571 | -0.098197 | -0.11187 | -0.026653 | -0.0076784 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00045419 | -7.5147e-05 | 4.7821e-05 | 4.1941e-05 | -0.00056248 | -0.00011261 | -8.3043e-05 | -4.9781e-05 | -6.4396e-05 | -6.1204e-05 | -6.3724e-05 | 7.6715e-05 | -1.4671e-05 | -4.8941e-05 | -9.329e-05 | 0.00010981 | -2.5758e-06 | 0.00050139 |  | 595.2753
|  | 0.00037907 | -0.00011646 | -6.0864e-05 | -0.00027201 | 0.00028467 | -0.00019625 | -9.6475e-05 | -0.0001205 | -0.00017391 | -0.00011999 | -4.0371e-05 | 5.4537e-05 | -2.1221e-05 | 6.9823e-05 | -2.4245e-05 | 6.0416e-05 | -5.5489e-05 | 0.00044928 |  | 595.3163
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -4004.6951
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 595.2398 | -4178.1546 | 5.1394 | -0.15424 | 1.0755 | 1.0167 | -5.0279 | -0.52888 | -0.2332 | 0.099435 | -0.046725 | -0.014805 | -0.040005 | 1.3645 | 0.45055 | 0.10783 | -0.33568 | 1.6954 | 0.57151 | 5.6115 | -3831.2356
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.083695 | 3.7934 | -0.23374 | 0.057279 | 0.0079452 | 0.2012 | -0.060898 | 0.37815 | 0.49349 | 0.023725 | -0.30003 | -0.65153 | -0.012332 | 0.24096 | 0.35724 | 0.26291 | 0.024656 | -0.23327 | -0.25858 | -0.29718 | -0.96733 |  | 0
db2 | -0.64903 | 3.6551 | -0.26429 | -0.10461 | 0.56294 | 0.52497 | -0.11801 | -0.081145 | 0.41531 | 0.25423 | -0.87196 | -0.53303 | -0.16538 | -0.33641 | 0.04704 | 0.40435 | 0.40192 | -0.50211 | 0.11633 | 0.24987 | -1.5643 |  | 0
db3 | -2.8715 | 45.8482 | 0.43319 | 0.74968 | -0.29771 | -0.14926 | -1.3517 | 0.082783 | 0.048132 | 0.39965 | -0.11157 | -1.1811 | 0.11021 | 0.79722 | -1.6579 | -1.3035 | 0.48141 | 0.93632 | 0.94493 | 1.0692 | 1.6246 |  | 2.2204e-16
db4 | -0.067205 | -0.18877 | -0.052851 | 0.0067085 | 0.035452 | 0.15337 | -0.05394 | -0.14385 | 0.16404 | 0.20924 | -0.25449 | 0.0067147 | -0.13336 | -0.015738 | 0.012872 | 0.039491 | 0.095891 | -0.032822 | -0.0046112 | -0.032111 | 0.31642 |  | 0
db5 | -0.42449 | -4.3616 | -0.49457 | -0.54414 | -0.081781 | 0.093994 | 0.064588 | 0.041664 | 0.037646 | -0.13545 | 0.025195 | 0.3107 | 0.12931 | 0.28642 | 0.3827 | 0.22818 | -0.12378 | 0.11361 | -0.090308 | -0.244 | -4.1125 |  | 0
db6 | -0.021381 | -0.19062 | 0.022059 | 0.021499 | -0.039714 | -0.078319 | -0.024131 | -0.015881 | 0.012675 | -0.0049644 | 0.060207 | -0.054441 | 0.010303 | 0.034368 | -0.035335 | -0.02634 | -0.025702 | 0.098903 | 0.021848 | 0.022966 | 0.10569 |  | 0
db7 | 1.1386 | 3.018 | -0.16405 | -0.078422 | 0.039892 | 0.042415 | 0.0071203 | 0.022201 | -0.0043995 | -0.0058768 | -0.018856 | 0.033466 | -0.0058622 | 0.042573 | 0.047939 | -0.02971 | 0.055964 | 0.0047923 | 0.00068475 | 0.010125 | 1.0083 |  | -5.5511e-17
db8 | -0.0044789 | -0.14801 | 0.0022066 | 0.00079807 | 0.026345 | 0.0020622 | -0.011134 | -0.028146 | 0.015773 | 0.0054907 | -0.022003 | -0.0053572 | -0.018679 | 0.0030706 | -0.019531 | 0.034426 | 0.0032297 | -0.021765 | 0.01676 | 0.016454 | -0.0043095 |  | 0
db9 | 0.45723 | 0.38936 | -0.0028301 | -0.0047333 | 0.0053292 | 0.0069759 | -0.021577 | 0.0088617 | 0.0039696 | 0.012932 | 0.0085137 | -0.019654 | -0.017848 | 0.009091 | -0.00070808 | -0.033722 | 0.0052184 | 0.010219 | 0.011503 | 0.018459 | 0.41714 |  | -6.1679e-18
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.7299 | 0.5927 | -0.013352 | -0.012367 | 0.0007856 | 0.0038619 | 0.0081621 | 0.00051807 | -5.6669e-05 | -0.000437 | 0.0026953 | 0.0097643 | -0.00151 | 0.0021193 | 0.0090776 | 0.0092617 | -0.0033377 | -0.00090954 | -0.0036254 | -0.01065 | 0.62503 |  | -3.7007e-17
db12 | -0.0010788 | -0.0079395 | -0.00068392 | -2.0736e-05 | 0.0019207 | -0.00079865 | 0.0013194 | 0.0026316 | -0.0011008 | 0.0015762 | -0.00058123 | -0.0035138 | -0.0028057 | -0.0038442 | 0.0010346 | 0.0010915 | 0.0007067 | -0.0010193 | 0.0021973 | 0.0018905 | 0.0032163 |  | 0
db13 | 0.068074 | 0.090679 | 0.0033048 | 0.0032065 | -0.001068 | 0.00064144 | -0.00031345 | -0.00082385 | -0.0009845 | -0.0012726 | -0.0020288 | -0.00064337 | -0.0011991 | -0.00010037 | -0.0023835 | -0.0022661 | 0.0026769 | -0.00055888 | 0.0013211 | 0.0024919 | 0.056862 |  | -4.6259e-18
db14 | 0.0017078 | 0.00052019 | -0.0013001 | -0.00052236 | 0.00027827 | 0.00025667 | 0.0010756 | 0.0012815 | 0.00089186 | 0.00010867 | -3.0358e-05 | -0.0010668 | -0.001114 | -0.00020796 | 0.00034773 | 0.00042131 | -0.00010652 | 0.00014945 | -0.00029558 | -0.00016739 | 0.002999 |  | -1.5661e-19
db15 | 0.026546 | -0.0067506 | 0.00043322 | -0.00032925 | -0.00096702 | 0.00024727 | 0.00013509 | -0.00072999 | -0.00068709 | 0.0010121 | -7.5199e-05 | -0.0023013 | -0.00021775 | -0.00065038 | 0.0029804 | -4.6529e-05 | 0.0016127 | 0.00013031 | 0.0012016 | -0.0017482 | 0.013208 |  | -5.7824e-19
da2 | 0.090393 | 13.5789 | -0.44549 | 0.26159 | -0.25775 | 0.71704 | -0.18784 | -0.78772 | 0.62992 | 1.1126 | -0.86769 | -0.55466 | -0.23149 | -0.60909 | 0.4425 | -0.78811 | 1.2126 | 0.54865 | 0.21215 | -0.40715 | 3.427 |  | 0
da3 | 0.47584 | -0.58872 | -0.12798 | 0.015829 | 0.2744 | 0.19601 | 0.12497 | 0.011824 | -0.0028395 | -0.085603 | 0.21565 | -0.14802 | -0.26719 | -0.21505 | 0.031487 | -0.12558 | 0.020735 | 0.27897 | -0.056028 | -0.14158 | -0.62681 |  | 0
da4 | -0.38503 | 3.3463 | -0.018207 | 0.22589 | -0.032507 | -0.25359 | -0.18158 | 0.11121 | -0.28236 | -0.065972 | -0.15192 | 0.043277 | -0.12027 | 0.094833 | 0.17463 | 0.31736 | -0.056468 | 0.1623 | 0.025692 | 0.0076905 | 0.020823 |  | -7.7099e-18
da5 | 0.12788 | 3.042 | -0.12834 | -0.10034 | -0.066582 | 0.34098 | -0.057327 | -0.15394 | -0.076738 | -0.029401 | 0.12839 | 0.19761 | -0.16124 | -0.12812 | 0.22543 | 0.14507 | 0.022415 | -0.060708 | -0.10093 | 0.0037764 | -0.30832 |  | 3.8549e-18
da6 | -0.037026 | 0.61662 | 0.080116 | 0.046163 | 0.016853 | -0.012076 | -0.049968 | -0.043249 | 0.02768 | 0.04664 | -0.027459 | 0.03677 | -0.015896 | -0.073977 | 0.069493 | -0.025044 | -0.10783 | -0.0024102 | 0.00033107 | 0.033863 | 0.070792 |  | 0
da7 | -0.035468 | 1.6853 | -0.057816 | -0.0061955 | -0.06082 | -0.034456 | 0.02939 | 0.043159 | -0.03352 | -0.050943 | 0.0067584 | 0.0044673 | 0.039011 | -0.02798 | -0.011461 | 0.053807 | 0.013079 | 0.065285 | 0.031659 | -0.0034262 | -0.019471 |  | 0
da8 | -0.030436 | 0.036273 | 0.00043499 | 0.014625 | -0.011737 | 0.0067095 | -0.0077725 | -0.035208 | 0.023442 | 0.027527 | 0.016837 | -0.016424 | 0.0005872 | -0.028716 | -0.0023088 | -0.022241 | 0.018066 | 0.000969 | 0.0068118 | 0.0083981 | 0.014568 |  | 4.2404e-18
da9 | -0.0066524 | -0.20963 | -0.0018418 | -0.0094479 | 0.013102 | 0.02336 | 0.013887 | 0.013858 | 0.013804 | -0.010608 | -0.0049408 | -0.0075187 | -0.0033302 | 0.0021614 | 0.0041791 | -0.026293 | -0.014941 | 0.0098277 | -0.0092725 | -0.005984 | -0.036119 |  | 0
da10 | 9.9385e-20 | 0 | 1.6895e-18 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | -9.9385e-20 | 0 |  | 0
da11 | -0.042718 | 0.17916 | -0.013332 | -0.0018967 | -0.0040256 | 0.0020765 | -0.0013983 | 0.0026063 | 0.0058566 | -0.002505 | -0.008134 | -0.0061938 | 0.0023615 | 0.0036809 | 0.005766 | 0.00646 | -0.0019051 | 0.0062885 | 0.0044452 | -0.00015091 | -0.055144 |  | -2.6985e-18
da12 | -0.001977 | -0.0029105 | 0.0014028 | 0.0028867 | 0.0020132 | 0.0019012 | 0.0025296 | -0.00082574 | 0.0035303 | 0.0017766 | -0.0029435 | -0.00043605 | -0.0019047 | -0.0053666 | 0.00062957 | -0.0053261 | -0.00079471 | 0.00047208 | 8.2153e-05 | 0.00037314 | 0.0026651 |  | 0
da13 | -0.003261 | -0.015405 | -2.1881e-05 | -0.0014695 | -0.00069862 | 0.0019898 | 0.0011206 | 0.0011899 | -0.00088576 | -0.0021281 | 0.00059037 | -0.0021067 | -0.0015413 | 0.00073454 | 0.0035404 | -0.00031458 | -0.0017942 | 0.0015715 | 0.00031626 | -9.2832e-05 | -0.0067042 |  | 4.0959e-19
da14 | -0.011365 | 0.0060066 | -0.001681 | -0.00019585 | -0.00037028 | 9.9783e-05 | -0.00018148 | -0.0014285 | 0.0019644 | 0.0012341 | -5.9564e-05 | -0.00045049 | -0.00071844 | -0.00020246 | -0.0015032 | -0.0010022 | 0.0017915 | 0.00075947 | 0.00088648 | 0.0010577 | -0.00121 |  | 5.7824e-19
da15 | 0.00091196 | 0.0056186 | -0.010037 | -0.0045188 | -0.0016718 | 0.00094583 | 0.0031243 | -0.0016757 | 0.001759 | -0.00095251 | 0.0042603 | 0.0023667 | 0.00026795 | 0.0021067 | 0.0034553 | -0.0044271 | -0.002696 | 0.0029422 | 0.0019253 | 0.0028256 | 0.0036079 |  | 0
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -4075.6495
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.2855 | -4346.1718 | 4.3086 | -0.64696 | -0.090923 | -2.2025 | 3.3646 | -1.4449 | -0.44705 | -0.68728 | -1.2215 | -0.68224 | 0.11402 | 1.0631 | 0.30553 | 1.216 | 0.27529 | 1.1219 | -0.037167 | 5.0108 | -3805.1272
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.16102 | 5.3728 | -0.55813 | -0.16766 | -0.23931 | 0.0016058 | 0.41508 | 0.24891 | 0.52827 | 0.087933 | -0.082487 | -0.4604 | 0.054011 | 0.21882 | 0.43401 | 0.28533 | -0.053149 | -0.17701 | -0.27232 | -0.26351 | 0.069831 |  | 0
db2 | -0.60561 | -8.8493 | 0.017138 | 0.24296 | -0.28126 | 0.26282 | 0.39823 | 0.09276 | -0.33876 | -0.23006 | 0.11946 | -0.27218 | 0.56749 | 0.54955 | -0.21807 | -0.37212 | -0.33415 | 0.022913 | 0.16187 | -0.3886 | 1.2397 |  | 0
db3 | -2.6354 | 43.4612 | -0.088283 | 0.705 | -0.30287 | 0.1739 | -1.399 | -0.015902 | 0.43409 | 0.50219 | -0.96158 | -1.697 | 0.16588 | 1.0497 | -1.805 | -1.4083 | 0.3599 | 1.5762 | 1.23 | 1.4811 | 2.1172 |  | -3.2073e-16
db4 | -0.072146 | -3.0996 | 0.091197 | 0.011513 | 0.019799 | 0.11516 | -0.095239 | -0.017788 | -0.058258 | -0.12027 | -0.13348 | -0.085344 | 0.019504 | 0.34448 | -0.071006 | -0.097531 | -0.15889 | 0.18144 | 0.090635 | -0.035922 | 0.05029 |  | 0
db5 | -0.34522 | -5.1916 | -0.39561 | -0.50906 | -0.07004 | -0.095908 | 0.25073 | 0.14196 | -0.032978 | -0.039136 | 0.12317 | 0.39198 | 0.19299 | 0.30156 | 0.2219 | 0.26164 | -0.05028 | -0.19203 | -0.16274 | -0.33814 | -4.1687 |  | 0
db6 | 0.012545 | -0.66952 | 0.022918 | 0.0027077 | -0.060274 | -0.094065 | -0.0057165 | -0.098533 | 0.059859 | 0.063525 | -0.061502 | 0.002715 | -0.13823 | 0.071751 | -0.076718 | 0.060802 | 0.14222 | 0.04465 | -0.0034355 | 0.067319 | 0.11459 |  | 0
db7 | 1.1334 | 2.9837 | -0.14258 | -0.072031 | 0.043564 | 0.039984 | 0.024962 | -0.0023427 | 0.0093219 | 0.022044 | -0.057015 | 0.0014531 | 0.042348 | 0.060978 | -0.06565 | -0.036803 | 0.064888 | 0.0044814 | 0.029564 | 0.03283 | 1.022 |  | 0
db8 | -0.0097502 | -0.13803 | -0.0015246 | -0.0056892 | -0.015836 | 0.015085 | 0.0049799 | 0.0058257 | -0.0090943 | -0.0055336 | 0.03132 | 0.013297 | 0.038693 | 0.026041 | -0.020709 | -0.03481 | -0.02722 | -0.010952 | -0.0026064 | -0.0012664 | 0.018505 |  | 7.7099e-19
db9 | 0.46088 | 0.40181 | 0.0069535 | -0.014351 | 0.00043923 | 0.013929 | -0.027656 | 0.0044477 | -0.0013909 | 0.0072413 | -0.014502 | -0.0072517 | -0.0031842 | 0.011547 | -0.014343 | -0.02462 | 0.010923 | 0.012679 | 0.01434 | 0.024798 | 0.44585 |  | 4.6259e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.72657 | 0.58804 | -0.0072377 | -0.011497 | -0.0019511 | -0.001352 | 0.0094968 | -0.0013068 | -0.0057759 | -0.0016207 | 0.0055075 | 0.011222 | 0.0049275 | 0.00046032 | 0.0099722 | 0.014472 | -0.0042043 | -0.0033578 | -0.0046168 | -0.013138 | 0.61134 |  | -6.1679e-18
db12 | -0.0012848 | 0.0044641 | -0.0019458 | -0.00071453 | 0.0010048 | 7.8143e-05 | 0.00081068 | -0.0017209 | 0.00026175 | -0.001006 | 0.00025822 | 0.0001825 | 0.0013405 | 0.00063182 | -0.0013629 | 0.0021356 | 0.0014977 | -0.0012606 | 0.00057289 | -0.00076379 | 0.0015165 |  | 0
db13 | 0.06724 | 0.090804 | 0.0011808 | 0.0029239 | -0.0013574 | -3.022e-05 | -0.0019481 | -0.0022262 | -0.00078687 | 0.0014357 | -0.0026008 | -0.003317 | 0.00057101 | 0.00016853 | -0.0032873 | -0.0016202 | 0.002533 | 0.0019846 | 0.0030219 | 0.0033544 | 0.054634 |  | 6.1679e-18
db14 | 0.0074475 | 0.0046762 | -0.00062254 | -0.0004558 | -2.9182e-05 | 0.0001099 | 0.00052053 | -0.0005563 | -0.00096389 | -0.00014007 | -0.00075828 | -0.00035837 | -0.00024057 | 7.9253e-06 | 6.0084e-05 | 0.00023496 | 0.0016428 | 0.00070363 | 0.00011143 | 0.00073378 | 0.00063482 |  | -3.8549e-19
db15 | 0.02161 | -0.010975 | -0.00044642 | -0.0046833 | -0.00013131 | 0.001772 | 0.0025797 | -0.0015644 | -0.00026894 | 0.0036443 | 0.0049813 | 0.0053245 | 0.00033236 | -0.0011799 | 0.0004523 | -0.0031002 | -0.00093842 | -0.00078242 | -0.0022083 | -0.003783 | 0.00059117 |  | 0
da2 | 0.46774 | 17.1129 | -0.080737 | -0.096347 | -0.80034 | 0.52417 | -0.085691 | -0.67459 | 0.39287 | 0.63461 | 0.018435 | -1.0869 | -0.52734 | -0.078008 | 0.38378 | -0.29677 | 0.18959 | 0.96481 | -0.13987 | 0.75836 | -5.0554 |  | 0
da3 | 0.26679 | -0.55083 | 0.070583 | -0.13743 | 0.20374 | 0.35824 | -0.11722 | -0.29203 | 0.27752 | -0.097726 | 0.33311 | 0.28882 | -0.16217 | -0.62506 | 0.099252 | -0.36928 | -0.05111 | 0.23307 | 0.28586 | -0.29817 | 0.04051 |  | 0
da4 | -0.12587 | 2.6521 | -0.074082 | 0.033286 | 0.25549 | -0.78136 | 0.19622 | 0.22888 | -0.096796 | 0.00094669 | 0.12704 | -0.15275 | 0.47739 | -0.3559 | 0.34101 | -0.10459 | 0.039436 | -0.18071 | -0.042453 | 0.088941 | -0.73019 |  | 2.313e-17
da5 | -0.029447 | 2.2566 | -0.028343 | -0.042382 | 0.018418 | 0.21621 | 0.03063 | -0.013342 | 0.02756 | 0.095757 | 0.10853 | 0.15663 | -0.069655 | -0.36521 | 0.049141 | -0.025939 | 0.049157 | -0.083048 | -0.027413 | -0.096707 | -0.48483 |  | 0
da6 | 0.10106 | 0.27436 | -0.05932 | -0.05852 | 0.09494 | 0.039862 | 0.0031951 | -0.014167 | 0.07146 | 0.028092 | 0.039257 | 0.084793 | 0.032977 | -0.12199 | 0.063309 | -0.086202 | 0.042599 | -0.0033581 | -0.056562 | -0.10037 | 0.040217 |  | 0
da7 | -0.022182 | 1.6077 | -0.041582 | -0.014938 | -0.072179 | 0.039692 | 0.02857 | 0.010505 | 0.046772 | 0.017266 | 0.055767 | -0.00094033 | 0.07405 | -0.089162 | 0.035558 | 0.030772 | -0.093554 | 0.021896 | -0.0087459 | -0.039748 | -0.010515 |  | 0
da8 | -0.015007 | -0.017315 | 0.021689 | 0.024062 | -0.021768 | 0.008798 | -0.012116 | -0.022543 | 0.029825 | 0.014885 | -0.012296 | -0.039335 | -0.0035354 | -0.046515 | 0.022598 | 0.0054863 | -0.0011066 | 0.036476 | -0.014568 | 0.009964 | 0.01602 |  | 0
da9 | -0.0092715 | -0.23929 | 0.025553 | -0.0097261 | 0.015021 | 0.011687 | 0.0092114 | 0.002099 | 0.019303 | -0.0059962 | 0.015181 | 0.0095271 | -0.0026707 | -0.047487 | 0.015058 | -0.0058645 | -0.012167 | -0.02116 | -0.0028134 | -0.014756 | -0.031564 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.014635 | 0.20152 | -0.013934 | -0.0024562 | -0.0031062 | 0.0048841 | 0.0057128 | 0.0049127 | 0.0097972 | 0.0070201 | -0.0025506 | -0.0031781 | 0.0035626 | -0.0082451 | 0.0070736 | 0.0060723 | -0.0041672 | -0.0026205 | -0.0048186 | -0.0039587 | -0.01603 |  | 1.0119e-18
da12 | -0.000745 | -0.0058634 | 0.0039705 | 0.00060069 | 0.0002001 | 0.0039325 | 0.0013623 | -0.0013171 | 0.0025394 | 0.00068176 | -0.0012998 | -0.00066809 | -0.0016473 | -0.0068499 | 0.0022641 | 0.0014667 | 0.00032296 | 0.00034019 | -0.0023139 | -0.0035852 | -0.00021008 |  | 0
da13 | 0.00066781 | -0.016769 | 0.0010544 | -0.0010276 | -0.000187 | 0.0011895 | -0.00016962 | -0.00082799 | 0.0042864 | 0.00017008 | 0.0035354 | 0.0012667 | 0.0004682 | -0.0040767 | 0.0046771 | -0.00043623 | -0.0024395 | -0.0026323 | -0.0014142 | -0.0034366 | -0.0028988 |  | 0
da14 | -0.0090313 | 0.0041357 | 0.0011197 | 0.00089202 | -0.00069361 | 0.00033798 | 0.00068951 | 0.00023486 | 0.00079284 | 0.00039274 | -0.0017981 | -0.0022179 | -0.00086708 | -0.0018277 | -0.00094893 | -0.0002978 | -0.00033688 | 0.0016382 | 0.00080888 | 0.0020813 | -0.0027076 |  | 1.7347e-18
da15 | -0.0079596 | -0.004208 | -0.0041052 | -0.0024509 | 0.000915 | 0.0048407 | 0.0016141 | -0.0013468 | 0.0047655 | 0.0027162 | 0.0062318 | 0.0009808 | 0.001149 | -0.0025541 | -0.00093628 | -0.002926 | -0.0011951 | -0.0026116 | -0.0018601 | -0.0032269 | -0.0026653 |  | -1.0601e-18
Field Colinearity | -0.3624
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.7003 | -0.31671 | -0.026917 | -0.075704 | 0.11753 | -0.14452 | 0.2947 | 0.40999 | -0.059969 | -0.38412 | -0.73617 | -0.096276 | 0.15679 | 0.27395 | 0.18006 | -0.059685 | -0.31665 | -0.34207 | -0.38127 | -1.0485
Twist (m2 mrad) | 27.7524 | -4.5131 | -0.36338 | -0.96523 | 1.4104 | -1.6258 | 3.0944 | 3.9975 | -0.53972 | -3.169 | -5.5213 | -0.64986 | 0.94072 | 1.4383 | 0.81025 | -0.22382 | -0.94994 | -0.76966 | -0.5719 | -0.39318
Twist Integral (I0) (m2 rad) | 0.014391
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.2253 | -0.71956 | -0.32887 | -0.40041 | -0.15854 | 0.25406 | 0.088068 | 0.36784 | -0.072298 | -0.24275 | -0.62192 | -0.10637 | 0.057442 | 0.27281 | 0.1242 | -0.21434 | -0.3387 | -0.43361 | -0.42528 | -0.089652
Twist (m2 mrad) | 39.1899 | -10.2537 | -4.4398 | -5.1052 | -1.9025 | 2.8582 | 0.92471 | 3.5864 | -0.65068 | -2.0027 | -4.6644 | -0.71801 | 0.34465 | 1.4322 | 0.55888 | -0.80376 | -1.0161 | -0.97561 | -0.63792 | -0.033619
Twist Integral (I0) (m2 rad) | 0.011768
Note on B154: 4.6s
Note on B189: 4.3s
Note on H154: 4.6s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.64903 | -2.8709 | -0.067205 | -0.42607 | 1.1376 | -0.83695 | 0.090393 | 0.47584 | -0.38503 | 15.6807 | 0.59528 | 0.60561 | -2.6333 | 0.072146 | -0.3461 | 1.1322 | -1.6102 | 0.46774 | 0.26679 | -0.12587 | 18.5919 | 0.59532 | -0.54792 | -1.2491 | -0.058566 | -0.66386 | 1.1904 | 0.0023326 | 0.60658 | 0.40813 | -0.25884 | 14419.222 | 0.34746 | -1.1215 | -0.022026 | -0.61233 | 1.1832 | -0.004456 | 0.77994 | 0.23537 | -0.06351 | 14420.588
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8