## Alarm sheet
File name | HCMB__A001-03000406_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 3000406
Date of test Ap 1 | 2005-08-16 |  |  |  | Date of test Ap 2 | 2005-08-16
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | yellow alarm | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | yellow alarm
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | yellow alarm | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
924 | HCMB__A001 | 3000406 |  | 3 |  |  | 3
| COMMENTS: | CBT is exchanged without de-collaring. CC has to be re-measured.
| HOLDING POINT | OK | Data at CERN |  | 2005-08-16 | Analysis by |  | D. Richter | Answer to MMS-LD |  |  | 2005-08-17 | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-19 | E. Todesco | New bounds computed with C. Vollinger over 270 collared coils - for allowed multipoles valid only for X-section 3 (100 coils)
| 2004-01-19 | E. Todesco | New evaluation of integrals in the template (worksheet) with feed down and not from the .txt. file
| 2004-01-19 | E. Todesco | The analysis is now pointing to data with feed-down (worksheet) and data from .txt file without feed-down (summary data)
| 2004-01-19 | E. Todesco | Correction of the coil waviness estimate (previous version was good for only cold mass, overestimate in cc of 18%)
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Dropped alarm colouring of "field colinearity". Commented out in macro CollaredCoilAnalysis
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2008-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2008-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2008-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-04 | E. Todesco | Computation of dx dy with complete first order formula dx=17/10*(b10b11+a11a10)/(b11**2+a11**2) and not dx=17/10*b10/b11
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims, fixed protection issue for Xsec 1 and 2
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -

## Original data
File name |  | HCMB__A001-03000406_cc.xls | [41 empty] | F1
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 3000406 | [11 empty] | Signs | [23 empty] | F1
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1 | [23 empty] | F1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 16-AUG-2005_3406_11-08cc1.txt |  |  |  | a odd | 1 | [23 empty] | F1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 16-AUG-2005_3406_11-08cc2.txt |  |  |  | a even | 1 | [23 empty] | F1
Test Operator |  | Kenklies |  |  | Test Controller |  |  | D. Richter | [35 empty] | F1
F1
Aperture 1 | [11 empty] | Aperture 2 | [31 empty] | F1
Date of test |  | 2005-08-16 |  |  | Time of test |  |  | 11:08 - 13:46 |  |  |  | Date of test |  | 2005-08-16 |  |  | Time of test |  |  | 11:09 - 13:47 | [23 empty] | F1
Mole name |  | Dipole 09 | [9 empty] | Mole name |  | Dipole 21 | [5 empty] | 2004-09-20 | [23 empty] | F1
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 22 | [23 empty] | F1
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw | [23 empty] | F1
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) | [23 empty] | F1
Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.62 |  |  | I. Gain |  |  | 50/500 | [23 empty] | F1
Dx average |  | 7.7044e-05 |  |  | Dy average |  |  | -1.5247e-05 |  |  |  | Dx average |  | 0.00016204 |  |  | Dy average |  |  | 6.1721e-05 | [23 empty] | F1
Magnetic length |  | 14.4197 |  |  | Tranfs. Function (mT/KA) |  |  | 595.663 |  |  |  | Magnetic length |  | 14.4204 |  |  | Tranfs. Function (mT/KA) |  |  | 595.8634 | [23 empty] | F1
Phase |  | 2.9316 |  |  | Integral field (T.m) |  |  | 85.798 |  |  |  | Phase |  | 0.4717 |  |  | Integral field (T.m) |  |  | 85.7776 | [23 empty] | F1
F1
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals | F1
C1 (T) | 3.4841 | 5.9585 | 5.9558 | 5.9561 | 5.9557 | 5.9561 | 5.9573 | 5.9558 | 5.9559 | 5.9561 | 5.9563 | 5.9563 | 5.9557 | 5.9562 | 5.9565 | 5.9566 | 5.9576 | 5.9569 | 5.96 | 3.7093 | 85.798 | C1 (T) | 3.3756 | 5.9604 | 5.9572 | 5.9583 | 5.958 | 5.9587 | 5.9586 | 5.9581 | 5.9582 | 5.9581 | 5.9575 | 5.9584 | 5.9578 | 5.9584 | 5.9584 | 5.9588 | 5.9596 | 5.9594 | 5.9617 | 3.7545 | 85.7776 | F1
Angle (mrad) | 5.1344 | -0.19963 | -0.18253 | -0.38467 | -0.22079 | 0.16211 | 0.31189 | 0.39965 | 0.12884 | -0.11731 | -0.13653 | 0.22504 | 0.52233 | 0.49129 | -0.010951 | -0.057441 | -0.69671 | -0.59526 | -1.1259 | -2.4339 | 2.9316 | Angle (mrad) | 6.5124 | -0.56735 | -0.67913 | -0.84997 | -0.84667 | -0.72979 | -0.4257 | -0.34784 | -0.37671 | -0.44264 | -0.46506 | 0.37426 | 0.90157 | 1.0897 | 0.79997 | 0.38773 | 0.27075 | -0.036305 | -0.12865 | -2.5669 | 0.4717 | F1
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | F1
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | F1
b2 | -2.4343 | -0.14862 | -0.42692 | -0.47243 | -0.61762 | -0.0096019 | 0.23347 | -0.37272 | -0.21607 | -0.3714 | -0.37489 | 0.1299 | -0.75971 | -0.24131 | -0.038723 | -0.3477 | -0.25914 | -0.51441 | -0.58516 | -1.331 | -0.39806 | b2 | -2.7893 | 0.38126 | -0.20097 | -0.4267 | -0.57827 | -0.092259 | 0.24862 | -0.19169 | -0.48187 | -0.29896 | 0.7839 | -0.58071 | 0.64146 | 0.39573 | -0.42632 | -0.38025 | -0.42534 | 0.13733 | -0.46043 | 0.39363 | -0.17127 | F1
b3 | 40.767 | -4.3131 | -3.8604 | -5.0259 | -4.4501 | -4.7386 | -5.0812 | -3.8538 | -4.0806 | -2.5537 | -4.7837 | -3.7684 | -4.5606 | -6.6771 | -5.2381 | -3.1823 | -3.0096 | -2.8144 | -2.8169 | 0.63311 | -2.6327 | b3 | 40.6096 | -4.6591 | -4.0751 | -6.031 | -3.9551 | -4.797 | -3.6391 | -3.811 | -3.0843 | -3.7413 | -4.339 | -3.8138 | -3.5047 | -6.0731 | -4.693 | -3.8447 | -3.0304 | -3.0754 | -2.3736 | 0.37773 | -2.568 | F1
b4 | -1.6155 | -0.023926 | -0.11176 | -0.070124 | -0.092535 | -0.12772 | -0.11982 | -0.10124 | -0.12143 | 0.043452 | -0.02224 | 0.053953 | -0.14483 | -0.1611 | -0.11466 | -0.046878 | -0.13093 | -0.045532 | -0.1642 | 0.67382 | -0.10552 | b4 | -3.2911 | 0.045468 | -0.029484 | 0.055515 | 0.015337 | 0.075626 | 0.23689 | 0.037499 | -0.1025 | -0.052379 | 0.22467 | -0.12927 | 0.2201 | 0.081078 | -0.01445 | 0.34011 | 0.011044 | -0.014571 | -0.05462 | 0.064173 | -0.045732 | F1
b5 | -4.961 | -0.36418 | -0.40277 | -0.12245 | 0.048663 | -0.013425 | 0.16821 | -0.0234 | 0.088916 | -0.13199 | 0.25544 | -0.060686 | 0.23847 | 0.27882 | 0.081805 | 0.028681 | -0.003999 | -0.12316 | -0.23353 | -3.867 | -0.29158 | b5 | -4.5971 | -0.38711 | -0.51549 | -0.074832 | -0.18972 | -0.02831 | -0.037076 | -0.048497 | -0.16137 | -0.26268 | 0.043589 | -0.088984 | 0.0030934 | 0.063026 | 0.0024429 | -0.23549 | -0.13405 | -0.20464 | -0.31907 | -3.6213 | -0.38867 | F1
b6 | -0.21926 | -0.0030371 | -0.058023 | -0.031572 | -0.045539 | 0.044843 | 0.069058 | -0.055082 | -0.00072645 | 0.12422 | 0.047914 | -0.01703 | 0.011582 | -0.069384 | -0.063541 | -0.002528 | -0.012488 | 0.067834 | -0.014221 | -0.12204 | -0.011035 | b6 | -0.50423 | 0.051756 | 0.0546 | -0.020963 | 0.042189 | -0.025348 | 0.02151 | -0.050736 | 0.068656 | 0.12464 | -0.012413 | 0.014149 | 0.064583 | 0.0076109 | 0.1261 | -0.10448 | -0.053541 | -0.018185 | -0.016906 | 0.023105 | 0.00010873 | F1
b7 | 2.961 | 0.95205 | 0.99205 | 1.0887 | 1.0968 | 1.1101 | 1.2111 | 1.1327 | 1.1415 | 1.1066 | 1.1732 | 1.1 | 1.0883 | 1.1644 | 1.0984 | 1.0984 | 1.0737 | 1.1198 | 1.1025 | 0.99954 | 1.156 | b7 | 3.1789 | 0.97276 | 1.0337 | 1.2697 | 1.2486 | 1.2061 | 1.2011 | 1.2298 | 1.1953 | 1.1922 | 1.2035 | 1.2303 | 1.1822 | 1.1918 | 1.1651 | 1.2615 | 1.1584 | 1.1735 | 1.1584 | 1.0345 | 1.236 | F1
b8 | -0.056337 | 0.036099 | -0.013612 | -0.050638 | -0.022976 | -0.024177 | -0.042744 | -0.029682 | 0.015379 | 0.0042066 | -0.036367 | 0.0096436 | -0.0063639 | 0.0062088 | 0.01325 | 0.023944 | 0.023218 | 0.03041 | 0.014249 | -0.052178 | -0.0060058 | b8 | -0.057072 | 0.0060315 | -0.049197 | -0.057914 | -0.050564 | -0.01704 | -0.0189 | -0.013484 | -0.017423 | -0.0047561 | -0.022925 | -0.037921 | -0.0075664 | -0.0095701 | -0.019159 | -0.0093227 | 0.031137 | 0.01903 | -0.0078107 | -0.019779 | -0.017301 | F1
b9 | 0.34633 | 0.45831 | 0.44565 | 0.44809 | 0.45761 | 0.43988 | 0.44336 | 0.46725 | 0.47061 | 0.48219 | 0.44362 | 0.44142 | 0.462 | 0.46959 | 0.44843 | 0.4775 | 0.46283 | 0.4817 | 0.47677 | 0.35615 | 0.45301 | b9 | 0.37362 | 0.46137 | 0.45103 | 0.43749 | 0.46353 | 0.43245 | 0.46388 | 0.45988 | 0.47163 | 0.46743 | 0.45196 | 0.44031 | 0.46768 | 0.44461 | 0.44881 | 0.4692 | 0.45813 | 0.47345 | 0.48047 | 0.41724 | 0.45414 | F1
b10 | 0.022818 | 0.029657 | -0.023146 | -0.031508 | -0.014691 | 0.00067424 | -0.041645 | -0.018783 | 0.036678 | 0.055302 | -0.0090867 | 0.0015794 | 0.0020724 | -0.014396 | -0.017659 | 0.015657 | 0.034146 | 0.06404 | 0.023622 | -0.17004 | 0 | b10 | 0.071062 | 0.015425 | -0.016595 | -0.051133 | -0.028684 | 0.014371 | 0.022853 | 0.0014615 | 0.017918 | 0.035714 | -0.00069113 | -0.017319 | -0.010049 | -0.0054567 | 0.0095216 | 0.015642 | 0.027739 | -0.013428 | -0.026059 | -0.049974 | 0 | F1
b11 | 0.60887 | 0.74011 | 0.74386 | 0.7518 | 0.75414 | 0.75328 | 0.75293 | 0.74468 | 0.74769 | 0.74188 | 0.75862 | 0.74588 | 0.75136 | 0.76241 | 0.76132 | 0.75491 | 0.74809 | 0.74661 | 0.73876 | 0.63371 | 0.74185 | b11 | 0.62291 | 0.73819 | 0.74269 | 0.74925 | 0.74617 | 0.75565 | 0.74917 | 0.74246 | 0.74829 | 0.74136 | 0.75598 | 0.74724 | 0.74581 | 0.75644 | 0.75682 | 0.75145 | 0.74993 | 0.75027 | 0.74287 | 0.64047 | 0.74109 | F1
b12 | 0.0040542 | 0.0053833 | -0.0033104 | -0.0048909 | -0.0080578 | -0.0011277 | -0.0045849 | -0.0005327 | 0.0031255 | 0.0032091 | -0.0052444 | -0.0025945 | -0.0015454 | -0.0030847 | -0.0013004 | 0.00080543 | 0.0043735 | 0.0046942 | 0.0014471 | -0.017371 | -0.0011285 | b12 | 0.021095 | 0.0013012 | -0.0069975 | -0.0087688 | -0.0039901 | 0.00066214 | -0.0018134 | -0.0029141 | -0.0033051 | 0.0032048 | -0.0012578 | -0.0047991 | -0.0026191 | -0.0041783 | -0.00083717 | -0.0062637 | 0.00063719 | -0.0010156 | -0.0046788 | -0.0068753 | -0.0020844 | F1
b13 | 0.085078 | 0.064134 | 0.066874 | 0.064915 | 0.06597 | 0.063892 | 0.064064 | 0.063798 | 0.060532 | 0.062789 | 0.063247 | 0.063851 | 0.060464 | 0.060287 | 0.063629 | 0.065214 | 0.065692 | 0.066878 | 0.067206 | 0.059378 | 0.064567 | b13 | 0.086675 | 0.064058 | 0.066891 | 0.06553 | 0.070778 | 0.066873 | 0.063783 | 0.065854 | 0.065925 | 0.06373 | 0.066364 | 0.067727 | 0.064459 | 0.063877 | 0.064466 | 0.066682 | 0.066741 | 0.06828 | 0.068413 | 0.053286 | 0.066319 | F1
b14 | -0.0058285 | 0.00034813 | -0.0030637 | -0.0035209 | -0.0038251 | -0.0015743 | -0.003032 | -0.0022755 | 0.00044603 | 0.0013274 | -0.0021472 | -0.0013341 | -0.0012813 | -0.0016352 | -0.002183 | -0.0015801 | -0.00032316 | 0.00084122 | -0.0011905 | -0.012373 | -0.0019324 | b14 | -0.0056198 | -0.0039015 | -0.0044042 | -0.0056586 | -0.0048232 | -0.0035101 | -0.0029882 | -0.0030016 | -0.0028674 | -0.0020337 | -0.0026616 | -0.004209 | -0.0033657 | -0.0031272 | -0.0028924 | -0.0045642 | -0.0026025 | -0.0042617 | -0.004412 | -0.0021621 | -0.0036377 | F1
b15 | -0.013718 | 0.025489 | 0.025118 | 0.023322 | 0.025061 | 0.023215 | 0.023209 | 0.022334 | 0.023915 | 0.023012 | 0.023113 | 0.023448 | 0.025275 | 0.024416 | 0.025382 | 0.025658 | 0.025262 | 0.024715 | 0.023213 | 0.020115 | 0.02289 | b15 | -0.014663 | 0.024057 | 0.022693 | 0.02027 | 0.020609 | 0.022864 | 0.022904 | 0.021212 | 0.022057 | 0.022263 | 0.020889 | 0.021945 | 0.021791 | 0.023477 | 0.022922 | 0.02128 | 0.023324 | 0.02331 | 0.020755 | 0.0028013 | 0.020425 | F1
b16 | [21 empty] | b16 | [21 empty] | F1
b17 | [21 empty] | b17 | [21 empty] | F1
a1 | 51.3447 | -1.9963 | -1.8253 | -3.8468 | -2.2079 | 1.6211 | 3.1189 | 3.9965 | 1.2884 | -1.1731 | -1.3653 | 2.2503 | 5.2233 | 4.9129 | -0.10951 | -0.57441 | -6.9671 | -5.9526 | -11.259 | -24.3392 | 0 | a1 | 65.1251 | -5.6735 | -6.7913 | -8.4997 | -8.4668 | -7.298 | -4.257 | -3.4783 | -3.7671 | -4.4264 | -4.6506 | 3.7426 | 9.0157 | 10.8967 | 7.9997 | 3.8773 | 2.7075 | -0.36305 | -1.2865 | -25.6695 | 0 | F1
a2 | 10.075 | 1.5058 | 0.58538 | 1.8596 | 0.63649 | 1.1801 | 1.96 | 1.1233 | 1.7723 | 0.049157 | -0.39004 | 1.0861 | 0.60098 | 0.93102 | 1.0656 | 0.13969 | 1.4245 | 1.6446 | 0.51753 | -1.1662 | 1.1901 | a2 | 21.8431 | 0.44079 | 0.29708 | 0.65095 | -0.44176 | 0.69015 | -0.1396 | 1.0254 | 0.42564 | 0.93525 | 1.7924 | 0.24845 | 0.6799 | 0.10266 | 0.32428 | -0.51625 | 0.75567 | 1.4688 | 0.5273 | -6.0018 | 0.93032 | F1
a3 | 0.57887 | 0.47586 | 0.63845 | 0.75186 | 0.98955 | 0.33563 | -0.012648 | 0.02833 | 0.12922 | 0.32406 | 0.56564 | 0.52854 | 0.18332 | 0.87213 | 0.81506 | 0.6113 | 0.79727 | 1.0281 | 0.81567 | -0.023976 | 0.5311 | a3 | 1.3641 | 0.9949 | 0.77125 | 0.41906 | 1.1046 | 0.79692 | 0.38909 | 0.66791 | 0.38545 | 1.0066 | 1.0953 | 0.69343 | 0.67237 | 0.88764 | 0.71501 | 0.84099 | 0.99257 | 1.2282 | 1.0634 | 0.8309 | 0.83459 | F1
a4 | 1.5205 | -0.4872 | -0.35775 | -0.37706 | -0.41337 | -0.73398 | -1.2915 | -1.0669 | -1.0258 | -0.4012 | 0.10133 | -0.33041 | -0.75221 | -0.32566 | -0.58271 | -0.45037 | -0.41346 | -0.24249 | 0.048965 | -1.2524 | -0.46814 | a4 | 2.5657 | 0.048459 | -0.20031 | 0.25173 | 0.26951 | -0.14558 | 0.15375 | -0.24292 | -0.66509 | 0.08532 | -0.0041998 | -0.13688 | -0.17069 | 0.41746 | -0.074157 | 0.0027451 | 0.085885 | -0.11886 | 0.19406 | -0.13913 | 0.058145 | F1
a5 | 3.3241 | -0.021727 | 0.066711 | 0.28034 | 0.26776 | 0.026294 | -0.12643 | 0.025881 | 0.037704 | -0.12478 | -0.010454 | 0.15417 | -0.067465 | 0.086094 | 0.16247 | 0.18894 | 0.15792 | 0.23975 | 0.22419 | -0.16673 | 0.17742 | a5 | 3.8386 | 0.10299 | 0.092594 | 0.11314 | 0.33489 | 0.092766 | 0.043515 | 0.093834 | 0.11058 | 0.232 | 0.096662 | -0.04531 | -0.033394 | 0.15331 | 0.073945 | 0.034383 | 0.089375 | 0.19429 | 0.20694 | 0.14595 | 0.22155 | F1
a6 | 0.34865 | 0.031276 | 0.084106 | 0.052194 | 0.13079 | 0.059455 | -0.077347 | 0.02825 | 0.089601 | 0.1135 | 0.04278 | 0.10088 | -0.01078 | -0.083295 | 0.03571 | 0.006287 | 0.11225 | 0.070373 | 0.10224 | 0.43036 | 0.070815 | a6 | 0.83994 | 0.016454 | 0.053933 | -0.042151 | 0.095048 | 0.006659 | 0.0093837 | -0.023024 | -0.14831 | 0.00050515 | -0.11602 | -0.14438 | -0.26939 | -0.083563 | -0.12955 | -0.078018 | -0.16903 | -0.073763 | -0.029857 | 0.025091 | -0.032998 | F1
a7 | 1.6642 | -0.070566 | -0.062332 | -0.17362 | -0.086172 | -0.10796 | -0.14112 | -0.096491 | -0.11762 | -0.12162 | -0.053375 | -0.082405 | -0.061674 | -0.083152 | -0.1035 | 0.011527 | -0.034314 | -0.039845 | -0.10688 | -0.16434 | -0.034365 | a7 | 1.7069 | -0.097316 | -0.099208 | -0.18923 | -0.13685 | -0.14245 | -0.054275 | -0.031555 | -0.082596 | -0.15983 | -0.0086173 | -0.0098931 | -0.10604 | -0.07881 | -0.083643 | 0.0067603 | -0.041608 | -0.094251 | -0.07763 | -0.11478 | -0.03086 | F1
a8 | -0.0094225 | -0.045277 | -0.045312 | -0.02027 | -0.04041 | -0.023656 | -0.035319 | -0.043253 | -0.021118 | -0.017891 | -0.054578 | 0.022813 | -0.0062983 | -0.015593 | -0.014952 | -0.030909 | -0.0087023 | -0.021767 | 0.032981 | 0.23087 | -0.013079 | a8 | -0.00017644 | 0.0013431 | -0.010782 | -0.0064237 | -0.010451 | -0.029319 | -0.048745 | 0.0020866 | -0.051069 | 0.0061411 | -0.0044254 | 0.047369 | -0.02694 | -0.022745 | -0.053008 | -0.062077 | 0.0070678 | -0.0039572 | 0.021663 | -0.0035759 | -0.012846 | F1
a9 | -0.23079 | -0.039331 | -0.039751 | -0.031249 | -0.012193 | -0.057957 | -0.096332 | -0.05774 | -0.049958 | -0.032236 | -0.046422 | -0.030222 | -0.058679 | -0.016044 | -0.0074098 | -0.041245 | -0.026668 | -0.031201 | -0.029345 | -0.19273 | -0.049927 | a9 | -0.20449 | -0.046529 | -0.069947 | -0.047042 | -0.027711 | -0.049567 | -0.07143 | -0.049027 | -0.078569 | -0.031253 | -0.04517 | -0.061375 | -0.07023 | -0.040157 | -0.054797 | -0.093058 | -0.065947 | -0.050043 | -0.046103 | -0.075952 | -0.060513 | F1
a10 | 0.024379 | -0.053555 | -0.0040898 | -0.016619 | 0.00158 | -0.017454 | -0.027859 | 0.0036776 | -0.020881 | -0.0015788 | -0.0050789 | 0.030379 | -0.012788 | -0.047717 | -0.045347 | -0.033839 | 0.0036573 | -0.023311 | 0.020744 | 0.37869 | 0 | a10 | 0.14278 | -0.00015874 | 0.0095258 | -0.003259 | 0.015384 | -0.032174 | 0.0047678 | 0.01779 | -0.038708 | 0.0157 | -0.028866 | 0.055556 | -0.030567 | -0.027617 | -0.03555 | -0.01127 | 0.0041331 | 0.012906 | 0.023201 | -0.050305 | 0 | F1
a11 | 0.16426 | -0.090793 | -0.085344 | -0.090485 | -0.085879 | -0.082954 | -0.081788 | -0.081655 | -0.082393 | -0.092588 | -0.073834 | -0.080936 | -0.079222 | -0.077368 | -0.087601 | -0.071499 | -0.083508 | -0.08364 | -0.095305 | -0.12568 | -0.077521 | a11 | 0.16581 | -0.11799 | -0.11382 | -0.12196 | -0.11455 | -0.11428 | -0.10836 | -0.10254 | -0.10655 | -0.11056 | -0.10009 | -0.10106 | -0.10428 | -0.099172 | -0.10313 | -0.095569 | -0.10426 | -0.10733 | -0.10555 | -0.11725 | -0.099549 | F1
a12 | -0.0067591 | -0.0039746 | 0.001203 | -0.00098223 | 0.0009073 | -0.0018113 | -0.0036248 | 0.00053715 | 0.0025251 | -0.00013663 | -0.00083512 | 0.003185 | -0.0004559 | -0.0058106 | -0.0039306 | -0.0069098 | 0.00067348 | -0.0021051 | 0.0020798 | 0.037537 | -2.3255e-06 | a12 | 0.0061979 | -0.0010442 | 0.0016253 | -0.001791 | 0.0022499 | -0.0033611 | -0.0012058 | -0.0024912 | -0.0084236 | 1.2497e-05 | -0.0036401 | 0.0035752 | -0.0095782 | -0.0062708 | -0.0054945 | -0.0024859 | -0.0059312 | -0.0039757 | -0.0020295 | -0.0053528 | -0.002611 | F1
a13 | -0.024299 | -0.0059834 | -0.0069176 | -0.0077722 | -0.0054635 | -0.0079593 | -0.010833 | -0.0063165 | -0.0071037 | -0.0069434 | -0.0066579 | -0.0062844 | -0.0077457 | -0.0041939 | -0.0052759 | -0.0053723 | -0.0045305 | -0.0044163 | -0.0048614 | -0.0091804 | -0.0070055 | a13 | -0.022499 | -0.007299 | -0.01097 | -0.010264 | -0.0072981 | -0.0093249 | -0.0097444 | -0.0068294 | -0.0099707 | -0.0096431 | -0.0076767 | -0.0095829 | -0.0098595 | -0.0066717 | -0.0079564 | -0.010112 | -0.0087568 | -0.007685 | -0.0071402 | -0.010281 | -0.0091687 | F1
a14 | -0.0058562 | -0.0041765 | -0.0025809 | -0.0040666 | -0.0041698 | -0.0057124 | -0.0062171 | -0.0036857 | -0.0050536 | -0.0022647 | -0.0018725 | 0.00066581 | -0.0019017 | -0.0043852 | -0.0032609 | -0.0044961 | -0.0027279 | -0.0025583 | 0.00046964 | -0.0038727 | -0.0033231 | a14 | -0.0077745 | -0.0014625 | -0.0018963 | -0.00096125 | -0.0013944 | -0.0050567 | -0.0036545 | -0.001267 | -0.0053063 | -0.00055226 | -0.0016962 | 0.0020491 | -0.0016249 | -0.0013393 | -0.00395 | -0.0027856 | 0.0016593 | 0.00013909 | 0.00027384 | 0.0005069 | -0.0017143 | F1
a15 | 0.0065521 | -0.0028113 | -0.0023314 | -0.0024813 | -0.0016647 | -0.0028136 | -0.0035186 | -0.0013683 | -0.0026309 | -0.0022751 | -0.0026693 | -0.0016057 | -0.0021291 | -0.0018862 | -0.0018327 | -0.0026864 | -0.0024822 | -0.0019373 | -0.0023664 | 0.027916 | -0.0010556 | a15 | 0.003958 | -0.0053864 | -0.0042997 | -0.005731 | -0.0023376 | -0.0043383 | -0.0042937 | -0.0024136 | -0.0046655 | -0.0027061 | -0.0034262 | -0.00089975 | -0.0041494 | -0.0046168 | -0.003805 | -0.0037873 | -0.0021508 | -0.0028041 | -0.0024722 | -0.00090168 | -0.0032615 | F1
a16 | [21 empty] | a16 | [21 empty] | F1
a17 | [21 empty] | a17 | [21 empty] | F1
Dx (m) | -7.6533e-05 | -8.1952e-05 | 5.1155e-05 | 6.5772e-05 | 3.3097e-05 | 0 | 8.6165e-05 | 4.3279e-05 | -8.7576e-05 | -0.00012523 | 1.9074e-05 | 0 | 0 | 2.1066e-05 | 2.741e-05 | -4.2082e-05 | -7.5729e-05 | -0.0001499 | -4.7408e-05 | 0.00061773 | 7.7044e-05 | Dx (m) | -0.00027802 | -3.4695e-05 | 4.0383e-05 | 0.00011188 | 6.9107e-05 | -4.2305e-05 | -4.9262e-05 | 0 | -5.2149e-05 | -7.4866e-05 | 0 | 5.5455e-05 | 1.2902e-05 | 0 | -3.1673e-05 | -3.801e-05 | -6.0418e-05 | 3.3919e-05 | 6.586e-05 | 0.00010459 | 0.00016204 | F1
Dy (m) | -4.7416e-05 | 0.00011299 | 1.5216e-05 | 4.5507e-05 | 0 | 3.8752e-05 | 7.228e-05 | 0 | 3.7824e-05 | -1.2012e-05 | 1.3239e-05 | -6.8816e-05 | 2.8126e-05 | 0.00010852 | 0.0001044 | 7.2209e-05 | -1.6769e-05 | 3.6309e-05 | -5.3852e-05 | -0.00089429 | -1.5247e-05 | Dy (m) | -0.00031692 | 0 | -1.562e-05 | 2.5631e-05 | -2.4448e-05 | 6.5984e-05 | -1.7941e-05 | -4.0427e-05 | 8.0488e-05 | -4.7173e-05 | 6.3991e-05 | -0.00011891 | 7.1463e-05 | 6.2587e-05 | 7.5527e-05 | 2.0658e-05 | -1.776e-05 | -2.4391e-05 | -4.3742e-05 | 0.00015273 | 6.1721e-05 | F1
Magnetic length (m) |  | 14.4197 | [19 empty] | Magnetic length (m) |  | 14.4204 | [19 empty] | F1
Transfer funct. (mT/KA) |  | 595.663 | [19 empty] | Transfer funct. (mT/KA) |  | 595.8634 | [19 empty] | F1
Coil positioning (mm) |  | 0.016953 | [19 empty] | Coil positioning (mm) |  | 0.017837 | [19 empty] | F1
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10 | [19 empty] | F1
F1
F1
Soft version | DIMM V2.1.6 | [20 empty] | Soft version | DIMM V2.1.6 | [20 empty] | F1
Level Average | 0.38073 | 0.43007 | 0.47876 | 0.45992 | 0.45161 | 0.36323 | 0.56193 | 0.25999 | 0.3061 | 0.32799 | 0.27221 | 0.10619 | 0.08025 | 0.16313 | 0.069092 | -0.046721 | -0.074952 | -0.092035 | -0.24209 | -0.54313 |  | Level Average | 0.020713 | 0.45457 | 0.17009 | 0.081654 | -0.06688 | 0.063678 | -0.211 | -0.092415 | -0.28391 | -0.23951 | -0.44602 | -0.15064 | -0.19331 | -0.3162 | -0.4682 | -0.15703 | -0.46697 | -0.2683 | -0.086939 | -0.0612 |  | F1
Level Sigma2 | 0.65351 | 0.32112 | 0.18859 | 0.25834 | 0.38507 | 0.16226 | 0.25573 | 0.27915 | 0.23568 | 0.008091 | 0.11086 | 0.07937 | 0.016192 | 0.004099 | 0.001147 | 0.044138 | 0.049737 | 0.20092 | 1.1681 | 4.0391 |  | Level Sigma2 | 3.7e-05 | 0.016116 | 0.019948 | 0.030563 | 0.041532 | 0.019794 | 0.067723 | 0.049896 | 0.058316 | 0.027373 | 0.10298 | 0.027368 | 0.015275 | 0.012481 | 0.092504 | 0.005909 | 0.094279 | 0.10773 | 0.42815 | 0.004621 |  | F1
Position (mm) | 0 | 747 | 1498 | 2248 | 3000 | 3751 | 4501 | 5252 | 6004 | 6754 | 7505 | 8256 | 9007 | 9758 | 10510 | 11261 | 12012 | 12762 | 13513 | 14264 |  | Position (mm) | 0 | 751 | 1502 | 2253 | 3005 | 3756 | 4507 | 5258 | 6010 | 6761 | 7512 | 8264 | 9014 | 9766 | 10517 | 11268 | 12020 | 12771 | 13522 | 14273 |  | F1
Fist move (mm) | First move Ap1 | 943.4915 | First move Ap2 | 948.4545 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2 | [18 empty] | F1
F1
F1
F1
F1
F1
F1
F1
F1
F1
F1
F1

## Summary Data
File | HCMB__A001-3000406 |  |  | Aperture 1 - Collared coils | [17 empty] | File | HCMB__A001-3000406 |  |  | Aperture 2 - Collared coils
C1/i (mT/KA) | 348.4073 | 595.8518 | 595.5838 | 595.6078 | 595.5697 | 595.6063 | 595.7286 | 595.5798 | 595.5939 | 595.6055 | 595.6348 | 595.6266 | 595.5674 | 595.6226 | 595.6496 | 595.657 | 595.7557 | 595.6926 | 596.0006 | 370.9308 | 8579.8028 | C1/i (mT/KA) | 337.556 | 596.0352 | 595.7192 | 595.8297 | 595.8045 | 595.868 | 595.8583 | 595.8085 | 595.8169 | 595.8069 | 595.7497 | 595.8396 | 595.7794 | 595.843 | 595.8399 | 595.8812 | 595.9613 | 595.9353 | 596.1653 | 375.4542 | 8577.7606
Angle (mrad) | 5.1344 | -0.19963 | -0.18253 | -0.38467 | -0.22079 | 0.16211 | 0.31189 | 0.39965 | 0.12884 | -0.11731 | -0.13653 | 0.22504 | 0.52233 | 0.49129 | -0.010951 | -0.057441 | -0.69671 | -0.59526 | -1.1259 | -2.4339 | 2.9316 | Angle (mrad) | 6.5124 | -0.56735 | -0.67913 | -0.84997 | -0.84667 | -0.72979 | -0.4257 | -0.34784 | -0.37671 | -0.44264 | -0.46506 | 0.37426 | 0.90157 | 1.0897 | 0.79997 | 0.38773 | 0.27075 | -0.036305 | -0.12865 | -2.5669 | 0.4717
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.4343 | -0.14862 | -0.42692 | -0.47243 | -0.61762 | -0.0096019 | 0.23347 | -0.37272 | -0.21607 | -0.3714 | -0.37489 | 0.1299 | -0.75971 | -0.24131 | -0.038723 | -0.3477 | -0.25914 | -0.51441 | -0.58516 | -1.331 |  | b2 | -2.7893 | 0.38126 | -0.20097 | -0.4267 | -0.57827 | -0.092259 | 0.24862 | -0.19169 | -0.48187 | -0.29896 | 0.7839 | -0.58071 | 0.64146 | 0.39573 | -0.42632 | -0.38025 | -0.42534 | 0.13733 | -0.46043 | 0.39363
b3 | 40.767 | -4.3131 | -3.8604 | -5.0259 | -4.4501 | -4.7386 | -5.0812 | -3.8538 | -4.0806 | -2.5537 | -4.7837 | -3.7684 | -4.5606 | -6.6771 | -5.2381 | -3.1823 | -3.0096 | -2.8144 | -2.8169 | 0.63311 |  | b3 | 40.6096 | -4.6591 | -4.0751 | -6.031 | -3.9551 | -4.797 | -3.6391 | -3.811 | -3.0843 | -3.7413 | -4.339 | -3.8138 | -3.5047 | -6.0731 | -4.693 | -3.8447 | -3.0304 | -3.0754 | -2.3736 | 0.37773
b4 | -1.6155 | -0.023926 | -0.11176 | -0.070124 | -0.092535 | -0.12772 | -0.11982 | -0.10124 | -0.12143 | 0.043452 | -0.02224 | 0.053953 | -0.14483 | -0.1611 | -0.11466 | -0.046878 | -0.13093 | -0.045532 | -0.1642 | 0.67382 |  | b4 | -3.2911 | 0.045468 | -0.029484 | 0.055515 | 0.015337 | 0.075626 | 0.23689 | 0.037499 | -0.1025 | -0.052379 | 0.22467 | -0.12927 | 0.2201 | 0.081078 | -0.01445 | 0.34011 | 0.011044 | -0.014571 | -0.05462 | 0.064173
b5 | -4.961 | -0.36418 | -0.40277 | -0.12245 | 0.048663 | -0.013425 | 0.16821 | -0.0234 | 0.088916 | -0.13199 | 0.25544 | -0.060686 | 0.23847 | 0.27882 | 0.081805 | 0.028681 | -0.003999 | -0.12316 | -0.23353 | -3.867 |  | b5 | -4.5971 | -0.38711 | -0.51549 | -0.074832 | -0.18972 | -0.02831 | -0.037076 | -0.048497 | -0.16137 | -0.26268 | 0.043589 | -0.088984 | 0.0030934 | 0.063026 | 0.0024429 | -0.23549 | -0.13405 | -0.20464 | -0.31907 | -3.6213
b6 | -0.21926 | -0.0030371 | -0.058023 | -0.031572 | -0.045539 | 0.044843 | 0.069058 | -0.055082 | -0.00072645 | 0.12422 | 0.047914 | -0.01703 | 0.011582 | -0.069384 | -0.063541 | -0.002528 | -0.012488 | 0.067834 | -0.014221 | -0.12204 |  | b6 | -0.50423 | 0.051756 | 0.0546 | -0.020963 | 0.042189 | -0.025348 | 0.02151 | -0.050736 | 0.068656 | 0.12464 | -0.012413 | 0.014149 | 0.064583 | 0.0076109 | 0.1261 | -0.10448 | -0.053541 | -0.018185 | -0.016906 | 0.023105
b7 | 2.961 | 0.95205 | 0.99205 | 1.0887 | 1.0968 | 1.1101 | 1.2111 | 1.1327 | 1.1415 | 1.1066 | 1.1732 | 1.1 | 1.0883 | 1.1644 | 1.0984 | 1.0984 | 1.0737 | 1.1198 | 1.1025 | 0.99954 |  | b7 | 3.1789 | 0.97276 | 1.0337 | 1.2697 | 1.2486 | 1.2061 | 1.2011 | 1.2298 | 1.1953 | 1.1922 | 1.2035 | 1.2303 | 1.1822 | 1.1918 | 1.1651 | 1.2615 | 1.1584 | 1.1735 | 1.1584 | 1.0345
b8 | -0.056337 | 0.036099 | -0.013612 | -0.050638 | -0.022976 | -0.024177 | -0.042744 | -0.029682 | 0.015379 | 0.0042066 | -0.036367 | 0.0096436 | -0.0063639 | 0.0062088 | 0.01325 | 0.023944 | 0.023218 | 0.03041 | 0.014249 | -0.052178 |  | b8 | -0.057072 | 0.0060315 | -0.049197 | -0.057914 | -0.050564 | -0.01704 | -0.0189 | -0.013484 | -0.017423 | -0.0047561 | -0.022925 | -0.037921 | -0.0075664 | -0.0095701 | -0.019159 | -0.0093227 | 0.031137 | 0.01903 | -0.0078107 | -0.019779
b9 | 0.34633 | 0.45831 | 0.44565 | 0.44809 | 0.45761 | 0.43988 | 0.44336 | 0.46725 | 0.47061 | 0.48219 | 0.44362 | 0.44142 | 0.462 | 0.46959 | 0.44843 | 0.4775 | 0.46283 | 0.4817 | 0.47677 | 0.35615 |  | b9 | 0.37362 | 0.46137 | 0.45103 | 0.43749 | 0.46353 | 0.43245 | 0.46388 | 0.45988 | 0.47163 | 0.46743 | 0.45196 | 0.44031 | 0.46768 | 0.44461 | 0.44881 | 0.4692 | 0.45813 | 0.47345 | 0.48047 | 0.41724
b10 | 0.022818 | 0.029657 | -0.023146 | -0.031508 | -0.014691 | 0.00067424 | -0.041645 | -0.018783 | 0.036678 | 0.055302 | -0.0090867 | 0.0015794 | 0.0020724 | -0.014396 | -0.017659 | 0.015657 | 0.034146 | 0.06404 | 0.023622 | -0.17004 |  | b10 | 0.071062 | 0.015425 | -0.016595 | -0.051133 | -0.028684 | 0.014371 | 0.022853 | 0.0014615 | 0.017918 | 0.035714 | -0.00069113 | -0.017319 | -0.010049 | -0.0054567 | 0.0095216 | 0.015642 | 0.027739 | -0.013428 | -0.026059 | -0.049974
b11 | 0.60887 | 0.74011 | 0.74386 | 0.7518 | 0.75414 | 0.75328 | 0.75293 | 0.74468 | 0.74769 | 0.74188 | 0.75862 | 0.74588 | 0.75136 | 0.76241 | 0.76132 | 0.75491 | 0.74809 | 0.74661 | 0.73876 | 0.63371 |  | b11 | 0.62291 | 0.73819 | 0.74269 | 0.74925 | 0.74617 | 0.75565 | 0.74917 | 0.74246 | 0.74829 | 0.74136 | 0.75598 | 0.74724 | 0.74581 | 0.75644 | 0.75682 | 0.75145 | 0.74993 | 0.75027 | 0.74287 | 0.64047
b12 | 0.0040542 | 0.0053833 | -0.0033104 | -0.0048909 | -0.0080578 | -0.0011277 | -0.0045849 | -0.0005327 | 0.0031255 | 0.0032091 | -0.0052444 | -0.0025945 | -0.0015454 | -0.0030847 | -0.0013004 | 0.00080543 | 0.0043735 | 0.0046942 | 0.0014471 | -0.017371 |  | b12 | 0.021095 | 0.0013012 | -0.0069975 | -0.0087688 | -0.0039901 | 0.00066214 | -0.0018134 | -0.0029141 | -0.0033051 | 0.0032048 | -0.0012578 | -0.0047991 | -0.0026191 | -0.0041783 | -0.00083717 | -0.0062637 | 0.00063719 | -0.0010156 | -0.0046788 | -0.0068753
b13 | 0.085078 | 0.064134 | 0.066874 | 0.064915 | 0.06597 | 0.063892 | 0.064064 | 0.063798 | 0.060532 | 0.062789 | 0.063247 | 0.063851 | 0.060464 | 0.060287 | 0.063629 | 0.065214 | 0.065692 | 0.066878 | 0.067206 | 0.059378 |  | b13 | 0.086675 | 0.064058 | 0.066891 | 0.06553 | 0.070778 | 0.066873 | 0.063783 | 0.065854 | 0.065925 | 0.06373 | 0.066364 | 0.067727 | 0.064459 | 0.063877 | 0.064466 | 0.066682 | 0.066741 | 0.06828 | 0.068413 | 0.053286
b14 | -0.0058285 | 0.00034813 | -0.0030637 | -0.0035209 | -0.0038251 | -0.0015743 | -0.003032 | -0.0022755 | 0.00044603 | 0.0013274 | -0.0021472 | -0.0013341 | -0.0012813 | -0.0016352 | -0.002183 | -0.0015801 | -0.00032316 | 0.00084122 | -0.0011905 | -0.012373 |  | b14 | -0.0056198 | -0.0039015 | -0.0044042 | -0.0056586 | -0.0048232 | -0.0035101 | -0.0029882 | -0.0030016 | -0.0028674 | -0.0020337 | -0.0026616 | -0.004209 | -0.0033657 | -0.0031272 | -0.0028924 | -0.0045642 | -0.0026025 | -0.0042617 | -0.004412 | -0.0021621
b15 | -0.013718 | 0.025489 | 0.025118 | 0.023322 | 0.025061 | 0.023215 | 0.023209 | 0.022334 | 0.023915 | 0.023012 | 0.023113 | 0.023448 | 0.025275 | 0.024416 | 0.025382 | 0.025658 | 0.025262 | 0.024715 | 0.023213 | 0.020115 |  | b15 | -0.014663 | 0.024057 | 0.022693 | 0.02027 | 0.020609 | 0.022864 | 0.022904 | 0.021212 | 0.022057 | 0.022263 | 0.020889 | 0.021945 | 0.021791 | 0.023477 | 0.022922 | 0.02128 | 0.023324 | 0.02331 | 0.020755 | 0.0028013
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 51.3447 | -1.9963 | -1.8253 | -3.8468 | -2.2079 | 1.6211 | 3.1189 | 3.9965 | 1.2884 | -1.1731 | -1.3653 | 2.2503 | 5.2233 | 4.9129 | -0.10951 | -0.57441 | -6.9671 | -5.9526 | -11.259 | -24.3392 |  | a1 | 65.1251 | -5.6735 | -6.7913 | -8.4997 | -8.4668 | -7.298 | -4.257 | -3.4783 | -3.7671 | -4.4264 | -4.6506 | 3.7426 | 9.0157 | 10.8967 | 7.9997 | 3.8773 | 2.7075 | -0.36305 | -1.2865 | -25.6695
a2 | 10.075 | 1.5058 | 0.58538 | 1.8596 | 0.63649 | 1.1801 | 1.96 | 1.1233 | 1.7723 | 0.049157 | -0.39004 | 1.0861 | 0.60098 | 0.93102 | 1.0656 | 0.13969 | 1.4245 | 1.6446 | 0.51753 | -1.1662 |  | a2 | 21.8431 | 0.44079 | 0.29708 | 0.65095 | -0.44176 | 0.69015 | -0.1396 | 1.0254 | 0.42564 | 0.93525 | 1.7924 | 0.24845 | 0.6799 | 0.10266 | 0.32428 | -0.51625 | 0.75567 | 1.4688 | 0.5273 | -6.0018
a3 | 0.57887 | 0.47586 | 0.63845 | 0.75186 | 0.98955 | 0.33563 | -0.012648 | 0.02833 | 0.12922 | 0.32406 | 0.56564 | 0.52854 | 0.18332 | 0.87213 | 0.81506 | 0.6113 | 0.79727 | 1.0281 | 0.81567 | -0.023976 |  | a3 | 1.3641 | 0.9949 | 0.77125 | 0.41906 | 1.1046 | 0.79692 | 0.38909 | 0.66791 | 0.38545 | 1.0066 | 1.0953 | 0.69343 | 0.67237 | 0.88764 | 0.71501 | 0.84099 | 0.99257 | 1.2282 | 1.0634 | 0.8309
a4 | 1.5205 | -0.4872 | -0.35775 | -0.37706 | -0.41337 | -0.73398 | -1.2915 | -1.0669 | -1.0258 | -0.4012 | 0.10133 | -0.33041 | -0.75221 | -0.32566 | -0.58271 | -0.45037 | -0.41346 | -0.24249 | 0.048965 | -1.2524 |  | a4 | 2.5657 | 0.048459 | -0.20031 | 0.25173 | 0.26951 | -0.14558 | 0.15375 | -0.24292 | -0.66509 | 0.08532 | -0.0041998 | -0.13688 | -0.17069 | 0.41746 | -0.074157 | 0.0027451 | 0.085885 | -0.11886 | 0.19406 | -0.13913
a5 | 3.3241 | -0.021727 | 0.066711 | 0.28034 | 0.26776 | 0.026294 | -0.12643 | 0.025881 | 0.037704 | -0.12478 | -0.010454 | 0.15417 | -0.067465 | 0.086094 | 0.16247 | 0.18894 | 0.15792 | 0.23975 | 0.22419 | -0.16673 |  | a5 | 3.8386 | 0.10299 | 0.092594 | 0.11314 | 0.33489 | 0.092766 | 0.043515 | 0.093834 | 0.11058 | 0.232 | 0.096662 | -0.04531 | -0.033394 | 0.15331 | 0.073945 | 0.034383 | 0.089375 | 0.19429 | 0.20694 | 0.14595
a6 | 0.34865 | 0.031276 | 0.084106 | 0.052194 | 0.13079 | 0.059455 | -0.077347 | 0.02825 | 0.089601 | 0.1135 | 0.04278 | 0.10088 | -0.01078 | -0.083295 | 0.03571 | 0.006287 | 0.11225 | 0.070373 | 0.10224 | 0.43036 |  | a6 | 0.83994 | 0.016454 | 0.053933 | -0.042151 | 0.095048 | 0.006659 | 0.0093837 | -0.023024 | -0.14831 | 0.00050515 | -0.11602 | -0.14438 | -0.26939 | -0.083563 | -0.12955 | -0.078018 | -0.16903 | -0.073763 | -0.029857 | 0.025091
a7 | 1.6642 | -0.070566 | -0.062332 | -0.17362 | -0.086172 | -0.10796 | -0.14112 | -0.096491 | -0.11762 | -0.12162 | -0.053375 | -0.082405 | -0.061674 | -0.083152 | -0.1035 | 0.011527 | -0.034314 | -0.039845 | -0.10688 | -0.16434 |  | a7 | 1.7069 | -0.097316 | -0.099208 | -0.18923 | -0.13685 | -0.14245 | -0.054275 | -0.031555 | -0.082596 | -0.15983 | -0.0086173 | -0.0098931 | -0.10604 | -0.07881 | -0.083643 | 0.0067603 | -0.041608 | -0.094251 | -0.07763 | -0.11478
a8 | -0.0094225 | -0.045277 | -0.045312 | -0.02027 | -0.04041 | -0.023656 | -0.035319 | -0.043253 | -0.021118 | -0.017891 | -0.054578 | 0.022813 | -0.0062983 | -0.015593 | -0.014952 | -0.030909 | -0.0087023 | -0.021767 | 0.032981 | 0.23087 |  | a8 | -0.00017644 | 0.0013431 | -0.010782 | -0.0064237 | -0.010451 | -0.029319 | -0.048745 | 0.0020866 | -0.051069 | 0.0061411 | -0.0044254 | 0.047369 | -0.02694 | -0.022745 | -0.053008 | -0.062077 | 0.0070678 | -0.0039572 | 0.021663 | -0.0035759
a9 | -0.23079 | -0.039331 | -0.039751 | -0.031249 | -0.012193 | -0.057957 | -0.096332 | -0.05774 | -0.049958 | -0.032236 | -0.046422 | -0.030222 | -0.058679 | -0.016044 | -0.0074098 | -0.041245 | -0.026668 | -0.031201 | -0.029345 | -0.19273 |  | a9 | -0.20449 | -0.046529 | -0.069947 | -0.047042 | -0.027711 | -0.049567 | -0.07143 | -0.049027 | -0.078569 | -0.031253 | -0.04517 | -0.061375 | -0.07023 | -0.040157 | -0.054797 | -0.093058 | -0.065947 | -0.050043 | -0.046103 | -0.075952
a10 | 0.024379 | -0.053555 | -0.0040898 | -0.016619 | 0.00158 | -0.017454 | -0.027859 | 0.0036776 | -0.020881 | -0.0015788 | -0.0050789 | 0.030379 | -0.012788 | -0.047717 | -0.045347 | -0.033839 | 0.0036573 | -0.023311 | 0.020744 | 0.37869 |  | a10 | 0.14278 | -0.00015874 | 0.0095258 | -0.003259 | 0.015384 | -0.032174 | 0.0047678 | 0.01779 | -0.038708 | 0.0157 | -0.028866 | 0.055556 | -0.030567 | -0.027617 | -0.03555 | -0.01127 | 0.0041331 | 0.012906 | 0.023201 | -0.050305
a11 | 0.16426 | -0.090793 | -0.085344 | -0.090485 | -0.085879 | -0.082954 | -0.081788 | -0.081655 | -0.082393 | -0.092588 | -0.073834 | -0.080936 | -0.079222 | -0.077368 | -0.087601 | -0.071499 | -0.083508 | -0.08364 | -0.095305 | -0.12568 |  | a11 | 0.16581 | -0.11799 | -0.11382 | -0.12196 | -0.11455 | -0.11428 | -0.10836 | -0.10254 | -0.10655 | -0.11056 | -0.10009 | -0.10106 | -0.10428 | -0.099172 | -0.10313 | -0.095569 | -0.10426 | -0.10733 | -0.10555 | -0.11725
a12 | -0.0067591 | -0.0039746 | 0.001203 | -0.00098223 | 0.0009073 | -0.0018113 | -0.0036248 | 0.00053715 | 0.0025251 | -0.00013663 | -0.00083512 | 0.003185 | -0.0004559 | -0.0058106 | -0.0039306 | -0.0069098 | 0.00067348 | -0.0021051 | 0.0020798 | 0.037537 |  | a12 | 0.0061979 | -0.0010442 | 0.0016253 | -0.001791 | 0.0022499 | -0.0033611 | -0.0012058 | -0.0024912 | -0.0084236 | 1.2497e-05 | -0.0036401 | 0.0035752 | -0.0095782 | -0.0062708 | -0.0054945 | -0.0024859 | -0.0059312 | -0.0039757 | -0.0020295 | -0.0053528
a13 | -0.024299 | -0.0059834 | -0.0069176 | -0.0077722 | -0.0054635 | -0.0079593 | -0.010833 | -0.0063165 | -0.0071037 | -0.0069434 | -0.0066579 | -0.0062844 | -0.0077457 | -0.0041939 | -0.0052759 | -0.0053723 | -0.0045305 | -0.0044163 | -0.0048614 | -0.0091804 |  | a13 | -0.022499 | -0.007299 | -0.01097 | -0.010264 | -0.0072981 | -0.0093249 | -0.0097444 | -0.0068294 | -0.0099707 | -0.0096431 | -0.0076767 | -0.0095829 | -0.0098595 | -0.0066717 | -0.0079564 | -0.010112 | -0.0087568 | -0.007685 | -0.0071402 | -0.010281
a14 | -0.0058562 | -0.0041765 | -0.0025809 | -0.0040666 | -0.0041698 | -0.0057124 | -0.0062171 | -0.0036857 | -0.0050536 | -0.0022647 | -0.0018725 | 0.00066581 | -0.0019017 | -0.0043852 | -0.0032609 | -0.0044961 | -0.0027279 | -0.0025583 | 0.00046964 | -0.0038727 |  | a14 | -0.0077745 | -0.0014625 | -0.0018963 | -0.00096125 | -0.0013944 | -0.0050567 | -0.0036545 | -0.001267 | -0.0053063 | -0.00055226 | -0.0016962 | 0.0020491 | -0.0016249 | -0.0013393 | -0.00395 | -0.0027856 | 0.0016593 | 0.00013909 | 0.00027384 | 0.0005069
a15 | 0.0065521 | -0.0028113 | -0.0023314 | -0.0024813 | -0.0016647 | -0.0028136 | -0.0035186 | -0.0013683 | -0.0026309 | -0.0022751 | -0.0026693 | -0.0016057 | -0.0021291 | -0.0018862 | -0.0018327 | -0.0026864 | -0.0024822 | -0.0019373 | -0.0023664 | 0.027916 |  | a15 | 0.003958 | -0.0053864 | -0.0042997 | -0.005731 | -0.0023376 | -0.0043383 | -0.0042937 | -0.0024136 | -0.0046655 | -0.0027061 | -0.0034262 | -0.00089975 | -0.0041494 | -0.0046168 | -0.003805 | -0.0037873 | -0.0021508 | -0.0028041 | -0.0024722 | -0.00090168
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.076504 | 0.081977 | -0.051152 | -0.065771 | -0.033094 | 0.0057893 | -0.086179 | -0.043279 | 0.087562 | 0.12522 | -0.019074 | -0.003868 | 0.0076544 | -0.021085 | -0.027419 | 0.042098 | 0.075723 | 0.14988 | 0.047411 | -0.63273 |  | Dx (mm) | 0.27797 | 0.034694 | -0.040379 | -0.11185 | -0.069103 | 0.042311 | 0.049263 | -0.0022366 | 0.052171 | 0.074863 | 0.0069189 | -0.055481 | -0.012911 | -0.0040567 | 0.031681 | 0.038015 | 0.060412 | -0.033916 | -0.065848 | -0.10469
Dy (mm) | 0.047429 | -0.11296 | -0.015216 | -0.045497 | -0.00020684 | -0.038754 | -0.072262 | 0.0036498 | -0.037828 | 0.01201 | -0.013238 | 0.068819 | -0.028127 | -0.10854 | -0.10441 | -0.072216 | 0.016764 | -0.036288 | 0.053851 | 0.89041 |  | Dy (mm) | 0.31569 | 0.0051797 | 0.015616 | -0.025601 | 0.024441 | -0.065985 | 0.017944 | 0.040424 | -0.08051 | 0.047167 | -0.063996 | 0.11889 | -0.07148 | -0.062596 | -0.075537 | -0.020661 | 0.017768 | 0.024392 | 0.043737 | -0.15269
Mag. len. (m) | 14.4197 | [20 empty] | Mag. len. (m) | 14.4204
TF (mT/KA) | 595.663 | [20 empty] | TF (mT/KA) | 595.8634
Coil pos. (mm) | 0.016953 | [20 empty] | Coil pos. (mm) | 0.017837
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.026058 | [20 empty] | Twist (m2 rad) | 0.0017074

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 595.663 | 0.11187 | -0.0013715 |  | -0.82702 | 4.3097 |  |  | 595.8634 | 0.10601 | -0.0096304 |  | -1.1506 | 6.1336 |  | 0 | 4
2 | -0.40738 |  | -0.29929 | 0.25696 | 1.1626 |  | 0.96345 | 0.65356 | -0.21057 |  | -0.10999 | 0.43615 | 0.88028 |  | 0.50973 | 0.58505 |  | 0 | 5
3 | -2.6323 |  | -4.1525 | 1.0525 | 0.52274 |  | 0.54855 | 0.32525 | -2.5587 |  | -4.0299 | 0.96807 | 0.83731 |  | 0.81916 | 0.25312 |  | 0 | 6
4 | -0.12102 |  | -0.08293 | 0.065434 | -0.44278 |  | -0.50562 | 0.35963 | -0.031741 |  | 0.052354 | 0.12746 | 0.057426 |  | -0.013062 | 0.24584 |  | 0 | 7
5 | -0.28865 |  | -0.016452 | 0.19517 | 0.18018 |  | 0.086292 | 0.12951 | -0.38477 |  | -0.14263 | 0.15763 | 0.22129 |  | 0.11071 | 0.091401 |  | 0 | 8
6 | -0.011486 |  | -0.0051522 | 0.047109 | 0.06738 |  | 0.06089 | 0.050252 | -0.0013662 |  | 0.015881 | 0.061313 | -0.044472 |  | -0.060127 | 0.085184 |  | 0 | 9
7 | 1.1583 |  | 1.1027 | 0.059027 | -0.031737 |  | -0.085048 | 0.043397 | 1.2363 |  | 1.1821 | 0.073607 | -0.030476 |  | -0.082391 | 0.054058 |  | 0 | 10
8 | -0.0085464 |  | -0.0057876 | 0.017896 | -0.013669 |  | -0.015139 | 0.020366 | -0.018555 |  | -0.015719 | 0.018657 | -0.011544 |  | -0.012195 | 0.020499 |  | 0 | 11
9 | 0.45671 |  | 0.45958 | 0.013908 | -0.04321 |  | -0.039053 | 0.020927 | 0.45468 |  | 0.45808 | 0.013271 | -0.061379 |  | -0.055091 | 0.016658 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.74227 |  | 0.74987 | 0.0070258 | -0.076698 |  | -0.083717 | 0.0062076 | 0.74105 |  | 0.74836 | 0.0056758 | -0.099711 |  | -0.10725 | 0.0070645 |  | 0 | 14
12 | -0.0001138 |  | -0.0012874 | 0.0020305 | -1.3163e-05 |  | 0.0002635 | 0.0015637 | -0.00024847 |  | -0.0025568 | 0.0021208 | -0.00026392 |  | -0.0025277 | 0.002253 |  | 0 | 15
13 | 0.0064357 |  | 0.06415 | 0.0020745 | -0.00067924 |  | -0.0063857 | 0.0017518 | 0.0066324 |  | 0.06617 | 0.0018456 | -0.00090659 |  | -0.008688 | 0.0013843 |  | 0 | 16
14 | -0.0001129 |  | -0.0016945 | 0.00055829 | -0.00027315 |  | -0.0025875 | 0.0014171 | -0.00034635 |  | -0.0035916 | 0.00062438 | -0.00015071 |  | -0.0014815 | 0.0014111 |  | 0 | 17
15 | 0.002289 |  | 0.024176 | 0.0010784 | -0.00010556 |  | -0.002305 | 0.00053033 | 0.0020425 |  | 0.022146 | 0.001121 | -0.00032615 |  | -0.0035713 | 0.0012581 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.89207 | -0.68292
3 | [12 empty] | 3 | 9 | -0.53602 | -0.70317
4 | [12 empty] | 4 | 16 | -1.8747 | -1.7315
5 | [12 empty] | 5 | 25 | -1.839 | -2.12
6 | [12 empty] | 6 | 36 | -3.023 | -2.6274
7 | [12 empty] | 7 | 49 | -2.9836 | -2.7634
8 | [12 empty] | 8 | 64 | -3.9585 | -3.9345
9 | [12 empty] | 9 | 81 | -4.071 | -4.2085
10 | [12 empty] | sigma (mm) |  | 0.016953 | 0.017837
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 5.4307 | 1.3577 | 0.32584 |  | Scaling law constants
2 | 3.341 | 0.83525 | 0.20046 |  | alpha | 87.202
3 | 2.0304 | 0.5076 | 0.12182 |  | beta | 0.6266
4 | 1.2189 | 0.30473 | 0.073135 |  | gamma | 0.9939
5 | 0.72285 | 0.18071 | 0.043371
6 | 0.42345 | 0.10586 | 0.025407
7 | 0.24505 | 0.061262 | 0.014703
8 | 0.14008 | 0.03502 | 0.0084049
9 | 0.079104 | 0.019776 | 0.0047462
10 | 0.044126 | 0.011032 | 0.0026476
11 | 0.024316 | 0.0060789 | 0.0014589
12 | 0.013236 | 0.003309 | 0.00079416
13 | 0.0071173 | 0.0017793 | 0.00042704
14 | 0.0037806 | 0.00094515 | 0.00022684
15 | 0.0019838 | 0.00049594 | 0.00011903
16 | 0.0010283 | 0.00025706 | 6.1695e-05
17 | 0.0005265 | 0.00013163 | 3.159e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 348.4073 | 595.8518 | 595.5838 | 595.6078 | 595.5697 | 595.6063 | 595.7286 | 595.5798 | 595.5939 | 595.6055 | 595.6348 | 595.6266 | 595.5674 | 595.6226 | 595.6496 | 595.657 | 595.7557 | 595.6926 | 596.0006 | 370.9308
2 | -2.798 | -0.11335 | -0.45129 | -0.51534 | -0.63497 | -0.0079046 | 0.18206 | -0.39233 | -0.17461 | -0.33332 | -0.3865 | 0.13246 | -0.75621 | -0.26901 | -0.065632 | -0.33713 | -0.23075 | -0.46917 | -0.56428 | -1.2864 | -0.40738
3 | 40.8015 | -4.3031 | -3.8605 | -5.0237 | -4.4506 | -4.7334 | -5.0666 | -3.8552 | -4.0719 | -2.5555 | -4.784 | -3.7724 | -4.5567 | -6.6715 | -5.2279 | -3.1762 | -3.009 | -2.8116 | -2.815 | 0.51156 | -2.6323
4 | -1.4891 | -0.016324 | -0.11684 | -0.075021 | -0.092169 | -0.12794 | -0.11426 | -0.10145 | -0.1236 | 0.046988 | -0.021061 | 0.056394 | -0.14482 | -0.16192 | -0.11812 | -0.050372 | -0.13024 | -0.043236 | -0.15875 | 0.06318 | -0.12102
5 | -4.9512 | -0.36515 | -0.40402 | -0.12376 | 0.048211 | -0.014179 | 0.17161 | -0.024071 | 0.087937 | -0.13617 | 0.25554 | -0.058664 | 0.23854 | 0.28105 | 0.080196 | 0.028578 | -0.0031674 | -0.1269 | -0.23171 | -3.777 | -0.28865
6 | -0.27135 | -0.027769 | -0.039778 | -0.0035123 | -0.032721 | 0.044052 | 0.10949 | -0.037906 | -0.034433 | 0.074799 | 0.056061 | -0.01753 | 0.0092544 | -0.057533 | -0.049097 | -0.019142 | -0.041386 | 0.0091087 | -0.034701 | 0.049523 | -0.011486
7 | 2.9626 | 0.95293 | 0.99204 | 1.0877 | 1.0965 | 1.1105 | 1.2106 | 1.1321 | 1.1413 | 1.1062 | 1.1732 | 1.1007 | 1.0884 | 1.1651 | 1.0992 | 1.0989 | 1.0729 | 1.1182 | 1.1029 | 1.0706 | 1.1583
8 | -0.073956 | 0.020509 | -0.0025997 | -0.0361 | -0.015848 | -0.024319 | -0.021487 | -0.020265 | -0.0031228 | -0.02439 | -0.032095 | 0.0094684 | -0.0072514 | 0.011688 | 0.019401 | 0.015886 | 0.0065153 | -0.0030326 | 0.0028678 | -0.026888 | -0.0085464
9 | 0.34602 | 0.46023 | 0.44505 | 0.44739 | 0.45736 | 0.44023 | 0.44253 | 0.46682 | 0.46933 | 0.47851 | 0.44357 | 0.44253 | 0.46219 | 0.47217 | 0.45068 | 0.47844 | 0.46149 | 0.47707 | 0.47677 | 0.4777 | 0.45671
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.60846 | 0.74012 | 0.74374 | 0.75162 | 0.75397 | 0.75333 | 0.75285 | 0.74466 | 0.74745 | 0.74162 | 0.75856 | 0.74602 | 0.75138 | 0.76278 | 0.76156 | 0.75521 | 0.74789 | 0.7462 | 0.73879 | 0.64823 | 0.74227
12 | -0.013538 | 0.021492 | -0.0082143 | -0.016276 | -0.06516 | -0.011711 | -0.0013518 | 0.014001 | -0.0042621 | -0.023999 | -0.043306 | -0.027254 | -0.017183 | -0.018661 | 0.0031992 | -0.0085862 | 0.0080859 | -0.022682 | -0.0098685 | 0.033794 | -0.001138 |  | ten times!
13 | 0.85206 | 0.64473 | 0.66784 | 0.64879 | 0.65874 | 0.64068 | 0.64207 | 0.63713 | 0.60648 | 0.62641 | 0.63235 | 0.63883 | 0.60513 | 0.60625 | 0.63843 | 0.65513 | 0.65676 | 0.66852 | 0.67268 | 0.50754 | 0.064357 |  | ten times!
14 | -0.047083 | -0.011111 | -0.019764 | -0.021647 | -0.031418 | -0.015952 | -0.011755 | -0.014836 | -0.011965 | -0.010682 | -0.01755 | -0.013504 | -0.013913 | -0.010426 | -0.014523 | -0.023099 | -0.019327 | -0.021514 | -0.022018 | 0.18579 | -0.001129 |  | ten times!
15 | -0.13718 | 0.25489 | 0.25118 | 0.23322 | 0.25061 | 0.23215 | 0.23209 | 0.22334 | 0.23915 | 0.23012 | 0.23113 | 0.23448 | 0.25275 | 0.24416 | 0.25382 | 0.25658 | 0.25262 | 0.24715 | 0.23213 | 0.20115 | 0.02289 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 51.3061 | -2.0046 | -1.824 | -3.8408 | -2.2067 | 1.6206 | 3.1298 | 3.9994 | 1.2788 | -1.1732 | -1.366 | 2.2501 | 5.2218 | 4.9125 | -0.10803 | -0.57623 | -6.9732 | -5.9682 | -11.2585 | -24.3129 | -0.0013715
2 | 9.8423 | 1.4439 | 0.58231 | 1.8385 | 0.64023 | 1.1583 | 1.9166 | 1.1251 | 1.7528 | 0.047991 | -0.39622 | 1.1168 | 0.58572 | 0.84793 | 1.0039 | 0.10963 | 1.4233 | 1.6145 | 0.53083 | -1.2343 | 1.1626
3 | 0.57187 | 0.48243 | 0.63493 | 0.74693 | 0.98713 | 0.3355 | -0.033817 | 0.020247 | 0.14426 | 0.33283 | 0.56593 | 0.52766 | 0.18362 | 0.86784 | 0.81013 | 0.61405 | 0.80318 | 1.0342 | 0.81682 | -0.26969 | 0.52274
4 | 1.516 | -0.49646 | -0.35839 | -0.37403 | -0.41128 | -0.73414 | -1.2912 | -1.0666 | -1.0258 | -0.39715 | 0.10208 | -0.32929 | -0.75051 | -0.31812 | -0.57965 | -0.45175 | -0.41625 | -0.252 | 0.049423 | -0.46703 | -0.44278
5 | 3.3194 | -0.022582 | 0.067717 | 0.28092 | 0.26903 | 0.026704 | -0.12692 | 0.0263 | 0.035388 | -0.1294 | -0.010028 | 0.15463 | -0.067345 | 0.083362 | 0.16081 | 0.18881 | 0.15548 | 0.23738 | 0.22299 | -0.05468 | 0.18018
6 | 0.25415 | 0.071273 | 0.088308 | 0.065645 | 0.12986 | 0.074859 | -0.050752 | 0.025317 | 0.10848 | 0.11418 | 0.047901 | 0.074053 | 0.00019085 | -0.039309 | 0.075185 | 0.034111 | 0.10681 | 0.086823 | 0.083073 | 0.079539 | 0.06738
7 | 1.6656 | -0.067358 | -0.063372 | -0.17512 | -0.086725 | -0.10829 | -0.14365 | -0.097217 | -0.11662 | -0.12072 | -0.054002 | -0.082642 | -0.061728 | -0.08301 | -0.1031 | 0.012775 | -0.034203 | -0.038048 | -0.10784 | -0.085062 | -0.031737
8 | -0.0088434 | -0.019397 | -0.043078 | -0.011643 | -0.040555 | -0.015476 | -0.024149 | -0.045231 | -0.010682 | -0.018717 | -0.052231 | 0.0084626 | 2.8141e-05 | 0.0082328 | 0.0069866 | -0.013865 | -0.011403 | -0.011341 | 0.021553 | 0.024256 | -0.013669
9 | -0.23235 | -0.035233 | -0.040048 | -0.032586 | -0.012167 | -0.05789 | -0.099196 | -0.05762 | -0.048255 | -0.032483 | -0.046537 | -0.030218 | -0.058596 | -0.017403 | -0.0090442 | -0.039893 | -0.027118 | -0.028121 | -0.030539 | 0.014282 | -0.04321
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.16447 | -0.090189 | -0.085337 | -0.09067 | -0.085861 | -0.082975 | -0.082205 | -0.081639 | -0.08246 | -0.092602 | -0.073889 | -0.080813 | -0.079247 | -0.077664 | -0.087759 | -0.071273 | -0.083588 | -0.083326 | -0.09542 | -0.1003 | -0.076698
12 | -0.082953 | 0.014853 | 0.016715 | 0.0074169 | 0.007893 | -0.00030977 | -0.01016 | 0.0017982 | 0.045805 | -0.00055195 | -0.0033377 | 0.00066061 | 0.0078641 | -0.012542 | 0.0065697 | -0.034258 | 0.0013828 | 0.00075192 | -0.0031218 | -0.038837 | -0.00013163 |  | ten times!
13 | -0.23745 | -0.056915 | -0.070542 | -0.080992 | -0.055696 | -0.079806 | -0.11411 | -0.064321 | -0.067524 | -0.067387 | -0.06707 | -0.062122 | -0.077622 | -0.044003 | -0.055186 | -0.053148 | -0.043684 | -0.040997 | -0.048294 | -0.026295 | -0.0067924 |  | ten times!
14 | -0.057332 | -0.016156 | -0.023644 | -0.033271 | -0.042109 | -0.049581 | -0.050856 | -0.038016 | -0.041188 | -0.022577 | -0.016625 | -0.0066822 | -0.013028 | -0.022356 | -0.011197 | -0.02877 | -0.029219 | -0.015807 | -0.0046744 | -0.04076 | -0.0027315 |  | ten times!
15 | 0.06552 | -0.028113 | -0.023314 | -0.024813 | -0.016647 | -0.028136 | -0.035186 | -0.013683 | -0.026309 | -0.022751 | -0.026693 | -0.016057 | -0.021291 | -0.018862 | -0.018327 | -0.026864 | -0.024822 | -0.019373 | -0.023664 | 0.27916 | -0.0010556 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 337.556 | 596.0352 | 595.7192 | 595.8297 | 595.8045 | 595.868 | 595.8583 | 595.8085 | 595.8169 | 595.8069 | 595.7497 | 595.8396 | 595.7794 | 595.843 | 595.8399 | 595.8812 | 595.9613 | 595.9353 | 596.1653 | 375.4542
2 | -4.0667 | 0.40088 | -0.21891 | -0.50732 | -0.60725 | -0.074567 | 0.27053 | -0.18952 | -0.46659 | -0.26042 | 0.77919 | -0.59591 | 0.63048 | 0.3863 | -0.41518 | -0.3651 | -0.40173 | 0.12858 | -0.47335 | 0.38335 | -0.21057
3 | 40.914 | -4.6594 | -4.0759 | -6.031 | -3.9538 | -4.7959 | -3.6407 | -3.8128 | -3.0739 | -3.7399 | -4.3392 | -3.818 | -3.502 | -6.0777 | -4.6919 | -3.847 | -3.0302 | -3.076 | -2.3727 | 0.38266 | -2.5587
4 | -2.7053 | 0.048754 | -0.034041 | 0.052864 | 0.014178 | 0.074468 | 0.2375 | 0.038366 | -0.10261 | -0.045177 | 0.22315 | -0.1317 | 0.22067 | 0.07888 | -0.015783 | 0.34205 | 0.013323 | -0.015089 | -0.057434 | -0.030277 | -0.031741
5 | -4.4778 | -0.38761 | -0.51459 | -0.075204 | -0.18818 | -0.028124 | -0.037338 | -0.048804 | -0.15891 | -0.26542 | 0.045798 | -0.093802 | 0.0090022 | 0.064574 | 0.0041461 | -0.23385 | -0.13399 | -0.20535 | -0.31978 | -3.6217 | -0.38477
6 | -0.62592 | 0.039666 | 0.068785 | 0.030871 | 0.071461 | -0.040041 | 0.00028288 | -0.050215 | 0.048993 | 0.090481 | -0.015158 | 0.037825 | 0.072645 | 0.011058 | 0.1153 | -0.12146 | -0.078502 | -0.0049493 | 0.0088171 | 0.067518 | -0.0013662
7 | 3.1854 | 0.97268 | 1.0328 | 1.2671 | 1.247 | 1.2072 | 1.2011 | 1.2298 | 1.1974 | 1.1924 | 1.2037 | 1.2318 | 1.1829 | 1.1923 | 1.167 | 1.2622 | 1.1577 | 1.1737 | 1.1586 | 1.0339 | 1.2363
8 | -0.13632 | -0.0016145 | -0.04114 | -0.03432 | -0.035809 | -0.024111 | -0.030257 | -0.013932 | -0.026026 | -0.021917 | -0.023036 | -0.029859 | -0.0023625 | -0.0075384 | -0.023902 | -0.016812 | 0.017562 | 0.026012 | 0.0061289 | 0.006235 | -0.018555
9 | 0.38703 | 0.46108 | 0.45076 | 0.43451 | 0.46268 | 0.43325 | 0.46333 | 0.46026 | 0.47279 | 0.4664 | 0.45294 | 0.44329 | 0.46877 | 0.44551 | 0.45007 | 0.46901 | 0.45728 | 0.47337 | 0.4801 | 0.41853 | 0.45468
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.62038 | 0.73816 | 0.74252 | 0.74865 | 0.74603 | 0.75577 | 0.74921 | 0.7424 | 0.74884 | 0.7412 | 0.75614 | 0.74734 | 0.74623 | 0.75668 | 0.75711 | 0.75164 | 0.74984 | 0.75019 | 0.74261 | 0.64053 | 0.74105
12 | -0.0092543 | -0.0029428 | -0.052119 | -0.034096 | -0.006636 | -0.009008 | -0.041549 | -0.03005 | -0.051662 | -0.0048396 | -0.012351 | -0.029509 | -0.015342 | -0.037005 | -0.018546 | -0.079056 | -0.023187 | 0.0048678 | -0.017194 | -0.018293 | -0.0024847 |  | ten times!
13 | 0.85993 | 0.64156 | 0.66732 | 0.65065 | 0.70497 | 0.67242 | 0.63846 | 0.65809 | 0.66366 | 0.63826 | 0.66461 | 0.67735 | 0.64515 | 0.63931 | 0.64765 | 0.66859 | 0.66884 | 0.68172 | 0.682 | 0.53054 | 0.066324 |  | ten times!
14 | -0.012343 | -0.046119 | -0.037049 | -0.036707 | -0.036974 | -0.04071 | -0.039809 | -0.030429 | -0.035057 | -0.035114 | -0.026001 | -0.032944 | -0.028897 | -0.028108 | -0.032538 | -0.051659 | -0.037944 | -0.036669 | -0.033755 | -0.018072 | -0.0034635 |  | ten times!
15 | -0.14663 | 0.24057 | 0.22693 | 0.2027 | 0.20609 | 0.22864 | 0.22904 | 0.21212 | 0.22057 | 0.22263 | 0.20889 | 0.21945 | 0.21791 | 0.23477 | 0.22922 | 0.2128 | 0.23324 | 0.2331 | 0.20755 | 0.028013 | 0.020425 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 64.8197 | -5.6745 | -6.7904 | -8.4961 | -8.4677 | -7.3 | -4.2569 | -3.4778 | -3.7707 | -4.4297 | -4.6484 | 3.7475 | 9.0189 | 10.8981 | 7.9972 | 3.878 | 2.7053 | -0.36032 | -1.2833 | -25.7029 | -0.0096304
2 | 20.2902 | 0.43957 | 0.30823 | 0.6383 | -0.42141 | 0.64894 | -0.13418 | 1.0437 | 0.39406 | 0.94714 | 1.7589 | 0.30632 | 0.65145 | 0.058358 | 0.27991 | -0.52935 | 0.75495 | 1.4825 | 0.54775 | -5.9848 | 0.88028
3 | 1.4216 | 0.99456 | 0.76991 | 0.42428 | 1.1078 | 0.79888 | 0.38701 | 0.66755 | 0.39011 | 1.0059 | 1.0978 | 0.6948 | 0.67475 | 0.88883 | 0.71523 | 0.84221 | 0.99162 | 1.2275 | 1.0661 | 0.83006 | 0.83731
4 | 2.6561 | 0.04809 | -0.19753 | 0.25426 | 0.27605 | -0.14694 | 0.1534 | -0.24241 | -0.66951 | 0.084148 | -0.0037008 | -0.13499 | -0.17074 | 0.41853 | -0.074665 | 0.0012928 | 0.085175 | -0.11613 | 0.20055 | -0.26563 | 0.057426
5 | 3.8168 | 0.10274 | 0.092984 | 0.1116 | 0.33651 | 0.092192 | 0.043266 | 0.094422 | 0.11449 | 0.23026 | 0.096664 | -0.04816 | -0.03306 | 0.15335 | 0.077954 | 0.03462 | 0.092658 | 0.19369 | 0.20658 | 0.14776 | 0.22129
6 | 0.31829 | 0.015868 | 0.046822 | -0.038148 | 0.08094 | 0.036874 | 0.0027205 | -0.040595 | -0.11282 | -0.015118 | -0.088812 | -0.1962 | -0.24005 | -0.057347 | -0.097551 | -0.06891 | -0.17541 | -0.084994 | -0.049543 | 0.076602 | -0.044472
7 | 1.7144 | -0.097348 | -0.099071 | -0.19013 | -0.13664 | -0.1424 | -0.053146 | -0.031329 | -0.082077 | -0.15993 | -0.0092088 | -0.0069546 | -0.10641 | -0.079095 | -0.083548 | 0.0076527 | -0.042011 | -0.094497 | -0.076902 | -0.11618 | -0.030476
8 | -0.028933 | 0.00097814 | -0.015426 | -0.003629 | -0.016684 | -0.014904 | -0.051006 | -0.0067134 | -0.031271 | -0.0031329 | 0.0093326 | 0.021132 | -0.011635 | -0.0097253 | -0.036238 | -0.05585 | 0.0051121 | -0.01019 | 0.010346 | 0.022663 | -0.011544
9 | -0.23737 | -0.046569 | -0.069607 | -0.047928 | -0.026777 | -0.048345 | -0.071772 | -0.049038 | -0.076737 | -0.032767 | -0.045087 | -0.058653 | -0.070819 | -0.040397 | -0.053819 | -0.09266 | -0.06634 | -0.049638 | -0.04469 | -0.08278 | -0.061379
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.16038 | -0.11797 | -0.11371 | -0.12224 | -0.11438 | -0.11416 | -0.1083 | -0.10247 | -0.10644 | -0.11066 | -0.10013 | -0.10057 | -0.10448 | -0.099358 | -0.10306 | -0.095591 | -0.10403 | -0.1074 | -0.10551 | -0.11829 | -0.099711
12 | -0.087019 | -0.010997 | 0.0057522 | -0.014172 | 0.0067284 | 0.00032186 | -0.016749 | -0.043811 | -0.043098 | -0.015997 | -0.0060476 | -0.024839 | -0.064156 | -0.034675 | -0.018792 | -0.01242 | -0.063949 | -0.053354 | -0.044735 | -0.0036929 | -0.0026392 |  | ten times!
13 | -0.1949 | -0.072447 | -0.10976 | -0.10457 | -0.072816 | -0.093384 | -0.095657 | -0.067388 | -0.099355 | -0.095381 | -0.07798 | -0.091133 | -0.1006 | -0.068255 | -0.080278 | -0.10104 | -0.087981 | -0.076019 | -0.069789 | -0.10493 | -0.0090659 |  | ten times!
14 | -0.048685 | -0.014112 | -0.023311 | -0.010618 | -0.019422 | -0.03663 | -0.038188 | -0.019776 | -0.036434 | -0.012502 | -0.0057581 | -0.0014063 | -0.0038625 | -0.0014447 | -0.024248 | -0.023049 | 0.01425 | -0.0040747 | -0.006078 | 0.0078141 | -0.0015071 |  | ten times!
15 | 0.03958 | -0.053864 | -0.042997 | -0.05731 | -0.023376 | -0.043383 | -0.042937 | -0.024136 | -0.046655 | -0.027061 | -0.034262 | -0.0089975 | -0.041494 | -0.046168 | -0.03805 | -0.037873 | -0.021508 | -0.028041 | -0.024722 | -0.0090168 | -0.0032615 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00031695 | -0.00013297 | -9.2679e-05 | -0.00015664 | -9.5197e-05 | 0.00011012 | -0.00013969 | -0.00011601 | -9.654e-05 | -4.7352e-05 | -6.1118e-05 | -0.0001605 | -6.7833e-05 | -2.2505e-05 | -1.0082e-05 | 0.00015562 | 4.9683e-05 | 0.00056675 |  | 595.663
|  | 0.00028826 | -0.00024207 | -5.6622e-05 | -9.8913e-05 | 7.6546e-06 | -8.6243e-06 | -9.22e-05 | -7.8103e-05 | -9.4886e-05 | -0.00019088 | -4.0007e-05 | -0.00014104 | -3.4301e-05 | -3.9504e-05 | 2.9807e-05 | 0.00016423 | 0.0001206 | 0.00050659 |  | 595.8634
Check Computation for Aperture1
Magnetic Length
dB/B Heads CS+NCS | -3961.5368
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
dB/B (units) | 595.6301 | -4150.6096 | 3.722 | -0.77743 | -0.3745 | -1.0142 | -0.39968 | 1.6536 | -0.84459 | -0.60786 | -0.41311 | 0.078803 | -0.058866 | -1.0528 | -0.12602 | 0.32728 | 0.45152 | 2.1086 | 1.0492 | 6.2202 | -3772.4639
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20
D Angle (mrad) | -0.082702 | 5.2133 | -0.11776 | -0.099694 | -0.30138 | -0.13796 | 0.24477 | 0.39568 | 0.48265 | 0.21058 | -0.034615 | -0.053897 | 0.30771 | 0.60488 | 0.57395 | 0.0719 | 0.025079 | -0.61462 | -0.51411 | -1.0432 | -2.3486 |  | 0
db2 | -0.29929 | -2.798 | 0.18595 | -0.152 | -0.21605 | -0.33568 | 0.29139 | 0.48136 | -0.093033 | 0.12469 | -0.03403 | -0.087209 | 0.43175 | -0.45691 | 0.030283 | 0.23366 | -0.037839 | 0.068541 | -0.16988 | -0.26498 | -1.2864 |  | 0
db3 | -4.156 | 40.8015 | -0.15058 | 0.29203 | -0.87117 | -0.29816 | -0.58096 | -0.91412 | 0.29727 | 0.080618 | 1.597 | -0.63155 | 0.38013 | -0.40421 | -2.519 | -1.0754 | 0.97626 | 1.1435 | 1.3409 | 1.3375 | 0.51156 |  | 0
db4 | -0.08293 | -1.4891 | 0.066606 | -0.033912 | 0.0079094 | -0.0092393 | -0.045011 | -0.031326 | -0.018523 | -0.040671 | 0.12992 | 0.061869 | 0.13932 | -0.061887 | -0.078986 | -0.035194 | 0.032558 | -0.047306 | 0.039694 | -0.075825 | 0.06318 |  | 0
db5 | -0.016144 | -4.9512 | -0.34869 | -0.38757 | -0.10731 | 0.064663 | 0.0022727 | 0.18806 | -0.0076189 | 0.10439 | -0.11971 | 0.27199 | -0.042212 | 0.25499 | 0.2975 | 0.096648 | 0.04503 | 0.013284 | -0.11045 | -0.21526 | -3.777 |  | 0
db6 | -0.0051522 | -0.27135 | -0.022617 | -0.034626 | 0.0016399 | -0.027569 | 0.049204 | 0.11465 | -0.032753 | -0.02928 | 0.079951 | 0.061213 | -0.012377 | 0.014407 | -0.052381 | -0.043945 | -0.013989 | -0.036234 | 0.014261 | -0.029549 | 0.049523 |  | 0
db7 | 1.1028 | 2.9626 | -0.14981 | -0.1107 | -0.015058 | -0.0062212 | 0.0077576 | 0.10786 | 0.029317 | 0.038511 | 0.0035041 | 0.070417 | -0.0020847 | -0.014302 | 0.062403 | -0.0035671 | -0.0038604 | -0.029835 | 0.015488 | 0.00018765 | 1.0706 |  | -1.7887e-16
db8 | -0.0057876 | -0.073956 | 0.026297 | 0.0031879 | -0.030313 | -0.010061 | -0.018531 | -0.0157 | -0.014477 | 0.0026647 | -0.018603 | -0.026308 | 0.015256 | -0.0014638 | 0.017475 | 0.025188 | 0.021674 | 0.012303 | 0.002755 | 0.0086554 | -0.026888 |  | 0
db9 | 0.45982 | 0.34602 | 0.00065265 | -0.014525 | -0.012183 | -0.0022194 | -0.019342 | -0.017047 | 0.007248 | 0.0097488 | 0.018936 | -0.01601 | -0.017041 | 0.0026103 | 0.012598 | -0.0088981 | 0.018865 | 0.0019178 | 0.017494 | 0.017195 | 0.4777 |  | 1.2336e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.74991 | 0.60846 | -0.0097569 | -0.0061388 | 0.0017445 | 0.004098 | 0.0034528 | 0.0029732 | -0.0052105 | -0.0024206 | -0.0082534 | 0.0086856 | -0.0038582 | 0.0015021 | 0.012904 | 0.01169 | 0.0053332 | -0.0019874 | -0.0036696 | -0.011088 | 0.64823 |  | -4.3175e-17
db12 | -0.0012874 | -0.0013538 | 0.0034367 | 0.000466 | -0.00034012 | -0.0052285 | 0.00011635 | 0.0011523 | 0.0026875 | 0.00086122 | -0.0011125 | -0.0030432 | -0.001438 | -0.00043089 | -0.00057871 | 0.0016074 | 0.00042881 | 0.002096 | -0.00098076 | 0.00030058 | 0.0033794 |  | 0
db13 | 0.06415 | 0.085206 | 0.00032355 | 0.0026346 | 0.00072922 | 0.0017239 | -8.133e-05 | 5.7679e-05 | -0.00043724 | -0.0035015 | -0.0015085 | -0.00091521 | -0.00026723 | -0.0036372 | -0.0035251 | -0.00030655 | 0.0013635 | 0.0015264 | 0.0027024 | 0.0031185 | 0.050754 |  | 1.542e-18
db14 | -0.0016945 | -0.0047083 | 0.00058334 | -0.00028195 | -0.00047024 | -0.0014473 | 9.9292e-05 | 0.00051901 | 0.00021089 | 0.00049794 | 0.00062627 | -6.0564e-05 | 0.00034405 | 0.00030317 | 0.00065184 | 0.00024222 | -0.00061539 | -0.00023828 | -0.00045697 | -0.00050734 | 0.018579 |  | 2.0479e-19
db15 | 0.024176 | -0.013718 | 0.0013137 | 0.00094293 | -0.00085309 | 0.00088561 | -0.00096046 | -0.0009662 | -0.0018412 | -0.00026028 | -0.0011635 | -0.0010629 | -0.00072745 | 0.0010999 | 0.00024036 | 0.0012067 | 0.0014827 | 0.0010862 | 0.00053906 | -0.00096209 | 0.020115 |  | -1.7347e-18
da2 | 0.96345 | 9.8423 | 0.48046 | -0.38114 | 0.87504 | -0.32322 | 0.19483 | 0.95317 | 0.16164 | 0.78933 | -0.91546 | -1.3597 | 0.15336 | -0.37773 | -0.11553 | 0.04045 | -0.85383 | 0.45988 | 0.65105 | -0.43263 | -1.2343 |  | 0
da3 | 0.54855 | 0.57187 | -0.066117 | 0.086377 | 0.19838 | 0.43858 | -0.21305 | -0.58236 | -0.5283 | -0.40428 | -0.21572 | 0.01738 | -0.020888 | -0.36493 | 0.31929 | 0.26158 | 0.065501 | 0.25464 | 0.48565 | 0.26827 | -0.26969 |  | 0
da4 | -0.50562 | 1.516 | 0.0091533 | 0.14723 | 0.13159 | 0.094337 | -0.22852 | -0.78555 | -0.56102 | -0.52015 | 0.10847 | 0.60769 | 0.17632 | -0.2449 | 0.1875 | -0.074035 | 0.053867 | 0.089363 | 0.25362 | 0.55504 | -0.46703 |  | 0
da5 | 0.086292 | 3.3194 | -0.10887 | -0.018575 | 0.19463 | 0.18274 | -0.059588 | -0.21321 | -0.059992 | -0.050904 | -0.21569 | -0.09632 | 0.068337 | -0.15364 | -0.0029297 | 0.074517 | 0.10252 | 0.069191 | 0.15108 | 0.1367 | -0.05468 |  | 0
da6 | 0.06089 | 0.25415 | 0.010383 | 0.027419 | 0.0047559 | 0.068974 | 0.013969 | -0.11164 | -0.035573 | 0.047587 | 0.053292 | -0.012988 | 0.013164 | -0.060699 | -0.1002 | 0.014296 | -0.026778 | 0.045921 | 0.025933 | 0.022183 | 0.079539 |  | 8.4809e-18
da7 | -0.085048 | 1.6656 | 0.017689 | 0.021676 | -0.090072 | -0.0016769 | -0.023241 | -0.058598 | -0.012169 | -0.031575 | -0.03567 | 0.031046 | 0.0024061 | 0.02332 | 0.0020382 | -0.018052 | 0.097823 | 0.050844 | 0.047 | -0.022788 | -0.085062 |  | 0
da8 | -0.015139 | -0.0088434 | -0.0042582 | -0.027939 | 0.0034959 | -0.025416 | -0.00033645 | -0.0090096 | -0.030092 | 0.0044572 | -0.0035778 | -0.037092 | 0.023602 | 0.015167 | 0.023372 | 0.022126 | 0.0012743 | 0.003736 | 0.0037985 | 0.036692 | 0.024256 |  | 0
da9 | -0.039053 | -0.23235 | 0.0038194 | -0.00099581 | 0.0064663 | 0.026886 | -0.018837 | -0.060144 | -0.018567 | -0.0092024 | 0.0065701 | -0.0074848 | 0.0088349 | -0.019543 | 0.021649 | 0.030008 | -0.00083995 | 0.011935 | 0.010931 | 0.0085133 | 0.014282 |  | 0
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.083717 | 0.16447 | -0.006471 | -0.0016198 | -0.0069528 | -0.0021432 | 0.00074242 | 0.0015129 | 0.0020786 | 0.0012579 | -0.0088842 | 0.0098288 | 0.0029048 | 0.00447 | 0.0060532 | -0.0040413 | 0.012445 | 0.00012944 | 0.0003915 | -0.011702 | -0.1003 |  | 6.9389e-18
da12 | 0.0002635 | -0.0082953 | 0.0012219 | 0.001408 | 0.00047819 | 0.00052581 | -0.00029447 | -0.0012795 | -8.3678e-05 | 0.004317 | -0.00031869 | -0.00059726 | -0.00019744 | 0.00052292 | -0.0015177 | 0.00039347 | -0.0036893 | -0.00012521 | -0.0001883 | -0.00057568 | -0.0038837 |  | 0
da13 | -0.0063857 | -0.023745 | 0.00069412 | -0.00066855 | -0.0017135 | 0.00081603 | -0.001595 | -0.0050251 | -4.6474e-05 | -0.00036678 | -0.00035305 | -0.00032131 | 0.00017341 | -0.0013765 | 0.0019853 | 0.00086705 | 0.0010709 | 0.0020173 | 0.0022859 | 0.0015563 | -0.0026295 |  | -9.6374e-19
da14 | -0.0025875 | -0.0057332 | 0.00097197 | 0.00022313 | -0.00073958 | -0.0016234 | -0.0023705 | -0.0024981 | -0.0012141 | -0.0015313 | 0.00032984 | 0.00092505 | 0.0019193 | 0.0012847 | 0.00035192 | 0.0014678 | -0.00028949 | -0.00033433 | 0.0010069 | 0.0021201 | -0.004076 |  | 0
da15 | -0.002305 | 0.0065521 | -0.00050629 | -2.6417e-05 | -0.00017626 | 0.00064037 | -0.00050862 | -0.0012136 | 0.00093671 | -0.00032588 | 2.9943e-05 | -0.00036424 | 0.00069928 | 0.00017596 | 0.0004188 | 0.00047237 | -0.00038137 | -0.00017717 | 0.00036773 | -6.1327e-05 | 0.027916 |  | -1.5661e-19
Check Computation for Aperture2
Magnetic Length
dB/B Heads CS+NCS | -4016.7027
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.8338 | -4334.7293 | 3.3795 | -1.924 | -0.069441 | -0.49238 | 0.57336 | 0.41056 | -0.42524 | -0.28427 | -0.4521 | -1.4121 | 0.096713 | -0.91364 | 0.15378 | 0.10175 | 0.79489 | 2.1392 | 1.7029 | 5.563 | -3698.6761
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.11506 | 6.597 | -0.45239 | -0.56398 | -0.73455 | -0.73171 | -0.61494 | -0.31063 | -0.23272 | -0.26201 | -0.32791 | -0.34978 | 0.48981 | 1.017 | 1.2049 | 0.91478 | 0.50286 | 0.38559 | 0.079028 | -0.013272 | -2.4552 |  | 0
db2 | 0.10999 | -4.0667 | 0.51088 | -0.10892 | -0.39733 | -0.49725 | 0.035426 | 0.38052 | -0.079525 | -0.35659 | -0.15043 | 0.88918 | -0.48592 | 0.74048 | 0.49629 | -0.30519 | -0.2551 | -0.29174 | 0.23858 | -0.36335 | 0.38335 |  | 0
db3 | -4.03 | 40.914 | -0.6295 | -0.046007 | -2.0011 | 0.076105 | -0.766 | 0.38921 | 0.21712 | 0.95596 | 0.29002 | -0.30934 | 0.2119 | 0.52785 | -2.0478 | -0.66201 | 0.18292 | 0.99967 | 0.95385 | 1.6572 | 0.38266 |  | 3.9475e-16
db4 | -0.052354 | -2.7053 | -0.0036 | -0.086395 | 0.00050989 | -0.038176 | 0.022114 | 0.18515 | -0.013987 | -0.15497 | -0.097531 | 0.17079 | -0.18405 | 0.16832 | 0.026526 | -0.068136 | 0.28969 | -0.039031 | -0.067442 | -0.10979 | -0.030277 |  | 0
db5 | -0.14306 | -4.4778 | -0.24498 | -0.37196 | 0.06743 | -0.045549 | 0.11451 | 0.1053 | 0.093831 | -0.016275 | -0.12278 | 0.18843 | 0.048833 | 0.15164 | 0.20721 | 0.14678 | -0.091212 | 0.0086484 | -0.062715 | -0.17714 | -3.6217 |  | 0
db6 | -0.015881 | -0.62592 | 0.023785 | 0.052904 | 0.01499 | 0.05558 | -0.055922 | -0.015598 | -0.066096 | 0.033112 | 0.0746 | -0.031039 | 0.021944 | 0.056764 | -0.0048228 | 0.099417 | -0.13734 | -0.094384 | -0.02083 | -0.007064 | 0.067518 |  | 0
db7 | 1.1819 | 3.1854 | -0.20941 | -0.14929 | 0.085056 | 0.064956 | 0.025077 | 0.019047 | 0.04775 | 0.015299 | 0.010331 | 0.021634 | 0.049688 | 0.00086323 | 0.010239 | -0.015077 | 0.080071 | -0.024358 | -0.0083535 | -0.023517 | 1.0339 |  | 1.3569e-16
db8 | 0.015719 | -0.13632 | 0.014104 | -0.025421 | -0.018602 | -0.020091 | -0.0083925 | -0.014538 | 0.0017864 | -0.010307 | -0.0061984 | -0.0073174 | -0.01414 | 0.013356 | 0.0081802 | -0.0081838 | -0.001093 | 0.03328 | 0.04173 | 0.021847 | 0.006235 |  | 0
db9 | 0.45796 | 0.38703 | 0.0030068 | -0.0073198 | -0.023569 | 0.0046038 | -0.02483 | 0.0052496 | 0.0021858 | 0.014709 | 0.0083249 | -0.0051408 | -0.014785 | 0.010691 | -0.012568 | -0.0080029 | 0.010927 | -0.00080079 | 0.015295 | 0.022021 | 0.41853 |  | 2.1588e-17
db10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
db11 | 0.74834 | 0.62038 | -0.010204 | -0.0058441 | 0.00028593 | -0.0023352 | 0.0074086 | 0.00085009 | -0.0059686 | 0.00047178 | -0.0071622 | 0.0077742 | -0.0010196 | -0.0021362 | 0.0083195 | 0.0087456 | 0.003275 | 0.0014727 | 0.0018222 | -0.0057555 | 0.64053 |  | -3.7007e-17
db12 | 0.0025568 | -0.00092543 | 0.0022625 | -0.0026551 | -0.00085279 | 0.0018932 | 0.001656 | -0.001598 | -0.0004482 | -0.0026094 | 0.0020728 | 0.0013217 | -0.00039413 | 0.0010226 | -0.0011437 | 0.0007022 | -0.0053488 | 0.00023808 | 0.0030436 | 0.00083743 | -0.0018293 |  | 3.1321e-19
db13 | 0.06617 | 0.085993 | -0.0020143 | 0.00056214 | -0.0011054 | 0.0043269 | 0.0010722 | -0.0023244 | -0.00036068 | 0.0001958 | -0.0023439 | 0.00029084 | 0.0015651 | -0.0016549 | -0.0022388 | -0.0014053 | 0.00068922 | 0.00071406 | 0.0020019 | 0.0020297 | 0.053054 |  | -4.6259e-18
db14 | 0.0035916 | -0.0012343 | -0.0010203 | -0.0001133 | -7.9112e-05 | -0.00010585 | -0.00047948 | -0.00038929 | 0.0005487 | 8.5869e-05 | 8.0183e-05 | 0.00099151 | 0.00029715 | 0.00070187 | 0.00078077 | 0.0003378 | -0.0015744 | -0.00020282 | -7.5378e-05 | 0.00021606 | -0.0018072 |  | -2.1684e-19
db15 | 0.022146 | -0.014663 | 0.0019113 | 0.00054683 | -0.0018757 | -0.0015365 | 0.00071851 | 0.00075822 | -0.00093337 | -8.8812e-05 | 0.00011769 | -0.0012571 | -0.00020092 | -0.00035431 | 0.0013313 | 0.00077603 | -0.00086541 | 0.0011784 | 0.0011645 | -0.0013906 | 0.0028013 |  | 1.3492e-18
da2 | 0.50973 | 20.2902 | -0.070163 | -0.2015 | 0.12857 | -0.93113 | 0.13921 | -0.64391 | 0.53396 | -0.11567 | 0.43741 | 1.2492 | -0.20341 | 0.14172 | -0.45137 | -0.22982 | -1.0391 | 0.24522 | 0.97279 | 0.038019 | -5.9848 |  | -2.1588e-17
da3 | 0.81916 | 1.4216 | 0.1754 | -0.049254 | -0.39488 | 0.28863 | -0.020278 | -0.43216 | -0.15161 | -0.42905 | 0.18673 | 0.27866 | -0.12436 | -0.14441 | 0.069672 | -0.10393 | 0.023052 | 0.17246 | 0.40837 | 0.24694 | 0.83006 |  | 0
da4 | -0.013062 | 2.6561 | 0.061152 | -0.18447 | 0.26732 | 0.28911 | -0.13388 | 0.16647 | -0.22934 | -0.65645 | 0.09721 | 0.009361 | -0.12192 | -0.15768 | 0.43159 | -0.061603 | 0.014355 | 0.098236 | -0.10307 | 0.21361 | -0.26563 |  | 0
da5 | 0.11071 | 3.8168 | -0.0079685 | -0.017725 | 0.00088901 | 0.22581 | -0.018517 | -0.067443 | -0.016286 | 0.0037773 | 0.11955 | -0.014045 | -0.15887 | -0.14377 | 0.042641 | -0.032755 | -0.076088 | -0.01805 | 0.082976 | 0.095875 | 0.14776 |  | 0
da6 | -0.060127 | 0.31829 | 0.075994 | 0.10695 | 0.021979 | 0.14107 | 0.097001 | 0.062847 | 0.019532 | -0.052698 | 0.045009 | -0.028686 | -0.13607 | -0.17993 | 0.0027797 | -0.037424 | -0.0087834 | -0.11529 | -0.024868 | 0.010583 | 0.076602 |  | 5.3969e-18
da7 | -0.082391 | 1.7144 | -0.014957 | -0.01668 | -0.10774 | -0.05425 | -0.060006 | 0.029245 | 0.051062 | 0.00031439 | -0.077538 | 0.073182 | 0.075437 | -0.024015 | 0.0032962 | -0.0011568 | 0.090044 | 0.04038 | -0.012106 | 0.0054891 | -0.11618 |  | 0
da8 | -0.012195 | -0.028933 | 0.013173 | -0.0032311 | 0.0085657 | -0.004489 | -0.0027094 | -0.038811 | 0.0054813 | -0.019077 | 0.0090618 | 0.021527 | 0.033327 | 0.00055941 | 0.0024694 | -0.024043 | -0.043655 | 0.017307 | 0.0020043 | 0.02254 | 0.022663 |  | 0
da9 | -0.055091 | -0.23737 | 0.0085226 | -0.014515 | 0.0071629 | 0.028314 | 0.0067467 | -0.01668 | 0.0060538 | -0.021645 | 0.022324 | 0.010004 | -0.0035618 | -0.015728 | 0.014694 | 0.0012718 | -0.037569 | -0.011249 | 0.005453 | 0.010401 | -0.08278 |  | -2.313e-18
da10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
da11 | -0.10725 | 0.16038 | -0.010723 | -0.0064619 | -0.014993 | -0.0071379 | -0.0069116 | -0.0010498 | 0.0047785 | 0.00080376 | -0.0034088 | 0.007119 | 0.006681 | 0.0027695 | 0.0078886 | 0.0041873 | 0.011655 | 0.0032145 | -0.00015224 | 0.001741 | -0.11829 |  | -2.313e-18
da12 | -0.0025277 | -0.0087019 | 0.001428 | 0.0031029 | 0.0011105 | 0.0032005 | 0.0025599 | 0.00085284 | -0.0018534 | -0.0017821 | 0.00092796 | 0.0019229 | 4.3852e-05 | -0.0038879 | -0.00093978 | 0.00064853 | 0.0012857 | -0.0038672 | -0.0028077 | -0.0019458 | -0.00036929 |  | 0
da13 | -0.008688 | -0.01949 | 0.0014433 | -0.0022884 | -0.0017695 | 0.0014063 | -0.00065045 | -0.00087778 | 0.0019492 | -0.0012476 | -0.00085015 | 0.00088993 | -0.00042535 | -0.0013716 | 0.0018624 | 0.00066014 | -0.0014156 | -0.00011012 | 0.0010861 | 0.0017091 | -0.010493 |  | 6.2643e-19
da14 | -0.0014815 | -0.0048685 | 7.023e-05 | -0.00084964 | 0.0004197 | -0.00046076 | -0.0021816 | -0.0023373 | -0.00049615 | -0.0021619 | 0.00023126 | 0.00090566 | 0.0013408 | 0.0010952 | 0.001337 | -0.00094332 | -0.00082346 | 0.0029065 | 0.001074 | 0.00087367 | 0.00078141 |  | 1.807e-19
da15 | -0.0035713 | 0.003958 | -0.0018151 | -0.00072838 | -0.0021597 | 0.0012337 | -0.00076705 | -0.00072243 | 0.0011577 | -0.0010942 | 0.00086523 | 0.00014506 | 0.0026716 | -0.00057811 | -0.0010455 | -0.00023368 | -0.00021597 | 0.0014205 | 0.00076724 | 0.0010991 | -0.00090168 |  | 2.8912e-19
Field Colinearity | 2.4599
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.1344 | -0.19963 | -0.18253 | -0.38467 | -0.22079 | 0.16211 | 0.31189 | 0.39965 | 0.12884 | -0.11731 | -0.13653 | 0.22504 | 0.52233 | 0.49129 | -0.010951 | -0.057441 | -0.69671 | -0.59526 | -1.1259 | -2.4339
Twist (m2 mrad) | 38.5082 | -2.8448 | -2.4642 | -4.9046 | -2.6495 | 1.8237 | 3.2748 | 3.8966 | 1.1595 | -0.96778 | -1.0239 | 1.519 | 3.134 | 2.5793 | -0.049279 | -0.2154 | -2.0901 | -1.3393 | -1.6888 | -0.91272
Twist Integral (I0) (m2 rad) | 0.026058
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 6.5124 | -0.56735 | -0.67913 | -0.84997 | -0.84667 | -0.72979 | -0.4257 | -0.34784 | -0.37671 | -0.44264 | -0.46506 | 0.37426 | 0.90157 | 1.0897 | 0.79997 | 0.38773 | 0.27075 | -0.036305 | -0.12865 | -2.5669
Twist (m2 mrad) | 48.8431 | -8.0847 | -9.1683 | -10.8371 | -10.1601 | -8.2102 | -4.4699 | -3.3914 | -3.3904 | -3.6518 | -3.488 | 2.5263 | 5.4094 | 5.7207 | 3.5999 | 1.454 | 0.81226 | -0.081686 | -0.19298 | -0.9626
Twist Integral (I0) (m2 rad) | 0.0017074
Note on V169: 6.4s
Note on V184: 4.4s
Note on V215: 4.2s

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.426 | 0.0024 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3755.4 | 92.7 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.28 | 0.37 | 4 | 8 | 0 | 1.6 | 4 | 8 | -3739.7 | 142.1 | 3.5 | 7 | -3771 | 160.3 | 3.5 | 7 | 8.5 | 1.75 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.47 | 4 | 8 | 3.9 | 1.1 | 3.5 | 7 | 0.04 | 0.78 | 3.5 | 7
b2 (units) | -0.38 | 0.71 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -0.68 | 2.2 | 3.5 | 7 | -0.038 | 1.65 | 3.5 | 7
b3 (units) | -4.3 | 1.26 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 42 | 3.3 | 3.5 | 7 | -5.66 | 2.8 | 3.5 | 7
b4 (units) | -0.05 | 0.13 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 0.19 | 0.87 | 3.5 | 7 | -0.2 | 0.32 | 3.5 | 7
b5 (units) | 0.13 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -2.1 | 0.63 | 3.5 | 7 | -3.13 | 0.45 | 3.5 | 7
b6 (units) | 0.005 | 0.036 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | 0.06 | 0.25 | 3.5 | 7 | -0.01 | 0.12 | 3.5 | 7
b7 (units) | 1.1 | 0.08 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.77 | 0.23 | 3.5 | 7 | 0.76 | 0.14 | 3.5 | 7
b8 (units) | -0.005 | 0.014 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | 0.028 | 0.1 | 3.5 | 7 | -0.003 | 0.045 | 3.5 | 7
b9 (units) | 0.57 | 0.028 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.51 | 0.08 | 3.5 | 7 | 0.51 | 0.044 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.011 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.66 | 0.022 | 3.5 | 7 | 0.66 | 0.017 | 3.5 | 7
b12 (units) | -0.0003 | 0.0019 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0095 | 3.5 | 7 | -0.0012 | 0.006 | 3.5 | 7
b13 (units) | 0.07 | 0.0037 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.087 | 0.008 | 3.5 | 7 | 0.054 | 0.0066 | 3.5 | 7
b14 (units) | -0.0001 | 0.0026 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0055 | 0.0042 | 3.5 | 7 | -0.0024 | 0.003 | 3.5 | 7
b15 (units) | 0.022 | 0.0041 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | -0.002 | 0.005 | 3.5 | 7 | 0.0129 | 0.0087 | 3.5 | 7
a2 (units) | 0.08 | 1.32 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | -1.5 | 3.6 | 3.5 | 7 | 0.26 | 2.9 | 3.5 | 7
a3 (units) | -0.11 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.5 | 1.2 | 3.5 | 7 | 0.58 | 0.6 | 3.5 | 7
a4 (units) | -0.03 | 0.29 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | -0.658 | 1.4 | 3.5 | 7 | 0 | 0.55 | 3.5 | 7
a5 (units) | -0.01 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 3.3 | 0.59 | 3.5 | 7 | -0.58 | 0.26 | 3.5 | 7
a6 (units) | 0 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | -0.21 | 0.38 | 3.5 | 7 | 0 | 0.14 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.56 | 0.22 | 3.5 | 7 | -0.067 | 0.081 | 3.5 | 7
a8 (units) | 0.001 | 0.025 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.03 | 0.09 | 3.5 | 7 | 0 | 0.04 | 3.5 | 7
a9 (units) | -0.012 | 0.031 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.13 | 0.05 | 3.5 | 7 | 0.018 | 0.037 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.021 | 0.043 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.18 | 0.04 | 3.5 | 7 | -0.017 | 0.034 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | -0.0012 | 0.0069 | 3.5 | 7
a13 (units) | -0.002 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.014 | 0.008 | 3.5 | 7 | 0.00076 | 0.0053 | 3.5 | 7
a14 (units) | -0.007 | 0.004 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.005 | 0.004 | 3.5 | 7 | -0.0012 | 0.0036 | 3.5 | 7
a15 (units) | -0.003 | 0.005 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.005 | 0.005 | 3.5 | 7 | -0.0067 | 0.0052 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4176 | 14.4344 | 14.4092 | 14.4428 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4079.85 | -3430.95 | -4404.3 | -3106.5
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 593.8 | 596.76 | 592.32 | 598.24 | -6.4 | 6.4 | -12.8 | 12.8 | -4237.05 | -3242.35 | -4734.4 | -2745 | -4332.05 | -3209.95 | -4893.1 | -2648.9 | 2.375 | 14.625 | -3.75 | 20.75
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -1.88 | 1.88 | -3.76 | 3.76 | 0.05 | 7.75 | -3.8 | 11.6 | -2.69 | 2.77 | -5.42 | 5.5
b2 (units) | -2.865 | 2.105 | -5.35 | 4.59 | -2.08 | 2.08 | -4.16 | 4.16 | -8.38 | 7.02 | -16.08 | 14.72 | -5.813 | 5.737 | -11.588 | 11.512
b3 (units) | -8.71 | 0.11 | -13.12 | 4.52 | -2.96 | 2.96 | -5.92 | 5.92 | 30.45 | 53.55 | 18.9 | 65.1 | -15.46 | 4.14 | -25.26 | 13.94
b4 (units) | -0.505 | 0.405 | -0.96 | 0.86 | -0.52 | 0.52 | -1.04 | 1.04 | -2.855 | 3.235 | -5.9 | 6.28 | -1.32 | 0.92 | -2.44 | 2.04
b5 (units) | -0.64 | 0.9 | -1.41 | 1.67 | -0.8 | 0.8 | -1.6 | 1.6 | -4.305 | 0.105 | -6.51 | 2.31 | -4.705 | -1.555 | -6.28 | 0.02
b6 (units) | -0.121 | 0.131 | -0.247 | 0.257 | -0.24 | 0.24 | -0.48 | 0.48 | -0.815 | 0.935 | -1.69 | 1.81 | -0.43 | 0.41 | -0.85 | 0.83
b7 (units) | 0.82 | 1.38 | 0.54 | 1.66 | -0.256 | 0.256 | -0.512 | 0.512 | 1.965 | 3.575 | 1.16 | 4.38 | 0.27 | 1.25 | -0.22 | 1.74
b8 (units) | -0.054 | 0.044 | -0.103 | 0.093 | -0.092 | 0.092 | -0.184 | 0.184 | -0.322 | 0.378 | -0.672 | 0.728 | -0.1605 | 0.1545 | -0.318 | 0.312
b9 (units) | 0.472 | 0.668 | 0.374 | 0.766 | -0.06 | 0.06 | -0.12 | 0.12 | 0.23 | 0.79 | -0.05 | 1.07 | 0.356 | 0.664 | 0.202 | 0.818
b10 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7085 | 0.7855 | 0.67 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.583 | 0.737 | 0.506 | 0.814 | 0.6005 | 0.7195 | 0.541 | 0.779
b12 (units) | -0.00695 | 0.00635 | -0.0136 | 0.013 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.03025 | 0.03625 | -0.0635 | 0.0695 | -0.0222 | 0.0198 | -0.0432 | 0.0408
b13 (units) | 0.05705 | 0.08295 | 0.0441 | 0.0959 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.059 | 0.115 | 0.031 | 0.143 | 0.0309 | 0.0771 | 0.0078 | 0.1002
b14 (units) | -0.0092 | 0.009 | -0.0183 | 0.0181 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0092 | 0.0202 | -0.0239 | 0.0349 | -0.0129 | 0.0081 | -0.0234 | 0.0186
b15 (units) | 0.00765 | 0.03635 | -0.0067 | 0.0507 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0195 | 0.0155 | -0.037 | 0.033 | -0.01755 | 0.04335 | -0.048 | 0.0738
a2 (units) | -4.54 | 4.7 | -9.16 | 9.32 | -2.8 | 2.8 | -5.6 | 5.6 | -14.1 | 11.1 | -26.7 | 23.7 | -9.89 | 10.41 | -20.04 | 20.56
a3 (units) | -1.055 | 0.835 | -2 | 1.78 | -1.28 | 1.28 | -2.56 | 2.56 | -5.7 | 2.7 | -9.9 | 6.9 | -1.52 | 2.68 | -3.62 | 4.78
a4 (units) | -1.045 | 0.985 | -2.06 | 2 | -1.12 | 1.12 | -2.24 | 2.24 | -5.558 | 4.242 | -10.458 | 9.142 | -1.925 | 1.925 | -3.85 | 3.85
a5 (units) | -0.36 | 0.34 | -0.71 | 0.69 | -0.52 | 0.52 | -1.04 | 1.04 | 1.235 | 5.365 | -0.83 | 7.43 | -1.49 | 0.33 | -2.4 | 1.24
a6 (units) | -0.245 | 0.245 | -0.49 | 0.49 | -0.28 | 0.28 | -0.56 | 0.56 | -1.54 | 1.12 | -2.87 | 2.45 | -0.49 | 0.49 | -0.98 | 0.98
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 0.79 | 2.33 | 0.02 | 3.1 | -0.3505 | 0.2165 | -0.634 | 0.5
a8 (units) | -0.0865 | 0.0885 | -0.174 | 0.176 | -0.096 | 0.096 | -0.192 | 0.192 | -0.345 | 0.285 | -0.66 | 0.6 | -0.14 | 0.14 | -0.28 | 0.28
a9 (units) | -0.1205 | 0.0965 | -0.229 | 0.205 | -0.08 | 0.08 | -0.16 | 0.16 | -0.305 | 0.045 | -0.48 | 0.22 | -0.1115 | 0.1475 | -0.241 | 0.277
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1715 | 0.1295 | -0.322 | 0.28 | -0.034 | 0.034 | -0.068 | 0.068 | 0.04 | 0.32 | -0.1 | 0.46 | -0.136 | 0.102 | -0.255 | 0.221
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.028 | 0.028 | -0.056 | 0.056 | -0.02535 | 0.02295 | -0.0495 | 0.0471
a13 (units) | -0.023 | 0.019 | -0.044 | 0.04 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.042 | 0.014 | -0.07 | 0.042 | -0.01779 | 0.01931 | -0.03634 | 0.03786
a14 (units) | -0.021 | 0.007 | -0.035 | 0.021 | -0.006 | 0.006 | -0.012 | 0.012 | -0.009 | 0.019 | -0.023 | 0.033 | -0.0138 | 0.0114 | -0.0264 | 0.024
a15 (units) | -0.0205 | 0.0145 | -0.038 | 0.032 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0125 | 0.0225 | -0.03 | 0.04 | -0.0249 | 0.0115 | -0.0431 | 0.0297
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.422 | 0.0025 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3733.8 | 101.5 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.56 | 0.2 | 4 | 8 | 0 | 1.2 | 4 | 8 | -3757.8 | 182.5 | 3.5 | 7 | -3709.8 | 127.5 | 3.5 | 7 | 7.4 | 1.5 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.61 | 4 | 8 | 2.6 | 1.3 | 3.5 | 7 | -0.7 | 0.89 | 3.5 | 7
b2 (units) | -0.17 | 0.46 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -2.5 | 2.6 | 3.5 | 7 | 0.01 | 1.17 | 3.5 | 7
b3 (units) | -4.1 | 0.66 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 29.5 | 4.4 | 3.5 | 7 | -1.84 | 3.7 | 3.5 | 7
b4 (units) | -0.02 | 0.09 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -1 | 0.6 | 3.5 | 7 | -0.09 | 0.29 | 3.5 | 7
b5 (units) | -0.35 | 0.27 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -5.7 | 0.82 | 3.5 | 7 | -3.31 | 0.81 | 3.5 | 7
b6 (units) | 0.031 | 0.034 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.24 | 0.17 | 3.5 | 7 | 0 | 0.09 | 3.5 | 7
b7 (units) | 0.73 | 0.084 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.49 | 0.22 | 3.5 | 7
b8 (units) | 0.002 | 0.02 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.05 | 3.5 | 7 | 0.001 | 0.049 | 3.5 | 7
b9 (units) | 0.47 | 0.018 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.37 | 0.11 | 3.5 | 7 | 0.42 | 0.046 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.754 | 0.01 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.61 | 0.024 | 3.5 | 7 | 0.645 | 0.025 | 3.5 | 7
b12 (units) | -0.001 | 0.0027 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.003 | 0.0068 | 3.5 | 7 | -0.0009 | 0.006 | 3.5 | 7
b13 (units) | 0.063 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.085 | 0.008 | 3.5 | 7 | 0.05 | 0.0094 | 3.5 | 7
b14 (units) | -0.001 | 0.0052 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0033 | 0.009 | 3.5 | 7 | -0.0034 | 0.0037 | 3.5 | 7
b15 (units) | 0.031 | 0.0033 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.0012 | 0.009 | 3.5 | 7 | 0.018 | 0.009 | 3.5 | 7
a2 (units) | 0.02 | 1.2 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 11.5 | 4.4 | 3.5 | 7 | -8.2 | 4.5 | 3.5 | 7
a3 (units) | -0.43 | 0.27 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | -1.4 | 1.4 | 3.5 | 7 | 0.26 | 0.65 | 3.5 | 7
a4 (units) | 0.34 | 0.35 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.84 | 0.79 | 3.5 | 7 | -0.35 | 0.6 | 3.5 | 7
a5 (units) | -0.08 | 0.08 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 1.7 | 0.42 | 3.5 | 7 | -0.11 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.1 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.12 | 0.19 | 3.5 | 7 | -0.13 | 0.17 | 3.5 | 7
a7 (units) | 0.04 | 0.05 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.69 | 0.15 | 3.5 | 7 | 0.057 | 0.074 | 3.5 | 7
a8 (units) | 0.029 | 0.03 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | 0.02 | 0.04 | 3.5 | 7 | 0.004 | 0.044 | 3.5 | 7
a9 (units) | 0.021 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.24 | 0.05 | 3.5 | 7 | 0.011 | 0.041 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | 0.02 | 0.054 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.2 | 0.05 | 3.5 | 7 | 0.022 | 0.053 | 3.5 | 7
a12 (units) | -0.001 | 0.004 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | -0.0037 | 0.0079 | 3.5 | 7
a13 (units) | 0.001 | 0.006 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.016 | 0.01 | 3.5 | 7 | 0.00053 | 0.0055 | 3.5 | 7
a14 (units) | -0.011 | 0.007 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.003 | 0.005 | 3.5 | 7 | -0.0031 | 0.0051 | 3.5 | 7
a15 (units) | -0.003 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0035 | 0.006 | 3.5 | 7 | -0.0055 | 0.0077 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4132 | 14.4307 | 14.4045 | 14.4395 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4089.05 | -3378.55 | -4444.3 | -3023.3
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.76 | 596.36 | 593.96 | 597.16 | -4.8 | 4.8 | -9.6 | 9.6 | -4396.55 | -3119.05 | -5035.3 | -2480.3 | -4156.05 | -3263.55 | -4602.3 | -2817.3 | 2.15 | 12.65 | -3.1 | 17.9
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.44 | 2.44 | -4.88 | 4.88 | -1.95 | 7.15 | -6.5 | 11.7 | -3.815 | 2.415 | -6.93 | 5.53
b2 (units) | -1.78 | 1.44 | -3.39 | 3.05 | -2.08 | 2.08 | -4.16 | 4.16 | -11.6 | 6.6 | -20.7 | 15.7 | -4.085 | 4.105 | -8.18 | 8.2
b3 (units) | -6.41 | -1.79 | -8.72 | 0.52 | -2.96 | 2.96 | -5.92 | 5.92 | 14.1 | 44.9 | -1.3 | 60.3 | -14.79 | 11.11 | -27.74 | 24.06
b4 (units) | -0.335 | 0.295 | -0.65 | 0.61 | -0.52 | 0.52 | -1.04 | 1.04 | -3.1 | 1.1 | -5.2 | 3.2 | -1.105 | 0.925 | -2.12 | 1.94
b5 (units) | -1.295 | 0.595 | -2.24 | 1.54 | -0.8 | 0.8 | -1.6 | 1.6 | -8.57 | -2.83 | -11.44 | 0.04 | -6.145 | -0.475 | -8.98 | 2.36
b6 (units) | -0.088 | 0.15 | -0.207 | 0.269 | -0.24 | 0.24 | -0.48 | 0.48 | -0.835 | 0.355 | -1.43 | 0.95 | -0.315 | 0.315 | -0.63 | 0.63
b7 (units) | 0.436 | 1.024 | 0.142 | 1.318 | -0.256 | 0.256 | -0.512 | 0.512 | 0.97 | 3.07 | -0.08 | 4.12 | -0.28 | 1.26 | -1.05 | 2.03
b8 (units) | -0.068 | 0.072 | -0.138 | 0.142 | -0.092 | 0.092 | -0.184 | 0.184 | -0.215 | 0.135 | -0.39 | 0.31 | -0.1705 | 0.1725 | -0.342 | 0.344
b9 (units) | 0.407 | 0.533 | 0.344 | 0.596 | -0.06 | 0.06 | -0.12 | 0.12 | -0.015 | 0.755 | -0.4 | 1.14 | 0.259 | 0.581 | 0.098 | 0.742
b10 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.719 | 0.789 | 0.684 | 0.824 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.526 | 0.694 | 0.442 | 0.778 | 0.5575 | 0.7325 | 0.47 | 0.82
b12 (units) | -0.01045 | 0.00845 | -0.0199 | 0.0179 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0208 | 0.0268 | -0.0446 | 0.0506 | -0.0219 | 0.0201 | -0.0429 | 0.0411
b13 (units) | 0.05145 | 0.07455 | 0.0399 | 0.0861 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.057 | 0.113 | 0.029 | 0.141 | 0.0171 | 0.0829 | -0.0158 | 0.1158
b14 (units) | -0.0192 | 0.0172 | -0.0374 | 0.0354 | -0.004 | 0.004 | -0.008 | 0.008 | -0.0282 | 0.0348 | -0.0597 | 0.0663 | -0.01635 | 0.00955 | -0.0293 | 0.0225
b15 (units) | 0.01945 | 0.04255 | 0.0079 | 0.0541 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.0303 | 0.0327 | -0.0618 | 0.0642 | -0.0135 | 0.0495 | -0.045 | 0.081
a2 (units) | -4.18 | 4.22 | -8.38 | 8.42 | -2.8 | 2.8 | -5.6 | 5.6 | -3.9 | 26.9 | -19.3 | 42.3 | -23.95 | 7.55 | -39.7 | 23.3
a3 (units) | -1.375 | 0.515 | -2.32 | 1.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.3 | 3.5 | -11.2 | 8.4 | -2.015 | 2.535 | -4.29 | 4.81
a4 (units) | -0.885 | 1.565 | -2.11 | 2.79 | -1.12 | 1.12 | -2.24 | 2.24 | -0.925 | 4.605 | -3.69 | 7.37 | -2.45 | 1.75 | -4.55 | 3.85
a5 (units) | -0.36 | 0.2 | -0.64 | 0.48 | -0.52 | 0.52 | -1.04 | 1.04 | 0.23 | 3.17 | -1.24 | 4.64 | -0.985 | 0.765 | -1.86 | 1.64
a6 (units) | -0.35 | 0.35 | -0.7 | 0.7 | -0.28 | 0.28 | -0.56 | 0.56 | -0.545 | 0.785 | -1.21 | 1.45 | -0.725 | 0.465 | -1.32 | 1.06
a7 (units) | -0.135 | 0.215 | -0.31 | 0.39 | -0.192 | 0.192 | -0.384 | 0.384 | 1.165 | 2.215 | 0.64 | 2.74 | -0.202 | 0.316 | -0.461 | 0.575
a8 (units) | -0.076 | 0.134 | -0.181 | 0.239 | -0.096 | 0.096 | -0.192 | 0.192 | -0.12 | 0.16 | -0.26 | 0.3 | -0.15 | 0.158 | -0.304 | 0.312
a9 (units) | -0.084 | 0.126 | -0.189 | 0.231 | -0.08 | 0.08 | -0.16 | 0.16 | -0.415 | -0.065 | -0.59 | 0.11 | -0.1325 | 0.1545 | -0.276 | 0.298
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.169 | 0.209 | -0.358 | 0.398 | -0.034 | 0.034 | -0.068 | 0.068 | 0.025 | 0.375 | -0.15 | 0.55 | -0.1635 | 0.2075 | -0.349 | 0.393
a12 (units) | -0.015 | 0.013 | -0.029 | 0.027 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.03135 | 0.02395 | -0.059 | 0.0516
a13 (units) | -0.02 | 0.022 | -0.041 | 0.043 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.051 | 0.019 | -0.086 | 0.054 | -0.01872 | 0.01978 | -0.03797 | 0.03903
a14 (units) | -0.0355 | 0.0135 | -0.06 | 0.038 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0145 | 0.0205 | -0.032 | 0.038 | -0.02095 | 0.01475 | -0.0388 | 0.0326
a15 (units) | -0.024 | 0.018 | -0.045 | 0.039 | -0.014 | 0.014 | -0.028 | 0.028 | -0.0175 | 0.0245 | -0.0385 | 0.0455 | -0.03245 | 0.02145 | -0.0594 | 0.0484
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.421 | 0.0028 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3714.8 | 156.6 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 596.06 | 0.18 | 4 | 8 | 0 | 1.1 | 4 | 8 | -3766.1 | 232.2 | 3.5 | 7 | -3766.1 | 105.8 | 3.5 | 7 | 5.8 | 1.6 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.56 | 4 | 8 | 3.5 | 1.3 | 3.5 | 7 | -0.55 | 0.88 | 3.5 | 7
b2 (units) | -0.64 | 0.74 | 3.5 | 7 | 0 | 0.52 | 4 | 8 | -1.6 | 2.5 | 3.5 | 7 | 0.32 | 1.33 | 3.5 | 7
b3 (units) | -3.2 | 0.5 | 3.5 | 7 | 0 | 0.74 | 4 | 8 | 37.1 | 3.3 | 3.5 | 7 | -1 | 2.15 | 3.5 | 7
b4 (units) | 0 | 0.11 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | -0.59 | 0.92 | 3.5 | 7 | -0.13 | 0.24 | 3.5 | 7
b5 (units) | -0.45 | 0.22 | 3.5 | 7 | 0 | 0.2 | 4 | 8 | -4.8 | 0.92 | 3.5 | 7 | -3.38 | 0.5 | 3.5 | 7
b6 (units) | 0.017 | 0.03 | 3.5 | 7 | 0 | 0.06 | 4 | 8 | -0.16 | 0.22 | 3.5 | 7 | -0.01 | 0.08 | 3.5 | 7
b7 (units) | 1.15 | 0.055 | 3.5 | 7 | 0 | 0.064 | 4 | 8 | 2.74 | 0.21 | 3.5 | 7 | 0.9 | 0.1 | 3.5 | 7
b8 (units) | 0.001 | 0.012 | 3.5 | 7 | 0 | 0.023 | 4 | 8 | -0.04 | 0.07 | 3.5 | 7 | 0.006 | 0.038 | 3.5 | 7
b9 (units) | 0.49 | 0.014 | 3.5 | 7 | 0 | 0.015 | 4 | 8 | 0.36 | 0.06 | 3.5 | 7 | 0.455 | 0.031 | 3.5 | 7
b10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
b11 (units) | 0.747 | 0.009 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0.62 | 0.024 | 3.5 | 7 | 0.653 | 0.016 | 3.5 | 7
b12 (units) | -0.0009 | 0.0017 | 3.5 | 7 | 0 | 0.0029 | 4 | 8 | 0.001 | 0.0056 | 3.5 | 7 | 0 | 0.005 | 3.5 | 7
b13 (units) | 0.066 | 0.0021 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0.089 | 0.008 | 3.5 | 7 | 0.054 | 0.0053 | 3.5 | 7
b14 (units) | 0.0005 | 0.0028 | 3.5 | 7 | 0 | 0.001 | 4 | 8 | 0.0042 | 0.0029 | 3.5 | 7 | 0 | 0.0029 | 3.5 | 7
b15 (units) | 0.034 | 0.0046 | 3.5 | 7 | 0 | 0.0026 | 4 | 8 | 0 | 0.008 | 3.5 | 7 | 0.02 | 0.011 | 3.5 | 7
a2 (units) | 0.08 | 1.08 | 3.5 | 7 | 0 | 0.7 | 4 | 8 | 5.5 | 6.9 | 3.5 | 7 | 1.8 | 6 | 3.5 | 7
a3 (units) | 0.45 | 0.43 | 3.5 | 7 | 0 | 0.32 | 4 | 8 | 0.1 | 2 | 3.5 | 7 | 0.3 | 0.54 | 3.5 | 7
a4 (units) | -0.05 | 0.28 | 3.5 | 7 | 0 | 0.28 | 4 | 8 | 1.11 | 1.4 | 3.5 | 7 | 0.12 | 0.62 | 3.5 | 7
a5 (units) | 0.06 | 0.1 | 3.5 | 7 | 0 | 0.13 | 4 | 8 | 2.7 | 0.56 | 3.5 | 7 | -0.15 | 0.2 | 3.5 | 7
a6 (units) | 0.02 | 0.07 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0.1 | 0.26 | 3.5 | 7 | 0.2 | 0.15 | 3.5 | 7
a7 (units) | -0.02 | 0.04 | 3.5 | 7 | 0 | 0.048 | 4 | 8 | 1.6 | 0.17 | 3.5 | 7 | 0.029 | 0.062 | 3.5 | 7
a8 (units) | 0 | 0.023 | 3.5 | 7 | 0 | 0.024 | 4 | 8 | -0.01 | 0.06 | 3.5 | 7 | 0.006 | 0.031 | 3.5 | 7
a9 (units) | -0.012 | 0.019 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | -0.23 | 0.05 | 3.5 | 7 | -0.07 | 0.029 | 3.5 | 7
a10 (units) | 0 | 1 | 3.5 | 7 | 0 | 1 | 4 | 8 | 0 | 1 | 3.5 | 7 | 0 | 1 | 3.5 | 7
a11 (units) | -0.032 | 0.023 | 3.5 | 7 | 0 | 0.0085 | 4 | 8 | 0.16 | 0.025 | 3.5 | 7 | -0.022 | 0.023 | 3.5 | 7
a12 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0031 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.0011 | 0.0055 | 3.5 | 7
a13 (units) | -0.002 | 0.003 | 3.5 | 7 | 0 | 0.0021 | 4 | 8 | -0.02 | 0.006 | 3.5 | 7 | -0.0036 | 0.0032 | 3.5 | 7
a14 (units) | -0.013 | 0.005 | 3.5 | 7 | 0 | 0.0015 | 4 | 8 | 0.006 | 0.005 | 3.5 | 7 | -0.0036 | 0.0037 | 3.5 | 7
a15 (units) | -0.004 | 0.004 | 3.5 | 7 | 0 | 0.0035 | 4 | 8 | 0.0045 | 0.005 | 3.5 | 7 | -0.0066 | 0.0078 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.4112 | 14.4308 | 14.4014 | 14.4406 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4262.9 | -3166.7 | -4811 | -2618.6
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 595.34 | 596.78 | 594.62 | 597.5 | -4.4 | 4.4 | -8.8 | 8.8 | -4578.8 | -2953.4 | -5391.5 | -2140.7 | -4136.4 | -3395.8 | -4506.7 | -3025.5 | 0.2 | 11.4 | -5.4 | 17
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2.24 | 2.24 | -4.48 | 4.48 | -1.05 | 8.05 | -5.6 | 12.6 | -3.63 | 2.53 | -6.71 | 5.61
b2 (units) | -3.23 | 1.95 | -5.82 | 4.54 | -2.08 | 2.08 | -4.16 | 4.16 | -10.35 | 7.15 | -19.1 | 15.9 | -4.335 | 4.975 | -8.99 | 9.63
b3 (units) | -4.95 | -1.45 | -6.7 | 0.3 | -2.96 | 2.96 | -5.92 | 5.92 | 25.55 | 48.65 | 14 | 60.2 | -8.525 | 6.525 | -16.05 | 14.05
b4 (units) | -0.385 | 0.385 | -0.77 | 0.77 | -0.52 | 0.52 | -1.04 | 1.04 | -3.81 | 2.63 | -7.03 | 5.85 | -0.97 | 0.71 | -1.81 | 1.55
b5 (units) | -1.22 | 0.32 | -1.99 | 1.09 | -0.8 | 0.8 | -1.6 | 1.6 | -8.02 | -1.58 | -11.24 | 1.64 | -5.13 | -1.63 | -6.88 | 0.12
b6 (units) | -0.088 | 0.122 | -0.193 | 0.227 | -0.24 | 0.24 | -0.48 | 0.48 | -0.93 | 0.61 | -1.7 | 1.38 | -0.29 | 0.27 | -0.57 | 0.55
b7 (units) | 0.9575 | 1.3425 | 0.765 | 1.535 | -0.256 | 0.256 | -0.512 | 0.512 | 2.005 | 3.475 | 1.27 | 4.21 | 0.55 | 1.25 | 0.2 | 1.6
b8 (units) | -0.041 | 0.043 | -0.083 | 0.085 | -0.092 | 0.092 | -0.184 | 0.184 | -0.285 | 0.205 | -0.53 | 0.45 | -0.127 | 0.139 | -0.26 | 0.272
b9 (units) | 0.441 | 0.539 | 0.392 | 0.588 | -0.06 | 0.06 | -0.12 | 0.12 | 0.15 | 0.57 | -0.06 | 0.78 | 0.3465 | 0.5635 | 0.238 | 0.672
b10 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
b11 (units) | 0.7155 | 0.7785 | 0.684 | 0.81 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | 0.536 | 0.704 | 0.452 | 0.788 | 0.597 | 0.709 | 0.541 | 0.765
b12 (units) | -0.00685 | 0.00505 | -0.0128 | 0.011 | -0.0116 | 0.0116 | -0.0232 | 0.0232 | -0.0186 | 0.0206 | -0.0382 | 0.0402 | -0.0175 | 0.0175 | -0.035 | 0.035
b13 (units) | 0.05865 | 0.07335 | 0.0513 | 0.0807 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | 0.061 | 0.117 | 0.033 | 0.145 | 0.03545 | 0.07255 | 0.0169 | 0.0911
b14 (units) | -0.0093 | 0.0103 | -0.0191 | 0.0201 | -0.004 | 0.004 | -0.008 | 0.008 | -0.00595 | 0.01435 | -0.0161 | 0.0245 | -0.01015 | 0.01015 | -0.0203 | 0.0203
b15 (units) | 0.0179 | 0.0501 | 0.0018 | 0.0662 | -0.0104 | 0.0104 | -0.0208 | 0.0208 | -0.028 | 0.028 | -0.056 | 0.056 | -0.0185 | 0.0585 | -0.057 | 0.097
a2 (units) | -3.7 | 3.86 | -7.48 | 7.64 | -2.8 | 2.8 | -5.6 | 5.6 | -18.65 | 29.65 | -42.8 | 53.8 | -19.2 | 22.8 | -40.2 | 43.8
a3 (units) | -1.055 | 1.955 | -2.56 | 3.46 | -1.28 | 1.28 | -2.56 | 2.56 | -6.9 | 7.1 | -13.9 | 14.1 | -1.59 | 2.19 | -3.48 | 4.08
a4 (units) | -1.03 | 0.93 | -2.01 | 1.91 | -1.12 | 1.12 | -2.24 | 2.24 | -3.79 | 6.01 | -8.69 | 10.91 | -2.05 | 2.29 | -4.22 | 4.46
a5 (units) | -0.29 | 0.41 | -0.64 | 0.76 | -0.52 | 0.52 | -1.04 | 1.04 | 0.74 | 4.66 | -1.22 | 6.62 | -0.85 | 0.55 | -1.55 | 1.25
a6 (units) | -0.225 | 0.265 | -0.47 | 0.51 | -0.28 | 0.28 | -0.56 | 0.56 | -0.81 | 1.01 | -1.72 | 1.92 | -0.325 | 0.725 | -0.85 | 1.25
a7 (units) | -0.16 | 0.12 | -0.3 | 0.26 | -0.192 | 0.192 | -0.384 | 0.384 | 1.005 | 2.195 | 0.41 | 2.79 | -0.188 | 0.246 | -0.405 | 0.463
a8 (units) | -0.0805 | 0.0805 | -0.161 | 0.161 | -0.096 | 0.096 | -0.192 | 0.192 | -0.22 | 0.2 | -0.43 | 0.41 | -0.1025 | 0.1145 | -0.211 | 0.223
a9 (units) | -0.0785 | 0.0545 | -0.145 | 0.121 | -0.08 | 0.08 | -0.16 | 0.16 | -0.405 | -0.055 | -0.58 | 0.12 | -0.1715 | 0.0315 | -0.273 | 0.133
a10 (units) | -3.5 | 3.5 | -7 | 7 | -4 | 4 | -8 | 8 | -3.5 | 3.5 | -7 | 7 | -3.5 | 3.5 | -7 | 7
a11 (units) | -0.1125 | 0.0485 | -0.193 | 0.129 | -0.034 | 0.034 | -0.068 | 0.068 | 0.0725 | 0.2475 | -0.015 | 0.335 | -0.1025 | 0.0585 | -0.183 | 0.139
a12 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0124 | 0.0124 | -0.0248 | 0.0248 | -0.021 | 0.021 | -0.042 | 0.042 | -0.02035 | 0.01815 | -0.0396 | 0.0374
a13 (units) | -0.0125 | 0.0085 | -0.023 | 0.019 | -0.0084 | 0.0084 | -0.0168 | 0.0168 | -0.041 | 0.001 | -0.062 | 0.022 | -0.0148 | 0.0076 | -0.026 | 0.0188
a14 (units) | -0.0305 | 0.0045 | -0.048 | 0.022 | -0.006 | 0.006 | -0.012 | 0.012 | -0.0115 | 0.0235 | -0.029 | 0.041 | -0.01655 | 0.00935 | -0.0295 | 0.0223
a15 (units) | -0.018 | 0.01 | -0.032 | 0.024 | -0.014 | 0.014 | -0.028 | 0.028 | -0.013 | 0.022 | -0.0305 | 0.0395 | -0.0339 | 0.0207 | -0.0612 | 0.048
Note on F3: voelling: Field colinearity is not considered any more. Jan 04

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.29929 | -4.1525 | -0.08293 | -0.016452 | 1.1027 | -0.82702 | 0.96345 | 0.54855 | -0.50562 | 16.9535 | 0.59566 | -0.10999 | -4.0299 | 0.052354 | -0.14263 | 1.1821 | -1.1506 | 0.50973 | 0.81916 | -0.013062 | 17.8372 | 0.59586 | -0.40738 | -2.6323 | -0.12102 | -0.28865 | 1.1583 | -0.0013715 | 1.1626 | 0.52274 | -0.44278 | 14419.699 | -0.21057 | -2.5587 | -0.031741 | -0.38477 | 1.2363 | -0.0096304 | 0.88028 | 0.83731 | 0.057426 | 14420.43
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8