## Alarm sheet
File name | HCMB__A001-01000067_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 1000067
Date of test Ap 1 | 2003-10-24 |  |  |  | Date of test Ap 2 | 2003-10-24
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | red alarm | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | yellow alarm | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | yellow alarm | status ok | b6 | status ok | yellow alarm | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | yellow alarm | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | yellow alarm | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | yellow alarm | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
211 | HCMB__A001 | 1000067 |  | 3 |  |  | 3
| COMMENTS: | Red alarm in ap. 2 average straight (20 sigma).
| HOLDING POINT | HOLD | Data at CERN |  | 2003-10-24 | Analysis by |  | V. Remondino | Answer to MMS-LD |  |  |  | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.175 | 0.2 | -0.025 | 0.75 | 0.8 | -0.05 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.75 | 0.2 | 0.2 | 0.75 | Up | 0.75 | 0.15 | 0.15 | 0.75
[8 empty] | 0.75 | 0.2 | 0.2 | 0.75 | Down | 0.75 | 0.15 | 0.15 | 0.75
Seen from connection side
| Collared coil template - version 20.12.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2
| 2002-10-20 | E. Todesco | Feed-down dx and dy in summary data, pre-stress data canceled (never used)
| 2002-12-01 | J. Beauquis | New upload from new measurement system
| 2002-12-20 | E. Todesco | Correction of the systematic difference betyween old and new system (mag len +0.020 m, main field +2.3 mT/KA)
| 2003-03-05 | E. Todesco | Rounding of the measuring current to one digit in getsetdata (10.005 A becomes 10 A, 8.4999 A becomes 8.5 A)
| 2003-05-25 | E. Todesco | Correction of the caption of figure (/30 dropped), removal of calibration date, removal of check on field colinearity

## Original data
File name |  | HCMB__A001-01000067_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 1000067 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Alstom | [9 empty] | File name for aperture 1 |  |  | 1067_cc1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | 1067_cc2 |  |  |  | a even | 1
Test Operator |  | C.MOUGENOT |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-10-24 |  |  | Time of test |  |  | 9:52 |  |  |  | Date of test |  | 2003-10-24 |  |  | Time of test |  |  | 9:52
Mole name |  | taupe_dipole#13 | [9 empty] | Mole name |  | taupe_dipole#10 | [5 empty] | 2001-01-11
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 20
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.51 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.47 |  |  | I. Gain |  |  | 50/500
Dx average |  | 9.3001e-05 |  |  | Dy average |  |  | 0.00033801 |  |  |  | Dx average |  | 0.0002708 |  |  | Dy average |  |  | 0.00014373
Magnetic length |  | 14.4272 |  |  | Tranfs. Function (mT/KA) |  |  | 594.3281 |  |  |  | Magnetic length |  | 14.4237 |  |  | Tranfs. Function (mT/KA) |  |  | 585.9417
Phase |  | 11.9321 |  |  | Integral field (T.m) |  |  | 85.6858 |  |  |  | Phase |  | 23.908 |  |  | Integral field (T.m) |  |  | 84.5563
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 3.3627 | 5.9493 | 5.943 | 5.9432 | 5.9437 | 5.9422 | 5.9412 | 5.9421 | 5.9423 | 5.9413 | 5.942 | 5.9422 | 5.9429 | 5.9428 | 5.9429 | 5.9423 | 5.9438 | 5.9433 | 5.9485 | 4.0584 | 85.6858 | C1 (T) | 3.7786 | 5.8641 | 5.8594 | 5.858 | 5.8582 | 5.8582 | 5.8581 | 5.8584 | 5.8593 | 5.8585 | 5.8579 | 5.8577 | 5.8585 | 5.8596 | 5.8597 | 5.8599 | 5.8582 | 5.8601 | 5.8656 | 3.5085 | 84.5563
Angle (mrad) | 2.6949 | -0.57666 | 0.25066 | -0.33146 | 0.34508 | 0.18126 | -0.807 | 0.13659 | 0.056699 | -0.98811 | -0.67048 | 0.56924 | 0.10068 | 0.19599 | -0.0644 | 0.30428 | -0.10051 | -0.000733 | 0.1431 | -0.39405 | 11.9321 | Angle (mrad) | 2.2849 | 0.004472 | 0.157 | 0.045712 | 0.32069 | 0.18378 | 0.038227 | 0.082267 | 0.21682 | -0.031745 | -0.13504 | -0.11842 | 0.033668 | -0.29677 | -0.52724 | -0.52355 | -0.021799 | -0.11071 | -0.20512 | -0.9775 | 23.908
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -0.70659 | 1.3354 | 2.1272 | 2.6657 | 1.2764 | 1.7067 | 2.1719 | 0.67704 | 0.24307 | 1.0362 | 1.808 | 1.1473 | 0.7177 | -0.010715 | 1.2686 | 2.4644 | 1.2451 | 0.54241 | 0.41355 | -0.87915 | 1.1344 | b2 | 0.75169 | 0.5201 | -0.96287 | -0.84756 | 0.055514 | -0.073905 | 0.14862 | -0.033415 | -0.2076 | -0.64255 | 0.39707 | 0.60718 | 0.0052025 | 0.090822 | 0.010949 | 0.24211 | 0.93455 | -0.14489 | 1.0181 | 0.79121 | 0.10794
b3 | 50.054 | -3.4152 | -5.2171 | -4.5903 | -4.265 | -5.2157 | -4.7239 | -4.7533 | -5.1228 | -4.5283 | -5.2845 | -4.7071 | -4.3663 | -4.9473 | -5.7094 | -5.2443 | -5.8863 | -3.9966 | -3.7589 | -5.935 | -3.193 | b3 | 39.4514 | -3.3758 | -5.4526 | -4.4489 | -4.3051 | -5.4885 | -5.5436 | -4.1361 | -5.9759 | -4.0263 | -3.8283 | -4.6593 | -3.8438 | -6.1693 | -4.0477 | -5.3446 | -4.9418 | -5.4368 | -4.1951 | -6.6984 | -3.3148
b4 | 1.7914 | -0.085061 | -0.15902 | -0.056315 | -0.11404 | 0.057101 | 0.14311 | 0.21193 | 0.17882 | -0.10577 | 0.022082 | 0.10558 | 0.16505 | -0.014115 | -0.29917 | 0.10623 | 0.33389 | -0.19675 | -0.15304 | -0.10182 | 0.05633 | b4 | 0.35402 | -0.29824 | -0.020147 | -0.14006 | -0.060007 | -0.093572 | 0.24417 | -0.026847 | -0.74389 | -0.48588 | -0.075458 | -0.088829 | -0.17225 | -0.18361 | -0.071965 | 0.12356 | 0.28688 | 0.011512 | 0.15453 | 0.10208 | -0.070185
b5 | -2.1417 | 0.33731 | 0.11274 | 0.15823 | 0.41849 | 0.49583 | 0.44265 | 0.647 | 0.49947 | 0.64832 | 0.51344 | 0.47225 | 0.44122 | 0.54664 | 0.50256 | 0.27744 | 0.72606 | 0.45672 | 0.6609 | -2.3283 | 0.28862 | b5 | -2.0361 | 0.44987 | 0.34657 | 0.20274 | 0.10337 | 0.22749 | 0.45533 | 0.50721 | 0.038195 | 0.50668 | 0.22922 | 0.28763 | 0.45549 | 0.43712 | 0.50243 | 0.4815 | 0.42225 | 0.6455 | 0.54099 | -3.8421 | 0.16766
b6 | 1.2145 | 0.077859 | 0.050112 | 0.020043 | 0.034535 | 0.12431 | -0.022473 | 0.08784 | 0.045872 | 0.055698 | 0.067635 | 0.01588 | -0.028197 | 0.10037 | 0.0079029 | 0.085288 | -0.10935 | -0.1202 | 0.018961 | -0.078001 | 0.059534 | b6 | 0.067243 | -0.0046806 | 0.00053669 | 0.03156 | -0.046328 | -0.048456 | -0.084988 | -0.089862 | -0.16374 | 0.027732 | -0.082732 | 0.0080737 | 0.0042217 | -0.20356 | -0.11654 | 0.048499 | 0.24636 | 0.035243 | -0.19463 | 0.16983 | -0.025381
b7 | 2.7431 | 1.0617 | 1.2281 | 1.2091 | 1.25 | 1.1819 | 1.1899 | 1.1744 | 1.1864 | 1.164 | 1.1888 | 1.1957 | 1.2288 | 1.2248 | 1.1219 | 1.1581 | 1.2091 | 1.2741 | 1.1189 | 0.8542 | 1.2209 | b7 | 2.7706 | 1.1703 | 1.2524 | 1.2243 | 1.2604 | 1.2649 | 1.2234 | 1.3297 | 1.2163 | 1.2557 | 1.3101 | 1.2912 | 1.2982 | 1.2582 | 1.226 | 1.3437 | 1.2894 | 1.3654 | 1.2176 | 0.71094 | 1.2996
b8 | 0.29336 | -0.020057 | 0.034002 | -0.033804 | -0.07143 | -0.03809 | 0.026926 | 0.00083649 | 0.051134 | 0.0095803 | 0.034991 | -0.0063091 | 0.012623 | -0.037797 | -0.091258 | 0.023971 | -0.032963 | -0.0023505 | -0.015076 | -0.036861 | -0.00074383 | b8 | -0.0035112 | -0.00934 | 0.022586 | -0.012023 | 0.00019227 | 0.046776 | 0.029533 | 0.01479 | -0.016023 | -0.0094486 | 0.00037441 | 0.057512 | 0.033057 | -0.011839 | -0.062753 | -0.0024328 | -0.059687 | -0.021527 | 0.068376 | 0.054034 | 0.0051054
b9 | 0.48941 | 0.61259 | 0.64188 | 0.65843 | 0.6284 | 0.64231 | 0.64877 | 0.64647 | 0.64023 | 0.63127 | 0.63645 | 0.63654 | 0.64709 | 0.62642 | 0.65074 | 0.63413 | 0.58653 | 0.63724 | 0.60007 | 0.50527 | 0.62484 | b9 | 0.52061 | 0.65075 | 0.62602 | 0.65597 | 0.65784 | 0.64779 | 0.64703 | 0.67876 | 0.61617 | 0.67498 | 0.65261 | 0.66157 | 0.64067 | 0.59008 | 0.62903 | 0.63373 | 0.6362 | 0.6383 | 0.62973 | 0.48533 | 0.63364
b10 | 0.17467 | -0.054481 | 0.034365 | -0.0075241 | -0.020638 | 0.028456 | 0.02854 | 0.025526 | 0.076149 | -0.0080276 | -0.012113 | -0.0099222 | 0.013053 | -0.01123 | -0.088333 | 0.0070719 | -0.022206 | -0.013204 | -0.028735 | -0.051943 | 0 | b10 | 0.078488 | -0.077487 | 0.020475 | -0.023453 | -0.0035594 | 0.030839 | 0.019688 | -0.0057195 | 0.030791 | -0.0029825 | 0.027465 | 0.064023 | 0.045026 | 0.021808 | -0.073696 | -0.015649 | 0.003173 | -0.037379 | -0.031597 | -0.070546 | 0
b11 | 0.6237 | 0.73783 | 0.72654 | 0.73127 | 0.73842 | 0.73532 | 0.72442 | 0.7302 | 0.73218 | 0.73415 | 0.73526 | 0.72891 | 0.72641 | 0.73151 | 0.7388 | 0.72899 | 0.73531 | 0.72649 | 0.73524 | 0.66812 | 0.72662 | b11 | 0.71287 | 0.79176 | 0.79076 | 0.78731 | 0.7888 | 0.7867 | 0.79917 | 0.80272 | 0.78891 | 0.80425 | 0.78816 | 0.78755 | 0.78962 | 0.7889 | 0.80163 | 0.7991 | 0.78514 | 0.79068 | 0.8039 | 0.70053 | 0.78749
b12 | 0.030077 | -0.0028639 | 0.0072308 | 0.0040343 | 0.0022144 | 0.0064516 | 0.0060698 | 0.0024868 | 0.0088428 | -0.0043972 | -0.0010761 | 0.00043602 | -0.00014177 | -0.0043667 | -0.013883 | -0.0023529 | -0.0068038 | 0.00018258 | 0.0021119 | -0.00033012 | 0.0010891 | b12 | 0.010252 | -0.003929 | 0.0016366 | -0.003734 | 0.0001494 | 0.0083135 | 0.0023089 | 0.0021156 | 0.0052248 | -0.0018662 | 0.0054262 | 0.0082822 | 0.0075559 | 0.0020545 | -0.010517 | -0.0071164 | -0.0042281 | -0.0092516 | -0.0056739 | 0.0009466 | 0.0002033
b13 | 0.090278 | 0.078095 | 0.075141 | 0.072418 | 0.077501 | 0.072187 | 0.072974 | 0.070278 | 0.071106 | 0.074286 | 0.075268 | 0.072759 | 0.074077 | 0.072779 | 0.069696 | 0.072773 | 0.072422 | 0.069926 | 0.072373 | 0.059279 | 0.073128 | b13 | 0.089363 | 0.080156 | 0.081174 | 0.076731 | 0.081029 | 0.076814 | 0.077414 | 0.078586 | 0.074125 | 0.07933 | 0.082186 | 0.075667 | 0.077784 | 0.075701 | 0.080833 | 0.077936 | 0.073248 | 0.077506 | 0.074073 | 0.065081 | 0.077786
b14 | 0.011113 | -0.00538 | 0.00021587 | 0.00056401 | -0.00091241 | 0.00077771 | -0.0012289 | -0.0010725 | -0.0001097 | -0.003473 | -0.0039373 | -0.0016287 | -0.0021746 | -0.0014151 | -0.0053577 | -0.0043099 | -0.0038714 | -0.0027771 | -0.0036084 | -0.0011332 | -0.0017762 | b14 | 0.0044324 | -0.0080187 | -0.0057148 | -0.0039619 | -0.0036409 | -0.0027044 | -0.0041909 | -0.0033345 | -0.0025133 | -0.0037697 | -0.0040794 | -0.002067 | -0.0030678 | -0.0044389 | -0.0064884 | -0.0061953 | -0.0033617 | -0.005877 | -0.0060318 | -0.0011373 | -0.0040162
b15 | -0.0068901 | 0.020366 | 0.015915 | 0.015455 | 0.017223 | 0.019875 | 0.017404 | 0.020283 | 0.020625 | 0.019919 | 0.018778 | 0.017858 | 0.017394 | 0.019101 | 0.01688 | 0.018522 | 0.017628 | 0.02063 | 0.022331 | 0.014969 | 0.017794 | b15 | -0.0093952 | 0.01383 | 0.019119 | 0.015774 | 0.017489 | 0.018926 | 0.018454 | 0.017154 | 0.019381 | 0.020147 | 0.018894 | 0.018858 | 0.019799 | 0.0215 | 0.016182 | 0.019726 | 0.017497 | 0.017077 | 0.017082 | 0.010073 | 0.016985
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 26.9493 | -5.7666 | 2.5066 | -3.3146 | 3.4508 | 1.8126 | -8.07 | 1.3659 | 0.56699 | -9.8811 | -6.7048 | 5.6924 | 1.0068 | 1.9599 | -0.644 | 3.0428 | -1.0051 | -0.0073297 | 1.431 | -3.9405 | 0 | a1 | 22.8485 | 0.04472 | 1.57 | 0.45712 | 3.2069 | 1.8378 | 0.38227 | 0.82267 | 2.1682 | -0.31745 | -1.3504 | -1.1842 | 0.33668 | -2.9677 | -5.2724 | -5.2355 | -0.21799 | -1.1071 | -2.0512 | -9.7751 | 0
a2 | -7.5073 | 0.070112 | -2.2191 | -0.14252 | 0.40903 | -1.1189 | -0.75093 | -0.34567 | 0.050851 | 0.47404 | -0.34661 | -0.49164 | 0.45257 | 0.38258 | 0.51961 | -1.5648 | -0.5412 | -0.60805 | 0.34875 | 1.6016 | -0.44547 | a2 | -2.9633 | 1.2486 | 0.97083 | -0.053474 | -0.49887 | 0.47165 | -0.52902 | -0.92761 | 1.9894 | -0.82926 | -0.04837 | -0.11219 | -0.90477 | 0.73085 | -0.81944 | -1.5465 | -0.52549 | 1.373 | -0.54116 | 0.65758 | -0.10746
a3 | -0.54902 | -0.66178 | -0.53993 | -0.40046 | -0.10215 | -0.42565 | 0.35888 | 0.69642 | 0.42316 | 0.27042 | 0.28762 | 0.13022 | -0.66002 | -0.32095 | -0.48856 | -0.76144 | 0.02371 | -0.28581 | 0.56356 | 1.5055 | -0.061081 | a3 | -2.0715 | -0.77286 | -0.65425 | -0.55463 | 0.13909 | 0.50065 | 0.60626 | 0.54525 | 0.54404 | -0.17711 | 0.1707 | -0.10242 | -0.064964 | 0.39348 | 0.38538 | -0.14333 | -0.78821 | 0.33633 | 0.72601 | 1.5004 | 0.033889
a4 | -2.2474 | -0.15505 | 0.31145 | -0.3267 | 0.093151 | 0.5162 | 0.22746 | 0.42144 | 0.32227 | 0.70181 | -0.11767 | -0.66708 | -0.4319 | -1.0278 | -0.7284 | -0.039143 | 0.082894 | 0.2338 | -0.13142 | 0.038995 | -0.10183 | a4 | -0.10912 | -0.20172 | -0.55108 | -0.29324 | -0.23191 | 0.06779 | -0.060024 | -0.37883 | 0.43804 | -0.017471 | -0.047035 | -0.11958 | 0.25612 | 0.51203 | 0.251 | 0.82689 | -0.47604 | -1.2784 | -0.60446 | 0.47012 | -0.088201
a5 | 4.434 | -0.021173 | -0.099084 | -0.018222 | -0.011422 | 0.037654 | 0.28835 | 0.18383 | 0.13619 | 0.075692 | 0.17137 | -0.17793 | -0.061085 | -0.26032 | 0.091295 | 0.066968 | -0.10119 | 0.022016 | 0.026526 | -0.2206 | 0.14066 | a5 | 3.8881 | 0.033163 | -0.15081 | 0.050838 | 0.0072686 | 0.077415 | 0.10651 | 0.068559 | 0.17095 | -0.078655 | 0.026509 | 0.07902 | 0.035161 | 0.066201 | 0.22513 | 0.22985 | 0.19224 | 0.15542 | -0.054944 | -0.49069 | 0.17945
a6 | -1.003 | -0.081347 | 0.06881 | 0.073378 | 0.033869 | 0.097982 | 0.063216 | -0.0052279 | -0.070473 | -0.16487 | -0.1044 | -0.06347 | -0.13175 | -0.14922 | -0.15247 | -0.16534 | -0.051019 | -0.12937 | 0.0040282 | 0.055077 | -0.075722 | a6 | -0.28052 | -0.04969 | 0.029904 | -0.054492 | -0.031955 | 0.024174 | -0.12619 | -0.0076648 | 0.10222 | -0.022763 | -0.012065 | -0.031577 | -0.0023836 | 0.062203 | 0.09728 | -0.053131 | -0.031753 | 0.03292 | 0.085278 | 0.14984 | -0.004196
a7 | 1.8255 | -0.068586 | -0.023386 | -0.047357 | -0.046181 | 0.025773 | -0.015439 | -0.038117 | -0.092476 | -0.0037205 | 0.044224 | -0.012205 | -0.028499 | -0.017357 | -0.084709 | -0.058386 | 0.043983 | -0.051686 | -0.10295 | -0.055277 | 0.021712 | a7 | 1.6377 | -0.14054 | -0.090497 | -0.13014 | -0.022791 | -0.0084984 | -0.017708 | -0.02393 | 0.033832 | -0.016632 | -0.007181 | 0.011252 | -0.0061272 | 0.014812 | -0.0080425 | 0.084113 | -0.087634 | -0.11487 | -0.030615 | -0.26209 | 0.017559
a8 | -0.19848 | 0.012799 | 0.016495 | 0.012935 | 0.075677 | 0.052667 | 0.075455 | 0.034014 | 0.05635 | 0.033765 | 0.018488 | 0.011058 | 0.0084378 | -0.015021 | 0.020419 | 0.034814 | -0.01633 | -0.049956 | -0.015365 | -0.047441 | 0.01153 | a8 | -0.08709 | 0.022639 | -0.012945 | 0.027237 | 0.031226 | 0.0028105 | -0.0096913 | -0.020448 | 0.0085344 | -0.034444 | 0.077018 | 0.0026043 | 0.021152 | -0.0080385 | 0.048915 | -0.010092 | -0.057073 | 0.032608 | 0.0041926 | -0.070479 | 0.0014474
a9 | -0.14711 | -0.018457 | -0.035112 | -0.029462 | 0.022441 | -0.019416 | -0.016647 | -0.0059678 | 0.0052512 | -0.01412 | -0.068144 | -0.033921 | -0.021426 | -0.010789 | 0.025132 | -0.067095 | -0.027638 | -0.0089762 | 0.0085872 | 0.035893 | -0.019453 | a9 | -0.14037 | -0.023161 | -0.028743 | -0.00572 | 0.0011007 | -0.037429 | -0.032007 | -0.023574 | -0.03502 | -0.04332 | -0.055701 | -0.062372 | -0.037221 | 0.0013536 | 0.0055443 | -0.036805 | -0.041543 | -0.029558 | -0.017141 | 0.05254 | -0.029119
a10 | -0.10058 | -0.026607 | 0.057123 | 0.040603 | 0.062168 | 0.024074 | 0.036759 | -0.0083512 | -0.017226 | 0.020366 | 0.044239 | 0.034632 | 0.031833 | -0.036023 | -0.002835 | 0.013925 | 0.01514 | -0.058276 | -0.080361 | -0.13787 | 0 | a10 | -0.048107 | 0.00017143 | 0.00057971 | 0.052333 | 0.056093 | -0.005596 | 0.0009463 | 0.015042 | 0.020001 | -0.026827 | 0.092951 | 0.0046029 | 0.033089 | -0.017666 | 0.029275 | -0.040518 | -0.055889 | -0.05006 | 0.015429 | -0.15515 | 0
a11 | 0.18403 | -0.03104 | -0.032966 | -0.040681 | -0.034995 | -0.023147 | -0.038217 | -0.026457 | -0.036225 | -0.042858 | -0.033195 | -0.025104 | -0.032266 | -0.03226 | -0.038689 | -0.022334 | -0.031596 | -0.026926 | -0.03802 | -0.021453 | -0.025846 | a11 | 0.18241 | -0.052949 | -0.049089 | -0.052484 | -0.042692 | -0.041856 | -0.03959 | -0.040789 | -0.034956 | -0.048026 | -0.049858 | -0.045621 | -0.048729 | -0.055387 | -0.05387 | -0.037388 | -0.037391 | -0.046211 | -0.065286 | -0.038722 | -0.038857
a12 | -0.021349 | 0.002316 | 0.0081388 | 0.0029591 | 0.0061097 | 0.0053202 | 0.0048314 | -0.00051011 | -0.0046792 | -0.0057989 | 0.00051046 | 0.0042218 | 0.0022705 | -0.0085379 | -0.0031817 | 2.9193e-05 | -0.00053472 | -0.0080238 | -0.0090954 | -0.0071597 | -0.0010717 | a12 | -0.0081328 | -0.0060366 | -0.0025337 | 0.0021489 | 0.0051493 | -0.0019236 | -0.0091322 | -0.0054143 | -0.0015189 | -0.0041216 | 0.0056533 | -0.0023554 | 0.0012909 | -0.00068211 | 0.0082765 | -0.0070968 | -0.0066859 | -0.0072917 | -0.001386 | -0.020723 | -0.0026667
a13 | -0.018214 | -0.0039924 | -0.0069115 | -0.0069406 | -0.0049372 | -0.0048426 | -0.0029812 | -0.0034027 | -0.0047547 | -0.0036008 | -0.0028653 | -0.0022612 | -0.0045482 | -0.0025459 | -0.00058574 | -0.0023275 | -0.0045384 | -0.0030904 | -0.0020969 | -0.0023653 | -0.0041117 | a13 | -0.018396 | -0.0089755 | -0.0059671 | -0.0052145 | -0.0037311 | -0.0043158 | -0.0053928 | -0.0033293 | -0.0042026 | -0.0076453 | -0.0079834 | -0.0067516 | -0.0066111 | -0.0034392 | -0.0031983 | -0.0023933 | -0.0047605 | -0.0051329 | -0.00083309 | 0.0028691 | -0.0051976
a14 | 0.0089566 | -0.0046222 | -0.0027726 | -0.00556 | -0.00099401 | -0.0040243 | -0.0022423 | -0.0020638 | -0.0016767 | -0.0023543 | -0.0038917 | -0.0042724 | -0.0036656 | -0.0068932 | -0.0046829 | -0.0048499 | -0.0039643 | -0.0060386 | -0.0068946 | -0.0036677 | -0.0035797 | a14 | 0.0025276 | 0.00076518 | -0.0039941 | 0.0011487 | -0.0018507 | -0.0018601 | -0.0032878 | -0.0042902 | -0.0023654 | -0.0041664 | 0.00073526 | -0.001185 | -0.00050115 | -0.0032682 | -0.00037454 | -0.00049336 | -0.004102 | -0.0056266 | -0.0075532 | -0.0018865 | -0.0021708
a15 | 0.0087232 | -0.008691 | -0.0026735 | -0.0030477 | -0.004101 | -0.0033184 | -0.0025167 | -0.0019885 | -0.00076529 | -0.0025506 | -0.0021997 | -0.0023831 | -0.0021112 | -0.002998 | -0.0072628 | -0.0060756 | -0.0040052 | -0.0057432 | -0.0053776 | -0.0077032 | -0.0035401 | a15 | 0.0061212 | -0.0079637 | -0.0059766 | -0.0037301 | -0.0020338 | -0.0028834 | -0.0030153 | -0.0029629 | -0.0022573 | -0.005248 | -0.0025438 | -0.0030404 | -0.0031084 | -0.0031348 | -0.0041304 | -0.0064655 | -0.0095234 | -0.0060963 | -0.0022423 | -0.010925 | -0.0041028
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.0003609 | 0.00012276 | -7.4103e-05 | 2.2685e-05 | 5.4172e-05 | -6.396e-05 | -6.2225e-05 | -6.0056e-05 | -0.00017841 | 2.1273e-05 | 3.256e-05 | 2.5889e-05 | -2.7185e-05 | 2.234e-05 | 0.00020253 | -1.5483e-05 | 5.2753e-05 | 2.5798e-05 | 5.6676e-05 | 0.00012077 | 9.3001e-05 | Dx (m) | -0.00014799 | 0.00016565 | -4.3773e-05 | 5.7906e-05 | 1.4169e-05 | -6.7107e-05 | -4.1682e-05 | 1.3699e-05 | -6.4323e-05 | 0 | -4.629e-05 | -0.00013721 | -9.2195e-05 | -4.9424e-05 | 0.00015977 | 2.9166e-05 | -1.2623e-05 | 7.3763e-05 | 6.9024e-05 | 0.00014938 | 0.0002708
Dy (m) | 0.000384 | 6.6455e-05 | -0.00013704 | -9.3116e-05 | -0.00014051 | -5.769e-05 | -8.9563e-05 | 1.7268e-05 | 3.1252e-05 | -4.5927e-05 | -0.00010082 | -7.9871e-05 | -7.5701e-05 | 8.4687e-05 | 1.7214e-05 | -3.2946e-05 | -3.2743e-05 | 0.00013727 | 0.00018862 | 0.00035448 | 0.00033801 | Dy (m) | 0.00015275 | 1.078e-05 | 0 | -0.00010915 | -0.00012011 | 0 | 0 | -3.1162e-05 | -4.5958e-05 | 5.6876e-05 | -0.00020342 | -1.7877e-05 | -7.6961e-05 | 3.4601e-05 | -5.1464e-05 | 8.756e-05 | 0.00012038 | 0.00011198 | -2.7026e-05 | 0.00038377 | 0.00014373
Magnetic length (m) |  | 14.4272 | [19 empty] | Magnetic length (m) |  | 14.4237
Transfer funct. (mT/KA) |  | 594.3281 | [19 empty] | Transfer funct. (mT/KA) |  | 585.9417
Coil positioning (mm) |  | 0.025606 | [19 empty] | Coil positioning (mm) |  | 0.027259
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMB__A001-1000067 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-1000067 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 338.5705 | 597.2261 | 596.5965 | 596.623 | 596.6748 | 596.5186 | 596.424 | 596.5133 | 596.5293 | 596.43 | 596.5011 | 596.5175 | 596.5928 | 596.5816 | 596.5881 | 596.5291 | 596.6849 | 596.6255 | 597.1496 | 408.1437 | 85.6858 | C1/i (mT/KA) | 380.1603 | 588.7107 | 588.2395 | 588.1017 | 588.1211 | 588.1222 | 588.1109 | 588.1391 | 588.2308 | 588.1496 | 588.0949 | 588.0726 | 588.1543 | 588.2552 | 588.2723 | 588.2948 | 588.1207 | 588.305 | 588.8557 | 353.1516 | 84.5563
Angle (mrad) | 2.6949 | -0.57666 | 0.25066 | -0.33146 | 0.34508 | 0.18126 | -0.807 | 0.13659 | 0.056699 | -0.98811 | -0.67048 | 0.56924 | 0.10068 | 0.19599 | -0.0644 | 0.30428 | -0.10051 | -0.000733 | 0.1431 | -0.39405 | 11.9321 | Angle (mrad) | 2.2849 | 0.004472 | 0.157 | 0.045712 | 0.32069 | 0.18378 | 0.038227 | 0.082267 | 0.21682 | -0.031745 | -0.13504 | -0.11842 | 0.033668 | -0.29677 | -0.52724 | -0.52355 | -0.021799 | -0.11071 | -0.20512 | -0.9775 | 23.908
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -0.70659 | 1.3354 | 2.1272 | 2.6657 | 1.2764 | 1.7067 | 2.1719 | 0.67704 | 0.24307 | 1.0362 | 1.808 | 1.1473 | 0.7177 | -0.010715 | 1.2686 | 2.4644 | 1.2451 | 0.54241 | 0.41355 | -0.87915 | 1.1344 | b2 | 0.75169 | 0.5201 | -0.96287 | -0.84756 | 0.055514 | -0.073905 | 0.14862 | -0.033415 | -0.2076 | -0.64255 | 0.39707 | 0.60718 | 0.0052025 | 0.090822 | 0.010949 | 0.24211 | 0.93455 | -0.14489 | 1.0181 | 0.79121 | 0.10794
b3 | 50.054 | -3.4152 | -5.2171 | -4.5903 | -4.265 | -5.2157 | -4.7239 | -4.7533 | -5.1228 | -4.5283 | -5.2845 | -4.7071 | -4.3663 | -4.9473 | -5.7094 | -5.2443 | -5.8863 | -3.9966 | -3.7589 | -5.935 | -3.193 | b3 | 39.4514 | -3.3758 | -5.4526 | -4.4489 | -4.3051 | -5.4885 | -5.5436 | -4.1361 | -5.9759 | -4.0263 | -3.8283 | -4.6593 | -3.8438 | -6.1693 | -4.0477 | -5.3446 | -4.9418 | -5.4368 | -4.1951 | -6.6984 | -3.3148
b4 | 1.7914 | -0.085061 | -0.15902 | -0.056315 | -0.11404 | 0.057101 | 0.14311 | 0.21193 | 0.17882 | -0.10577 | 0.022082 | 0.10558 | 0.16505 | -0.014115 | -0.29917 | 0.10623 | 0.33389 | -0.19675 | -0.15304 | -0.10182 | 0.05633 | b4 | 0.35402 | -0.29824 | -0.020147 | -0.14006 | -0.060007 | -0.093572 | 0.24417 | -0.026847 | -0.74389 | -0.48588 | -0.075458 | -0.088829 | -0.17225 | -0.18361 | -0.071965 | 0.12356 | 0.28688 | 0.011512 | 0.15453 | 0.10208 | -0.070185
b5 | -2.1417 | 0.33731 | 0.11274 | 0.15823 | 0.41849 | 0.49583 | 0.44265 | 0.647 | 0.49947 | 0.64832 | 0.51344 | 0.47225 | 0.44122 | 0.54664 | 0.50256 | 0.27744 | 0.72606 | 0.45672 | 0.6609 | -2.3283 | 0.28862 | b5 | -2.0361 | 0.44987 | 0.34657 | 0.20274 | 0.10337 | 0.22749 | 0.45533 | 0.50721 | 0.038195 | 0.50668 | 0.22922 | 0.28763 | 0.45549 | 0.43712 | 0.50243 | 0.4815 | 0.42225 | 0.6455 | 0.54099 | -3.8421 | 0.16766
b6 | 1.2145 | 0.077859 | 0.050112 | 0.020043 | 0.034535 | 0.12431 | -0.022473 | 0.08784 | 0.045872 | 0.055698 | 0.067635 | 0.01588 | -0.028197 | 0.10037 | 0.0079029 | 0.085288 | -0.10935 | -0.1202 | 0.018961 | -0.078001 | 0.059534 | b6 | 0.067243 | -0.0046806 | 0.00053669 | 0.03156 | -0.046328 | -0.048456 | -0.084988 | -0.089862 | -0.16374 | 0.027732 | -0.082732 | 0.0080737 | 0.0042217 | -0.20356 | -0.11654 | 0.048499 | 0.24636 | 0.035243 | -0.19463 | 0.16983 | -0.025381
b7 | 2.7431 | 1.0617 | 1.2281 | 1.2091 | 1.25 | 1.1819 | 1.1899 | 1.1744 | 1.1864 | 1.164 | 1.1888 | 1.1957 | 1.2288 | 1.2248 | 1.1219 | 1.1581 | 1.2091 | 1.2741 | 1.1189 | 0.8542 | 1.2209 | b7 | 2.7706 | 1.1703 | 1.2524 | 1.2243 | 1.2604 | 1.2649 | 1.2234 | 1.3297 | 1.2163 | 1.2557 | 1.3101 | 1.2912 | 1.2982 | 1.2582 | 1.226 | 1.3437 | 1.2894 | 1.3654 | 1.2176 | 0.71094 | 1.2996
b8 | 0.29336 | -0.020057 | 0.034002 | -0.033804 | -0.07143 | -0.03809 | 0.026926 | 0.00083649 | 0.051134 | 0.0095803 | 0.034991 | -0.0063091 | 0.012623 | -0.037797 | -0.091258 | 0.023971 | -0.032963 | -0.0023505 | -0.015076 | -0.036861 | -0.00074383 | b8 | -0.0035112 | -0.00934 | 0.022586 | -0.012023 | 0.00019227 | 0.046776 | 0.029533 | 0.01479 | -0.016023 | -0.0094486 | 0.00037441 | 0.057512 | 0.033057 | -0.011839 | -0.062753 | -0.0024328 | -0.059687 | -0.021527 | 0.068376 | 0.054034 | 0.0051054
b9 | 0.48941 | 0.61259 | 0.64188 | 0.65843 | 0.6284 | 0.64231 | 0.64877 | 0.64647 | 0.64023 | 0.63127 | 0.63645 | 0.63654 | 0.64709 | 0.62642 | 0.65074 | 0.63413 | 0.58653 | 0.63724 | 0.60007 | 0.50527 | 0.62484 | b9 | 0.52061 | 0.65075 | 0.62602 | 0.65597 | 0.65784 | 0.64779 | 0.64703 | 0.67876 | 0.61617 | 0.67498 | 0.65261 | 0.66157 | 0.64067 | 0.59008 | 0.62903 | 0.63373 | 0.6362 | 0.6383 | 0.62973 | 0.48533 | 0.63364
b10 | 0.17467 | -0.054481 | 0.034365 | -0.0075241 | -0.020638 | 0.028456 | 0.02854 | 0.025526 | 0.076149 | -0.0080276 | -0.012113 | -0.0099222 | 0.013053 | -0.01123 | -0.088333 | 0.0070719 | -0.022206 | -0.013204 | -0.028735 | -0.051943 | 0 | b10 | 0.078488 | -0.077487 | 0.020475 | -0.023453 | -0.0035594 | 0.030839 | 0.019688 | -0.0057195 | 0.030791 | -0.0029825 | 0.027465 | 0.064023 | 0.045026 | 0.021808 | -0.073696 | -0.015649 | 0.003173 | -0.037379 | -0.031597 | -0.070546 | 0
b11 | 0.6237 | 0.73783 | 0.72654 | 0.73127 | 0.73842 | 0.73532 | 0.72442 | 0.7302 | 0.73218 | 0.73415 | 0.73526 | 0.72891 | 0.72641 | 0.73151 | 0.7388 | 0.72899 | 0.73531 | 0.72649 | 0.73524 | 0.66812 | 0.72662 | b11 | 0.71287 | 0.79176 | 0.79076 | 0.78731 | 0.7888 | 0.7867 | 0.79917 | 0.80272 | 0.78891 | 0.80425 | 0.78816 | 0.78755 | 0.78962 | 0.7889 | 0.80163 | 0.7991 | 0.78514 | 0.79068 | 0.8039 | 0.70053 | 0.78749
b12 | 0.030077 | -0.0028639 | 0.0072308 | 0.0040343 | 0.0022144 | 0.0064516 | 0.0060698 | 0.0024868 | 0.0088428 | -0.0043972 | -0.0010761 | 0.00043602 | -0.00014177 | -0.0043667 | -0.013883 | -0.0023529 | -0.0068038 | 0.00018258 | 0.0021119 | -0.00033012 | 0.0010891 | b12 | 0.010252 | -0.003929 | 0.0016366 | -0.003734 | 0.0001494 | 0.0083135 | 0.0023089 | 0.0021156 | 0.0052248 | -0.0018662 | 0.0054262 | 0.0082822 | 0.0075559 | 0.0020545 | -0.010517 | -0.0071164 | -0.0042281 | -0.0092516 | -0.0056739 | 0.0009466 | 0.0002033
b13 | 0.090278 | 0.078095 | 0.075141 | 0.072418 | 0.077501 | 0.072187 | 0.072974 | 0.070278 | 0.071106 | 0.074286 | 0.075268 | 0.072759 | 0.074077 | 0.072779 | 0.069696 | 0.072773 | 0.072422 | 0.069926 | 0.072373 | 0.059279 | 0.073128 | b13 | 0.089363 | 0.080156 | 0.081174 | 0.076731 | 0.081029 | 0.076814 | 0.077414 | 0.078586 | 0.074125 | 0.07933 | 0.082186 | 0.075667 | 0.077784 | 0.075701 | 0.080833 | 0.077936 | 0.073248 | 0.077506 | 0.074073 | 0.065081 | 0.077786
b14 | 0.011113 | -0.00538 | 0.00021587 | 0.00056401 | -0.00091241 | 0.00077771 | -0.0012289 | -0.0010725 | -0.0001097 | -0.003473 | -0.0039373 | -0.0016287 | -0.0021746 | -0.0014151 | -0.0053577 | -0.0043099 | -0.0038714 | -0.0027771 | -0.0036084 | -0.0011332 | -0.0017762 | b14 | 0.0044324 | -0.0080187 | -0.0057148 | -0.0039619 | -0.0036409 | -0.0027044 | -0.0041909 | -0.0033345 | -0.0025133 | -0.0037697 | -0.0040794 | -0.002067 | -0.0030678 | -0.0044389 | -0.0064884 | -0.0061953 | -0.0033617 | -0.005877 | -0.0060318 | -0.0011373 | -0.0040162
b15 | -0.0068901 | 0.020366 | 0.015915 | 0.015455 | 0.017223 | 0.019875 | 0.017404 | 0.020283 | 0.020625 | 0.019919 | 0.018778 | 0.017858 | 0.017394 | 0.019101 | 0.01688 | 0.018522 | 0.017628 | 0.02063 | 0.022331 | 0.014969 | 0.017794 | b15 | -0.0093952 | 0.01383 | 0.019119 | 0.015774 | 0.017489 | 0.018926 | 0.018454 | 0.017154 | 0.019381 | 0.020147 | 0.018894 | 0.018858 | 0.019799 | 0.0215 | 0.016182 | 0.019726 | 0.017497 | 0.017077 | 0.017082 | 0.010073 | 0.016985
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 26.9493 | -5.7666 | 2.5066 | -3.3146 | 3.4508 | 1.8126 | -8.07 | 1.3659 | 0.56699 | -9.8811 | -6.7048 | 5.6924 | 1.0068 | 1.9599 | -0.644 | 3.0428 | -1.0051 | -0.0073297 | 1.431 | -3.9405 | 0 | a1 | 22.8485 | 0.04472 | 1.57 | 0.45712 | 3.2069 | 1.8378 | 0.38227 | 0.82267 | 2.1682 | -0.31745 | -1.3504 | -1.1842 | 0.33668 | -2.9677 | -5.2724 | -5.2355 | -0.21799 | -1.1071 | -2.0512 | -9.7751 | 0
a2 | -7.5073 | 0.070112 | -2.2191 | -0.14252 | 0.40903 | -1.1189 | -0.75093 | -0.34567 | 0.050851 | 0.47404 | -0.34661 | -0.49164 | 0.45257 | 0.38258 | 0.51961 | -1.5648 | -0.5412 | -0.60805 | 0.34875 | 1.6016 | -0.44547 | a2 | -2.9633 | 1.2486 | 0.97083 | -0.053474 | -0.49887 | 0.47165 | -0.52902 | -0.92761 | 1.9894 | -0.82926 | -0.04837 | -0.11219 | -0.90477 | 0.73085 | -0.81944 | -1.5465 | -0.52549 | 1.373 | -0.54116 | 0.65758 | -0.10746
a3 | -0.54902 | -0.66178 | -0.53993 | -0.40046 | -0.10215 | -0.42565 | 0.35888 | 0.69642 | 0.42316 | 0.27042 | 0.28762 | 0.13022 | -0.66002 | -0.32095 | -0.48856 | -0.76144 | 0.02371 | -0.28581 | 0.56356 | 1.5055 | -0.061081 | a3 | -2.0715 | -0.77286 | -0.65425 | -0.55463 | 0.13909 | 0.50065 | 0.60626 | 0.54525 | 0.54404 | -0.17711 | 0.1707 | -0.10242 | -0.064964 | 0.39348 | 0.38538 | -0.14333 | -0.78821 | 0.33633 | 0.72601 | 1.5004 | 0.033889
a4 | -2.2474 | -0.15505 | 0.31145 | -0.3267 | 0.093151 | 0.5162 | 0.22746 | 0.42144 | 0.32227 | 0.70181 | -0.11767 | -0.66708 | -0.4319 | -1.0278 | -0.7284 | -0.039143 | 0.082894 | 0.2338 | -0.13142 | 0.038995 | -0.10183 | a4 | -0.10912 | -0.20172 | -0.55108 | -0.29324 | -0.23191 | 0.06779 | -0.060024 | -0.37883 | 0.43804 | -0.017471 | -0.047035 | -0.11958 | 0.25612 | 0.51203 | 0.251 | 0.82689 | -0.47604 | -1.2784 | -0.60446 | 0.47012 | -0.088201
a5 | 4.434 | -0.021173 | -0.099084 | -0.018222 | -0.011422 | 0.037654 | 0.28835 | 0.18383 | 0.13619 | 0.075692 | 0.17137 | -0.17793 | -0.061085 | -0.26032 | 0.091295 | 0.066968 | -0.10119 | 0.022016 | 0.026526 | -0.2206 | 0.14066 | a5 | 3.8881 | 0.033163 | -0.15081 | 0.050838 | 0.0072686 | 0.077415 | 0.10651 | 0.068559 | 0.17095 | -0.078655 | 0.026509 | 0.07902 | 0.035161 | 0.066201 | 0.22513 | 0.22985 | 0.19224 | 0.15542 | -0.054944 | -0.49069 | 0.17945
a6 | -1.003 | -0.081347 | 0.06881 | 0.073378 | 0.033869 | 0.097982 | 0.063216 | -0.0052279 | -0.070473 | -0.16487 | -0.1044 | -0.06347 | -0.13175 | -0.14922 | -0.15247 | -0.16534 | -0.051019 | -0.12937 | 0.0040282 | 0.055077 | -0.075722 | a6 | -0.28052 | -0.04969 | 0.029904 | -0.054492 | -0.031955 | 0.024174 | -0.12619 | -0.0076648 | 0.10222 | -0.022763 | -0.012065 | -0.031577 | -0.0023836 | 0.062203 | 0.09728 | -0.053131 | -0.031753 | 0.03292 | 0.085278 | 0.14984 | -0.004196
a7 | 1.8255 | -0.068586 | -0.023386 | -0.047357 | -0.046181 | 0.025773 | -0.015439 | -0.038117 | -0.092476 | -0.0037205 | 0.044224 | -0.012205 | -0.028499 | -0.017357 | -0.084709 | -0.058386 | 0.043983 | -0.051686 | -0.10295 | -0.055277 | 0.021712 | a7 | 1.6377 | -0.14054 | -0.090497 | -0.13014 | -0.022791 | -0.0084984 | -0.017708 | -0.02393 | 0.033832 | -0.016632 | -0.007181 | 0.011252 | -0.0061272 | 0.014812 | -0.0080425 | 0.084113 | -0.087634 | -0.11487 | -0.030615 | -0.26209 | 0.017559
a8 | -0.19848 | 0.012799 | 0.016495 | 0.012935 | 0.075677 | 0.052667 | 0.075455 | 0.034014 | 0.05635 | 0.033765 | 0.018488 | 0.011058 | 0.0084378 | -0.015021 | 0.020419 | 0.034814 | -0.01633 | -0.049956 | -0.015365 | -0.047441 | 0.01153 | a8 | -0.08709 | 0.022639 | -0.012945 | 0.027237 | 0.031226 | 0.0028105 | -0.0096913 | -0.020448 | 0.0085344 | -0.034444 | 0.077018 | 0.0026043 | 0.021152 | -0.0080385 | 0.048915 | -0.010092 | -0.057073 | 0.032608 | 0.0041926 | -0.070479 | 0.0014474
a9 | -0.14711 | -0.018457 | -0.035112 | -0.029462 | 0.022441 | -0.019416 | -0.016647 | -0.0059678 | 0.0052512 | -0.01412 | -0.068144 | -0.033921 | -0.021426 | -0.010789 | 0.025132 | -0.067095 | -0.027638 | -0.0089762 | 0.0085872 | 0.035893 | -0.019453 | a9 | -0.14037 | -0.023161 | -0.028743 | -0.00572 | 0.0011007 | -0.037429 | -0.032007 | -0.023574 | -0.03502 | -0.04332 | -0.055701 | -0.062372 | -0.037221 | 0.0013536 | 0.0055443 | -0.036805 | -0.041543 | -0.029558 | -0.017141 | 0.05254 | -0.029119
a10 | -0.10058 | -0.026607 | 0.057123 | 0.040603 | 0.062168 | 0.024074 | 0.036759 | -0.0083512 | -0.017226 | 0.020366 | 0.044239 | 0.034632 | 0.031833 | -0.036023 | -0.002835 | 0.013925 | 0.01514 | -0.058276 | -0.080361 | -0.13787 | 0 | a10 | -0.048107 | 0.00017143 | 0.00057971 | 0.052333 | 0.056093 | -0.005596 | 0.0009463 | 0.015042 | 0.020001 | -0.026827 | 0.092951 | 0.0046029 | 0.033089 | -0.017666 | 0.029275 | -0.040518 | -0.055889 | -0.05006 | 0.015429 | -0.15515 | 0
a11 | 0.18403 | -0.03104 | -0.032966 | -0.040681 | -0.034995 | -0.023147 | -0.038217 | -0.026457 | -0.036225 | -0.042858 | -0.033195 | -0.025104 | -0.032266 | -0.03226 | -0.038689 | -0.022334 | -0.031596 | -0.026926 | -0.03802 | -0.021453 | -0.025846 | a11 | 0.18241 | -0.052949 | -0.049089 | -0.052484 | -0.042692 | -0.041856 | -0.03959 | -0.040789 | -0.034956 | -0.048026 | -0.049858 | -0.045621 | -0.048729 | -0.055387 | -0.05387 | -0.037388 | -0.037391 | -0.046211 | -0.065286 | -0.038722 | -0.038857
a12 | -0.021349 | 0.002316 | 0.0081388 | 0.0029591 | 0.0061097 | 0.0053202 | 0.0048314 | -0.00051011 | -0.0046792 | -0.0057989 | 0.00051046 | 0.0042218 | 0.0022705 | -0.0085379 | -0.0031817 | 2.9193e-05 | -0.00053472 | -0.0080238 | -0.0090954 | -0.0071597 | -0.0010717 | a12 | -0.0081328 | -0.0060366 | -0.0025337 | 0.0021489 | 0.0051493 | -0.0019236 | -0.0091322 | -0.0054143 | -0.0015189 | -0.0041216 | 0.0056533 | -0.0023554 | 0.0012909 | -0.00068211 | 0.0082765 | -0.0070968 | -0.0066859 | -0.0072917 | -0.001386 | -0.020723 | -0.0026667
a13 | -0.018214 | -0.0039924 | -0.0069115 | -0.0069406 | -0.0049372 | -0.0048426 | -0.0029812 | -0.0034027 | -0.0047547 | -0.0036008 | -0.0028653 | -0.0022612 | -0.0045482 | -0.0025459 | -0.00058574 | -0.0023275 | -0.0045384 | -0.0030904 | -0.0020969 | -0.0023653 | -0.0041117 | a13 | -0.018396 | -0.0089755 | -0.0059671 | -0.0052145 | -0.0037311 | -0.0043158 | -0.0053928 | -0.0033293 | -0.0042026 | -0.0076453 | -0.0079834 | -0.0067516 | -0.0066111 | -0.0034392 | -0.0031983 | -0.0023933 | -0.0047605 | -0.0051329 | -0.00083309 | 0.0028691 | -0.0051976
a14 | 0.0089566 | -0.0046222 | -0.0027726 | -0.00556 | -0.00099401 | -0.0040243 | -0.0022423 | -0.0020638 | -0.0016767 | -0.0023543 | -0.0038917 | -0.0042724 | -0.0036656 | -0.0068932 | -0.0046829 | -0.0048499 | -0.0039643 | -0.0060386 | -0.0068946 | -0.0036677 | -0.0035797 | a14 | 0.0025276 | 0.00076518 | -0.0039941 | 0.0011487 | -0.0018507 | -0.0018601 | -0.0032878 | -0.0042902 | -0.0023654 | -0.0041664 | 0.00073526 | -0.001185 | -0.00050115 | -0.0032682 | -0.00037454 | -0.00049336 | -0.004102 | -0.0056266 | -0.0075532 | -0.0018865 | -0.0021708
a15 | 0.0087232 | -0.008691 | -0.0026735 | -0.0030477 | -0.004101 | -0.0033184 | -0.0025167 | -0.0019885 | -0.00076529 | -0.0025506 | -0.0021997 | -0.0023831 | -0.0021112 | -0.002998 | -0.0072628 | -0.0060756 | -0.0040052 | -0.0057432 | -0.0053776 | -0.0077032 | -0.0035401 | a15 | 0.0061212 | -0.0079637 | -0.0059766 | -0.0037301 | -0.0020338 | -0.0028834 | -0.0030153 | -0.0029629 | -0.0022573 | -0.005248 | -0.0025438 | -0.0030404 | -0.0031084 | -0.0031348 | -0.0041304 | -0.0064655 | -0.0095234 | -0.0060963 | -0.0022423 | -0.010925 | -0.0041028
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.3609 | -0.12276 | 0.074103 | -0.022685 | -0.054172 | 0.06396 | 0.062225 | 0.060056 | 0.17841 | -0.021273 | -0.03256 | -0.025889 | 0.027185 | -0.02234 | -0.20253 | 0.015483 | -0.052753 | -0.025798 | -0.056676 | -0.12077 |  | Dx (mm) | 0.14799 | -0.16565 | 0.043772 | -0.057906 | -0.014169 | 0.067107 | 0.041682 | -0.013699 | 0.064323 | 0 | 0.04629 | 0.13721 | 0.092195 | 0.049424 | -0.15977 | -0.029166 | 0.012623 | -0.073763 | -0.069024 | -0.14938
Dy (mm) | -0.384 | -0.066455 | 0.13704 | 0.093116 | 0.14051 | 0.05769 | 0.089563 | -0.017268 | -0.031252 | 0.045927 | 0.10082 | 0.079871 | 0.075701 | -0.084687 | -0.017214 | 0.032946 | 0.032743 | -0.13727 | -0.18862 | -0.35448 |  | Dy (mm) | -0.15275 | -0.01078 | 0 | 0.10915 | 0.12011 | 0 | 0 | 0.031162 | 0.045958 | -0.056876 | 0.20342 | 0.017877 | 0.076961 | -0.034601 | 0.051464 | -0.08756 | -0.12038 | -0.11198 | 0.027026 | -0.38377
Mag. len. (m) | 14.4272 | [20 empty] | Mag. len. (m) | 14.4237
TF (mT/KA) | 596.6281 | [20 empty] | TF (mT/KA) | 588.2417
Coil pos. (mm) | 0.025606 | [20 empty] | Coil pos. (mm) | 0.027259
Current (A) | 10 | [20 empty] | Current (A) | 10

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 1.1344 | 0.96672 | 1.2636 | 0.77029 | -0.44547 | 1.8354 | -0.28923 | 0.74239 | 0.10794 | 0.55084 | 0.065169 | 0.53666 | -0.10746 | 1.1344 | -0.026378 | 0.95646 |  | 0
3 | -3.193 | 12.2893 | -4.7632 | 0.64771 | -0.061081 | 0.57712 | -0.10908 | 0.46085 | -3.3148 | 9.9438 | -4.7335 | 0.82352 | 0.033889 | 0.75231 | 0.059892 | 0.4931 |  | 0
4 | 0.05633 | 0.43079 | 0.0093553 | 0.16195 | -0.10183 | 0.65961 | -0.040889 | 0.45875 | -0.070185 | 0.25552 | -0.090067 | 0.24453 | -0.088201 | 0.47343 | -0.10522 | 0.47974 |  | 0
5 | 0.28862 | 0.84578 | 0.46461 | 0.1634 | 0.14066 | 0.99942 | 0.018207 | 0.13327 | 0.16766 | 1.074 | 0.37999 | 0.16307 | 0.17945 | 0.87501 | 0.070202 | 0.10367 |  | 0
6 | 0.059534 | 0.27472 | 0.032068 | 0.068115 | -0.075722 | 0.23185 | -0.060279 | 0.084951 | -0.025381 | 0.11192 | -0.034895 | 0.11185 | -0.004196 | 0.09233 | -0.0063405 | 0.070193 |  | 0
7 | 1.2209 | 0.36281 | 1.1869 | 0.051874 | 0.021712 | 0.41763 | -0.032455 | 0.042776 | 1.2996 | 0.36775 | 1.2667 | 0.050816 | 0.017559 | 0.38353 | -0.031186 | 0.059093 |  | 0
8 | -0.00074383 | 0.077005 | -0.0067381 | 0.025961 | 0.01153 | 0.059194 | 0.014501 | 0.028027 | 0.0051054 | 0.035378 | 0.0041789 | 0.031935 | 0.0014474 | 0.039711 | 0.0028381 | 0.022153 |  | 0
9 | 0.62484 | 0.045386 | 0.63404 | 0.017529 | -0.019453 | 0.039801 | -0.018199 | 0.025245 | 0.63364 | 0.047734 | 0.6427 | 0.021786 | -0.029119 | 0.036454 | -0.028194 | 0.020679 |  | 0
10 | 0 | 0.054139 | 5.7782e-06 | 2.1709e-05 | 0 | 0.054624 | -5.9553e-07 | 1.9124e-05 | 0 | 0.043462 | -7.8655e-05 | 0.00033011 | 0 | 0.0522 | -7.8343e-06 | 0.001154 |  | 0
11 | 0.72662 | 0.027805 | 0.73209 | 0.0043129 | -0.025846 | 0.048707 | -0.032732 | 0.0060328 | 0.78749 | 0.027388 | 0.79298 | 0.0065917 | -0.038857 | 0.051701 | -0.046892 | 0.007613 |  | 0
12 | 0.0010891 | 0.0085241 | 0.0006052 | 0.0032473 | -0.0010717 | 0.0070116 | -0.0012272 | 0.003022 | 0.0002033 | 0.0060736 | -0.00012208 | 0.0030106 | -0.0026667 | 0.0063584 | -0.0027166 | 0.0031942 |  | 0
13 | 0.073128 | 0.0055118 | 0.073029 | 0.0022701 | -0.0041117 | 0.003616 | -0.0038245 | 0.0015215 | 0.077786 | 0.0046656 | 0.077756 | 0.0025525 | -0.0051976 | 0.0040642 | -0.0049633 | 0.0020106 |  | 0
14 | -0.0017762 | 0.0035092 | -0.0020766 | 0.0013179 | -0.0035797 | 0.0033169 | -0.0042885 | 0.0014505 | -0.0040162 | 0.002576 | -0.0044349 | 0.0010137 | -0.0021708 | 0.0025027 | -0.0026089 | 0.0020033 |  | 0
15 | 0.017794 | 0.0059937 | 0.018677 | 0.0018467 | -0.0035401 | 0.0035625 | -0.0037672 | 0.0020765 | 0.016985 | 0.006562 | 0.01816 | 0.0018301 | -0.0041028 | 0.0034826 | -0.004242 | 0.0021565 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.27943 | -0.33345
3 | [12 empty] | 3 | 9 | -0.6045 | -0.45061
4 | [12 empty] | 4 | 16 | -1.2999 | -1.0715
5 | [12 empty] | 5 | 25 | -1.9135 | -2.0401
6 | [12 empty] | 6 | 36 | -2.5761 | -2.4235
7 | [12 empty] | 7 | 49 | -3.0554 | -2.9041
8 | [12 empty] | 8 | 64 | -3.6129 | -3.6269
9 | [12 empty] | 9 | 81 | -3.8615 | -3.8526
10 | [12 empty] | sigma (mm) |  | 0.025606 | 0.027259
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.61258 | 0.15315 | 0.036755 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.35886 | 0.089715 | 0.021532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.20767 | 0.051917 | 0.01246 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.11871 | 0.029678 | 0.0071228 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.067037 | 0.016759 | 0.0040222 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.037395 | 0.0093488 | 0.0022437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.020607 | 0.0051516 | 0.0012364 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.011217 | 0.0028042 | 0.00067302 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0060316 | 0.0015079 | 0.0003619 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0032039 | 0.00080097 | 0.00019223 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0016811 | 0.00042029 | 0.00010087 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0008714 | 0.00021785 | 5.2284e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -2.807 | 1.2913 | 2.164 | 2.649 | 1.2475 | 1.7431 | 2.2102 | 0.70921 | 0.34904 | 1.0263 | 1.7911 | 1.1342 | 0.72579 | -0.020521 | 1.1335 | 2.471 | 1.2087 | 0.53489 | 0.37598 | -1.0263 | 1.1344
3 | 50.0922 | -3.4153 | -5.2075 | -4.5959 | -4.2638 | -5.2111 | -4.7219 | -4.7568 | -5.1302 | -4.523 | -5.2865 | -4.716 | -4.3728 | -4.932 | -5.7179 | -5.2449 | -5.8827 | -4.0032 | -3.7561 | -5.9397 | -3.193
4 | 1.5726 | -0.074987 | -0.16418 | -0.05587 | -0.10909 | 0.05015 | 0.1427 | 0.20204 | 0.15685 | -0.10171 | 0.030081 | 0.10512 | 0.16114 | -0.0060547 | -0.27559 | 0.10573 | 0.34212 | -0.19468 | -0.1454 | -0.14959 | 0.05633
5 | -2.1573 | 0.34171 | 0.11442 | 0.16038 | 0.42044 | 0.49515 | 0.44473 | 0.64548 | 0.49771 | 0.64644 | 0.51099 | 0.47088 | 0.43851 | 0.55102 | 0.50381 | 0.27545 | 0.72387 | 0.46103 | 0.66099 | -2.3368 | 0.28862
6 | 0.61765 | 0.12547 | 0.01686 | 0.028168 | 0.056144 | 0.098152 | -0.049094 | 0.06318 | -0.027817 | 0.064378 | 0.08287 | 0.026462 | -0.040749 | 0.11054 | 0.088611 | 0.07828 | -0.086325 | -0.1061 | 0.048195 | -0.034674 | 0.059534
7 | 2.7309 | 1.0603 | 1.228 | 1.2093 | 1.2528 | 1.1842 | 1.192 | 1.1741 | 1.182 | 1.1648 | 1.19 | 1.196 | 1.2289 | 1.225 | 1.1141 | 1.1584 | 1.2082 | 1.2769 | 1.1197 | 0.85929 | 1.2209
8 | 0.23683 | 0.015909 | 0.0093537 | -0.028066 | -0.053927 | -0.057949 | 0.0072267 | -0.017385 | -0.0026954 | 0.015595 | 0.041509 | 0.00017084 | 0.0035811 | -0.030782 | -0.029439 | 0.01831 | -0.018829 | 0.0059655 | 0.00016636 | -0.014132 | -0.00074383
9 | 0.47649 | 0.60998 | 0.64467 | 0.66034 | 0.63243 | 0.64208 | 0.64957 | 0.64573 | 0.63332 | 0.63168 | 0.6386 | 0.63787 | 0.64818 | 0.6279 | 0.6413 | 0.63431 | 0.58617 | 0.64129 | 0.60723 | 0.52782 | 0.62484
10 | 0.00069554 | 1.1726e-05 | 3.6974e-05 | 5.5517e-06 | -3.7553e-07 | 5.4262e-06 | 1.0984e-05 | -8.7594e-07 | -2.6181e-05 | 1.1778e-06 | 6.3477e-07 | -1.3264e-06 | 1.0786e-07 | -1.0109e-05 | 7.7967e-05 | -5.2716e-07 | 2.6917e-06 | -5.3402e-06 | -4.498e-06 | -4.2087e-06 | 0
11 | 0.62198 | 0.7375 | 0.72692 | 0.7315 | 0.73905 | 0.73525 | 0.72446 | 0.73011 | 0.73126 | 0.73391 | 0.73527 | 0.72913 | 0.72653 | 0.73192 | 0.73702 | 0.72902 | 0.73507 | 0.7272 | 0.73643 | 0.66973 | 0.72662
12 | 0.12016 | 0.040906 | 0.026317 | 0.047377 | 0.046883 | 0.029953 | 0.026761 | -0.0045102 | -7.2462e-05 | -0.033985 | 0.0044992 | 0.016382 | -0.018063 | -0.030668 | -0.03912 | -0.032024 | -0.042119 | 0.017554 | 0.052865 | 0.053154 | 0.010891 | ten times!
13 | 0.84581 | 0.77825 | 0.74838 | 0.72032 | 0.77356 | 0.71971 | 0.72879 | 0.70355 | 0.71161 | 0.74147 | 0.7487 | 0.72465 | 0.7391 | 0.73202 | 0.68928 | 0.72702 | 0.72166 | 0.70505 | 0.73211 | 0.60168 | 0.73128 | ten times!
14 | 0.10402 | -0.028454 | -0.010571 | 0.0061903 | -0.0061859 | -0.0042681 | -0.023064 | -0.020474 | -0.031204 | -0.032205 | -0.036164 | -0.014047 | -0.026956 | -0.0085461 | -0.024393 | -0.047109 | -0.032136 | -0.016896 | -0.017308 | 0.026044 | -0.017762 | ten times!
15 | -0.068901 | 0.20366 | 0.15915 | 0.15455 | 0.17223 | 0.19875 | 0.17404 | 0.20283 | 0.20625 | 0.19919 | 0.18778 | 0.17858 | 0.17394 | 0.19101 | 0.1688 | 0.18522 | 0.17628 | 0.2063 | 0.22331 | 0.14969 | 0.17794 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -5.2227 | 0.033853 | -2.1303 | -0.093301 | 0.47888 | -1.0803 | -0.70378 | -0.36024 | 0.023134 | 0.49918 | -0.28282 | -0.44701 | 0.49357 | 0.33244 | 0.49641 | -1.5431 | -0.51838 | -0.67346 | 0.2691 | 1.3755 | -0.44547
3 | -0.2845 | -0.66614 | -0.54016 | -0.40084 | -0.098432 | -0.43206 | 0.35412 | 0.6926 | 0.414 | 0.27391 | 0.28655 | 0.12568 | -0.66015 | -0.32521 | -0.51551 | -0.76195 | 0.022552 | -0.28952 | 0.55715 | 1.5 | -0.061081
4 | -2.8175 | -0.15039 | 0.30954 | -0.33026 | 0.079169 | 0.5089 | 0.21391 | 0.42147 | 0.32023 | 0.69519 | -0.12854 | -0.67704 | -0.43937 | -1.0183 | -0.72201 | -0.041537 | 0.076044 | 0.24869 | -0.10173 | -0.16147 | -0.10183
5 | 4.6776 | -0.022588 | -0.1026 | -0.018282 | -0.01231 | 0.033702 | 0.28779 | 0.18437 | 0.14031 | 0.073908 | 0.16837 | -0.17878 | -0.059403 | -0.2588 | 0.082253 | 0.066895 | -0.10093 | 0.016181 | 0.027645 | -0.22677 | 0.14066
6 | -0.86377 | -0.059417 | 0.01002 | 0.033261 | -0.029003 | 0.073335 | 0.02594 | 0.0027374 | -0.051563 | -0.18377 | -0.14619 | -0.097288 | -0.1643 | -0.11275 | -0.15171 | -0.17849 | -0.064173 | -0.068116 | 0.076455 | 0.15959 | -0.075722
7 | 1.9014 | -0.068488 | -0.025808 | -0.04594 | -0.04036 | 0.025291 | -0.018365 | -0.038952 | -0.095958 | -0.0036059 | 0.043019 | -0.01188 | -0.028987 | -0.018814 | -0.083653 | -0.058934 | 0.044073 | -0.05235 | -0.10448 | -0.063017 | 0.021712
8 | -0.08506 | 0.03089 | -0.023675 | -0.016231 | 0.034698 | 0.035813 | 0.048599 | 0.039436 | 0.065325 | 0.01998 | -0.012753 | -0.01328 | -0.01434 | 0.0098301 | 0.028086 | 0.025471 | -0.026054 | -0.0088998 | 0.038128 | 0.038887 | 0.01153
9 | -0.092387 | -0.022103 | -0.039846 | -0.028603 | 0.02576 | -0.0211 | -0.019211 | -0.0054689 | 0.0081381 | -0.013695 | -0.066735 | -0.033027 | -0.022408 | -0.011719 | 0.024023 | -0.067333 | -0.026831 | -0.010732 | 0.0033064 | 0.01733 | -0.019453
10 | 0.0012323 | -5.2387e-06 | -7.2831e-06 | 5.5102e-06 | 1.9934e-05 | -8.0839e-06 | -7.8423e-06 | 7.0544e-07 | 3.5745e-05 | -3.5967e-06 | -2.9489e-06 | 3.7432e-06 | 1.4768e-06 | -6.1799e-06 | 3.6732e-05 | 4.263e-07 | -3.5307e-06 | -2.0961e-05 | -4.9326e-05 | -7.3518e-05 | 0
11 | 0.19649 | -0.030979 | -0.033998 | -0.04088 | -0.034983 | -0.023608 | -0.038763 | -0.026409 | -0.035506 | -0.042808 | -0.033114 | -0.025056 | -0.032299 | -0.032623 | -0.03926 | -0.022284 | -0.03147 | -0.027044 | -0.038095 | -0.022089 | -0.025846 | ten times!
12 | 0.077613 | 0.056335 | 0.012317 | -0.01912 | -0.017659 | 0.025992 | 0.0034887 | 0.0049079 | -0.025117 | -0.082613 | -0.049122 | 0.00078385 | -0.016006 | -0.042274 | -0.024186 | -0.016378 | -0.023776 | -0.013043 | 0.0045695 | 0.074716 | -0.010717 | ten times!
13 | -0.17423 | -0.046997 | -0.06777 | -0.070773 | -0.048803 | -0.046801 | -0.027903 | -0.033221 | -0.045286 | -0.035171 | -0.026586 | -0.022463 | -0.043461 | -0.027553 | -0.013815 | -0.021615 | -0.046014 | -0.035011 | -0.029162 | -0.030112 | -0.041117 | ten times!
14 | 0.041851 | -0.043862 | -0.044056 | -0.068021 | -0.0317 | -0.047938 | -0.03397 | -0.01677 | -0.010334 | -0.031524 | -0.055098 | -0.054978 | -0.047027 | -0.056162 | -0.05655 | -0.052749 | -0.046137 | -0.038285 | -0.036767 | -0.00064122 | -0.035797 | ten times!
15 | 0.087232 | -0.08691 | -0.026735 | -0.030477 | -0.04101 | -0.033184 | -0.025167 | -0.019885 | -0.0076529 | -0.025506 | -0.021997 | -0.023831 | -0.021112 | -0.02998 | -0.072628 | -0.060756 | -0.040052 | -0.057432 | -0.053776 | -0.077032 | -0.035401 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 0.10203 | 0.45529 | -0.93479 | -0.88499 | 0.050303 | -0.030574 | 0.17581 | -0.038082 | -0.15943 | -0.64137 | 0.422 | 0.68218 | 0.046306 | 0.12509 | -0.062799 | 0.22525 | 0.95305 | -0.1965 | 0.9863 | 0.60575 | 0.10794
3 | 39.4451 | -3.3841 | -5.4525 | -4.456 | -4.3101 | -5.4874 | -5.5454 | -4.1383 | -5.9639 | -4.0261 | -3.8293 | -4.6576 | -3.8376 | -6.1708 | -4.0475 | -5.3567 | -4.9323 | -5.4114 | -4.1962 | -6.7276 | -3.3148
4 | 0.28517 | -0.28079 | -0.023717 | -0.13599 | -0.059457 | -0.097164 | 0.2397 | -0.024709 | -0.74262 | -0.48483 | -0.076686 | -0.097783 | -0.18149 | -0.18923 | -0.050351 | 0.12213 | 0.28018 | 0.01862 | 0.16297 | 0.011347 | -0.070185
5 | -2.0264 | 0.4498 | 0.34656 | 0.20153 | 0.10205 | 0.22845 | 0.45638 | 0.50678 | 0.042674 | 0.50706 | 0.22962 | 0.28713 | 0.45532 | 0.43945 | 0.49843 | 0.48329 | 0.42246 | 0.64518 | 0.53771 | -3.8516 | 0.16766
6 | -0.16576 | 0.064278 | -0.018812 | 0.051569 | -0.040991 | -0.078414 | -0.10299 | -0.083697 | -0.1908 | 0.028066 | -0.10465 | -0.054385 | -0.038186 | -0.22569 | -0.047554 | 0.059731 | 0.24434 | 0.075329 | -0.16526 | 0.24281 | -0.025381
7 | 2.7763 | 1.1696 | 1.252 | 1.2253 | 1.2619 | 1.2636 | 1.2229 | 1.3295 | 1.2169 | 1.2565 | 1.3166 | 1.2879 | 1.2976 | 1.2586 | 1.2229 | 1.344 | 1.2926 | 1.3632 | 1.2196 | 0.7254 | 1.2996
8 | -0.029678 | 0.041506 | 0.0096907 | 0.0055579 | 0.0046409 | 0.026319 | 0.016841 | 0.01882 | -0.035431 | -0.0082892 | -0.019174 | 0.014269 | 0.003913 | -0.025586 | -0.015327 | 0.0077818 | -0.061112 | 0.0021876 | 0.088613 | 0.078661 | 0.0051054
9 | 0.51835 | 0.64395 | 0.62554 | 0.65827 | 0.66138 | 0.64669 | 0.6466 | 0.67896 | 0.6156 | 0.67579 | 0.66195 | 0.65696 | 0.63982 | 0.58983 | 0.62359 | 0.63537 | 0.63974 | 0.63981 | 0.62879 | 0.51127 | 0.63364
10 | 3.8972e-05 | -1.0397e-07 | 0.00011394 | -4.8547e-06 | -9.9349e-07 | -0.0002156 | 9.3729e-05 | 1.3254e-06 | -4.1049e-06 | -0.0013757 | 3.747e-05 | -2.2743e-05 | -2.7571e-06 | -2.6325e-07 | 1.0304e-05 | -1.3825e-05 | -9.096e-06 | -2.7285e-05 | 4.7631e-06 | -0.00025085 | 0
11 | 0.71269 | 0.79138 | 0.79072 | 0.78732 | 0.7892 | 0.78634 | 0.79911 | 0.80263 | 0.78864 | 0.8044 | 0.78874 | 0.78679 | 0.78924 | 0.78885 | 0.80082 | 0.79937 | 0.78569 | 0.79077 | 0.80362 | 0.70576 | 0.78749
12 | 0.028999 | 0.055119 | -0.0087149 | -0.009994 | 0.0064352 | 0.046749 | 0.00031213 | 0.028023 | 0.017229 | -0.015593 | 0.015944 | 0.0086816 | 0.021346 | -0.0050252 | -0.015174 | -0.05364 | -0.044763 | -0.048103 | -0.020808 | 0.070316 | 0.002033 | ten times!
13 | 0.88566 | 0.79134 | 0.81365 | 0.76652 | 0.80819 | 0.76952 | 0.77547 | 0.78448 | 0.74165 | 0.79511 | 0.82445 | 0.75867 | 0.77971 | 0.75955 | 0.80026 | 0.77831 | 0.73658 | 0.77656 | 0.73598 | 0.65505 | 0.77786 | ten times!
14 | 0.048075 | -0.060614 | -0.06404 | -0.03545 | -0.03638 | -0.037503 | -0.048244 | -0.03217 | -0.036254 | -0.035239 | -0.052258 | -0.042427 | -0.04768 | -0.052246 | -0.045344 | -0.052553 | -0.025995 | -0.042775 | -0.051107 | 0.035547 | -0.040162 | ten times!
15 | -0.093952 | 0.1383 | 0.19119 | 0.15774 | 0.17489 | 0.18926 | 0.18454 | 0.17154 | 0.19381 | 0.20147 | 0.18894 | 0.18858 | 0.19799 | 0.215 | 0.16182 | 0.19726 | 0.17497 | 0.17077 | 0.17082 | 0.10073 | 0.16985 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -2.2183 | 1.2293 | 0.9742 | -0.0001207 | -0.43781 | 0.4677 | -0.53199 | -0.91157 | 2.0176 | -0.85621 | 0.042318 | -0.10074 | -0.86926 | 0.70345 | -0.78769 | -1.602 | -0.59431 | 1.3043 | -0.52193 | 0.38152 | -0.10746
3 | -2.0591 | -0.77933 | -0.65 | -0.55493 | 0.13979 | 0.49985 | 0.6067 | 0.54448 | 0.5451 | -0.18198 | 0.17379 | -0.099247 | -0.066792 | 0.38789 | 0.39311 | -0.13717 | -0.78105 | 0.31992 | 0.71791 | 1.5198 | 0.033889
4 | -0.31769 | -0.19929 | -0.54953 | -0.29776 | -0.2348 | 0.066567 | -0.061069 | -0.38233 | 0.43504 | -0.010691 | -0.058294 | -0.12334 | 0.24711 | 0.51482 | 0.25338 | 0.83839 | -0.46465 | -1.2587 | -0.60879 | 0.10594 | -0.088201
5 | 3.9034 | 0.030728 | -0.1512 | 0.048897 | 0.0087719 | 0.076938 | 0.10806 | 0.069352 | 0.17123 | -0.078191 | 0.031623 | 0.080252 | 0.03513 | 0.063225 | 0.23146 | 0.23065 | 0.20108 | 0.1573 | -0.051666 | -0.46494 | 0.17945
6 | -0.21669 | -0.053454 | 0.031302 | -0.10432 | -0.085498 | 0.024375 | -0.12593 | -0.022405 | 0.081721 | 0.0024428 | -0.10601 | -0.040269 | -0.037446 | 0.07731 | 0.074558 | -0.010742 | 0.023421 | 0.083892 | 0.072918 | 0.23232 | -0.004196
7 | 1.6428 | -0.13904 | -0.090264 | -0.12895 | -0.022618 | -0.0085761 | -0.017542 | -0.024235 | 0.03391 | -0.016853 | -0.0086804 | 0.010682 | -0.0079777 | 0.014807 | -0.0034947 | 0.083904 | -0.090296 | -0.11487 | -0.031256 | -0.25789 | 0.017559
8 | -0.039892 | 0.024135 | -0.012353 | -0.0066138 | -0.0059483 | 0.0039925 | -0.0090635 | -0.030554 | -0.0037315 | -0.016378 | 0.015758 | 0.001066 | -0.00043686 | 0.0015382 | 0.034098 | 0.015515 | -0.020786 | 0.065219 | -0.004373 | 0.020863 | 0.0014474
9 | -0.13025 | -0.023588 | -0.028757 | -0.0027605 | 0.0017478 | -0.03723 | -0.032028 | -0.02337 | -0.03645 | -0.04341 | -0.060936 | -0.063312 | -0.04067 | 0.0022153 | 0.010028 | -0.038156 | -0.040967 | -0.033729 | -0.016126 | 0.025937 | -0.029119
10 | 6.598e-05 | 3.2689e-05 | 0.0018437 | -6.624e-06 | 7.9256e-06 | -0.0039438 | 0.001917 | -5.1825e-07 | -3.69e-06 | 8.0306e-05 | -1.3828e-06 | 3.2283e-06 | -1.5542e-05 | 7.8705e-07 | -5.4786e-05 | -7.9845e-07 | -1.5029e-05 | 1.7236e-05 | -1.7356e-06 | -0.00041562 | 0
11 | 0.18421 | -0.053623 | -0.049018 | -0.05214 | -0.042656 | -0.041772 | -0.039344 | -0.040879 | -0.035048 | -0.048094 | -0.050742 | -0.045508 | -0.049182 | -0.055319 | -0.052664 | -0.037925 | -0.037666 | -0.047229 | -0.065248 | -0.04049 | -0.038857
12 | 0.034243 | -0.064762 | -0.023493 | -0.039764 | -0.017577 | -0.017191 | -0.089736 | -0.071751 | -0.037327 | -0.0093666 | -0.05887 | -0.026564 | -0.025045 | 0.012868 | 0.049793 | -0.023291 | -0.0041941 | -0.014325 | -0.028397 | -0.027899 | -0.026667 | ten times!
13 | -0.18164 | -0.089447 | -0.058334 | -0.04833 | -0.034168 | -0.042204 | -0.05288 | -0.032948 | -0.03998 | -0.078093 | -0.073749 | -0.06599 | -0.063952 | -0.034331 | -0.029887 | -0.028191 | -0.050303 | -0.059535 | -0.011071 | 0.023199 | -0.051976 | ten times!
14 | 0.0059969 | -0.0019844 | -0.037786 | -0.0044716 | -0.036043 | -0.017007 | -0.031843 | -0.047639 | -0.029794 | -0.032228 | -0.02333 | -0.011191 | -0.0152 | -0.025279 | -0.016038 | 0.0077375 | -0.022684 | -0.044221 | -0.080609 | -0.00047003 | -0.021708 | ten times!
15 | 0.061212 | -0.079637 | -0.059766 | -0.037301 | -0.020338 | -0.028834 | -0.030153 | -0.029629 | -0.022573 | -0.05248 | -0.025438 | -0.030404 | -0.031084 | -0.031348 | -0.041304 | -0.064655 | -0.095234 | -0.060963 | -0.022423 | -0.10925 | -0.041028 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.0010023 | -5.2964e-05 | -8.548e-06 | 7.8273e-05 | -0.00018353 | -0.00034209 | -0.00019241 | -0.0001656 | -0.00033203 | -0.00021286 | -0.00018538 | -5.9166e-05 | -7.7938e-05 | -6.7043e-05 | -0.00016593 | 9.5202e-05 | -4.3578e-06 | 0.00087408 |  | 596.6281
|  | 0.00079724 | -3.7872e-06 | -0.00023804 | -0.00020506 | -0.00020319 | -0.0002224 | -0.00017447 | -1.8577e-05 | -0.00015662 | -0.0002496 | -0.00028751 | -0.00014863 | 2.2903e-05 | 5.1972e-05 | 9.0222e-05 | -0.00020574 | 0.00010756 | 0.0010437 |  | 588.2417
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3741.48
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 596.678 | -4324.6017 | 11.197 | 0.64317 | 1.0874 | 1.9557 | -0.66266 | -2.2484 | -0.7515 | -0.48329 | -2.1478 | -0.956 | -0.68109 | 0.58115 | 0.3934 | 0.50236 | -0.48665 | 2.125 | 1.1293 | 9.9147 | -3158.3583
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.069766 | 2.7647 | -0.50689 | 0.32043 | -0.2617 | 0.41484 | 0.25103 | -0.73723 | 0.20636 | 0.12646 | -0.91835 | -0.60072 | 0.63901 | 0.17044 | 0.26575 | 0.0053658 | 0.37405 | -0.030742 | 0.069033 | 0.21286 | -0.32429
db2 | 1.2687 | -0.70659 | 0.066763 | 0.85855 | 1.397 | 0.0077237 | 0.43805 | 0.90321 | -0.59162 | -1.0256 | -0.23246 | 0.5393 | -0.12134 | -0.55096 | -1.2794 | -8.9313e-05 | 1.1957 | -0.023521 | -0.72625 | -0.85511 | -0.87915
db3 | -3.9529 | 50.054 | 1.3477 | -0.45416 | 0.17263 | 0.49788 | -0.45277 | 0.039 | 0.0096184 | -0.35988 | 0.23462 | -0.52158 | 0.055781 | 0.39664 | -0.18442 | -0.94654 | -0.48143 | -1.1234 | 0.76631 | 1.004 | -5.935
db4 | 0.0078072 | 1.7914 | -0.092868 | -0.16682 | -0.064122 | -0.12185 | 0.049294 | 0.1353 | 0.20413 | 0.17102 | -0.11358 | 0.014275 | 0.097778 | 0.15725 | -0.021923 | -0.30698 | 0.098419 | 0.32608 | -0.20455 | -0.16084 | -0.10182
db5 | 0.42579 | -2.1417 | -0.12699 | -0.35155 | -0.30606 | -0.045797 | 0.031533 | -0.021642 | 0.18271 | 0.035178 | 0.18403 | 0.049143 | 0.0079584 | -0.023076 | 0.082348 | 0.038272 | -0.18685 | 0.26176 | -0.007569 | 0.19661 | -2.3283
db6 | 0.028449 | 1.2145 | 0.04941 | 0.021663 | -0.0084058 | 0.0060863 | 0.095858 | -0.050922 | 0.059391 | 0.017423 | 0.027249 | 0.039186 | -0.012569 | -0.056647 | 0.071918 | -0.020546 | 0.056839 | -0.13779 | -0.14865 | -0.0094882 | -0.078001
db7 | 1.176 | 2.7431 | -0.12527 | 0.041161 | 0.022163 | 0.062994 | -0.0050772 | 0.0029558 | -0.01262 | -0.00054017 | -0.022938 | 0.0017848 | 0.0087278 | 0.041797 | 0.037815 | -0.065122 | -0.028918 | 0.02212 | 0.087101 | -0.068129 | 0.8542
db8 | -0.0086151 | 0.29336 | -0.011442 | 0.042617 | -0.025188 | -0.062815 | -0.029474 | 0.035541 | 0.0094516 | 0.05975 | 0.018195 | 0.043606 | 0.002306 | 0.021238 | -0.029182 | -0.082643 | 0.032586 | -0.024348 | 0.0062646 | -0.0064609 | -0.036861
db9 | 0.63519 | 0.48941 | -0.021054 | 0.0082359 | 0.024788 | -0.0052434 | 0.0086685 | 0.015128 | 0.012827 | 0.0065893 | -0.0023686 | 0.0028069 | 0.0028977 | 0.013453 | -0.0072224 | 0.0171 | 0.00048723 | -0.047115 | 0.0035942 | -0.033573 | 0.50527
db10 | -0.0035142 | 0.17467 | -0.050967 | 0.037879 | -0.00401 | -0.017124 | 0.03197 | 0.032055 | 0.02904 | 0.079663 | -0.0045134 | -0.008599 | -0.006408 | 0.016567 | -0.0077161 | -0.084819 | 0.010586 | -0.018692 | -0.00969 | -0.025221 | -0.051943
db11 | 0.73207 | 0.6237 | 0.0057598 | -0.0055262 | -0.00080502 | 0.006351 | 0.0032471 | -0.0076462 | -0.0018686 | 0.00011198 | 0.002076 | 0.0031892 | -0.0031625 | -0.0056571 | -0.00055652 | 0.0067341 | -0.0030783 | 0.0032412 | -0.0055833 | 0.0031736 | 0.66812
db12 | 0.00023196 | 0.030077 | -0.0030959 | 0.0069988 | 0.0038023 | 0.0019825 | 0.0062197 | 0.0058379 | 0.0022548 | 0.0086108 | -0.0046292 | -0.0013081 | 0.00020406 | -0.00037373 | -0.0045987 | -0.014115 | -0.0025848 | -0.0070357 | -4.9384e-05 | 0.0018799 | -0.00033012
db13 | 0.073114 | 0.090278 | 0.004981 | 0.0020264 | -0.0006964 | 0.0043863 | -0.00092756 | -0.00014058 | -0.0028363 | -0.0020082 | 0.0011713 | 0.0021541 | -0.0003558 | 0.00096214 | -0.00033512 | -0.0034181 | -0.0003411 | -0.00069258 | -0.0031885 | -0.00074107 | 0.059279
db14 | -0.0022055 | 0.011113 | -0.0031745 | 0.0024214 | 0.0027695 | 0.0012931 | 0.0029832 | 0.0009766 | 0.001133 | 0.0020958 | -0.0012675 | -0.0017318 | 0.00057684 | 3.0924e-05 | 0.00079038 | -0.0031522 | -0.0021044 | -0.0016659 | -0.00057161 | -0.0014029 | -0.0011332
db15 | 0.018677 | -0.0068901 | 0.0016893 | -0.0027621 | -0.0032221 | -0.0014537 | 0.0011978 | -0.0012732 | 0.0016059 | 0.0019483 | 0.0012418 | 0.00010103 | -0.00081935 | -0.0012828 | 0.00042413 | -0.0017973 | -0.00015536 | -0.0010491 | 0.0019526 | 0.0036541 | 0.014969
da2 | -0.30122 | -7.5073 | 0.37133 | -1.9179 | 0.1587 | 0.71024 | -0.8177 | -0.44971 | -0.044451 | 0.35207 | 0.77525 | -0.04539 | -0.19042 | 0.75379 | 0.68379 | 0.82083 | -1.2636 | -0.23998 | -0.30683 | 0.64997 | 1.6016
da3 | -0.10515 | -0.54902 | -0.55663 | -0.43478 | -0.29531 | 0.0030031 | -0.3205 | 0.46404 | 0.80157 | 0.52832 | 0.37557 | 0.39277 | 0.23537 | -0.55487 | -0.21579 | -0.38341 | -0.65628 | 0.12886 | -0.18066 | 0.66871 | 1.5055
da4 | -0.039704 | -2.2474 | -0.11534 | 0.35116 | -0.28699 | 0.13285 | 0.5559 | 0.26717 | 0.46114 | 0.36198 | 0.74152 | -0.077967 | -0.62738 | -0.3922 | -0.9881 | -0.68869 | 0.00056109 | 0.1226 | 0.27351 | -0.091711 | 0.038995
da5 | 0.019415 | 4.434 | -0.040588 | -0.1185 | -0.037638 | -0.030838 | 0.018239 | 0.26893 | 0.16442 | 0.11677 | 0.056277 | 0.15196 | -0.19734 | -0.0805 | -0.27974 | 0.07188 | 0.047553 | -0.1206 | 0.0026003 | 0.0071106 | -0.2206
da6 | -0.051537 | -1.003 | -0.02981 | 0.12035 | 0.12492 | 0.085406 | 0.14952 | 0.11475 | 0.046309 | -0.018935 | -0.11333 | -0.052863 | -0.011932 | -0.080208 | -0.097683 | -0.10093 | -0.1138 | 0.00051877 | -0.077837 | 0.055566 | 0.055077
da7 | -0.03206 | 1.8255 | -0.036527 | 0.0086737 | -0.015298 | -0.014121 | 0.057833 | 0.016621 | -0.0060576 | -0.060416 | 0.028339 | 0.076283 | 0.019854 | 0.0035606 | 0.014702 | -0.05265 | -0.026327 | 0.076043 | -0.019627 | -0.070887 | -0.055277
da8 | 0.020372 | -0.19848 | -0.0075736 | -0.0038775 | -0.0074372 | 0.055305 | 0.032294 | 0.055083 | 0.013642 | 0.035978 | 0.013392 | -0.001884 | -0.0093141 | -0.011934 | -0.035393 | 4.6535e-05 | 0.014442 | -0.036703 | -0.070328 | -0.035738 | -0.047441
da9 | -0.017542 | -0.14711 | -0.00091502 | -0.01757 | -0.01192 | 0.039984 | -0.0018739 | 0.00089513 | 0.011574 | 0.022793 | 0.0034226 | -0.050602 | -0.016379 | -0.0038841 | 0.006753 | 0.042674 | -0.049553 | -0.010096 | 0.008566 | 0.026129 | 0.035893
da10 | 0.008399 | -0.10058 | -0.035006 | 0.048724 | 0.032204 | 0.053769 | 0.015675 | 0.02836 | -0.01675 | -0.025625 | 0.011967 | 0.03584 | 0.026233 | 0.023434 | -0.044422 | -0.011234 | 0.0055258 | 0.0067405 | -0.066675 | -0.08876 | -0.13787
da11 | -0.03261 | 0.18403 | 0.0015698 | -0.00035654 | -0.0080711 | -0.0023857 | 0.0094632 | -0.0056072 | 0.0061529 | -0.0036148 | -0.010249 | -0.00058555 | 0.0075059 | 0.00034343 | 0.00034932 | -0.0060789 | 0.010276 | 0.0010142 | 0.0056839 | -0.0054099 | -0.021453
da12 | -0.00020304 | -0.021349 | 0.0025191 | 0.0083418 | 0.0031621 | 0.0063128 | 0.0055232 | 0.0050344 | -0.00030707 | -0.0044761 | -0.0055959 | 0.0007135 | 0.0044249 | 0.0024735 | -0.0083349 | -0.0029787 | 0.00023223 | -0.00033168 | -0.0078208 | -0.0088924 | -0.0071597
da13 | -0.0037346 | -0.018214 | -0.00025781 | -0.0031769 | -0.003206 | -0.0012025 | -0.001108 | 0.00075347 | 0.00033189 | -0.0010201 | 0.00013383 | 0.00086937 | 0.0014735 | -0.00081358 | 0.0011887 | 0.0031489 | 0.0014071 | -0.0008038 | 0.00064419 | 0.0016377 | -0.0023653
da14 | -0.0039702 | 0.0089566 | -0.00065198 | 0.0011976 | -0.0015898 | 0.0029762 | -5.4153e-05 | 0.0017279 | 0.0019064 | 0.0022935 | 0.0016159 | 7.8506e-05 | -0.00030222 | 0.00030458 | -0.002923 | -0.0007127 | -0.00087967 | 5.876e-06 | -0.0020684 | -0.0029244 | -0.0036677
da15 | -0.0037672 | 0.0087232 | -0.0049238 | 0.0010936 | 0.00071945 | -0.00033383 | 0.00044876 | 0.0012504 | 0.0017787 | 0.0030019 | 0.0012166 | 0.0015674 | 0.0013841 | 0.001656 | 0.00076915 | -0.0034956 | -0.0023084 | -0.000238 | -0.001976 | -0.0016104 | -0.0077032
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3766.1998
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 588.4658 | -3536.6019 | 9.1241 | 1.1129 | -1.23 | -0.90014 | -0.88144 | -1.0736 | -0.59411 | 0.96496 | -0.41559 | -1.3456 | -1.7247 | -0.33568 | 1.3798 | 1.6705 | 2.0531 | -0.90694 | 2.2265 | 11.5894 | -3995.7976
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.04932 | 2.3342 | 0.053792 | 0.20632 | 0.095032 | 0.37001 | 0.2331 | 0.087547 | 0.13159 | 0.26614 | 0.017575 | -0.085724 | -0.069099 | 0.082988 | -0.24745 | -0.47792 | -0.47423 | 0.027521 | -0.061387 | -0.1558 | -0.92818
db2 | 0.062077 | 0.75169 | 0.45802 | -1.0249 | -0.90964 | -0.0065635 | -0.13598 | 0.086544 | -0.095492 | -0.26967 | -0.70463 | 0.33499 | 0.5451 | -0.056875 | 0.028745 | -0.051128 | 0.18004 | 0.87247 | -0.20697 | 0.95598 | 0.79121
db3 | -2.8344 | 39.4514 | 1.3587 | -0.71818 | 0.2855 | 0.42934 | -0.75408 | -0.80919 | 0.5983 | -1.2415 | 0.70811 | 0.90615 | 0.075097 | 0.89059 | -1.4349 | 0.68671 | -0.61014 | -0.20739 | -0.7024 | 0.53928 | -6.6984
db4 | -0.091116 | 0.35402 | -0.20712 | 0.070969 | -0.048945 | 0.031109 | -0.0024558 | 0.33529 | 0.064269 | -0.65277 | -0.39476 | 0.015658 | 0.0022869 | -0.081131 | -0.092495 | 0.019151 | 0.21468 | 0.378 | 0.10263 | 0.24565 | 0.10208
db5 | 0.14148 | -2.0361 | 0.069896 | -0.033412 | -0.17724 | -0.27661 | -0.15248 | 0.075357 | 0.12723 | -0.34178 | 0.1267 | -0.15076 | -0.092351 | 0.075513 | 0.057147 | 0.12245 | 0.10153 | 0.042278 | 0.26552 | 0.16101 | -3.8421
db6 | -0.035183 | 0.067243 | 0.030502 | 0.035719 | 0.066743 | -0.011145 | -0.013273 | -0.049805 | -0.05468 | -0.12855 | 0.062915 | -0.04755 | 0.043256 | 0.039404 | -0.16838 | -0.081355 | 0.083681 | 0.28154 | 0.070426 | -0.15945 | 0.16983
db7 | 1.3305 | 2.7706 | -0.09617 | -0.014108 | -0.042178 | -0.0061099 | -0.0016189 | -0.043126 | 0.063162 | -0.050231 | -0.010854 | 0.043638 | 0.024667 | 0.031648 | -0.0082619 | -0.040535 | 0.077153 | 0.022925 | 0.098862 | -0.048863 | 0.71094
db8 | 0.0037846 | -0.0035112 | -0.013125 | 0.018801 | -0.015808 | -0.0035923 | 0.042991 | 0.025748 | 0.011005 | -0.019807 | -0.013233 | -0.0034102 | 0.053727 | 0.029273 | -0.015624 | -0.066538 | -0.0062174 | -0.063471 | -0.025311 | 0.064592 | 0.054034
db9 | 0.61617 | 0.52061 | 0.0081263 | -0.016604 | 0.013343 | 0.015213 | 0.0051671 | 0.0044101 | 0.036135 | -0.026458 | 0.032355 | 0.0099886 | 0.018946 | -0.001953 | -0.052545 | -0.013597 | -0.0088899 | -0.0064207 | -0.0043204 | -0.012896 | 0.48533
db10 | -0.00045739 | 0.078488 | -0.077029 | 0.020932 | -0.022995 | -0.003102 | 0.031296 | 0.020146 | -0.0052621 | 0.031248 | -0.0025251 | 0.027922 | 0.064481 | 0.045484 | 0.022265 | -0.073238 | -0.015192 | 0.0036304 | -0.036921 | -0.03114 | -0.070546
db11 | 0.79721 | 0.71287 | -0.0012962 | -0.0022965 | -0.0057461 | -0.0042562 | -0.0063609 | 0.0061102 | 0.0096616 | -0.004151 | 0.011189 | -0.0049005 | -0.0055082 | -0.0034339 | -0.00416 | 0.0085693 | 0.0060413 | -0.0079198 | -0.0023791 | 0.010837 | 0.70053
db12 | -0.00018049 | 0.010252 | -0.0037486 | 0.0018171 | -0.0035535 | 0.00032989 | 0.008494 | 0.0024894 | 0.0022961 | 0.0054053 | -0.0016857 | 0.0056067 | 0.0084627 | 0.0077363 | 0.002235 | -0.010337 | -0.0069359 | -0.0040476 | -0.0090711 | -0.0054934 | 0.0009466
db13 | 0.077794 | 0.089363 | 0.0023622 | 0.0033801 | -0.0010628 | 0.0032349 | -0.0009804 | -0.00038024 | 0.00079166 | -0.0036694 | 0.0015358 | 0.0043921 | -0.0021275 | -9.5039e-06 | -0.0020928 | 0.0030393 | 0.00014178 | -0.0045461 | -0.00028817 | -0.0037211 | 0.065081
db14 | -0.0044142 | 0.0044324 | -0.0036045 | -0.0013005 | 0.00045236 | 0.00077336 | 0.0017099 | 0.0002233 | 0.0010797 | 0.0019009 | 0.00064453 | 0.00033487 | 0.0023473 | 0.0013465 | -2.4644e-05 | -0.0020742 | -0.0017811 | 0.0010525 | -0.0014628 | -0.0016175 | -0.0011373
db15 | 0.01816 | -0.0093952 | -0.0043305 | 0.00095841 | -0.0023865 | -0.00067122 | 0.00076549 | 0.00029369 | -0.0010064 | 0.0012204 | 0.0019864 | 0.00073387 | 0.00069732 | 0.0016383 | 0.0033392 | -0.0019787 | 0.0015656 | -0.00066313 | -0.0010837 | -0.0010786 | 0.010073
da2 | -0.030654 | -2.9633 | 1.2793 | 1.0015 | -0.022819 | -0.46822 | 0.50231 | -0.49836 | -0.89695 | 2.0201 | -0.79861 | -0.017715 | -0.081538 | -0.87411 | 0.76151 | -0.78879 | -1.5158 | -0.49484 | 1.4036 | -0.51051 | 0.65758
da3 | 0.060523 | -2.0715 | -0.83339 | -0.71478 | -0.61515 | 0.078571 | 0.44013 | 0.54574 | 0.48472 | 0.48352 | -0.23763 | 0.11017 | -0.16295 | -0.12549 | 0.33296 | 0.32486 | -0.20385 | -0.84873 | 0.27581 | 0.66549 | 1.5004
da4 | -0.10599 | -0.10912 | -0.095726 | -0.44509 | -0.18725 | -0.12591 | 0.17378 | 0.04597 | -0.27283 | 0.54403 | 0.088523 | 0.058959 | -0.013585 | 0.36212 | 0.61803 | 0.357 | 0.93289 | -0.37004 | -1.1724 | -0.49846 | 0.47012
da5 | 0.06888 | 3.8881 | -0.035716 | -0.21969 | -0.018041 | -0.061611 | 0.0085353 | 0.037632 | -0.00032073 | 0.10208 | -0.14753 | -0.042371 | 0.010141 | -0.033718 | -0.0026784 | 0.15625 | 0.16097 | 0.12336 | 0.086542 | -0.12382 | -0.49069
da6 | 0.00057291 | -0.28052 | -0.050263 | 0.029331 | -0.055065 | -0.032528 | 0.023601 | -0.12676 | -0.0082377 | 0.10164 | -0.023336 | -0.012638 | -0.03215 | -0.0029565 | 0.06163 | 0.096707 | -0.053704 | -0.032326 | 0.032347 | 0.084705 | 0.14984
da7 | -0.031177 | 1.6377 | -0.10936 | -0.05932 | -0.098963 | 0.0083867 | 0.022679 | 0.013469 | 0.0072478 | 0.06501 | 0.014545 | 0.023996 | 0.04243 | 0.02505 | 0.045989 | 0.023135 | 0.11529 | -0.056457 | -0.083689 | 0.00056279 | -0.26209
da8 | 0.0070114 | -0.08709 | 0.015628 | -0.019957 | 0.020226 | 0.024214 | -0.0042008 | -0.016703 | -0.027459 | 0.0015231 | -0.041455 | 0.070007 | -0.0044071 | 0.01414 | -0.01505 | 0.041904 | -0.017104 | -0.064084 | 0.025597 | -0.0028187 | -0.070479
da9 | -0.027851 | -0.14037 | 0.0046903 | -0.00089257 | 0.022131 | 0.028952 | -0.0095779 | -0.0041566 | 0.0042772 | -0.0071689 | -0.015469 | -0.02785 | -0.034521 | -0.0093699 | 0.029204 | 0.033395 | -0.0089539 | -0.013692 | -0.0017071 | 0.010709 | 0.05254
da10 | 0.0068866 | -0.048107 | -0.0067152 | -0.0063069 | 0.045447 | 0.049207 | -0.012483 | -0.0059403 | 0.0081559 | 0.013115 | -0.033714 | 0.086065 | -0.0022837 | 0.026203 | -0.024553 | 0.022389 | -0.047404 | -0.062776 | -0.056947 | 0.0085424 | -0.15515
da11 | -0.046787 | 0.18241 | -0.0061615 | -0.002302 | -0.0056969 | 0.0040955 | 0.0049313 | 0.0071974 | 0.0059986 | 0.011831 | -0.0012385 | -0.0030711 | 0.001166 | -0.0019414 | -0.0085994 | -0.0070829 | 0.009399 | 0.009396 | 0.0005767 | -0.018498 | -0.038722
da12 | -0.00187 | -0.0081328 | -0.0041666 | -0.00066367 | 0.0040189 | 0.0070193 | -5.3557e-05 | -0.0072623 | -0.0035443 | 0.00035113 | -0.0022516 | 0.0075233 | -0.00048545 | 0.0031609 | 0.0011879 | 0.010146 | -0.0052268 | -0.0048159 | -0.0054217 | 0.00048397 | -0.020723
da13 | -0.0049932 | -0.018396 | -0.0039823 | -0.0009739 | -0.00022133 | 0.0012621 | 0.00067738 | -0.00039963 | 0.0016639 | 0.00079057 | -0.0026521 | -0.0029902 | -0.0017584 | -0.0016179 | 0.001554 | 0.0017949 | 0.0025999 | 0.00023272 | -0.00013966 | 0.0041601 | 0.0028691
da14 | -0.0023483 | 0.0025276 | 0.0031135 | -0.0016458 | 0.003497 | 0.00049757 | 0.00048825 | -0.00093951 | -0.0019419 | -1.7117e-05 | -0.0018181 | 0.0030836 | 0.0011633 | 0.0018472 | -0.00091985 | 0.0019738 | 0.001855 | -0.0017536 | -0.0032783 | -0.0052049 | -0.0018865
da15 | -0.004242 | 0.0061212 | -0.0037217 | -0.0017346 | 0.00051188 | 0.0022082 | 0.0013586 | 0.0012268 | 0.0012791 | 0.0019847 | -0.0010059 | 0.0016982 | 0.0012016 | 0.0011336 | 0.0011073 | 0.00011166 | -0.0022235 | -0.0052814 | -0.0018543 | 0.0019997 | -0.010925
Field Colinearity | -11.976
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
1.2687 | -4.7629 | 0.0078072 | 0.46429 | 1.187 | -0.69766 | -0.30122 | -0.10515 | -0.039704 | 25.6059 | 0.59663 | 0.062077 | -4.7344 | -0.091116 | 0.37998 | 1.2665 | -0.4932 | -0.030654 | 0.060523 | -0.10599 | 27.2588 | 0.58824 | 1.1344 | -3.193 | 0.05633 | 0.28862 | 1.2209 | 0 | -0.44547 | -0.061081 | -0.10183 | 14427.247 | 0.10794 | -3.3148 | -0.070185 | 0.16766 | 1.2996 | 0 | -0.10746 | 0.033889 | -0.088201 | 14423.72
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.75 | 0.75 | 0.2 | 0.2 | 0.75 | 0.75 | 0.15 | 0.15 | 0.75 | 0.75 | 0.15 | 0.15 | 0.75 | 0.75