## Alarm sheet
File name | HCMBBRA001-01000006_cm.xls
Component ID | HCMBBRA001 |  |  |  | Serial Number | 1000006
Date of test Ap 1 | 2002-04-25 |  |  |  | Date of test Ap 2 | 2002-04-25
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 |  |  |  |  | b10
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 |  |  |  |  | a10
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Magnet name
| HCMBBRA001 | 1000006 |  | Additional lamin. (mm) |  | 0
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2002-04-26 | Analysis by | E. Wildner | Answer to MMS-LD |  | 2002-04-26 | Answer to manufacturer | 2002-04-26
| Cold mass template - version 11.10.2002 | [7 empty] | Version : | cern
1 | Remplir les champs jaunes dans Assembly Resume et Original Data
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
| Last modified | by
| 2002-10-21 | E. Todesco | Global feed-down added in Summary data - cell for additional laminations in Assembly data

## Original data
File name |  | HCMBBRA001-01000006_cm.xls
Component ID |  | HCMBBRA001 |  |  | Serial Number |  |  | 1000006 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Alstom | [9 empty] | File name for aperture 1 |  |  | try1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 34 |  |  |  | File name for aperture 2 |  |  | try2 |  |  |  | a even | 1
Test Operator |  | Galbraith + Molinari |  |  | Test Controller |  |  | E. Todesco
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2002-04-25 |  |  | Time of test |  |  | 08.00-12.00 |  |  |  | Date of test |  | 2002-04-25 |  |  | Time of test |  |  | 13.00-17.00
Mole name |  | Dipole 6 |  |  | Calibration date |  |  | 20.07.99 |  |  |  | Mole name |  | Dipole 6 |  |  | Calibration date |  |  | 20.07.99
Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 19 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 22
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 9.7168e-05 |  |  | Dy average |  |  | 0.00042819 |  |  |  | Dx average |  | 0.00026484 |  |  | Dy average |  |  | 0.00056591
Magnetic length |  | 14.3887 |  |  | Tranfs. Function (mT/KA) |  |  | 704.0588 |  |  |  | Magnetic length |  | 14.3864 |  |  | Tranfs. Function (mT/KA) |  |  | 704.098
Phase |  | 2.0009 |  |  | Integral field (T.m) |  |  | 0.08611 |  |  |  | Phase |  | 2.1186 |  |  | Integral field (T.m) |  |  | 0.086101
File | Aperture 1 - Cold mass | [19 empty] | Integrals | File | Aperture 2 - Cold mass | [19 empty] | Integrals
C1 (T) | 0.003525 | 0.005984 | 0.005986 | 0.005987 | 0.005986 | 0.005985 | 0.005985 | 0.005985 | 0.005982 | 0.005982 | 0.005983 | 0.005983 | 0.005984 | 0.005985 | 0.005986 | 0.005985 | 0.005986 | 0.005986 | 0.005981 | 0.003581 | 0.08611 | C1 (T) | 0.0035 | 0.005983 | 0.005985 | 0.005987 | 0.005986 | 0.005986 | 0.005985 | 0.005986 | 0.005985 | 0.005984 | 0.005983 | 0.005983 | 0.005984 | 0.005984 | 0.005984 | 0.005985 | 0.005987 | 0.005987 | 0.005983 | 0.003589 | 0.086101
Angle (mrad) | 2.8241 | 0.38305 | 0.16522 | -0.065335 | 0.36968 | 0.28449 | 0.90424 | 0.60258 | 0.27637 | -0.26131 | -0.35483 | -0.36048 | -0.79167 | -0.52994 | -0.84379 | -0.09485 | 0.62651 | 0.39575 | -1.9418 | -0.42771 | 2.0009 | Angle (mrad) | 3.1103 | -0.10283 | -0.45137 | -0.43167 | -0.18588 | 0.23097 | 0.74811 | 0.6146 | 0.31673 | -0.30848 | -0.61097 | -0.62387 | -0.65662 | -0.36691 | -1.2114 | -0.27104 | 0.24302 | 0.11722 | -0.12329 | 1.8016 | 2.1186
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 24.1508 | -0.20815 | 1.1753 | 1.803 | 3.0137 | 1.3872 | 0.50984 | 1.2083 | 1.3831 | 1.8394 | 1.7036 | 2.0695 | 2.346 | 1.2265 | 3.4709 | 1.4898 | 2.4885 | 3.5409 | 6.3562 | 27.1192 | 3.5053 | b2 | -26.0644 | -1.3265 | -0.40607 | -1.3056 | -1.6491 | -1.119 | -1.9196 | -0.37853 | -0.34876 | 0.037425 | -0.46549 | -0.52589 | -1.0722 | -0.59646 | -0.013477 | -0.11053 | -1.367 | -1.4095 | -0.79316 | -23.1144 | -2.2869
b3 | 36.0487 | 7.6097 | 5.3324 | 5.0777 | 5.0305 | 3.6252 | 6.9812 | 7.846 | 7.1086 | 7.4885 | 6.5138 | 7.0792 | 6.1196 | 4.4545 | 0.17147 | 0.43061 | -1.6483 | 0.12133 | 5.3183 | 2.2786 | 5.5896 | b3 | 37.0654 | 7.3716 | 6.3455 | 5.8105 | 4.9088 | 1.1977 | 5.1592 | 6.2113 | 4.5914 | 3.8907 | 5.2972 | 6.4675 | 5.5141 | 4.6656 | 4.9033 | 3.2703 | -0.46698 | -0.068249 | 7.7297 | 2.7974 | 5.5329
b4 | 0.98466 | 0.16286 | 0.13134 | 0.66011 | 0.13157 | -0.20235 | 0.18924 | 0.20594 | 0.34848 | 0.37467 | 0.44313 | 0.4572 | 0.97486 | -0.012393 | 0.52981 | -0.29115 | 0.058245 | 0.078621 | 0.12472 | 0.15413 | 0.26252 | b4 | -1.6186 | -0.22131 | 0.094848 | 0.14108 | 0.86102 | 0.19847 | 0.14801 | 0.13287 | 0.129 | 0.61616 | -0.18425 | 0.0020129 | -0.20052 | -0.22276 | -0.06087 | 0.17036 | -0.96533 | -0.84461 | -0.30051 | -2.2707 | -0.14671
b5 | -3.0874 | 1.6662 | 1.2826 | 1.6046 | 2.1397 | 2.3919 | 2.0899 | 1.6494 | 1.5815 | 1.9261 | 1.952 | 1.8846 | 1.277 | 1.4614 | 2.5406 | 2.839 | 3.4607 | 2.1357 | 1.4473 | -1.7232 | 1.6928 | b5 | -3.0619 | 1.2528 | 1.1281 | 1.0317 | 1.6837 | 2.7033 | 1.4239 | 0.78945 | 0.74722 | 1.456 | 1.4285 | 1.5346 | 1.3462 | 1.7552 | 2.1645 | 3.2562 | 3.3445 | 3.2132 | 1.1426 | -2.5445 | 1.464
b6 | -0.27153 | 0.072049 | -0.093815 | -0.31729 | -0.11855 | 0.15947 | 0.045135 | -0.00089776 | 0.0074868 | -0.1351 | 0.055129 | 0.0086981 | -0.40338 | -0.059487 | -0.48611 | 0.094001 | -0.23708 | -0.25889 | 0.2134 | 0.34309 | -0.073514 | b6 | -0.65512 | 0.024446 | 0.028209 | -0.10006 | -0.33709 | -0.2718 | -0.080745 | -0.18272 | -0.079444 | -0.34199 | -0.18005 | -0.19473 | -0.045883 | -0.064501 | -0.27697 | -0.28534 | 0.15473 | 0.072876 | -0.083397 | -0.061005 | -0.13886
b7 | 1.8777 | 0.45365 | 0.91037 | 0.6989 | 0.47281 | 0.50953 | 0.51842 | 0.68568 | 0.79119 | 0.62798 | 0.52742 | 0.64137 | 0.76469 | 0.84045 | 0.80678 | 0.78993 | 0.93128 | 1.3616 | 0.43948 | 0.2525 | 0.73117 | b7 | 1.9892 | 0.64918 | 0.96333 | 1.0178 | 0.99734 | 0.68414 | 0.9184 | 0.75808 | 1.0177 | 1.0946 | 0.70135 | 0.72181 | 0.71892 | 0.69578 | 0.50507 | 0.084218 | 0.4743 | 0.52068 | 0.64311 | 0.45875 | 0.76125
b8 | -0.0001683 | -0.074361 | 0.0081964 | 0.0078305 | 0.034296 | 0.029351 | -0.034034 | 0.0099702 | -0.03179 | -0.0054032 | -0.085037 | -0.074655 | 0.04023 | 0.041239 | 0.10004 | 0.044135 | 0.072614 | 0.093559 | -0.057305 | 0.0015108 | 0.0062466 | b8 | 0.026914 | -0.0087197 | 0.014258 | -0.01313 | -0.049602 | -0.0079205 | -0.042566 | 0.0072428 | -0.0067001 | -0.02761 | 0.0010594 | 0.0096868 | 0.028598 | -0.0034353 | 0.043467 | -0.087237 | 0.069704 | 0.096065 | -0.0025858 | 0.051845 | 0.003512
b9 | 0.28302 | 0.32703 | 0.20605 | 0.23407 | 0.28768 | 0.29981 | 0.34194 | 0.34479 | 0.33933 | 0.32244 | 0.33145 | 0.33551 | 0.31469 | 0.25608 | 0.19138 | 0.20194 | 0.14668 | 0.063917 | 0.32879 | 0.32231 | 0.27273 | b9 | 0.28402 | 0.33945 | 0.24954 | 0.23445 | 0.19382 | 0.24106 | 0.30118 | 0.2898 | 0.31263 | 0.24621 | 0.28285 | 0.34014 | 0.32643 | 0.30003 | 0.30963 | 0.27793 | 0.22063 | 0.23642 | 0.30635 | 0.24318 | 0.27732
b10 | -0.16733 | -0.013663 | -0.014909 | 0.005096 | 0.0034642 | -0.046835 | -0.0044726 | 0.010344 | 0.018116 | -0.0021329 | 0.030286 | 0.023435 | 0.083949 | 0.00021469 | 0.015549 | 0.01674 | 0.0029175 | -0.044363 | -0.026073 | 0.068389 | 0 | b10 | -0.11098 | -0.051664 | -0.008319 | 0.0089966 | 0.070489 | 0.028237 | -0.013464 | 0.024925 | -0.039408 | 0.0081981 | -0.0087809 | 0.00037841 | 0.030831 | -0.037204 | 0.024459 | 0.061377 | -0.015367 | -0.050574 | -0.0027712 | 0.057593 | 0
b11 | 0.55044 | 0.62874 | 0.583 | 0.60918 | 0.62515 | 0.61245 | 0.63809 | 0.62571 | 0.61802 | 0.63424 | 0.63637 | 0.63632 | 0.61468 | 0.60518 | 0.62833 | 0.60479 | 0.58361 | 0.54812 | 0.62031 | 0.58827 | 0.61126 | b11 | 0.55231 | 0.63037 | 0.5961 | 0.58891 | 0.58529 | 0.62221 | 0.60733 | 0.6064 | 0.57897 | 0.58899 | 0.6107 | 0.62672 | 0.62744 | 0.6221 | 0.65185 | 0.65703 | 0.62291 | 0.61988 | 0.63247 | 0.57548 | 0.61214
b12 | -0.01142 | -0.0080737 | 0.0031704 | 0.0016997 | 0.013583 | 0.0039407 | -0.0038323 | 0.00032501 | 0.0070871 | -0.0032988 | 0.00026372 | -0.0004637 | 0.0087017 | 0.0076533 | 0.025629 | 0.013891 | 0.01562 | 0.0091105 | -0.0029457 | 0.011236 | 0.0047984 | b12 | -0.0015679 | -0.0065699 | 0.0025872 | 0.0011685 | -0.0089666 | -0.0055461 | -0.0021978 | 0.0013329 | 0.0044641 | -0.01592 | 0.001828 | 0.0028506 | 0.0022066 | -0.0040462 | 0.0077264 | 0.0080624 | 0.011569 | 0.0034909 | -0.0026912 | -0.0010356 | -9.168e-06
b13 | 0.082358 | 0.073669 | 0.097576 | 0.080184 | 0.068739 | 0.068091 | 0.070423 | 0.077962 | 0.078016 | 0.078984 | 0.074418 | 0.077254 | 0.085301 | 0.091059 | 0.082933 | 0.078372 | 0.077205 | 0.10841 | 0.091046 | 0.064534 | 0.080615 | b13 | 0.087597 | 0.081222 | 0.091648 | 0.096217 | 0.094846 | 0.076443 | 0.082321 | 0.087343 | 0.094444 | 0.098102 | 0.085405 | 0.082155 | 0.079943 | 0.077726 | 0.069227 | 0.057 | 0.06385 | 0.061225 | 0.08103 | 0.051314 | 0.080384
b14 | 0.0064699 | -0.0020959 | -0.0035646 | -0.0083115 | 0.0039222 | 0.0017928 | -0.00014018 | -0.00078612 | -0.001389 | -0.00099061 | -0.0029939 | -0.003597 | -0.006099 | -0.0023108 | -0.0026809 | 0.00022588 | -3.9664e-05 | -0.0062779 | -0.00047743 | -0.0037052 | -0.0017842 | b14 | 0.0075583 | -0.0043702 | -0.0037237 | -0.0084919 | -0.013503 | -0.011747 | -0.0039581 | -0.0055942 | -0.0077963 | -0.010872 | -0.0096122 | -0.0070958 | -0.0049868 | -0.0076268 | -0.0044783 | -0.0053962 | -0.0026749 | -0.0029594 | -0.0039126 | -0.00979 | -0.006267
b15 | -0.012909 | 0.027385 | 0.032714 | 0.031467 | 0.030468 | 0.03304 | 0.027087 | 0.021318 | 0.023175 | 0.024971 | 0.027474 | 0.022954 | 0.02228 | 0.024226 | 0.030125 | 0.034686 | 0.039048 | 0.034709 | 0.013729 | 0.0033055 | 0.025811 | b15 | -0.012675 | 0.026777 | 0.027025 | 0.029717 | 0.031565 | 0.035146 | 0.026789 | 0.027044 | 0.032452 | 0.02741 | 0.029694 | 0.024052 | 0.023413 | 0.025732 | 0.031557 | 0.040479 | 0.048834 | 0.044698 | 0.02471 | 0.015138 | 0.029127
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 28.2407 | 3.8305 | 1.6522 | -0.65335 | 3.6968 | 2.8449 | 9.0424 | 6.0258 | 2.7637 | -2.6131 | -3.5483 | -3.6048 | -7.9167 | -5.2994 | -8.4379 | -0.9485 | 6.2651 | 3.9575 | -19.4177 | -4.277 | 0 | a1 | 31.1029 | -1.0283 | -4.5137 | -4.3167 | -1.8588 | 2.3097 | 7.4811 | 6.146 | 3.1673 | -3.0848 | -6.1097 | -6.2387 | -6.5662 | -3.6691 | -12.1136 | -2.7105 | 2.4302 | 1.1722 | -1.2328 | 18.0162 | 0
a2 | 1.102 | -1.9764 | -0.51357 | 0.67209 | -5.4289 | -6.3146 | -2.5097 | -0.36768 | -1.6615 | -0.95842 | -0.65559 | -0.94169 | 1.2519 | 2.901 | 2.9076 | -2.9558 | -9.9387 | -5.5722 | 5.8728 | 3.422 | -1.2249 | a2 | 1.7171 | -0.32518 | -1.877 | 1.1561 | 1.2937 | -3.644 | -2.2207 | -0.21482 | 2.4458 | 1.7635 | -1.0296 | -0.5649 | -0.78188 | -0.45295 | 1.4259 | 0.96198 | -2.3088 | 1.4739 | 0.11858 | -1.0366 | -0.12505
a3 | -2.6159 | 0.41408 | 0.092681 | -0.3227 | -0.20746 | -0.81759 | 0.23535 | -0.033738 | -0.095659 | 0.14234 | -0.52986 | 0.52498 | -0.45191 | 0.064574 | 0.67505 | -0.23686 | -0.6082 | -0.96866 | 0.79393 | 1.3847 | -0.10663 | a3 | -2.7169 | -0.24564 | -0.24859 | -1.3457 | 0.20295 | -0.92718 | -0.11083 | 0.8934 | 0.2025 | 0.12981 | -0.42665 | 0.31458 | 0.40543 | 0.2369 | -0.20849 | 0.27098 | 0.18595 | 0.22302 | 0.055281 | 0.80688 | -0.077913
a4 | -1.3584 | -0.21593 | 0.012872 | -0.3297 | 1.0034 | 0.88697 | 0.18288 | 0.035392 | -0.31062 | -0.27711 | -0.1138 | 0.2202 | 0.17 | -0.41711 | -0.1086 | 1.7146 | 3.8564 | 2.0608 | -0.25755 | -0.92121 | 0.35253 | a4 | 0.050389 | 0.34694 | 0.36132 | -0.40051 | -0.46305 | 0.997 | 0.03868 | 0.0094122 | -0.35994 | -0.566 | -0.34473 | 0.084544 | 0.34493 | 0.54626 | -0.56842 | 0.097601 | 1.8706 | -0.088431 | -0.099333 | 0.98935 | 0.12667
a5 | 1.5745 | -0.22766 | 0.15282 | 0.97502 | -0.11324 | 0.3876 | 0.14291 | 0.070961 | 0.23817 | 0.26877 | 0.18852 | 0.33531 | 0.6791 | 0.23719 | 0.62951 | 0.35822 | -0.10798 | 0.2244 | -0.27376 | -0.59697 | 0.24678 | a5 | 1.3902 | 0.23621 | 0.45595 | 0.71907 | 0.22146 | 0.081142 | -0.095683 | -0.036537 | 0.10241 | 0.026685 | -0.1247 | 0.18174 | 0.10884 | 0.044105 | 0.11324 | 0.736 | 0.31862 | 0.74809 | -0.079636 | -0.28708 | 0.22934
a6 | 0.062984 | -0.24177 | -0.33725 | -0.43954 | -0.34934 | -0.34785 | -0.3658 | -0.0085003 | 0.058398 | -0.071961 | -0.23648 | -0.19416 | -0.30862 | -0.46492 | -0.21925 | -0.66965 | -0.60499 | -0.039584 | 0.35759 | 0.31781 | -0.22187 | a6 | 0.27582 | -0.069312 | 0.27463 | 0.012436 | -0.27621 | 0.037349 | 0.33236 | 0.12296 | -0.17796 | -0.18888 | -0.030848 | 0.0089052 | -0.22568 | -0.16457 | 0.037801 | -0.48685 | -0.87848 | -0.38544 | 0.070041 | -0.058691 | -0.097066
a7 | 1.597 | -0.046323 | 0.00099876 | -0.10765 | 0.14955 | 0.064777 | 0.030474 | 0.010156 | 0.0052879 | -0.065297 | -0.049615 | -0.082798 | -0.056479 | -0.025514 | -0.21615 | -0.01709 | 0.22383 | 0.15042 | 0.025612 | 0.0099078 | 0.048995 | a7 | 1.6872 | 0.0064331 | -0.0024158 | -0.10864 | -0.095954 | 0.10887 | 0.11849 | 0.031587 | 0.006473 | 0.074156 | 0.14468 | 0.042777 | 0.12076 | 0.092283 | 0.051555 | -0.14855 | 0.0012697 | -0.068127 | 0.037107 | -0.021911 | 0.0723
a8 | 0.13204 | -0.0046498 | 0.030621 | 0.088655 | -0.11233 | -0.082864 | 0.027494 | -0.011591 | -0.058782 | 0.0043329 | -0.0018218 | 0.048836 | 0.095772 | 0.23138 | 0.14229 | 0.086503 | -0.27932 | -0.26357 | -0.12481 | -0.15344 | -0.010316 | a8 | 0.055376 | 0.025121 | -0.11713 | 0.10228 | 0.11169 | -0.080087 | -0.097926 | -0.065253 | 0.050033 | 0.063843 | -0.017303 | -0.0095263 | 0.028034 | 0.021483 | 0.089085 | 0.17702 | 0.16886 | 0.14173 | 0.037845 | -0.024061 | 0.033769
a9 | -0.15667 | 0.099564 | 0.018938 | 0.0041527 | 0.0075538 | -0.032249 | 0.042857 | 0.030678 | 0.012983 | 0.038756 | 0.021555 | 0.038148 | -0.071751 | 0.0086179 | 0.097886 | -0.01083 | -0.02031 | -0.072526 | 0.005712 | -0.026192 | 0.0058004 | a9 | -0.1685 | 0.032139 | -0.023684 | -0.055575 | 0.073512 | 0.026062 | 0.012771 | 0.057074 | 0.072646 | 0.032324 | 0.031153 | 0.042288 | 0.029693 | 0.01946 | 0.037889 | 0.012079 | 0.017132 | 0.030797 | 0.038845 | -0.016064 | 0.019747
a10 | 0.15654 | -0.0053072 | -0.073221 | -0.011518 | -0.020915 | 0.0015205 | -0.014452 | 0.029588 | 0.0035141 | -0.0020642 | -0.01019 | 0.033303 | 0.016292 | -0.0051143 | 0.027856 | -0.029429 | -0.042865 | -0.084132 | 0.1059 | -0.01812 | 0 | a10 | 0.15308 | -0.0095896 | -0.019127 | -0.026773 | -0.0064168 | 0.00010277 | -0.017203 | 0.0022745 | 0.00071388 | -0.0029441 | -0.024857 | 0.026795 | 0.022777 | -0.0037054 | 0.00021427 | 0.026638 | 0.0021622 | -0.0085185 | -0.0099905 | -0.070148 | 0
a11 | 0.21532 | 0.030407 | 0.046463 | 0.054096 | 0.065413 | 0.093922 | 0.06948 | 0.057424 | 0.057913 | 0.038696 | 0.047527 | 0.045696 | 0.075734 | 0.052894 | 0.015477 | 0.044972 | 0.064429 | 0.071583 | 0.019959 | 0.027329 | 0.057024 | a11 | 0.21918 | 0.053256 | 0.060743 | 0.062264 | 0.014623 | 0.071502 | 0.085531 | 0.052679 | 0.045368 | 0.055197 | 0.062341 | 0.054785 | 0.057181 | 0.058706 | 0.035495 | 0.066298 | 0.04023 | 0.035127 | 0.039221 | 0.038069 | 0.05748
a12 | 0.025578 | -0.0032595 | -0.015081 | -0.013061 | 0.0030239 | -0.00027251 | -0.015525 | 0.0041711 | 0.0052659 | -0.00076441 | -0.0098234 | -0.0072969 | -0.018758 | -0.034509 | -0.0086583 | -0.035711 | 0.0058594 | 0.0045041 | 0.020054 | -0.00022448 | -0.0054692 | a12 | 0.024987 | -0.00035401 | 0.015143 | -0.011097 | -0.020798 | 0.0061692 | 0.017756 | 0.012528 | -0.016949 | -0.013434 | 0.0040954 | 0.0056935 | -0.0030668 | -0.0014426 | -0.0040741 | -0.017144 | -0.025542 | -0.017726 | 0.00094982 | -0.018895 | -0.0034419
a13 | -0.0010713 | 0.0099459 | 0.0067215 | 0.0078396 | 0.0062051 | 0.0077752 | 0.009199 | 0.0086918 | 0.0072051 | 0.010059 | 0.0058397 | 0.0079123 | 0.012053 | 0.0077402 | 0.011902 | 0.0076815 | 0.0011201 | 0.013474 | 0.0050532 | 0.00611 | 0.0077792 | a13 | -0.00083257 | 0.008314 | 0.013359 | 0.012914 | 0.01643 | 0.010742 | 0.013906 | 0.010767 | 0.010456 | 0.014471 | 0.0085241 | 0.0059347 | 0.0084452 | 0.0081405 | 0.0095509 | 0.010318 | 0.01401 | 0.015422 | 0.0093397 | 0.0058273 | 0.010646
a14 | -0.0045769 | -0.0062403 | -0.0082282 | -0.008005 | -0.0075692 | -0.0050506 | -0.0061884 | -0.0046826 | -0.0081583 | -0.0081394 | -0.0096573 | -0.0036662 | -0.0053869 | -0.0010684 | -0.0067033 | -3.5186e-05 | -0.0015401 | -0.01451 | 0.0013432 | 0.0064712 | -0.0053321 | a14 | -0.005709 | -0.0068758 | -0.0090086 | -0.0047912 | -0.0096052 | -0.012945 | -0.012286 | -0.010165 | -0.0092096 | -0.0087253 | -0.012801 | -0.0088812 | -0.0093283 | -0.0085915 | -0.0097974 | -0.0068345 | -0.0014312 | -0.011464 | -0.0098544 | -0.0075486 | -0.008886
a15 | 0.0058299 | -0.0024925 | 0.0001557 | -0.0034958 | 0.0022437 | 0.00020556 | 0.00097321 | 0.0009128 | 5.4988e-05 | 0.00050428 | -0.00051652 | 7.5786e-05 | 0.0007135 | 0.0024021 | 0.0055625 | 0.00296 | 0.0027176 | -0.0032808 | -0.00056932 | -0.0021878 | 0.000583 | a15 | 0.0038405 | -0.011216 | -0.0073172 | -0.010737 | -0.0047289 | -0.005537 | -0.0058444 | -0.0032725 | -0.0070549 | -0.0087171 | -0.0057503 | -0.0031338 | -0.001605 | -0.0054487 | -0.002033 | -0.0012753 | -0.00011059 | -0.0019096 | -0.0046063 | -0.0044113 | -0.0047235
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | 0.00028129 | 3.7553e-05 | 6.0038e-05 | -1.1273e-05 | 0 | 0.00012634 | 1.5908e-05 | -3.5178e-05 | -5.0303e-05 | 0 | -7.8407e-05 | -6.87e-05 | -0.00023429 | 0 | -4.3889e-05 | -4.0516e-05 | 0 | 0.00016946 | 6.1755e-05 | -0.00019494 | 9.7168e-05 | Dx (m) | 0.00013347 | 0.00014055 | 2.9013e-05 | -1.759e-05 | -0.00020371 | -7.6161e-05 | 4.3641e-05 | -6.9904e-05 | 0.00011483 | -2.266e-05 | 3.1198e-05 | 0 | -8.8446e-05 | 0.00010173 | -6.364e-05 | -0.0001643 | 4.1387e-05 | 0.00013949 | 0 | -0.00015519 | 0.00026484
Dy (m) | -0.00059247 | 1.2539e-05 | 0.00020844 | 3.3137e-05 | 5.724e-05 | -2.3573e-05 | 3.6765e-05 | -7.7152e-05 | 0 | 0 | 3.3098e-05 | -8.4041e-05 | -1.5638e-05 | 1.4306e-05 | -7.439e-05 | 8.5679e-05 | 0.00012433 | 0.00023839 | -0.00029164 | 6.152e-05 | 0.00042819 | Dy (m) | -0.00052194 | 1.3993e-05 | 5.1603e-05 | 7.9114e-05 | 2.4129e-05 | 0 | 4.2008e-05 | 0 | -1.096e-05 | 1.062e-05 | 6.6015e-05 | -7.2033e-05 | -5.3643e-05 | 0 | 0 | -5.2177e-05 | 0 | 1.5615e-05 | 2.6291e-05 | 0.00021715 | 0.00056591
Magnetic length (m) |  | 14.3887 | [19 empty] | Magnetic length (m) |  | 14.3864
Transfer funct. (mT/KA) |  | 704.0588 | [19 empty] | Transfer funct. (mT/KA) |  | 704.098
Coil positioning (mm) |  | 0.071626 | [19 empty] | Coil positioning (mm) |  | 0.0566
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMBBRA001-01000006 |  |  | Aperture 1 - Cold mass | [16 empty] | Integrals | File | HCMBBRA001-1000006 |  |  | Aperture 2 - Cold mass | [16 empty] | Integrals
C1/i (mT/KA) | 414.7059 | 704 | 704.2353 | 704.3529 | 704.2353 | 704.1176 | 704.1176 | 704.1176 | 703.7647 | 703.7647 | 703.8824 | 703.8824 | 704 | 704.1176 | 704.2353 | 704.1176 | 704.2353 | 704.2353 | 703.6471 | 421.2941 | 0.08611 | C1/i (mT/KA) | 411.7647 | 703.8824 | 704.1176 | 704.3529 | 704.2353 | 704.2353 | 704.1176 | 704.2353 | 704.1176 | 704 | 703.8824 | 703.8824 | 704 | 704 | 704 | 704.1176 | 704.3529 | 704.3529 | 703.8824 | 422.2353 | 0.086101
Angle (mrad) | 2.8241 | 0.38305 | 0.16522 | -0.065335 | 0.36968 | 0.28449 | 0.90424 | 0.60258 | 0.27637 | -0.26131 | -0.35483 | -0.36048 | -0.79167 | -0.52994 | -0.84379 | -0.09485 | 0.62651 | 0.39575 | -1.9418 | -0.42771 | 2.0009 | Angle (mrad) | 3.1103 | -0.10283 | -0.45137 | -0.43167 | -0.18588 | 0.23097 | 0.74811 | 0.6146 | 0.31673 | -0.30848 | -0.61097 | -0.62387 | -0.65662 | -0.36691 | -1.2114 | -0.27104 | 0.24302 | 0.11722 | -0.12329 | 1.8016 | 2.1186
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 24.1508 | -0.20815 | 1.1753 | 1.803 | 3.0137 | 1.3872 | 0.50984 | 1.2083 | 1.3831 | 1.8394 | 1.7036 | 2.0695 | 2.346 | 1.2265 | 3.4709 | 1.4898 | 2.4885 | 3.5409 | 6.3562 | 27.1192 | 3.5053 | b2 | -26.0644 | -1.3265 | -0.40607 | -1.3056 | -1.6491 | -1.119 | -1.9196 | -0.37853 | -0.34876 | 0.037425 | -0.46549 | -0.52589 | -1.0722 | -0.59646 | -0.013477 | -0.11053 | -1.367 | -1.4095 | -0.79316 | -23.1144 | -2.2869
b3 | 36.0487 | 7.6097 | 5.3324 | 5.0777 | 5.0305 | 3.6252 | 6.9812 | 7.846 | 7.1086 | 7.4885 | 6.5138 | 7.0792 | 6.1196 | 4.4545 | 0.17147 | 0.43061 | -1.6483 | 0.12133 | 5.3183 | 2.2786 | 5.5896 | b3 | 37.0654 | 7.3716 | 6.3455 | 5.8105 | 4.9088 | 1.1977 | 5.1592 | 6.2113 | 4.5914 | 3.8907 | 5.2972 | 6.4675 | 5.5141 | 4.6656 | 4.9033 | 3.2703 | -0.46698 | -0.068249 | 7.7297 | 2.7974 | 5.5329
b4 | 0.98466 | 0.16286 | 0.13134 | 0.66011 | 0.13157 | -0.20235 | 0.18924 | 0.20594 | 0.34848 | 0.37467 | 0.44313 | 0.4572 | 0.97486 | -0.012393 | 0.52981 | -0.29115 | 0.058245 | 0.078621 | 0.12472 | 0.15413 | 0.26252 | b4 | -1.6186 | -0.22131 | 0.094848 | 0.14108 | 0.86102 | 0.19847 | 0.14801 | 0.13287 | 0.129 | 0.61616 | -0.18425 | 0.0020129 | -0.20052 | -0.22276 | -0.06087 | 0.17036 | -0.96533 | -0.84461 | -0.30051 | -2.2707 | -0.14671
b5 | -3.0874 | 1.6662 | 1.2826 | 1.6046 | 2.1397 | 2.3919 | 2.0899 | 1.6494 | 1.5815 | 1.9261 | 1.952 | 1.8846 | 1.277 | 1.4614 | 2.5406 | 2.839 | 3.4607 | 2.1357 | 1.4473 | -1.7232 | 1.6928 | b5 | -3.0619 | 1.2528 | 1.1281 | 1.0317 | 1.6837 | 2.7033 | 1.4239 | 0.78945 | 0.74722 | 1.456 | 1.4285 | 1.5346 | 1.3462 | 1.7552 | 2.1645 | 3.2562 | 3.3445 | 3.2132 | 1.1426 | -2.5445 | 1.464
b6 | -0.27153 | 0.072049 | -0.093815 | -0.31729 | -0.11855 | 0.15947 | 0.045135 | -0.00089776 | 0.0074868 | -0.1351 | 0.055129 | 0.0086981 | -0.40338 | -0.059487 | -0.48611 | 0.094001 | -0.23708 | -0.25889 | 0.2134 | 0.34309 | -0.073514 | b6 | -0.65512 | 0.024446 | 0.028209 | -0.10006 | -0.33709 | -0.2718 | -0.080745 | -0.18272 | -0.079444 | -0.34199 | -0.18005 | -0.19473 | -0.045883 | -0.064501 | -0.27697 | -0.28534 | 0.15473 | 0.072876 | -0.083397 | -0.061005 | -0.13886
b7 | 1.8777 | 0.45365 | 0.91037 | 0.6989 | 0.47281 | 0.50953 | 0.51842 | 0.68568 | 0.79119 | 0.62798 | 0.52742 | 0.64137 | 0.76469 | 0.84045 | 0.80678 | 0.78993 | 0.93128 | 1.3616 | 0.43948 | 0.2525 | 0.73117 | b7 | 1.9892 | 0.64918 | 0.96333 | 1.0178 | 0.99734 | 0.68414 | 0.9184 | 0.75808 | 1.0177 | 1.0946 | 0.70135 | 0.72181 | 0.71892 | 0.69578 | 0.50507 | 0.084218 | 0.4743 | 0.52068 | 0.64311 | 0.45875 | 0.76125
b8 | -0.0001683 | -0.074361 | 0.0081964 | 0.0078305 | 0.034296 | 0.029351 | -0.034034 | 0.0099702 | -0.03179 | -0.0054032 | -0.085037 | -0.074655 | 0.04023 | 0.041239 | 0.10004 | 0.044135 | 0.072614 | 0.093559 | -0.057305 | 0.0015108 | 0.0062466 | b8 | 0.026914 | -0.0087197 | 0.014258 | -0.01313 | -0.049602 | -0.0079205 | -0.042566 | 0.0072428 | -0.0067001 | -0.02761 | 0.0010594 | 0.0096868 | 0.028598 | -0.0034353 | 0.043467 | -0.087237 | 0.069704 | 0.096065 | -0.0025858 | 0.051845 | 0.003512
b9 | 0.28302 | 0.32703 | 0.20605 | 0.23407 | 0.28768 | 0.29981 | 0.34194 | 0.34479 | 0.33933 | 0.32244 | 0.33145 | 0.33551 | 0.31469 | 0.25608 | 0.19138 | 0.20194 | 0.14668 | 0.063917 | 0.32879 | 0.32231 | 0.27273 | b9 | 0.28402 | 0.33945 | 0.24954 | 0.23445 | 0.19382 | 0.24106 | 0.30118 | 0.2898 | 0.31263 | 0.24621 | 0.28285 | 0.34014 | 0.32643 | 0.30003 | 0.30963 | 0.27793 | 0.22063 | 0.23642 | 0.30635 | 0.24318 | 0.27732
b10 | -0.16733 | -0.013663 | -0.014909 | 0.005096 | 0.0034642 | -0.046835 | -0.0044726 | 0.010344 | 0.018116 | -0.0021329 | 0.030286 | 0.023435 | 0.083949 | 0.00021469 | 0.015549 | 0.01674 | 0.0029175 | -0.044363 | -0.026073 | 0.068389 | 0 | b10 | -0.11098 | -0.051664 | -0.008319 | 0.0089966 | 0.070489 | 0.028237 | -0.013464 | 0.024925 | -0.039408 | 0.0081981 | -0.0087809 | 0.00037841 | 0.030831 | -0.037204 | 0.024459 | 0.061377 | -0.015367 | -0.050574 | -0.0027712 | 0.057593 | 0
b11 | 0.55044 | 0.62874 | 0.583 | 0.60918 | 0.62515 | 0.61245 | 0.63809 | 0.62571 | 0.61802 | 0.63424 | 0.63637 | 0.63632 | 0.61468 | 0.60518 | 0.62833 | 0.60479 | 0.58361 | 0.54812 | 0.62031 | 0.58827 | 0.61126 | b11 | 0.55231 | 0.63037 | 0.5961 | 0.58891 | 0.58529 | 0.62221 | 0.60733 | 0.6064 | 0.57897 | 0.58899 | 0.6107 | 0.62672 | 0.62744 | 0.6221 | 0.65185 | 0.65703 | 0.62291 | 0.61988 | 0.63247 | 0.57548 | 0.61214
b12 | -0.01142 | -0.0080737 | 0.0031704 | 0.0016997 | 0.013583 | 0.0039407 | -0.0038323 | 0.00032501 | 0.0070871 | -0.0032988 | 0.00026372 | -0.0004637 | 0.0087017 | 0.0076533 | 0.025629 | 0.013891 | 0.01562 | 0.0091105 | -0.0029457 | 0.011236 | 0.0047984 | b12 | -0.0015679 | -0.0065699 | 0.0025872 | 0.0011685 | -0.0089666 | -0.0055461 | -0.0021978 | 0.0013329 | 0.0044641 | -0.01592 | 0.001828 | 0.0028506 | 0.0022066 | -0.0040462 | 0.0077264 | 0.0080624 | 0.011569 | 0.0034909 | -0.0026912 | -0.0010356 | -9.168e-06
b13 | 0.082358 | 0.073669 | 0.097576 | 0.080184 | 0.068739 | 0.068091 | 0.070423 | 0.077962 | 0.078016 | 0.078984 | 0.074418 | 0.077254 | 0.085301 | 0.091059 | 0.082933 | 0.078372 | 0.077205 | 0.10841 | 0.091046 | 0.064534 | 0.080615 | b13 | 0.087597 | 0.081222 | 0.091648 | 0.096217 | 0.094846 | 0.076443 | 0.082321 | 0.087343 | 0.094444 | 0.098102 | 0.085405 | 0.082155 | 0.079943 | 0.077726 | 0.069227 | 0.057 | 0.06385 | 0.061225 | 0.08103 | 0.051314 | 0.080384
b14 | 0.0064699 | -0.0020959 | -0.0035646 | -0.0083115 | 0.0039222 | 0.0017928 | -0.00014018 | -0.00078612 | -0.001389 | -0.00099061 | -0.0029939 | -0.003597 | -0.006099 | -0.0023108 | -0.0026809 | 0.00022588 | -3.9664e-05 | -0.0062779 | -0.00047743 | -0.0037052 | -0.0017842 | b14 | 0.0075583 | -0.0043702 | -0.0037237 | -0.0084919 | -0.013503 | -0.011747 | -0.0039581 | -0.0055942 | -0.0077963 | -0.010872 | -0.0096122 | -0.0070958 | -0.0049868 | -0.0076268 | -0.0044783 | -0.0053962 | -0.0026749 | -0.0029594 | -0.0039126 | -0.00979 | -0.006267
b15 | -0.012909 | 0.027385 | 0.032714 | 0.031467 | 0.030468 | 0.03304 | 0.027087 | 0.021318 | 0.023175 | 0.024971 | 0.027474 | 0.022954 | 0.02228 | 0.024226 | 0.030125 | 0.034686 | 0.039048 | 0.034709 | 0.013729 | 0.0033055 | 0.025811 | b15 | -0.012675 | 0.026777 | 0.027025 | 0.029717 | 0.031565 | 0.035146 | 0.026789 | 0.027044 | 0.032452 | 0.02741 | 0.029694 | 0.024052 | 0.023413 | 0.025732 | 0.031557 | 0.040479 | 0.048834 | 0.044698 | 0.02471 | 0.015138 | 0.029127
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 28.2407 | 3.8305 | 1.6522 | -0.65335 | 3.6968 | 2.8449 | 9.0424 | 6.0258 | 2.7637 | -2.6131 | -3.5483 | -3.6048 | -7.9167 | -5.2994 | -8.4379 | -0.9485 | 6.2651 | 3.9575 | -19.4177 | -4.277 | 0 | a1 | 31.1029 | -1.0283 | -4.5137 | -4.3167 | -1.8588 | 2.3097 | 7.4811 | 6.146 | 3.1673 | -3.0848 | -6.1097 | -6.2387 | -6.5662 | -3.6691 | -12.1136 | -2.7105 | 2.4302 | 1.1722 | -1.2328 | 18.0162 | 0
a2 | 1.102 | -1.9764 | -0.51357 | 0.67209 | -5.4289 | -6.3146 | -2.5097 | -0.36768 | -1.6615 | -0.95842 | -0.65559 | -0.94169 | 1.2519 | 2.901 | 2.9076 | -2.9558 | -9.9387 | -5.5722 | 5.8728 | 3.422 | -1.2249 | a2 | 1.7171 | -0.32518 | -1.877 | 1.1561 | 1.2937 | -3.644 | -2.2207 | -0.21482 | 2.4458 | 1.7635 | -1.0296 | -0.5649 | -0.78188 | -0.45295 | 1.4259 | 0.96198 | -2.3088 | 1.4739 | 0.11858 | -1.0366 | -0.12505
a3 | -2.6159 | 0.41408 | 0.092681 | -0.3227 | -0.20746 | -0.81759 | 0.23535 | -0.033738 | -0.095659 | 0.14234 | -0.52986 | 0.52498 | -0.45191 | 0.064574 | 0.67505 | -0.23686 | -0.6082 | -0.96866 | 0.79393 | 1.3847 | -0.10663 | a3 | -2.7169 | -0.24564 | -0.24859 | -1.3457 | 0.20295 | -0.92718 | -0.11083 | 0.8934 | 0.2025 | 0.12981 | -0.42665 | 0.31458 | 0.40543 | 0.2369 | -0.20849 | 0.27098 | 0.18595 | 0.22302 | 0.055281 | 0.80688 | -0.077913
a4 | -1.3584 | -0.21593 | 0.012872 | -0.3297 | 1.0034 | 0.88697 | 0.18288 | 0.035392 | -0.31062 | -0.27711 | -0.1138 | 0.2202 | 0.17 | -0.41711 | -0.1086 | 1.7146 | 3.8564 | 2.0608 | -0.25755 | -0.92121 | 0.35253 | a4 | 0.050389 | 0.34694 | 0.36132 | -0.40051 | -0.46305 | 0.997 | 0.03868 | 0.0094122 | -0.35994 | -0.566 | -0.34473 | 0.084544 | 0.34493 | 0.54626 | -0.56842 | 0.097601 | 1.8706 | -0.088431 | -0.099333 | 0.98935 | 0.12667
a5 | 1.5745 | -0.22766 | 0.15282 | 0.97502 | -0.11324 | 0.3876 | 0.14291 | 0.070961 | 0.23817 | 0.26877 | 0.18852 | 0.33531 | 0.6791 | 0.23719 | 0.62951 | 0.35822 | -0.10798 | 0.2244 | -0.27376 | -0.59697 | 0.24678 | a5 | 1.3902 | 0.23621 | 0.45595 | 0.71907 | 0.22146 | 0.081142 | -0.095683 | -0.036537 | 0.10241 | 0.026685 | -0.1247 | 0.18174 | 0.10884 | 0.044105 | 0.11324 | 0.736 | 0.31862 | 0.74809 | -0.079636 | -0.28708 | 0.22934
a6 | 0.062984 | -0.24177 | -0.33725 | -0.43954 | -0.34934 | -0.34785 | -0.3658 | -0.0085003 | 0.058398 | -0.071961 | -0.23648 | -0.19416 | -0.30862 | -0.46492 | -0.21925 | -0.66965 | -0.60499 | -0.039584 | 0.35759 | 0.31781 | -0.22187 | a6 | 0.27582 | -0.069312 | 0.27463 | 0.012436 | -0.27621 | 0.037349 | 0.33236 | 0.12296 | -0.17796 | -0.18888 | -0.030848 | 0.0089052 | -0.22568 | -0.16457 | 0.037801 | -0.48685 | -0.87848 | -0.38544 | 0.070041 | -0.058691 | -0.097066
a7 | 1.597 | -0.046323 | 0.00099876 | -0.10765 | 0.14955 | 0.064777 | 0.030474 | 0.010156 | 0.0052879 | -0.065297 | -0.049615 | -0.082798 | -0.056479 | -0.025514 | -0.21615 | -0.01709 | 0.22383 | 0.15042 | 0.025612 | 0.0099078 | 0.048995 | a7 | 1.6872 | 0.0064331 | -0.0024158 | -0.10864 | -0.095954 | 0.10887 | 0.11849 | 0.031587 | 0.006473 | 0.074156 | 0.14468 | 0.042777 | 0.12076 | 0.092283 | 0.051555 | -0.14855 | 0.0012697 | -0.068127 | 0.037107 | -0.021911 | 0.0723
a8 | 0.13204 | -0.0046498 | 0.030621 | 0.088655 | -0.11233 | -0.082864 | 0.027494 | -0.011591 | -0.058782 | 0.0043329 | -0.0018218 | 0.048836 | 0.095772 | 0.23138 | 0.14229 | 0.086503 | -0.27932 | -0.26357 | -0.12481 | -0.15344 | -0.010316 | a8 | 0.055376 | 0.025121 | -0.11713 | 0.10228 | 0.11169 | -0.080087 | -0.097926 | -0.065253 | 0.050033 | 0.063843 | -0.017303 | -0.0095263 | 0.028034 | 0.021483 | 0.089085 | 0.17702 | 0.16886 | 0.14173 | 0.037845 | -0.024061 | 0.033769
a9 | -0.15667 | 0.099564 | 0.018938 | 0.0041527 | 0.0075538 | -0.032249 | 0.042857 | 0.030678 | 0.012983 | 0.038756 | 0.021555 | 0.038148 | -0.071751 | 0.0086179 | 0.097886 | -0.01083 | -0.02031 | -0.072526 | 0.005712 | -0.026192 | 0.0058004 | a9 | -0.1685 | 0.032139 | -0.023684 | -0.055575 | 0.073512 | 0.026062 | 0.012771 | 0.057074 | 0.072646 | 0.032324 | 0.031153 | 0.042288 | 0.029693 | 0.01946 | 0.037889 | 0.012079 | 0.017132 | 0.030797 | 0.038845 | -0.016064 | 0.019747
a10 | 0.15654 | -0.0053072 | -0.073221 | -0.011518 | -0.020915 | 0.0015205 | -0.014452 | 0.029588 | 0.0035141 | -0.0020642 | -0.01019 | 0.033303 | 0.016292 | -0.0051143 | 0.027856 | -0.029429 | -0.042865 | -0.084132 | 0.1059 | -0.01812 | 0 | a10 | 0.15308 | -0.0095896 | -0.019127 | -0.026773 | -0.0064168 | 0.00010277 | -0.017203 | 0.0022745 | 0.00071388 | -0.0029441 | -0.024857 | 0.026795 | 0.022777 | -0.0037054 | 0.00021427 | 0.026638 | 0.0021622 | -0.0085185 | -0.0099905 | -0.070148 | 0
a11 | 0.21532 | 0.030407 | 0.046463 | 0.054096 | 0.065413 | 0.093922 | 0.06948 | 0.057424 | 0.057913 | 0.038696 | 0.047527 | 0.045696 | 0.075734 | 0.052894 | 0.015477 | 0.044972 | 0.064429 | 0.071583 | 0.019959 | 0.027329 | 0.057024 | a11 | 0.21918 | 0.053256 | 0.060743 | 0.062264 | 0.014623 | 0.071502 | 0.085531 | 0.052679 | 0.045368 | 0.055197 | 0.062341 | 0.054785 | 0.057181 | 0.058706 | 0.035495 | 0.066298 | 0.04023 | 0.035127 | 0.039221 | 0.038069 | 0.05748
a12 | 0.025578 | -0.0032595 | -0.015081 | -0.013061 | 0.0030239 | -0.00027251 | -0.015525 | 0.0041711 | 0.0052659 | -0.00076441 | -0.0098234 | -0.0072969 | -0.018758 | -0.034509 | -0.0086583 | -0.035711 | 0.0058594 | 0.0045041 | 0.020054 | -0.00022448 | -0.0054692 | a12 | 0.024987 | -0.00035401 | 0.015143 | -0.011097 | -0.020798 | 0.0061692 | 0.017756 | 0.012528 | -0.016949 | -0.013434 | 0.0040954 | 0.0056935 | -0.0030668 | -0.0014426 | -0.0040741 | -0.017144 | -0.025542 | -0.017726 | 0.00094982 | -0.018895 | -0.0034419
a13 | -0.0010713 | 0.0099459 | 0.0067215 | 0.0078396 | 0.0062051 | 0.0077752 | 0.009199 | 0.0086918 | 0.0072051 | 0.010059 | 0.0058397 | 0.0079123 | 0.012053 | 0.0077402 | 0.011902 | 0.0076815 | 0.0011201 | 0.013474 | 0.0050532 | 0.00611 | 0.0077792 | a13 | -0.00083257 | 0.008314 | 0.013359 | 0.012914 | 0.01643 | 0.010742 | 0.013906 | 0.010767 | 0.010456 | 0.014471 | 0.0085241 | 0.0059347 | 0.0084452 | 0.0081405 | 0.0095509 | 0.010318 | 0.01401 | 0.015422 | 0.0093397 | 0.0058273 | 0.010646
a14 | -0.0045769 | -0.0062403 | -0.0082282 | -0.008005 | -0.0075692 | -0.0050506 | -0.0061884 | -0.0046826 | -0.0081583 | -0.0081394 | -0.0096573 | -0.0036662 | -0.0053869 | -0.0010684 | -0.0067033 | -3.5186e-05 | -0.0015401 | -0.01451 | 0.0013432 | 0.0064712 | -0.0053321 | a14 | -0.005709 | -0.0068758 | -0.0090086 | -0.0047912 | -0.0096052 | -0.012945 | -0.012286 | -0.010165 | -0.0092096 | -0.0087253 | -0.012801 | -0.0088812 | -0.0093283 | -0.0085915 | -0.0097974 | -0.0068345 | -0.0014312 | -0.011464 | -0.0098544 | -0.0075486 | -0.008886
a15 | 0.0058299 | -0.0024925 | 0.0001557 | -0.0034958 | 0.0022437 | 0.00020556 | 0.00097321 | 0.0009128 | 5.4988e-05 | 0.00050428 | -0.00051652 | 7.5786e-05 | 0.0007135 | 0.0024021 | 0.0055625 | 0.00296 | 0.0027176 | -0.0032808 | -0.00056932 | -0.0021878 | 0.000583 | a15 | 0.0038405 | -0.011216 | -0.0073172 | -0.010737 | -0.0047289 | -0.005537 | -0.0058444 | -0.0032725 | -0.0070549 | -0.0087171 | -0.0057503 | -0.0031338 | -0.001605 | -0.0054487 | -0.002033 | -0.0012753 | -0.00011059 | -0.0019096 | -0.0046063 | -0.0044113 | -0.0047235
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | -0.28129 | -0.037553 | -0.060038 | 0.011273 | 0 | -0.12634 | -0.015908 | 0.035178 | 0.050303 | 0 | 0.078407 | 0.0687 | 0.23429 | 0 | 0.043889 | 0.040516 | 0 | -0.16946 | -0.061755 | 0.19494 | -0.097168 | Dx (mm) | -0.13347 | -0.14055 | -0.029013 | 0.01759 | 0.20371 | 0.076161 | -0.043641 | 0.069904 | -0.11483 | 0.02266 | -0.031198 | 0 | 0.088446 | -0.10173 | 0.06364 | 0.1643 | -0.041387 | -0.13949 | 0 | 0.15519 | -0.26484
Dy (mm) | 0.59247 | -0.012539 | -0.20844 | -0.033137 | -0.05724 | 0.023573 | -0.036765 | 0.077152 | 0 | 0 | -0.033098 | 0.084041 | 0.015638 | -0.014306 | 0.07439 | -0.085679 | -0.12433 | -0.23839 | 0.29164 | -0.06152 | -0.42819 | Dy (mm) | 0.52194 | -0.013993 | -0.051603 | -0.079114 | -0.024129 | 0 | -0.042008 | 0 | 0.01096 | -0.01062 | -0.066015 | 0.072033 | 0.053643 | 0 | 0 | 0.052177 | 0 | -0.015615 | -0.026291 | -0.21715 | -0.56591
Mag. len. (m) | 14.3617 | [20 empty] | Mag. len. (m) | 14.3594
TF (mT/KA) | 704.0588 | [20 empty] | TF (mT/KA) | 704.098
Coil pos. (mm) | 0.071626 | [20 empty] | Coil pos. (mm) | 0.0566
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5
Twist (m2 rad) | 0.025719 | [20 empty] | Twist (m2 rad) | 0.0038566
|  | -0.022486 | -0.0064308 | 0.0012023 | 0.012519 | 0.041029 | 0.016587 | 0.0045303 | 0.0050194 | -0.014198 | -0.0053667 | -0.0071971 | 0.02284 | 2.5889e-07 | -0.037417 | -0.0079218 | 0.011536 | 0.018689 | -0.032934

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 3.5053 | 7.4025 | 2.0382 | 1.4451 | -1.2249 | 3.7441 | -1.4614 | 3.7324 | -2.2869 | 7.3528 | -0.82225 | 0.59105 | -0.12505 | 1.6297 | -0.14935 | 1.647 |  | 0
3 | 5.5896 | 7.5983 | 4.6884 | 2.9931 | -0.10663 | 0.8185 | -0.071036 | 0.48668 | 5.5329 | 7.6139 | 4.5966 | 2.3197 | -0.077913 | 0.80052 | -0.020126 | 0.50999 |  | 0
4 | 0.26252 | 0.3327 | 0.24043 | 0.28486 | 0.35253 | 1.1596 | 0.45915 | 1.1477 | -0.14671 | 0.72535 | -0.032166 | 0.40128 | 0.12667 | 0.60872 | 0.10116 | 0.60641 |  | 0
5 | 1.6928 | 1.4654 | 1.9699 | 0.56991 | 0.24678 | 0.46921 | 0.23044 | 0.32168 | 1.464 | 1.6116 | 1.7466 | 0.84189 | 0.22934 | 0.3958 | 0.2092 | 0.28419 |  | 0
6 | -0.073514 | 0.21194 | -0.082448 | 0.20181 | -0.22187 | 0.27253 | -0.23829 | 0.2389 | -0.13886 | 0.18176 | -0.12473 | 0.16125 | -0.097066 | 0.28251 | -0.10632 | 0.28735 |  | 0
7 | 0.73117 | 0.35702 | 0.71115 | 0.23374 | 0.048995 | 0.37002 | -0.0013684 | 0.10337 | 0.76125 | 0.37443 | 0.73234 | 0.24768 | 0.0723 | 0.38123 | 0.022419 | 0.085679 |  | 0
8 | 0.0062466 | 0.053637 | 0.0053786 | 0.059116 | -0.010316 | 0.13128 | -0.011032 | 0.13205 | 0.003512 | 0.041728 | 0.0013633 | 0.048138 | 0.033769 | 0.084872 | 0.0355 | 0.087485 |  | 0
9 | 0.27273 | 0.077106 | 0.27197 | 0.079293 | 0.0058004 | 0.057831 | 0.011438 | 0.047245 | 0.27732 | 0.041409 | 0.27688 | 0.043604 | 0.019747 | 0.05279 | 0.026602 | 0.030632 |  | 0
10 | 0 | 0.049938 | -0.00016617 | 0.00072868 | 0 | 0.053807 | -1.1355e-05 | 0.00067147 | 0 | 0.043646 | -2.1788e-05 | 0.0010543 | 0 | 0.041405 | -4.5591e-05 | 0.0011079 |  | 0
11 | 0.61126 | 0.026605 | 0.61431 | 0.023004 | 0.057024 | 0.041448 | 0.053373 | 0.019904 | 0.61214 | 0.026377 | 0.61531 | 0.021588 | 0.05748 | 0.040475 | 0.052857 | 0.016001 |  | 0
12 | 0.0047984 | 0.0089111 | 0.0048833 | 0.0091106 | -0.0054692 | 0.015238 | -0.0055879 | 0.014434 | -9.168e-06 | 0.00631 | -0.00024733 | 0.009177 | -0.0034419 | 0.014297 | -0.0033225 | 0.013584 |  | 0
13 | 0.080615 | 0.010509 | 0.081379 | 0.010767 | 0.0077792 | 0.0034043 | 0.008037 | 0.0027235 | 0.080384 | 0.013393 | 0.081294 | 0.012298 | 0.010646 | 0.0040018 | 0.011151 | 0.0030035 |  | 0
14 | -0.0017842 | 0.0034076 | -0.0019511 | 0.0037043 | -0.0053321 | 0.0045298 | -0.0050067 | 0.0035135 | -0.006267 | 0.0044607 | -0.006587 | 0.0050224 | -0.008886 | 0.0027782 | -0.0089062 | 0.0028291 |  | 0
15 | 0.025811 | 0.011885 | 0.027825 | 0.0061172 | 0.000583 | 0.0025227 | 0.00050704 | 0.0022449 | 0.029127 | 0.012244 | 0.03095 | 0.0071614 | -0.0047235 | 0.0035939 | -0.0050165 | 0.0031722 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | 0.84262 | -0.013433
3 | [12 empty] | 3 | 9 | 0.18809 | 0.08403
4 | [12 empty] | 4 | 16 | -0.559 | -0.70665
5 | [12 empty] | 5 | 25 | -0.84823 | -0.7151
6 | [12 empty] | 6 | 36 | -1.5161 | -1.5359
7 | [12 empty] | 7 | 49 | -1.8615 | -1.9264
8 | [12 empty] | 8 | 64 | -2.4264 | -2.735
9 | [12 empty] | 9 | 81 | -2.7935 | -3.3092
10 | [12 empty] | sigma (mm) |  | 0.071626 | 0.0566
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05
17 | 0.00044619 | 0.00011155 | 2.6771e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 25.1614 | -0.17514 | 1.2107 | 1.7975 | 3.0151 | 1.4388 | 0.52189 | 1.1755 | 1.341 | 1.8394 | 1.6456 | 2.0174 | 2.1765 | 1.2264 | 3.4759 | 1.4901 | 2.4974 | 3.5705 | 6.4221 | 27.0569 | 3.5053
3 | 35.9556 | 7.6113 | 5.3333 | 5.0783 | 5.0204 | 3.6244 | 6.9805 | 7.8452 | 7.1055 | 7.4885 | 6.5083 | 7.077 | 6.0798 | 4.4556 | 0.16594 | 0.40677 | -1.733 | 0.036989 | 5.3064 | 2.2833 | 5.5896
4 | 0.99981 | 0.17825 | 0.14196 | 0.64826 | 0.1331 | -0.1291 | 0.19583 | 0.19358 | 0.32977 | 0.37467 | 0.40565 | 0.43337 | 0.90696 | -0.013192 | 0.5146 | -0.32544 | 0.061404 | 0.15119 | 0.12697 | 0.24181 | 0.26252
5 | -3.0989 | 1.6679 | 1.3016 | 1.61 | 2.1456 | 2.3954 | 2.0941 | 1.6492 | 1.5814 | 1.9261 | 1.953 | 1.8796 | 1.3034 | 1.4634 | 2.5421 | 2.8548 | 3.4829 | 2.1255 | 1.4818 | -1.7486 | 1.6928
6 | 0.24882 | 0.078267 | -0.074598 | -0.31881 | -0.12157 | 0.18272 | 0.047651 | -0.0091344 | -0.0065599 | -0.1351 | 0.041113 | -0.0093091 | -0.46692 | -0.059358 | -0.50428 | 0.083222 | -0.24691 | -0.19011 | 0.22562 | 0.3255 | -0.073514
7 | 1.9099 | 0.45252 | 0.90794 | 0.69765 | 0.47546 | 0.51025 | 0.51778 | 0.68517 | 0.79184 | 0.62798 | 0.53019 | 0.64517 | 0.76143 | 0.83909 | 0.80933 | 0.78615 | 0.94558 | 1.394 | 0.42303 | 0.25626 | 0.73117
8 | -0.0063861 | -0.069169 | 0.01216 | 0.006524 | 0.034093 | 0.046818 | -0.032216 | 0.0053763 | -0.039823 | -0.0054032 | -0.097602 | -0.083993 | 0.0050057 | 0.041181 | 0.099512 | 0.040722 | 0.073802 | 0.10679 | -0.046966 | -0.027298 | 0.0062466
9 | 0.3072 | 0.32679 | 0.21365 | 0.23424 | 0.28831 | 0.2967 | 0.34219 | 0.34581 | 0.33885 | 0.32244 | 0.33037 | 0.33613 | 0.30441 | 0.25612 | 0.19212 | 0.20291 | 0.1495 | 0.070555 | 0.34429 | 0.31584 | 0.27273
10 | -0.001208 | 1.2453e-06 | -1.6545e-05 | 2.0339e-06 | 0.0012617 | -1.7132e-05 | -4.0752e-06 | 2.633e-06 | -0.00017145 | -0.0021329 | 1.0347e-05 | -2.0367e-05 | -6.9533e-05 | -0.00023043 | 4.4347e-06 | 5.9284e-05 | -0.0017944 | 0.00023891 | -0.00011485 | -5.5308e-05 | 0
11 | 0.55817 | 0.62857 | 0.58516 | 0.60945 | 0.62504 | 0.61277 | 0.63842 | 0.62591 | 0.61779 | 0.63424 | 0.63657 | 0.63594 | 0.61317 | 0.6055 | 0.62719 | 0.60641 | 0.58314 | 0.54842 | 0.62398 | 0.58686 | 0.61126
12 | 0.044842 | -0.062089 | 0.063167 | 0.0087832 | 0.13333 | 0.10142 | -0.032803 | -0.011375 | 0.043169 | -0.032988 | -0.039915 | -0.037407 | -0.052727 | 0.075751 | 0.23685 | 0.11185 | 0.15522 | 0.19811 | 0.020635 | 0.020908 | 0.047984 | ten times!
13 | 0.81676 | 0.73668 | 0.98724 | 0.80458 | 0.69071 | 0.68173 | 0.70595 | 0.77707 | 0.78069 | 0.78984 | 0.74842 | 0.77207 | 0.8633 | 0.91071 | 0.82642 | 0.78367 | 0.77351 | 1.1024 | 0.91323 | 0.64782 | 0.80615 | ten times!
14 | 0.063241 | -0.012232 | -0.019738 | -0.085082 | 0.038165 | 0.052344 | 0.0018522 | -0.013457 | -0.02349 | -0.0099061 | -0.047538 | -0.048904 | -0.10389 | -0.023391 | -0.03429 | -0.011403 | -0.0031791 | -0.0078985 | 0.00084051 | -0.04125 | -0.017842 | ten times!
15 | -0.12909 | 0.27385 | 0.32714 | 0.31467 | 0.30468 | 0.3304 | 0.27087 | 0.21318 | 0.23175 | 0.24971 | 0.27474 | 0.22954 | 0.2228 | 0.24226 | 0.30125 | 0.34686 | 0.39048 | 0.34709 | 0.13729 | 0.033055 | 0.25811 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -1.4973 | -1.9634 | -0.38215 | 0.69232 | -5.395 | -6.3368 | -2.479 | -0.43876 | -1.6609 | -0.95842 | -0.62534 | -1.0159 | 1.2531 | 2.9085 | 2.9026 | -2.9504 | -9.9628 | -5.5881 | 5.6961 | 3.4067 | -1.2249
3 | -2.7863 | 0.41301 | 0.097649 | -0.31818 | -0.20613 | -0.79697 | 0.23709 | -0.036762 | -0.092901 | 0.14234 | -0.5257 | 0.51553 | -0.46163 | 0.064542 | 0.66894 | -0.25352 | -0.60692 | -0.90372 | 0.7847 | 1.418 | -0.10663
4 | -0.82379 | -0.21303 | 0.077936 | -0.31977 | 1.0322 | 0.88523 | 0.20149 | 0.0048623 | -0.31344 | -0.27711 | -0.10207 | 0.17752 | 0.12787 | -0.41219 | -0.15957 | 1.7684 | 3.9577 | 2.1896 | -0.36084 | -0.91877 | 0.35253
5 | 1.627 | -0.23007 | 0.14111 | 0.97339 | -0.11524 | 0.37357 | 0.14169 | 0.07107 | 0.23731 | 0.26877 | 0.19451 | 0.33902 | 0.70223 | 0.23694 | 0.64297 | 0.36856 | -0.11665 | 0.20427 | -0.28557 | -0.60899 | 0.24678
6 | -0.17111 | -0.24038 | -0.27026 | -0.43094 | -0.33979 | -0.3492 | -0.35891 | -0.027298 | 0.058304 | -0.071961 | -0.22895 | -0.21118 | -0.30817 | -0.46067 | -0.23708 | -0.64552 | -0.56412 | 0.083977 | 0.31291 | 0.32261 | -0.22187
7 | 1.6123 | -0.046779 | 0.0024592 | -0.10796 | 0.15035 | 0.060181 | 0.030139 | 0.010008 | 0.0065054 | -0.065297 | -0.050715 | -0.081596 | -0.065978 | -0.025271 | -0.22179 | -0.016976 | 0.22754 | 0.14121 | 0.029319 | 0.022262 | 0.048995
8 | 0.032397 | -0.00096052 | 0.051367 | 0.092283 | -0.10458 | -0.088107 | 0.03373 | -0.024617 | -0.059089 | 0.0043329 | 0.0025454 | 0.034334 | 0.10137 | 0.23311 | 0.13356 | 0.094851 | -0.27074 | -0.26219 | -0.16977 | -0.14171 | -0.010316
9 | -0.080876 | 0.099368 | 0.014965 | 0.0043108 | 0.0076588 | -0.031563 | 0.042648 | 0.029705 | 0.012889 | 0.038756 | 0.022509 | 0.035894 | -0.074466 | 0.0086196 | 0.096626 | -0.0094396 | -0.020118 | -0.085673 | 0.0132 | -0.022095 | 0.0058004
10 | 0.00033414 | 2.0398e-06 | -9.7419e-05 | -2.0982e-06 | 0.00013462 | 8.0417e-06 | -2.3012e-06 | 2.5722e-06 | 0.0018004 | -0.0020642 | 7.8488e-06 | -5.7605e-07 | 0.00019968 | -2.1513e-05 | -3.8066e-05 | -1.9937e-05 | -0.00018357 | -0.00013573 | 0.00020583 | 3.4553e-05 | 0
11 | 0.22435 | 0.030263 | 0.046305 | 0.054228 | 0.065916 | 0.09384 | 0.069229 | 0.057313 | 0.057742 | 0.038696 | 0.048031 | 0.046046 | 0.078489 | 0.052965 | 0.014489 | 0.046678 | 0.065686 | 0.073482 | 0.021317 | 0.027805 | 0.057024 | ten times!
12 | -0.090775 | -0.023438 | -0.0043897 | -0.11248 | 0.058013 | -0.0071212 | -0.13594 | -0.002906 | 0.0501 | -0.0076441 | -0.08408 | -0.12263 | -0.21693 | -0.33589 | -0.13382 | -0.31191 | 0.12635 | 0.24358 | 0.015314 | 0.017373 | -0.054692 | ten times!
13 | -0.049871 | 0.097466 | 0.057756 | 0.07698 | 0.063768 | 0.07255 | 0.091198 | 0.088641 | 0.075189 | 0.10059 | 0.06343 | 0.083361 | 0.13091 | 0.077149 | 0.1228 | 0.076974 | 0.011164 | 0.10449 | 0.052231 | 0.04971 | 0.077792 | ten times!
14 | 0.03072 | -0.060346 | -0.026049 | -0.071139 | -0.06133 | -0.056706 | -0.053555 | -0.060636 | -0.081605 | -0.081394 | -0.088751 | -0.052592 | -0.058115 | -0.0078302 | -0.087499 | 0.023135 | 0.024579 | -0.081543 | -0.019832 | 0.069899 | -0.053321 | ten times!
15 | 0.058299 | -0.024925 | 0.001557 | -0.034958 | 0.022437 | 0.0020556 | 0.0097321 | 0.009128 | 0.00054988 | 0.0050428 | -0.0051652 | 0.00075786 | 0.007135 | 0.024021 | 0.055625 | 0.0296 | 0.027176 | -0.032808 | -0.0056932 | -0.021878 | 0.00583 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -25.6492 | -1.2042 | -0.3829 | -1.3051 | -1.7674 | -1.1297 | -1.8925 | -0.42961 | -0.28647 | 0.026891 | -0.44273 | -0.52323 | -1.127 | -0.54062 | -0.050189 | -0.17208 | -1.3693 | -1.4111 | -0.79333 | -23.1861 | -2.2869
3 | 37.0319 | 7.3653 | 6.3427 | 5.8156 | 4.8798 | 1.195 | 5.1601 | 6.2096 | 4.5933 | 3.8893 | 5.3002 | 6.4685 | 5.5205 | 4.6616 | 4.904 | 3.2662 | -0.47403 | -0.088796 | 7.7301 | 2.8217 | 5.5329
4 | -1.544 | -0.18065 | 0.097013 | 0.12342 | 0.77906 | 0.15002 | 0.16357 | 0.11988 | 0.14945 | 0.60833 | -0.17183 | 0.0050932 | -0.22716 | -0.18075 | -0.093282 | 0.053513 | -0.93276 | -0.7419 | -0.30002 | -2.1632 | -0.14671
5 | -3.0453 | 1.2541 | 1.1242 | 1.032 | 1.7059 | 2.7094 | 1.4188 | 0.79321 | 0.74397 | 1.4589 | 1.4274 | 1.5348 | 1.3438 | 1.7532 | 2.1697 | 3.2625 | 3.3463 | 3.218 | 1.1421 | -2.5379 | 1.464
6 | -0.2506 | 0.056617 | 0.038117 | -0.10334 | -0.40797 | -0.29019 | -0.068356 | -0.20142 | -0.038175 | -0.35102 | -0.17569 | -0.19364 | -0.066039 | -0.03952 | -0.28832 | -0.29296 | 0.16166 | 0.098885 | -0.083741 | -0.084452 | -0.13886
7 | 2.0026 | 0.64853 | 0.96598 | 1.0146 | 1.0004 | 0.68439 | 0.91933 | 0.75787 | 1.0176 | 1.0946 | 0.70184 | 0.72153 | 0.7185 | 0.69563 | 0.50393 | 0.093923 | 0.47548 | 0.52529 | 0.6427 | 0.45759 | 0.76125
8 | 0.003365 | 0.013521 | 0.01824 | -0.013002 | -0.069017 | -0.01656 | -0.036633 | -0.0022904 | 0.010568 | -0.030397 | 0.0042444 | 0.01112 | 0.015762 | 0.010928 | 0.034195 | -0.10843 | 0.074001 | 0.11136 | -0.0030664 | 0.035728 | 0.003512
9 | 0.31848 | 0.33568 | 0.24993 | 0.23549 | 0.1863 | 0.23992 | 0.30126 | 0.28887 | 0.31024 | 0.24612 | 0.28358 | 0.34116 | 0.32563 | 0.29802 | 0.30881 | 0.27333 | 0.22029 | 0.23275 | 0.30649 | 0.24651 | 0.27732
10 | -0.00032617 | 1.4675e-05 | 1.0373e-05 | 5.5573e-06 | 0.00014541 | 0.00036141 | 1.3103e-05 | -1.0207e-05 | -9.4106e-06 | 2.4716e-06 | 5.7975e-06 | 0.0026998 | -8.2966e-06 | 2.2672e-05 | 5.605e-05 | -8.7171e-05 | -0.00020265 | -3.3987e-05 | -0.0033778 | 0.00019721 | 0
11 | 0.56061 | 0.62977 | 0.59564 | 0.58947 | 0.5868 | 0.62248 | 0.60679 | 0.60634 | 0.57918 | 0.58932 | 0.61056 | 0.62699 | 0.6272 | 0.62184 | 0.65154 | 0.6556 | 0.62322 | 0.62038 | 0.63245 | 0.57824 | 0.61214
12 | 0.063781 | 0.014062 | 0.039775 | -0.0074736 | -0.22885 | -0.096558 | -0.00074165 | -0.02977 | 0.122 | -0.17598 | 0.033116 | 0.031524 | -0.024646 | 0.015351 | 0.046165 | 0.018319 | 0.13434 | 0.093492 | -0.028645 | -0.075499 | -9.168e-05 | ten times!
13 | 0.86089 | 0.80826 | 0.91921 | 0.96621 | 0.97127 | 0.77127 | 0.82584 | 0.87642 | 0.93682 | 0.98361 | 0.85822 | 0.81666 | 0.79898 | 0.77132 | 0.69445 | 0.57405 | 0.63765 | 0.61046 | 0.81228 | 0.53729 | 0.80384 | ten times!
14 | 0.07816 | -0.011416 | -0.02767 | -0.082229 | -0.18704 | -0.13952 | -0.027932 | -0.071511 | -0.047912 | -0.11307 | -0.085367 | -0.072817 | -0.06763 | -0.054711 | -0.061322 | -0.10928 | -0.010105 | 0.021997 | -0.038129 | -0.10936 | -0.06267 | ten times!
15 | -0.12675 | 0.26777 | 0.27025 | 0.29717 | 0.31565 | 0.35146 | 0.26789 | 0.27044 | 0.32452 | 0.2741 | 0.29694 | 0.24052 | 0.23413 | 0.25732 | 0.31557 | 0.40479 | 0.48834 | 0.44698 | 0.2471 | 0.15138 | 0.29127 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -0.60156 | -0.31711 | -1.8393 | 1.213 | 1.3028 | -3.6357 | -2.1957 | -0.22217 | 2.4426 | 1.7681 | -0.99001 | -0.61971 | -0.8209 | -0.45011 | 1.4274 | 0.93667 | -2.3079 | 1.4774 | 0.14249 | -0.97987 | -0.12505
3 | -2.5667 | -0.23758 | -0.24587 | -1.3425 | 0.22326 | -0.94058 | -0.10944 | 0.89328 | 0.19496 | 0.13323 | -0.4307 | 0.31455 | 0.40195 | 0.24671 | -0.20211 | 0.26658 | 0.19961 | 0.21852 | 0.053886 | 0.69277 | -0.077913
4 | 0.47008 | 0.35888 | 0.37813 | -0.38428 | -0.46411 | 0.99554 | 0.051772 | 0.010013 | -0.3591 | -0.56251 | -0.32345 | 0.058534 | 0.32568 | 0.54731 | -0.57012 | 0.029172 | 1.8737 | -0.052072 | -0.092265 | 0.86983 | 0.12667
5 | 1.5016 | 0.23344 | 0.45872 | 0.71668 | 0.23562 | 0.080305 | -0.092414 | -0.039065 | 0.096655 | 0.026876 | -0.12848 | 0.18586 | 0.11544 | 0.039181 | 0.11253 | 0.76391 | 0.30792 | 0.73262 | -0.080281 | -0.2883 | 0.22934
6 | -0.011144 | -0.065786 | 0.29215 | 0.04153 | -0.26082 | 0.034422 | 0.3478 | 0.12218 | -0.18163 | -0.18537 | -0.012914 | -0.0094457 | -0.24306 | -0.16126 | 0.036643 | -0.47979 | -0.87846 | -0.38592 | 0.076008 | -0.022331 | -0.097066
7 | 1.6845 | 0.0078367 | -0.0035121 | -0.10981 | -0.10582 | 0.11138 | 0.11599 | 0.033465 | 0.0088689 | 0.07344 | 0.14449 | 0.04249 | 0.11911 | 0.093183 | 0.04922 | -0.15865 | 0.0041472 | -0.059369 | 0.037079 | -0.015738 | 0.0723
8 | -0.024968 | 0.029482 | -0.11139 | 0.11147 | 0.10684 | -0.081021 | -0.09171 | -0.067131 | 0.052346 | 0.064728 | -0.0080589 | -0.021056 | 0.018558 | 0.022415 | 0.08795 | 0.16926 | 0.16919 | 0.14549 | 0.041635 | 0.0019627 | 0.033769
9 | -0.12702 | 0.031043 | -0.024205 | -0.054949 | 0.075104 | 0.026058 | 0.012074 | 0.05699 | 0.072918 | 0.032405 | 0.030435 | 0.042274 | 0.02775 | 0.01926 | 0.037882 | 0.0080666 | 0.017179 | 0.02975 | 0.038806 | -0.00368 | 0.019747
10 | 0.00071615 | 2.2338e-06 | 3.9273e-06 | -1.0311e-05 | 0.00013844 | -0.0031006 | 5.4005e-07 | 0.00010837 | 4.5422e-05 | -4.9877e-07 | 1.7094e-06 | 0.00023932 | 2.8333e-06 | -0.00019245 | -0.0011145 | 6.4752e-05 | 0.0031416 | 5.7641e-05 | -0.00020905 | -0.00011351 | 0
11 | 0.22187 | 0.053164 | 0.061114 | 0.06245 | 0.017225 | 0.071198 | 0.085973 | 0.052112 | 0.044077 | 0.055284 | 0.062501 | 0.054652 | 0.05728 | 0.058611 | 0.035663 | 0.067849 | 0.039546 | 0.033562 | 0.039175 | 0.039821 | 0.05748
12 | -0.073644 | 0.012731 | 0.18755 | -0.058844 | -0.21545 | 0.055917 | 0.20625 | 0.11997 | -0.16833 | -0.1293 | 0.08263 | 0.015162 | -0.066211 | -0.0085808 | -0.045031 | -0.2044 | -0.25133 | -0.15533 | 0.024536 | -0.11668 | -0.034419 | ten times!
13 | -0.04432 | 0.075282 | 0.13012 | 0.12464 | 0.17677 | 0.11496 | 0.13368 | 0.11311 | 0.097125 | 0.14533 | 0.077334 | 0.063255 | 0.092807 | 0.074722 | 0.10028 | 0.11392 | 0.13965 | 0.14164 | 0.09261 | 0.050975 | 0.10646 | ten times!
14 | 0.0016109 | -0.078654 | -0.08035 | -0.026996 | -0.081846 | -0.12598 | -0.11569 | -0.099771 | -0.1017 | -0.08323 | -0.11334 | -0.10308 | -0.10246 | -0.09048 | -0.096908 | -0.084013 | -0.01435 | -0.11108 | -0.093194 | -0.042777 | -0.08886 | ten times!
15 | 0.038405 | -0.11216 | -0.073172 | -0.10737 | -0.047289 | -0.05537 | -0.058444 | -0.032725 | -0.070549 | -0.087171 | -0.057503 | -0.031338 | -0.01605 | -0.054487 | -0.02033 | -0.012753 | -0.0011059 | -0.019096 | -0.046063 | -0.044113 | -0.047235 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -8.3549e-05 | 0.00025065 | 0.00041775 | 0.00025065 | 8.3549e-05 | 8.3549e-05 | 8.3549e-05 | -0.00041775 | -0.00041775 | -0.00025065 | -0.00025065 | -8.3549e-05 | 8.3549e-05 | 0.00025065 | 8.3549e-05 | 0.00025065 | 0.00025065 | -0.00058484 |  | 704.0588
|  | -0.00030633 | 2.7848e-05 | 0.00036203 | 0.00019494 | 0.00019494 | 2.7848e-05 | 0.00019494 | 2.7848e-05 | -0.00013924 | -0.00030633 | -0.00030633 | -0.00013924 | -0.00013924 | -0.00013924 | 2.7848e-05 | 0.00036203 | 0.00036203 | -0.00030633 |  | 704.098
Magnetic Length | -0.066045
mbh dB/B | -4062.9961
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.3897 | -4109.7836 | -0.83549 | 2.5065 | 4.1775 | 2.5065 | 0.83549 | 0.83549 | 0.83549 | -4.1775 | -4.1775 | -2.5065 | -2.5065 | -0.83549 | 0.83549 | 2.5065 | 0.83549 | 2.5065 | 2.5065 | -5.8484 | -4016.2085
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 2.8927 | 0.45172 | 0.23389 | 0.0033368 | 0.43835 | 0.35316 | 0.97291 | 0.67125 | 0.34504 | -0.19264 | -0.28616 | -0.29181 | -0.723 | -0.46127 | -0.77512 | -0.026178 | 0.69518 | 0.46442 | -1.8731 | -0.35903
db2 | 1.3813 | 22.4365 | -0.016996 | 0.19005 | -0.11678 | 0.018768 | 0.0091161 | -0.12709 | 0.044326 | -0.0072904 | 0.075467 | -0.06566 | -0.0070199 | -0.10212 | -0.017304 | 0.069928 | -0.17678 | 0.16597 | -0.13368 | 0.1971 | 27.0569
db3 | 4.5694 | -3.7139 | 0.042947 | -0.20711 | 0.10695 | -0.14689 | 0.14222 | -0.092027 | -0.026334 | -0.02089 | -0.05433 | -0.011341 | -0.054435 | -0.032301 | 0.091082 | -0.39796 | 0.42665 | -0.18289 | 0.34854 | 0.068115 | 8.7016
db4 | 0.13315 | 0.99981 | 0.012492 | 0.033217 | -0.035827 | -0.0031141 | -0.017043 | 0.0092683 | 0.025611 | 0.034805 | 0.017888 | 0.0032537 | 0.028149 | 0.014299 | -0.0081285 | 0.043081 | -0.089513 | -0.034383 | -0.025039 | -0.0090153 | 0.24181
db5 | 0.03136 | -1.7049 | 0.005476 | -0.0010255 | -0.027156 | 0.029738 | -0.01803 | 0.0084554 | -0.020581 | -0.004662 | -0.0043407 | 0.015653 | -0.01295 | -0.014396 | 0.0072695 | 0.061485 | -0.020314 | 0.0148 | -0.034747 | 0.015324 | -0.028943
db6 | -0.0028641 | 0.24882 | -0.002641 | -0.026287 | 0.027861 | -0.011057 | 0.013186 | 0.0048707 | -0.0087792 | -0.018809 | -0.012649 | 0.016592 | -0.015988 | -0.0021264 | 0.018825 | -0.011198 | 0.0564 | -0.031445 | 0.0046626 | -0.0014174 | 0.00011448
db7 | -0.013675 | -0.48307 | 0.0080252 | 0.0062767 | -0.008534 | 5.6505e-05 | -0.0069196 | -0.00012432 | 0.00042344 | 0.00054088 | -0.006088 | 0.0025105 | 0.0088082 | 0.0079304 | -0.0078162 | 0.023801 | -0.017984 | 0.025402 | -0.02449 | -0.011819 | -0.0077279
db8 | 0.00051043 | -0.0063861 | 0.00092212 | 0.0049982 | -0.011209 | 0.0083626 | -0.0027185 | -0.0015717 | 0.00046393 | -0.00028398 | -0.00051817 | -0.0030202 | 0.0059157 | -0.0028863 | -0.010692 | -0.0081718 | -0.0065124 | 0.023514 | 0.00069255 | 0.0027154 | 0.0070459
db9 | 0.01342 | 0.022202 | -0.0022939 | -0.0023582 | 0.0060308 | -0.0034636 | 0.004984 | -0.0036873 | -0.00078165 | -0.001337 | -0.0039178 | -0.0023618 | -0.0022313 | -0.0049135 | 0.0014047 | -0.00916 | 0.017336 | -0.009266 | 0.0061473 | 0.0098694 | 0.030699
db10 | -0.00014769 | -0.001208 | 0.00015372 | 0.00014994 | 0.00016233 | 0.0013659 | 0.00014028 | 0.00017505 | 0.00015109 | -2.2914e-05 | -0.001986 | 0.00016654 | 0.00011481 | 0.00011071 | -5.4983e-05 | 0.00021517 | 0.00045116 | -0.0016614 | 0.00030344 | 6.5202e-05 | -5.5308e-05
db11 | 0.00013021 | -0.0019236 | 0.0022561 | 0.00086049 | -1.9612e-05 | -0.0031173 | -0.004206 | 0.00026272 | -0.0021463 | -1.4924e-05 | 0.0015645 | -7.6492e-05 | -0.00056529 | -0.0029743 | 0.00067494 | 0.0036908 | 0.003328 | -0.0045788 | -0.00020313 | 0.0052647 | 0.017016
db12 | 0.0008022 | 0.0044842 | 0.00011589 | -0.00060089 | -0.00013444 | -0.0006199 | 0.00086241 | 0.00060459 | -4.4882e-05 | -0.00057225 | -0.0019148 | 0.00024994 | 0.00061694 | 0.00041827 | 0.0018398 | 0.0012232 | 0.0024897 | -0.0022407 | -0.0029175 | 0.00062459 | 0.0020908
db13 | -0.00017911 | 0.0070227 | -0.00072752 | 0.0026905 | -0.0004129 | -0.0011471 | -0.0013205 | -0.00019299 | -0.0013345 | 0.00035925 | 0.00079328 | 9.0841e-06 | 0.00018444 | 0.0018478 | -0.0013417 | -0.00075223 | 0.00048619 | 0.002205 | -0.00059657 | -0.00074949 | -0.0020952
db14 | -0.0059312 | 0.0063241 | -0.0003251 | 0.00046441 | -0.00064015 | -0.00020702 | -0.0010782 | -8.3101e-05 | 0.00089015 | 0.0013977 | 0.0014825 | 2.0399e-05 | 0.0010197 | 0.0004249 | 0.0010602 | -0.00032301 | -0.0024571 | -0.0039134 | 0.00016344 | 0.0021038 | -0.004125
db15 | 0.0013211 | -0.012909 | -0.00076871 | 0.0054859 | -0.00066749 | 0.00076511 | -0.00052626 | -0.00074797 | -0.0017012 | -0.00071385 | -0.00011737 | 0.0010928 | -0.0026398 | -0.0006967 | 0.0013532 | -0.0033595 | 0.005646 | -0.0012716 | 0.003649 | -0.0047815 | 0.0014782
da2 | 0.074918 | -1.4973 | -0.23362 | -0.02148 | -0.018046 | -0.14754 | 0.28521 | -0.099969 | 0.080767 | -0.050378 | -0.01573 | -0.070751 | 0.0049538 | 0.12479 | -0.018406 | 0.03285 | -0.32799 | -0.15214 | 0.19412 | 0.43335 | -0.45407
da3 | 0.087377 | -2.7863 | 0.086054 | 0.0022048 | 0.037545 | -0.064463 | -0.01727 | 0.12543 | 0.065146 | 0.077443 | 0.073787 | -0.0029537 | 0.036153 | -0.0074153 | 0.055604 | -0.059777 | -0.17923 | -0.16032 | -0.096894 | 0.028948 | 0.052604
da4 | 0.011998 | -0.82379 | 0.03112 | -0.011479 | -0.021495 | 0.058316 | -0.086706 | 0.04047 | -0.011346 | -0.015796 | 0.0020149 | 0.015559 | 0.0039156 | 0.0080322 | 0.0011596 | 0.014278 | 0.070091 | 0.06541 | -0.12884 | -0.0347 | -0.14034
da5 | 0.062597 | 1.627 | 0.029908 | -0.02546 | 0.00026454 | -0.03207 | 0.028718 | 0.011575 | 0.0012467 | -0.011298 | 0.0016916 | 0.0062685 | -0.022744 | -0.016266 | -0.031528 | 0.0098247 | 0.073668 | 0.014539 | -0.001242 | -0.037095 | -0.17877
da6 | -0.0060821 | -0.17111 | 0.0070603 | 0.0064035 | -0.027521 | 0.0049555 | 0.0061735 | 0.0022177 | 0.0089751 | 0.0083576 | -0.0042248 | -0.00925 | 0.0042746 | -0.014643 | 0.015345 | -0.017386 | -0.010311 | 0.0072734 | -0.0059926 | 0.018293 | 0.027574
da7 | 0.030056 | 1.6123 | -0.028528 | 0.017907 | -0.0034877 | -0.00045638 | -0.013462 | -0.0071464 | 0.0018307 | 0.0060409 | -0.0043518 | -0.015938 | 0.0013771 | -0.0048841 | 0.003966 | 0.01607 | 0.0075448 | 0.024928 | 0.016076 | -0.017486 | 0.018657
da8 | -0.0012389 | 0.032397 | -0.0067288 | 0.0010284 | -0.0021743 | -0.0020672 | 0.01157 | -0.004254 | 0.0012943 | -0.00085066 | -0.00080644 | -0.0075153 | -0.0028891 | 0.0039175 | 0.0008496 | 0.0087108 | -0.0033676 | -0.012171 | 0.013938 | 0.0015159 | -0.0048715
da9 | 0.016415 | -0.080876 | 0.001163 | -0.011626 | 0.0057618 | 0.00049444 | 0.0028809 | 0.0057868 | 0.0050304 | 0.0049492 | 0.0035489 | 0.0066244 | -0.0002181 | 0.0015402 | 0.0036883 | -0.00044075 | -0.0094834 | -0.0056747 | -0.017883 | 0.0038567 | -0.022095
da10 | 0.00012539 | 0.00033414 | 0.0013415 | -0.00026091 | -0.0015884 | 0.0018347 | -0.00012716 | -0.0010595 | -0.00011577 | 0.0016751 | -0.0021905 | -0.00013119 | -0.00010593 | 1.4901e-05 | -0.00018499 | -0.00019013 | 0.001493 | -0.00024634 | -0.00023275 | 7.4441e-05 | 3.4553e-05
da11 | 0.0042417 | 0.22435 | -0.0027831 | -0.0026311 | 0.00054563 | 0.0010459 | -0.00018582 | 0.0040214 | 0.0041673 | 0.0022655 | -0.0021637 | -0.0028303 | -0.0021856 | -0.00098742 | -0.0034865 | 0.003942 | 0.011406 | 0.00092853 | -0.0085782 | -0.0024907 | 0.027805
da12 | -0.00026183 | -0.0090775 | -0.00080772 | 0.0011361 | -0.0017869 | 0.0014555 | -0.00049698 | -0.0017191 | -0.00059396 | 0.00038051 | -0.0018119 | -0.0011298 | -0.00079981 | -0.00054112 | 0.00010668 | 0.0012025 | 0.00013962 | 0.0037232 | 0.0018007 | -0.0002576 | 0.0017373
da13 | 0.008037 | -0.0042807 | -0.00045085 | -0.0010332 | 0.00089094 | -6.9227e-05 | -0.0014664 | -6.8457e-05 | 4.7238e-05 | 0.00085129 | -0.00019579 | -0.00057948 | -0.00068019 | 0.0011793 | -0.00039176 | 0.0012258 | -0.0010325 | -0.00074965 | 0.0019927 | 0.00053025 | 0.003575
da14 | -0.0050067 | 0.003072 | 0.00026114 | 0.00019983 | -0.00021524 | 0.00073766 | -0.00023092 | -0.00093266 | 0.00044494 | 0.00015297 | -0.00033466 | 0.00021209 | -0.00027088 | -0.00026886 | -0.00021185 | -0.0015517 | 0.00080584 | -5.9149e-05 | -7.6199e-05 | 0.0013377 | 0.0069899
da15 | -0.0025409 | 0.0058299 | 0.00013802 | -0.0045322 | 0.0016052 | 0.0026666 | 0.00012038 | -0.00044847 | 0.0018418 | 0.00076005 | -0.00087362 | -0.0023073 | 0.0013626 | 0.00034887 | -0.00077073 | 0.0015191 | -0.0019122 | 0.002707 | -0.0031265 | 0.00090141 | -0.0021878
Check Computation for Aperture2
Magnetic Length | -0.069641
mbh dB/B | -4077.5293
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.375 | -4151.8839 | -3.0633 | 0.27848 | 3.6203 | 1.9494 | 1.9494 | 0.27848 | 1.9494 | 0.27848 | -1.3924 | -3.0633 | -3.0633 | -1.3924 | -1.3924 | -1.3924 | 0.27848 | 3.6203 | 3.6203 | -3.0633 | -4003.1747
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 3.281 | 0.067928 | -0.28062 | -0.26091 | -0.015124 | 0.40172 | 0.91886 | 0.78536 | 0.48749 | -0.13772 | -0.44021 | -0.45311 | -0.48587 | -0.19615 | -1.0406 | -0.10029 | 0.41378 | 0.28798 | 0.047472 | 1.9724
db2 | 1.172 | 24.4397 | 0.13065 | -0.17008 | 0.026457 | 0.060698 | -0.14286 | -0.093989 | -0.12075 | -0.028262 | 0.12535 | -0.075359 | 0.10252 | 0.013197 | -0.005561 | 0.060435 | 0.17473 | -0.01152 | -0.11298 | 0.067313 | 23.1861
db3 | 4.5285 | -3.5257 | -0.0059097 | -0.13646 | 0.051241 | -0.31946 | 0.13804 | -0.12246 | -0.0102 | -0.084109 | 0.0017438 | 0.19616 | -0.012967 | -0.072602 | 0.21415 | -0.22998 | 0.24646 | -0.19152 | 0.22913 | 0.10874 | 8.2777
db4 | 0.13692 | 1.544 | 0.038107 | 0.00097102 | -0.018535 | -0.06977 | 0.083934 | -0.029459 | 0.029554 | -0.024601 | -0.026632 | 0.0029082 | -0.00079312 | -0.0049185 | 0.025273 | -0.058488 | 0.018001 | 0.077384 | -0.068941 | 0.026007 | 2.1632
db5 | 0.03878 | -1.9637 | -0.013981 | -0.024867 | -0.013538 | 0.044675 | -0.027331 | -0.0043913 | 0.001506 | 0.012629 | 0.048896 | -0.0066865 | 0.014671 | 0.01562 | -0.0051823 | 0.084173 | -0.047387 | 0.047069 | -0.092931 | -0.032945 | -0.065485
db6 | -0.015346 | 0.2506 | -0.010634 | 0.014149 | -0.0099746 | 0.015487 | -0.0058759 | 0.014533 | 0.018299 | -0.0014992 | -0.0036213 | 0.013985 | -4.829e-05 | -0.016369 | 0.0037764 | 0.029178 | -0.042372 | -0.041354 | 0.010516 | 0.011822 | 0.032702
db7 | -0.013757 | -0.51707 | 0.0089233 | 0.010568 | -0.0081736 | 0.01179 | -0.0098424 | 0.0021095 | -0.02049 | 0.020603 | -0.014632 | -0.012887 | 0.0043616 | -0.001122 | -0.0076101 | -0.025234 | 0.012387 | 0.020822 | 0.0076071 | 0.00082035 | 0.00045351
db8 | 0.00033548 | -0.003365 | -0.0016888 | -0.0079931 | 0.0031577 | 0.016918 | -0.023202 | -0.0056074 | -0.0060284 | 0.00095285 | 0.0052493 | -0.0079055 | 0.001122 | -0.00031162 | -0.0059497 | 0.0097804 | 0.010196 | -0.00095313 | 0.011225 | 0.0010386 | -0.0029041
db9 | 0.014149 | 0.0010535 | -0.00064177 | 0.00082317 | 0.010648 | -0.0092607 | 0.0086057 | -0.00086704 | -0.0058738 | -0.0045589 | 0.00082878 | 0.0033661 | 0.0059774 | 0.0010916 | 0.00072322 | -0.0025716 | 0.0038358 | -0.0061714 | -0.0037893 | -0.0021651 | 0.025699
db10 | -6.6866e-06 | 0.00032617 | -8.7379e-06 | 1.9891e-05 | -6.2071e-06 | -0.00013344 | -0.00033983 | -1.6854e-05 | 1.3661e-05 | -2.3513e-06 | 2.2556e-05 | 7.5422e-06 | -0.0027119 | 9.1169e-06 | -0.00037164 | -5.4047e-05 | -3.3428e-05 | 0.00023229 | -1.0002e-05 | 0.0033834 | -0.00019721
db11 | 0.0014172 | -0.00081091 | 0.00038135 | -0.0017274 | 0.0022954 | 0.001337 | -0.0017012 | -0.0045226 | 0.002506 | -0.0026276 | 0.0014349 | -0.0017951 | 0.00077153 | 0.00021665 | -0.00085671 | 0.0046846 | -0.0033366 | 0.0010187 | 0.00011178 | 0.0018093 | 0.013723
db12 | 0.00013779 | -0.0063781 | -0.00011771 | 0.00059281 | -0.0015035 | 8.9822e-05 | -0.00042205 | 0.0027334 | -0.00051402 | -0.00085836 | -0.0010553 | 0.0015906 | 0.00024905 | -0.00017745 | 0.00094075 | 0.00060831 | -0.00074435 | -0.002889 | -0.0004368 | 0.0019136 | 0.0075499
db13 | 4.2362e-05 | 9.7369e-05 | -0.0010907 | -0.0015011 | 5.2906e-06 | 0.0037871 | 0.0010671 | -0.00085396 | 0.00034043 | 0.00019504 | 0.0014551 | -0.00010909 | 0.0014133 | 0.00056607 | -0.0005201 | -0.0016578 | 0.0016927 | -0.001482 | -0.0031542 | -0.00015324 | -0.0013857
db14 | 0.012238 | -0.007816 | -0.0017438 | -0.0020771 | 0.00049664 | 0.00076663 | 0.0012033 | -0.002091 | -0.0016164 | 0.00097796 | -0.00098306 | -0.00099804 | -0.0019832 | -0.0028462 | -0.0018584 | -0.00021279 | 0.0054244 | 0.0058372 | 0.0048388 | -0.0031351 | 0.010936
db15 | 0.0026687 | -0.012675 | 8.6019e-06 | 0.0016889 | -0.00069552 | 0.0018586 | -0.00064278 | -4.0652e-05 | -0.0023335 | -0.0007258 | -0.0020877 | 0.0021468 | -0.0034164 | 0.00012659 | -0.0023339 | -0.00086537 | 0.0027111 | 0.0023761 | 0.0014562 | 0.00076871 | 0.0053782
da2 | -0.16593 | -0.60156 | 0.033498 | -0.0065558 | 0.13116 | -0.11397 | -0.0049354 | -0.0010335 | 0.17311 | 0.029506 | -0.036894 | -0.068137 | -0.032616 | -0.11188 | 0.088792 | 0.052326 | -0.22798 | 0.040474 | 0.13288 | -0.077748 | -1.1232
da3 | 0.10979 | -2.5667 | 0.054815 | 0.061214 | 0.085173 | 0.051524 | -0.13144 | 0.07749 | -0.0043942 | -0.024437 | -0.036893 | 0.011597 | 0.047421 | 0.00060914 | -0.024124 | 0.013751 | -0.038185 | -0.12759 | -0.095671 | 0.07915 | 0.12719
da4 | -0.018637 | 0.47008 | -0.035369 | 0.028751 | -0.018124 | 0.069062 | -0.020229 | 0.023582 | -0.00076949 | -0.009334 | 0.016259 | 0.010947 | 0.00726 | 0.061807 | -0.078575 | -0.028856 | 0.094097 | -0.0044054 | -0.093668 | -0.022434 | -0.3349
da5 | 0.059163 | 1.5016 | 0.011811 | -0.01402 | -0.046878 | 0.0069703 | 0.029795 | -0.0030132 | -0.013335 | -0.039923 | -0.009689 | -0.018575 | -0.0025475 | -0.0088547 | -0.0026822 | 0.0020919 | 0.053393 | 0.054688 | 0.0193 | -0.018531 | -0.21242
da6 | -0.0057671 | -0.011144 | 0.028396 | 0.018435 | -0.0083604 | -0.016224 | 0.0014537 | 0.0075178 | -0.032192 | 0.0011906 | 0.0036635 | -0.016968 | 0.0030845 | -0.0064811 | 0.010832 | -0.0053023 | -0.042423 | 0.00097289 | 0.050499 | 0.0019058 | -0.034122
da7 | 0.035584 | 1.6845 | -0.014131 | 0.0093804 | 0.014804 | 0.0097972 | 0.013052 | 0.01063 | 0.0066658 | 0.018175 | 0.021779 | -0.010648 | -0.00048062 | -0.0027499 | -0.0077141 | -0.008823 | -0.038037 | -0.0067693 | -0.011472 | -0.0034602 | 0.012739
da8 | -0.00050104 | -0.024968 | -0.0035502 | -0.0025269 | 0.012052 | -0.011782 | -0.0037908 | -0.0085099 | 0.0053654 | 0.0013732 | -0.0025869 | -0.00068183 | 0.0025933 | 0.0014562 | 0.0045194 | 0.0051747 | -0.0034962 | 0.0072677 | -0.0032884 | 0.00041171 | -0.013987
da9 | 0.021905 | -0.12702 | 0.00063932 | -0.0017545 | 0.00020534 | 0.003214 | -0.0075697 | 0.003486 | 0.0053534 | 0.0055178 | -0.0030043 | -0.0011522 | 0.0020706 | -0.0013975 | 0.00017731 | 0.0021812 | 0.0016306 | -0.010431 | 0.00079671 | 3.6471e-05 | -0.00368
da10 | 3.3724e-05 | 0.00071615 | -3.1663e-05 | 0.0010393 | -7.1864e-05 | -7.5483e-05 | -0.0030937 | -1.4878e-05 | 7.1324e-05 | 0.00058206 | -5.4544e-05 | -2.3325e-05 | 0.00025458 | -3.2466e-05 | -0.00021801 | -0.0011472 | -5.8565e-05 | 0.0031021 | 1.5087e-05 | -0.00024281 | -0.00011351
da11 | 0.052857 | 0.22187 | 0.0017967 | 0.00085029 | -0.0020953 | -0.0027409 | 0.010341 | 0.0048322 | 0.0018992 | 0.0024549 | -0.0023866 | -0.0060263 | -0.0042606 | -0.0025819 | -0.002493 | -0.0040937 | 0.0026089 | -0.00056678 | 0.0032679 | -0.00080602 | 0.039821
da12 | -0.00070138 | -0.0073644 | 0.0012137 | 0.0014251 | -0.0018042 | -0.00058924 | -0.0013027 | 0.00096179 | -0.001036 | 0.0021731 | -0.0013849 | 0.00061879 | 0.00051336 | -0.00080657 | 0.00021174 | -0.0019754 | -0.00044535 | 0.0010513 | 0.0011863 | -1.0877e-05 | -0.011668
da13 | 0.011151 | 0.0016421 | -0.0011423 | 0.0023023 | 0.00020273 | 0.0028137 | 0.00099005 | 0.0025411 | 0.0016266 | 0.0025884 | 0.00091643 | -0.00053198 | 0.00016802 | 0.0015244 | -0.0014812 | -0.00066755 | -0.0022409 | -0.0045497 | -0.0057081 | 0.00064786 | 0.0029536
da14 | -0.0089062 | 0.00016109 | 0.0009117 | 0.0020426 | 0.0028317 | 0.00034973 | -0.00030593 | -0.00037443 | 0.00024311 | -0.00075529 | 0.0011552 | -0.00045013 | 0.00070099 | 0.0006304 | -3.2866e-05 | 0.00056966 | 8.3164e-05 | -0.0019179 | -0.0046986 | -0.00098313 | -0.0042777
da15 | -0.0099391 | 0.0038405 | -0.0047421 | -0.0010762 | 0.0013999 | -0.00045446 | -0.00075315 | 0.001492 | 0.0020336 | -0.0025938 | -0.0031127 | -0.0018734 | 0.004056 | 0.0011086 | -0.0013699 | 0.0014235 | 0.0011835 | 0.0038448 | -2.8775e-05 | -0.00053735 | -0.0044113
Field Colinearity | -0.11768
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.8241 | 0.38305 | 0.16522 | -0.065335 | 0.36968 | 0.28449 | 0.90424 | 0.60258 | 0.27637 | -0.26131 | -0.35483 | -0.36048 | -0.79167 | -0.52994 | -0.84379 | -0.09485 | 0.62651 | 0.39575 | -1.9418 | -0.42771
Twist (m2 mrad) | 21.1805 | 5.4584 | 2.2305 | -0.83302 | 4.4362 | 3.2005 | 9.4945 | 5.8751 | 2.4874 | -2.1558 | -2.6612 | -2.4332 | -4.75 | -2.7822 | -3.7971 | -0.35569 | 1.8795 | 0.89043 | -2.9127 | -0.16039
Twist Integral (I0) (m2 rad) | 0.025719
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.1103 | -0.10283 | -0.45137 | -0.43167 | -0.18588 | 0.23097 | 0.74811 | 0.6146 | 0.31673 | -0.30848 | -0.61097 | -0.62387 | -0.65662 | -0.36691 | -1.2114 | -0.27104 | 0.24302 | 0.11722 | -0.12329 | 1.8016
Twist (m2 mrad) | 23.3271 | -1.4653 | -6.0935 | -5.5037 | -2.2306 | 2.5984 | 7.8551 | 5.9924 | 2.8506 | -2.5449 | -4.5822 | -4.2111 | -3.9397 | -1.9263 | -5.4511 | -1.0164 | 0.72906 | 0.26375 | -0.18493 | 0.67561
Twist Integral (I0) (m2 rad) | 0.0038566

## Work sheet diff
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 3.5053 | 7.4025 | 1.3813 | 0.1111 | -1.2249 | 3.7441 | 0.074918 | 0.18085 | -2.2869 | 7.3528 | -1.172 | 0.10253 | -0.12505 | 1.6297 | -0.16593 | 0.099925 |  | 0
3 | 5.5896 | 7.5983 | 4.5694 | 0.19142 | -0.10663 | 0.8185 | 0.087377 | 0.084289 | 5.5329 | 7.6139 | 4.5285 | 0.16609 | -0.077913 | 0.80052 | 0.10979 | 0.067837 |  | 0
4 | 0.26252 | 0.3327 | 0.13315 | 0.032493 | 0.35253 | 1.1596 | 0.011998 | 0.049719 | -0.14671 | 0.72535 | -0.13692 | 0.043791 | 0.12667 | 0.60872 | -0.018637 | 0.047056 |  | 0
5 | 1.6928 | 1.4654 | 0.03136 | 0.023061 | 0.24678 | 0.46921 | 0.062597 | 0.027213 | 1.464 | 1.6116 | 0.03878 | 0.040664 | 0.22934 | 0.3958 | 0.059163 | 0.026946 |  | 0
6 | -0.073514 | 0.21194 | -0.0028641 | 0.021135 | -0.22187 | 0.27253 | -0.0060821 | 0.012053 | -0.13886 | 0.18176 | 0.015346 | 0.019302 | -0.097066 | 0.28251 | -0.0057671 | 0.020904 |  | 0
7 | 0.73117 | 0.35702 | -0.013675 | 0.012706 | 0.048995 | 0.37002 | 0.030056 | 0.013793 | 0.76125 | 0.37443 | -0.013757 | 0.013468 | 0.0723 | 0.38123 | 0.035584 | 0.014613 |  | 0
8 | 0.0062466 | 0.053637 | 0.00051043 | 0.0079283 | -0.010316 | 0.13128 | -0.0012389 | 0.0065221 | 0.003512 | 0.041728 | -0.00033548 | 0.0091535 | 0.033769 | 0.084872 | -0.00050104 | 0.0057382 |  | 0
9 | 0.27273 | 0.077106 | 0.01342 | 0.0066754 | 0.0058004 | 0.057831 | 0.016415 | 0.0068214 | 0.27732 | 0.041409 | 0.014149 | 0.0051919 | 0.019747 | 0.05279 | 0.021905 | 0.0040285 |  | 0
10 | 0 | 0.049938 | -0.00014769 | 0.00073239 | 0 | 0.053807 | 0.00012539 | 0.0010561 | 0 | 0.043646 | 6.6866e-06 | 0.0010609 | 0 | 0.041405 | 3.3724e-05 | 0.0011408 |  | 0
11 | 0.61126 | 0.026605 | 0.00013021 | 0.0026943 | 0.057024 | 0.041448 | 0.042417 | 0.0042817 | 0.61214 | 0.026377 | 0.0014172 | 0.0023108 | 0.05748 | 0.040475 | 0.046956 | 0.003941 |  | 0
12 | 0.0047984 | 0.0089111 | 0.0008022 | 0.0013613 | -0.0054692 | 0.015238 | -0.00026183 | 0.0014368 | -9.168e-06 | 0.00631 | -0.00013779 | 0.0013068 | -0.0034419 | 0.014297 | -0.00070138 | 0.0012223 |  | 0
13 | 0.080615 | 0.010509 | -0.00017911 | 0.0012225 | 0.0077792 | 0.0034043 | 0.0064463 | 0.00093038 | 0.080384 | 0.013393 | 4.2362e-05 | 0.0015775 | 0.010646 | 0.0040018 | 0.0083387 | 0.0023786 |  | 0
14 | -0.0017842 | 0.0034076 | -0.0059312 | 0.0014378 | -0.0053321 | 0.0045298 | -0.0014029 | 0.00064382 | -0.006267 | 0.0044607 | -0.012238 | 0.0027845 | -0.008886 | 0.0027782 | -0.003536 | 0.0015994 |  | 0
15 | 0.025811 | 0.011885 | 0.0013211 | 0.0027376 | 0.000583 | 0.0025227 | -0.0025409 | 0.0019848 | 0.029127 | 0.012244 | 0.0026687 | 0.0018107 | -0.0047235 | 0.0035939 | -0.0099391 | 0.0023101 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -1.9537 | -2.2905
3 | [12 empty] | 3 | 9 | -2.0634 | -2.2429
4 | [12 empty] | 4 | 16 | -3.214 | -3.0924
5 | [12 empty] | 5 | 25 | -3.6868 | -3.4082
6 | [12 empty] | 6 | 36 | -4.1376 | -3.9077
7 | [12 empty] | 7 | 49 | -4.3246 | -4.2667
8 | [12 empty] | 8 | 64 | -4.9349 | -4.9271
9 | [12 empty] | 9 | 81 | -4.9985 | -5.3875
10 | [12 empty] | sigma (mm) |  | 0.0056351 | 0.0054896
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05 | [5 empty] | Aperture 1 | Aperture 2
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | Magnetic length | -0.066045 | -0.069641
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 57.0588 | 107.6471 | 108.2353 | 108.3529 | 108.4706 | 108.3529 | 108.4706 | 108.4706 | 108.3529 | 108.4706 | 108.4706 | 108.4706 | 108.4706 | 108.4706 | 108.3529 | 108.3529 | 108.2353 | 108.2353 | 107.7647 | 19.0588
2 | 22.4365 | 1.3643 | 1.5713 | 1.2645 | 1.4 | 1.3904 | 1.2542 | 1.4256 | 1.374 | 1.4567 | 1.3156 | 1.3742 | 1.2791 | 1.364 | 1.4512 | 1.2045 | 1.5472 | 1.2476 | 1.5784 | 23.4246 | 3.5053
3 | -3.7139 | 4.6124 | 4.3623 | 4.6764 | 4.4225 | 4.7117 | 4.4774 | 4.5431 | 4.5485 | 4.5151 | 4.5581 | 4.515 | 4.5371 | 4.6605 | 4.1715 | 4.9961 | 4.3866 | 4.918 | 4.6376 | 8.7016 | 5.5896
4 | 0.37457 | 0.14564 | 0.16637 | 0.097322 | 0.13004 | 0.11611 | 0.14242 | 0.15876 | 0.16795 | 0.15104 | 0.1364 | 0.1613 | 0.14745 | 0.12502 | 0.17623 | 0.043636 | 0.098766 | 0.10811 | 0.12413 | 0.86441 | 0.26252
5 | -1.7049 | 0.036836 | 0.030334 | 0.0042037 | 0.061098 | 0.013329 | 0.039815 | 0.010779 | 0.026698 | 0.027019 | 0.047013 | 0.01841 | 0.016964 | 0.038629 | 0.092845 | 0.011046 | 0.04616 | -0.0033868 | 0.046684 | -0.028943 | 1.6928
6 | -0.24316 | -0.0055051 | -0.029151 | 0.024997 | -0.013922 | 0.010322 | 0.0020066 | -0.011643 | -0.021673 | -0.015513 | 0.013728 | -0.018852 | -0.0049905 | 0.015961 | -0.014062 | 0.053536 | -0.034309 | 0.0017984 | -0.0042815 | 0.00011448 | -0.073514
7 | -0.48307 | -0.0056495 | -0.0073979 | -0.022209 | -0.013618 | -0.020594 | -0.013799 | -0.013251 | -0.013134 | -0.019763 | -0.011164 | -0.0048665 | -0.0057442 | -0.021491 | 0.010127 | -0.031659 | 0.011727 | -0.038165 | -0.025494 | -0.0077279 | 0.73117
8 | -0.015723 | 0.0014326 | 0.0055086 | -0.010699 | 0.008873 | -0.0022081 | -0.0010613 | 0.00097437 | 0.00022646 | -7.7398e-06 | -0.0025098 | 0.0064262 | -0.0023759 | -0.010182 | -0.0076614 | -0.006002 | 0.024025 | 0.001203 | 0.0032258 | 0.0070459 | 0.0062466
9 | 0.022202 | 0.011126 | 0.011062 | 0.019451 | 0.0099566 | 0.018404 | 0.0097329 | 0.012639 | 0.012083 | 0.0095024 | 0.011058 | 0.011189 | 0.0085067 | 0.014825 | 0.0042602 | 0.030756 | 0.0041541 | 0.019567 | 0.02329 | 0.030699 | 0.27273
10 | -0.0069874 | 6.0312e-06 | 2.2459e-06 | 1.4641e-05 | 0.0012182 | -7.4146e-06 | 2.7355e-05 | 3.3922e-06 | -0.00017061 | -0.0021337 | 1.8849e-05 | -3.2887e-05 | -3.6981e-05 | -0.00020268 | 6.748e-05 | 0.00030347 | -0.0018091 | 0.00015575 | -8.2491e-05 | -0.00016456 | 0
11 | -0.0019236 | 0.0023863 | 0.00099071 | 0.0001106 | -0.0029871 | -0.0040758 | 0.00039293 | -0.0020161 | 0.00011529 | 0.0016947 | 5.3722e-05 | -0.00043507 | -0.0028441 | 0.00080516 | 0.003821 | 0.0034582 | -0.0044486 | -7.2914e-05 | 0.0053949 | 0.017016 | 0.61126
12 | 0.21739 | 0.0091809 | 0.0020131 | 0.0066776 | 0.0018229 | 0.016646 | 0.014068 | 0.0075732 | 0.0022995 | -0.011126 | 0.010521 | 0.014191 | 0.012205 | 0.02642 | 0.020254 | 0.032919 | -0.014385 | -0.021153 | 0.014268 | -0.017162 | 0.047984 | ten times!
13 | 0.070227 | -0.0090663 | 0.025114 | -0.0059202 | -0.013262 | -0.014996 | -0.003721 | -0.015136 | 0.0018014 | 0.0061416 | -0.0017003 | 5.3265e-05 | 0.016687 | -0.015208 | -0.0093134 | 0.0030708 | 0.020259 | -0.0077569 | -0.009286 | -0.020952 | 0.80615 | ten times!
14 | 0.22801 | -0.062563 | -0.054668 | -0.065714 | -0.061382 | -0.070094 | -0.060143 | -0.050411 | -0.045335 | -0.044487 | -0.059108 | -0.049115 | -0.055063 | -0.04871 | -0.062542 | -0.083883 | -0.098447 | -0.057678 | -0.038274 | -0.036702 | -0.017842 | ten times!
15 | -0.22852 | 0.0055243 | 0.06807 | 0.0065365 | 0.020862 | 0.0079488 | 0.0057317 | -0.0038006 | 0.0060729 | 0.012038 | 0.024139 | -0.013186 | 0.0062444 | 0.026743 | -0.020384 | 0.069671 | 0.00049552 | 0.049701 | -0.034604 | 0.014782 | 0.25811 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -2.7652 | -0.4523 | -0.07559 | 0.22575 | 0.087962 | -0.21072 | 0.5359 | 0.47998 | 0.30581 | -0.20783 | -0.59182 | -0.47079 | -0.30277 | -0.26071 | 0.75386 | 1.1286 | 0.19262 | -0.32532 | -0.36195 | 1.9965
2 | -1.2758 | -0.1587 | 0.053438 | 0.056872 | -0.072624 | 0.36013 | -0.025051 | 0.15568 | 0.02454 | 0.059188 | 0.0041676 | 0.079872 | 0.19971 | 0.056512 | 0.10777 | -0.25307 | -0.077225 | 0.26904 | 0.50827 | -0.45407 | -1.2249
3 | -1.6395 | 0.17343 | 0.089582 | 0.12492 | 0.022914 | 0.070107 | 0.21281 | 0.15252 | 0.16482 | 0.16116 | 0.084423 | 0.12353 | 0.079962 | 0.14298 | 0.0276 | -0.091852 | -0.072939 | -0.0095172 | 0.11633 | 0.052604 | -0.10663
4 | -1.7572 | 0.043118 | 0.00051851 | -0.009497 | 0.070313 | -0.074709 | 0.052468 | 0.00065135 | -0.0037986 | 0.014013 | 0.027557 | 0.015913 | 0.02003 | 0.013157 | 0.026275 | 0.082088 | 0.077407 | -0.11685 | -0.022702 | -0.14034 | 0.35253
5 | -1.5627 | 0.092505 | 0.037137 | 0.062862 | 0.030527 | 0.091315 | 0.074172 | 0.063844 | 0.051299 | 0.064289 | 0.068866 | 0.039853 | 0.046331 | 0.031069 | 0.072422 | 0.13627 | 0.077136 | 0.061355 | 0.025502 | -0.17877 | 0.24678
6 | -0.14755 | 0.00097824 | 0.00032139 | -0.033603 | -0.0011266 | 9.1424e-05 | -0.0038644 | 0.002893 | 0.0022755 | -0.010307 | -0.015332 | -0.0018075 | -0.020725 | 0.0092631 | -0.023468 | -0.016393 | 0.0011913 | -0.012075 | 0.012211 | 0.027574 | -0.22187
7 | -0.36329 | 0.001528 | 0.047963 | 0.026569 | 0.0296 | 0.016595 | 0.02291 | 0.031887 | 0.036097 | 0.025705 | 0.014118 | 0.031433 | 0.025172 | 0.034022 | 0.046127 | 0.037601 | 0.054985 | 0.046132 | 0.01257 | 0.018657 | 0.048995
8 | -0.12693 | -0.0079677 | -0.00021055 | -0.0034133 | -0.0033061 | 0.010331 | -0.0054929 | 5.5422e-05 | -0.0020896 | -0.0020454 | -0.0087542 | -0.0041281 | 0.0026785 | -0.00038932 | 0.0074718 | -0.0046066 | -0.01341 | 0.012699 | 0.00027695 | -0.0048715 | -0.010316
9 | -0.12109 | 0.017578 | 0.004789 | 0.022177 | 0.016909 | 0.019296 | 0.022201 | 0.021445 | 0.021364 | 0.019964 | 0.023039 | 0.016197 | 0.017955 | 0.020103 | 0.015974 | 0.0069314 | 0.01074 | -0.0014678 | 0.020271 | 0.014363 | 0.0058004
10 | -0.002275 | 0.0014669 | -0.00013552 | -0.001463 | 0.0019601 | -1.7679e-06 | -0.00093416 | 9.6156e-06 | 0.0018004 | -0.0020651 | -5.8064e-06 | 1.9456e-05 | 0.00014029 | -5.9598e-05 | -6.4741e-05 | 0.0016183 | -0.00012095 | -0.00010736 | 0.00019983 | 0.00022377 | 0
11 | -0.0079401 | 0.039634 | 0.039786 | 0.042963 | 0.043463 | 0.042231 | 0.046438 | 0.046584 | 0.044683 | 0.040253 | 0.039587 | 0.040231 | 0.04143 | 0.038931 | 0.046359 | 0.053823 | 0.043346 | 0.033839 | 0.039926 | 0.04523 | 0.057024 | ten times!
12 | -0.089811 | -0.010695 | 0.0087432 | -0.020487 | 0.011937 | -0.0075881 | -0.019809 | -0.0085579 | 0.0011869 | -0.020738 | -0.013916 | -0.010616 | -0.0080295 | -0.0015515 | 0.0094064 | -0.0012221 | 0.034614 | 0.015389 | -0.0051942 | -0.019222 | -0.054692 | ten times!
13 | -0.042807 | 0.059954 | 0.054131 | 0.073372 | 0.063771 | 0.049799 | 0.063778 | 0.064935 | 0.072976 | 0.062505 | 0.058668 | 0.057661 | 0.076256 | 0.060545 | 0.076721 | 0.054138 | 0.056966 | 0.08439 | 0.069765 | 0.03575 | 0.077792 | ten times!
14 | -0.054785 | -0.011417 | -0.01203 | -0.016181 | -0.0066519 | -0.016338 | -0.023355 | -0.0095792 | -0.012499 | -0.017375 | -0.011908 | -0.016737 | -0.016717 | -0.016147 | -0.029546 | -0.0059701 | -0.01462 | -0.014791 | -0.00065162 | 0.0017242 | -0.053321 | ten times!
15 | -0.40613 | -0.024029 | -0.070732 | -0.0093572 | 0.0012571 | -0.024205 | -0.029894 | -0.0069911 | -0.017809 | -0.034145 | -0.048483 | -0.011783 | -0.021921 | -0.033117 | -0.010219 | -0.044531 | 0.0016609 | -0.056674 | -0.016395 | 0.024612 | 0.00583 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 54.1176 | 107.6471 | 108.2353 | 108.3529 | 108.3529 | 108.3529 | 108.3529 | 108.3529 | 108.3529 | 108.3529 | 108.4706 | 108.4706 | 108.4706 | 108.3529 | 108.4706 | 108.3529 | 108.3529 | 108.3529 | 107.7647 | 18.9412
2 | -24.4397 | -1.3027 | -1.0019 | -1.1985 | -1.2327 | -1.0291 | -1.078 | -1.0513 | -1.1437 | -1.2973 | -1.0966 | -1.2745 | -1.1852 | -1.1664 | -1.2324 | -1.3467 | -1.1605 | -1.059 | -1.2393 | -22.8835 | -2.2869
3 | -3.5257 | 4.5226 | 4.392 | 4.5797 | 4.209 | 4.6665 | 4.406 | 4.5183 | 4.4444 | 4.5302 | 4.7247 | 4.5155 | 4.4559 | 4.7426 | 4.2985 | 4.775 | 4.337 | 4.7576 | 4.6372 | 8.2777 | 5.5329
4 | -1.7657 | -0.17503 | -0.1379 | -0.11839 | -0.067155 | -0.22086 | -0.10747 | -0.16648 | -0.11232 | -0.11029 | -0.13983 | -0.13613 | -0.13201 | -0.1622 | -0.078436 | -0.15493 | -0.21431 | -0.067984 | -0.16293 | -1.3806 | -0.14671
5 | -1.9637 | 0.0248 | 0.013914 | 0.025243 | 0.083455 | 0.01145 | 0.034389 | 0.040286 | 0.051409 | 0.087676 | 0.032094 | 0.053451 | 0.0544 | 0.033598 | 0.12295 | -0.0086071 | 0.085849 | -0.054151 | 0.005835 | -0.065485 | 1.464
6 | -0.31925 | 0.02598 | 0.0011962 | 0.02532 | -0.00014138 | 0.021222 | 0.0008122 | -0.0029538 | 0.016845 | 0.018967 | 0.0013603 | 0.015394 | 0.031715 | 0.011569 | -0.013833 | 0.057718 | 0.056699 | 0.0048293 | 0.0035232 | -0.032702 | -0.13886
7 | -0.51707 | -0.0048342 | -0.0031895 | -0.021931 | -0.001967 | -0.0236 | -0.011648 | -0.034247 | 0.0068451 | -0.02839 | -0.026645 | -0.0093959 | -0.014879 | -0.021368 | -0.038992 | -0.0013702 | 0.0070642 | -0.0061504 | -0.012937 | 0.00045351 | 0.76125
8 | -0.13054 | 0.0013533 | 0.0076576 | -0.0034931 | -0.017254 | 0.022866 | 0.0052719 | 0.005693 | -0.0012883 | -0.0055848 | 0.0075701 | -0.0014575 | -2.3853e-05 | 0.0056143 | -0.010116 | -0.010531 | 0.00061766 | -0.01156 | -0.0013741 | 0.0029041 | 0.003512
9 | 0.0010535 | 0.013508 | 0.014973 | 0.024797 | 0.0048887 | 0.022755 | 0.013282 | 0.0082756 | 0.0095905 | 0.014978 | 0.017515 | 0.020127 | 0.015241 | 0.014873 | 0.011578 | 0.017985 | 0.007978 | 0.01036 | 0.011984 | 0.025699 | 0.27732
10 | -0.00069915 | 1.5424e-05 | -1.3205e-05 | 1.2894e-05 | 0.00014012 | 0.00034651 | 2.354e-05 | -6.9748e-06 | 9.0379e-06 | -1.587e-05 | -8.5564e-07 | 0.0027186 | -2.4303e-06 | 0.00037833 | 6.0734e-05 | 4.0114e-05 | -0.0002256 | 1.6689e-05 | -0.0033767 | 0.00013052 | 0
11 | -0.00081091 | 0.0017985 | -0.00031022 | 0.0037125 | 0.0027541 | -0.00028406 | -0.0031054 | 0.0039231 | -0.0012104 | 0.0028521 | -0.00037792 | 0.0021887 | 0.0016338 | 0.00056045 | 0.0061017 | -0.0019195 | 0.0024358 | 0.0015289 | 0.0032265 | 0.013723 | 0.61214
12 | -0.061075 | -0.00020078 | -0.007306 | 0.013657 | -0.0022761 | 0.0028427 | -0.028712 | 0.0037624 | 0.0072057 | 0.0091748 | -0.017284 | -0.0038684 | 0.00039667 | -0.010785 | -0.0074609 | 0.0060657 | 0.027512 | 0.0029901 | -0.020514 | -0.026724 | -9.168e-05 | ten times!
13 | 0.00097369 | -0.010483 | -0.014588 | 0.00047653 | 0.038294 | 0.011095 | -0.008116 | 0.0038279 | 0.002374 | 0.014975 | -0.00066725 | 0.014556 | 0.0060844 | -0.0047774 | -0.016154 | 0.017351 | -0.014397 | -0.031118 | -0.0011088 | -0.013857 | 0.80384 | ten times!
14 | -0.00077885 | -0.10494 | -0.10161 | -0.12735 | -0.13005 | -0.13441 | -0.10147 | -0.10622 | -0.13216 | -0.11255 | -0.1124 | -0.10255 | -0.093917 | -0.10379 | -0.12025 | -0.17662 | -0.18075 | -0.17077 | -0.091028 | -0.10026 | -0.06267 | ten times!
15 | -0.13486 | 0.026773 | 0.043576 | 0.019732 | 0.045274 | 0.02026 | 0.026281 | 0.0033528 | 0.019429 | 0.0058106 | 0.048155 | -0.0074763 | 0.027953 | 0.0033482 | 0.018034 | 0.053798 | 0.050448 | 0.04125 | 0.034375 | 0.053782 | 0.29127 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -1.9516 | 5.9e-05 | 0.068822 | -0.045941 | 0.31103 | 0.80918 | 0.63676 | 0.49349 | 0.55124 | -0.12434 | -0.8418 | -0.60607 | -0.58642 | -0.36573 | -0.91036 | -0.10533 | 0.20769 | 0.37419 | -0.0277 | 2.0466
2 | 2.2118 | -0.13243 | -0.17249 | -0.034769 | -0.2799 | -0.17087 | -0.16696 | 0.007184 | -0.13642 | -0.20282 | -0.23407 | -0.19855 | -0.27781 | -0.077137 | -0.1136 | -0.39391 | -0.12546 | -0.033049 | -0.24368 | -1.1232 | -0.12505
3 | -2.207 | 0.16461 | 0.17101 | 0.19497 | 0.16132 | -0.021652 | 0.18728 | 0.1054 | 0.085356 | 0.072899 | 0.12139 | 0.15721 | 0.1104 | 0.085668 | 0.12354 | 0.071608 | -0.017801 | 0.014121 | 0.18894 | 0.12719 | -0.077913
4 | 0.80081 | -0.054006 | 0.010114 | -0.03676 | 0.050425 | -0.038865 | 0.0049454 | -0.019406 | -0.027971 | -0.0023774 | -0.0076897 | -0.011377 | 0.04317 | -0.097212 | -0.047493 | 0.07546 | -0.023042 | -0.11231 | -0.041071 | -0.3349 | 0.12667
5 | -1.7159 | 0.070974 | 0.045143 | 0.012286 | 0.066134 | 0.088959 | 0.05615 | 0.045828 | 0.01924 | 0.049474 | 0.040588 | 0.056616 | 0.050309 | 0.056481 | 0.061255 | 0.11256 | 0.11385 | 0.078463 | 0.040633 | -0.21242 | 0.22934
6 | 0.40523 | 0.022629 | 0.012668 | -0.014128 | -0.021991 | -0.0043134 | 0.0017507 | -0.03796 | -0.0045765 | -0.0021036 | -0.022735 | -0.0026826 | -0.012248 | 0.0050651 | -0.011069 | -0.048191 | -0.0047942 | 0.044732 | -0.0038613 | -0.034122 | -0.097066
7 | -0.41362 | 0.021453 | 0.044964 | 0.050388 | 0.045381 | 0.048635 | 0.046214 | 0.042249 | 0.053759 | 0.057362 | 0.024936 | 0.035103 | 0.032834 | 0.027869 | 0.02676 | -0.0024531 | 0.028814 | 0.024112 | 0.032123 | 0.012739 | 0.0723
8 | 0.044265 | -0.0040512 | -0.0030279 | 0.011551 | -0.012283 | -0.0042918 | -0.0090109 | 0.0048644 | 0.00087212 | -0.003088 | -0.0011829 | 0.0020922 | 0.00095519 | 0.0040184 | 0.0046736 | -0.0039973 | 0.0067666 | -0.0037895 | -8.933e-05 | -0.013987 | 0.033769
9 | -0.10343 | 0.022544 | 0.020151 | 0.02211 | 0.025119 | 0.014335 | 0.025391 | 0.027259 | 0.027423 | 0.018901 | 0.020753 | 0.023976 | 0.020508 | 0.022082 | 0.024086 | 0.023536 | 0.011475 | 0.022702 | 0.021942 | 0.02293 | 0.019747
10 | 0.00084092 | 2.0606e-06 | 0.0010731 | -3.814e-05 | -4.1759e-05 | -0.00306 | 1.8846e-05 | 0.00010505 | 0.00061579 | -2.082e-05 | 1.0399e-05 | 0.00028831 | 1.2584e-06 | -0.00018429 | -0.0011135 | -2.4841e-05 | 0.0031359 | 4.8811e-05 | -0.00020908 | -0.00030674 | 0
11 | 0.0087956 | 0.048753 | 0.047806 | 0.044861 | 0.044215 | 0.057297 | 0.051788 | 0.048855 | 0.049411 | 0.044569 | 0.04093 | 0.042695 | 0.044374 | 0.044463 | 0.042862 | 0.049565 | 0.046389 | 0.050224 | 0.04615 | 0.048901 | 0.05748
12 | -0.13112 | 0.0051237 | 0.0072377 | -0.025056 | -0.012906 | -0.020041 | 0.0026041 | -0.017374 | 0.014717 | -0.020863 | -0.00082587 | -0.0018802 | -0.015079 | -0.0048964 | -0.026768 | -0.011467 | 0.0034995 | 0.0048494 | -0.0071225 | -0.073176 | -0.034419 | ten times!
13 | 0.016421 | 0.071964 | 0.10641 | 0.085414 | 0.11152 | 0.093288 | 0.1088 | 0.099653 | 0.10927 | 0.092551 | 0.078067 | 0.085067 | 0.098631 | 0.068575 | 0.076712 | 0.060978 | 0.037891 | 0.026306 | 0.089866 | 0.029536 | 0.10646 | ten times!
14 | -0.10677 | -0.026243 | -0.014934 | -0.0070437 | -0.031863 | -0.03842 | -0.039105 | -0.032929 | -0.042913 | -0.023809 | -0.039862 | -0.02835 | -0.029056 | -0.035689 | -0.029664 | -0.034529 | -0.054539 | -0.082346 | -0.045192 | -0.042519 | -0.08886 | ten times!
15 | -0.098828 | -0.14681 | -0.11015 | -0.085392 | -0.10394 | -0.10692 | -0.084472 | -0.079056 | -0.12533 | -0.13052 | -0.11812 | -0.058832 | -0.088306 | -0.11309 | -0.085157 | -0.087556 | -0.060943 | -0.099679 | -0.10476 | -0.050176 | -0.047235 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
2.0447 | 4.7034 | 0.24249 | 1.9628 | 0.70953 | -0.68672 | -1.455 | -0.07387 | 0.45073 | 71.6261 | 0.70406 | -0.82053 | 4.6 | -0.02813 | 1.7445 | 0.73144 | -1.7076 | -0.15446 | -0.021795 | 0.10038 | 56.6003 | 0.7041 | 3.5053 | 5.5896 | 0.26252 | 1.6928 | 0.73117 | 0 | -1.2249 | -0.10663 | 0.35253 | 14361.717 | -2.2869 | 5.5329 | -0.14671 | 1.464 | 0.76125 | 0 | -0.12505 | -0.077913 | 0.12667 | 14359.409
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0