## Alarm sheet
File name | HCMBALA001-2000097_cm.xls
Component ID | HCMBALA001 |  |  |  | Serial Number | 2000097
Date of test Ap 1 | 2004-09-04 |  |  |  | Date of test Ap 2 | 2004-09-04
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 |  |  |  |  | b10
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | yellow alarm | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 |  |  |  |  | a10
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Magnet name
| HCMBALA001 | 2000097 |  | Additional lamin. (mm) |  | 0
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2004-09-04 | Analysis by | P. Hagen | Answer to MMS-LD |  | 2004-09-05 | Answer to manufacturer |  | 2004-09-06
1 | Fill in the yellow fields in Original Data | [10 empty] | English
2 | CTRL-g to load the measurement files .txt
3 | Save as … with the "official" name in Original Data C1
4 | Open the file of the collared coil measurement of the same magnet
5 | Go back to cold_mass file. Tools - Macro - Macros - Make difference - Run (to make the difference between collared coil and cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (to perform the analysis - results in Alarm sheet)
7 | Save the file
| Cold mass template - version 11.10.2002 | [7 empty] | Version : | cern
1 | Remplir les champs jaunes dans Original Data | [10 empty] | Francais
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec le nom "officiel" dans Original Data C1
4 | Ouvrir le fichier avec les mesures de la collared coil
5 | Revenir au ficher cold_mass. Tools - Macro - Macros - Make difference - Run (pour faire la difference entre collared coil et cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (pour faire l'analyse - resultats dans Alarm sheet)
7 | Sauver le fichier
1 | Riempire I campi gialli in Original Data | [10 empty] | Italiano
2 | CTRL-g per caricare I file di misura .txt
3 | Salvare il file con il nome "ufficiale" in Original Data C1
4 | Aprire il file delle misure della collared coil corrispondente
5 | Tornare sul file cold_mass. Tools - Macro - Macros - Make difference - Run (per fare la differenza tra collared coil e cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (per fare l'analisi - risultati in Alarm sheet)
7 | Salvare il file
1 | Die gelben Felder in Original Data ausfüllen | [10 empty] | Deutsch
2 | Ctrl-g um die Messdatei ( Datei .txt) zu laden.
3 | Die Datei mit der ”offiziellen” bezeichnung im Original Data C1 speichern mit “Save as…”.
4 | Öffnen “Collared Coils” Messdatei (gleiche Magnete versteht sich)
5 | Zurück zum cold mass Datei. Tools - Macro - Macros - Make difference - Run (Zeigt allen unterschieden in Messdatei zw. cc und cm)
6 | Tools - Macro - Macros - Collared coil analysis - Run (Um Maßdaten zu analysieren - Resultaten ins Alarm Sheet)
7 | Datei speichern
| Last modif on | by | Modifications:
| 2002-10-20 | E. Todesco | Cell for iron additional iron laminations - check on magnetic length includes additional iron laminations
| 2002-11-13 | E. Wildner | German version of instructions added
| 2003-10-01 | E. Wildner | Update of bounds: all heads tested on differences
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2004-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2004-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2004-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-03 | E. Todesco | Correction of integrals in worksheet, worksheet diff, MTF, and computation of dx dy from multipoles
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -
| 2004-07-02 | P. Hagen | Added Inclinometer data into separate sheet if present in ASCII files

## Original data
File name |  | HCMBALA001-02000097_cm.xls | [41 empty] | T
Component ID |  | HCMBALA001 |  |  | Serial Number |  |  | 2000097 | [11 empty] | Signs | [23 empty] | T
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1 | [23 empty] | T
Firm Name |  | Ansaldo | [9 empty] | File name for aperture 1 |  |  | 04-SEP-2004_2097_08-06cm1.txt |  |  |  | a odd | 1 | [23 empty] | T
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 34 |  |  |  | File name for aperture 2 |  |  | 04-SEP-2004_2097_08-06cm2.txt |  |  |  | a even | 1 | [23 empty] | T
Test Operator |  | Federico_Galeazzi |  |  | Test Controller |  |  | P. Hagen | [35 empty] | T
T
Aperture 1 | [11 empty] | Aperture 2 | [31 empty] | T
Date of test |  | 2004-09-04 |  |  | Time of test |  |  | 08:06 - 09:56 |  |  |  | Date of test |  | 2004-09-04 |  |  | Time of test |  |  | 08:07 - 09:56 | [23 empty] | T
Mole name |  | Dipole 23 | [9 empty] | Mole name |  | Dipole 24 | [29 empty] | T
Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 24 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 24 | [23 empty] | T
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw | [23 empty] | T
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) | [23 empty] | T
Rotation speed |  | 0.46 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.49 |  |  | I. Gain |  |  | 50/500 | [23 empty] | T
Dx average |  | 7.1186e-05 |  |  | Dy average |  |  | 0.0008336 |  |  |  | Dx average |  | 0.0003079 |  |  | Dy average |  |  | 0.00045779 | [23 empty] | T
Magnetic length |  | 14.3596 |  |  | Tranfs. Function (mT/KA) |  |  | 703.8547 |  |  |  | Magnetic length |  | 14.3599 |  |  | Tranfs. Function (mT/KA) |  |  | 703.4502 | [23 empty] | T
Phase |  | 5.2077 |  |  | Integral field (T.m) |  |  | 101.0436 |  |  |  | Phase |  | 5.2543 |  |  | Integral field (T.m) |  |  | 100.9091 | [23 empty] | T
T
File | Aperture 1 - Cold mass | [19 empty] | Integrals | File | Aperture 2 - Cold mass | [19 empty] | Integrals | T
C1 (mT) | 3.9855 | 7.0358 | 7.0385 | 7.0391 | 7.0382 | 7.0394 | 7.0398 | 7.0385 | 7.0387 | 7.0382 | 7.0388 | 7.0388 | 7.0398 | 7.0392 | 7.0396 | 7.0388 | 7.0383 | 7.0385 | 7.0357 | 4.0632 | 101.0436 | C1 (mT) | 3.9194 | 7.0314 | 7.0359 | 7.0363 | 7.0356 | 7.0357 | 7.0348 | 7.0338 | 7.0341 | 7.0345 | 7.0342 | 7.0347 | 7.0344 | 7.0345 | 7.0353 | 7.035 | 7.0338 | 7.0352 | 7.0318 | 4.0236 | 100.9091 | T
Angle (mrad) | 3.8663 | 0.25887 | -0.15578 | -0.46243 | -0.1963 | -0.41275 | -0.094205 | -0.19912 | -0.57856 | -0.44606 | -0.59487 | -0.52361 | -0.20725 | 0.078754 | 0.13968 | 0.32219 | 0.67313 | 0.78394 | 0.66978 | -2.1551 | 5.2077 | Angle (mrad) | 4.4572 | 0.74796 | -0.14853 | -0.48719 | 0.026097 | 0.36074 | 0.20841 | -0.13738 | -0.84003 | -0.82059 | -1.0256 | -0.89337 | -0.39779 | -0.005406 | 0.021157 | 0.15313 | 0.15292 | 0.65022 | 0.35934 | -0.71143 | 5.2543 | T
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | T
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | T
b2 | 26.3676 | 1.0905 | 1.2803 | 0.50981 | 0.35325 | 0.33209 | 0.67934 | 1.2559 | 0.39625 | -0.35371 | -0.14149 | 0.37119 | 1.2333 | 0.10943 | 0.96442 | 0.59567 | 1.2482 | 1.7404 | 1.3665 | 27.4403 | 2.2881 | b2 | -29.9165 | -1.6009 | -0.43294 | 0.23973 | -0.24622 | 0.45217 | -0.53973 | -0.099193 | 0.25632 | -0.10051 | -0.89408 | -1.0373 | -1.0085 | -0.77679 | -0.93597 | -1.2964 | -1.163 | -0.18284 | -1.1099 | -26.7668 | -2.2192 | T
b3 | 32.2936 | 0.93379 | -0.56622 | -0.12519 | -1.423 | -1.2341 | -1.3961 | -1.9201 | -1.9516 | -1.6948 | -1.8877 | -1.2214 | -2.0007 | -1.6084 | -2.1609 | -1.7251 | -1.53 | -1.9373 | -0.34622 | 6.1493 | -0.1024 | b3 | 31.3875 | 0.13952 | -1.2379 | -1.5297 | -2.458 | -3.4113 | -2.5415 | -2.9974 | -3.1878 | -2.4831 | -1.9461 | -1.7191 | -2.1414 | -2.5132 | -1.5791 | -1.7056 | -1.4879 | -2.2337 | -0.83701 | 4.5041 | -0.82641 | T
b4 | 1.1078 | -0.10347 | 0.17939 | 0.1611 | 0.044196 | 0.078086 | 0.023794 | 0.19689 | 0.22131 | 0.18507 | 0.19833 | 0.10645 | 0.30981 | 0.30418 | 0.30956 | 0.095704 | 0.075446 | 0.0035816 | 0.087799 | 0.76795 | 0.18534 | b4 | -0.66653 | -0.13769 | -0.28539 | -0.20292 | -0.2151 | -0.42804 | -0.16683 | -0.22289 | -0.12991 | -0.10315 | -0.20724 | -0.10108 | -0.22013 | -0.30583 | 0.048394 | -0.067275 | -0.077024 | -0.35619 | -0.15584 | -1.1732 | -0.22879 | T
b5 | -5.5883 | -0.22757 | -0.29257 | -0.21282 | -0.1353 | 0.013799 | -0.17981 | -0.071482 | 0.038538 | 0.072523 | -0.047192 | -0.20302 | 0.22907 | 0.21305 | 0.43668 | 0.1246 | 0.07642 | 0.16511 | -0.1476 | -3.3162 | -0.273 | b5 | -5.9635 | -0.24783 | -0.075231 | 0.23952 | 0.38587 | 0.86509 | 0.243 | 0.62779 | 0.59048 | 0.46903 | 0.50308 | 0.39617 | 0.34331 | 0.63857 | 0.65703 | 0.32791 | 0.21276 | 0.33691 | 0.044201 | -3.7467 | 0.057117 | T
b6 | 0.34031 | 0.11584 | -0.029705 | 0.03173 | -0.0089573 | -0.050932 | 0.038863 | 0.072269 | 0.14473 | 0.063033 | 0.070458 | -0.047892 | -0.16245 | 0.022743 | 0.01616 | 0.020769 | -0.054205 | -0.065204 | -0.045572 | -0.00091002 | 0.016915 | b6 | 0.08792 | -0.052631 | -0.14225 | -0.11791 | -0.094256 | -0.28538 | -0.16205 | -0.15766 | -0.20192 | -0.15024 | -0.022 | -0.0056955 | -0.11198 | -0.095224 | 0.006497 | -0.038764 | -0.032971 | -0.0086015 | -0.090102 | -0.14644 | -0.093991 | T
b7 | 1.9269 | 0.6159 | 0.74694 | 0.77206 | 0.73222 | 0.61454 | 0.70637 | 0.64233 | 0.65644 | 0.572 | 0.60321 | 0.63529 | 0.62692 | 0.57492 | 0.49883 | 0.66959 | 0.60204 | 0.65106 | 0.71146 | 0.43007 | 0.67759 | b7 | 1.795 | 0.65982 | 0.74713 | 0.65398 | 0.63282 | 0.42886 | 0.5319 | 0.46315 | 0.5489 | 0.51682 | 0.51672 | 0.60127 | 0.54856 | 0.55476 | 0.50768 | 0.64876 | 0.64708 | 0.6426 | 0.69121 | 0.25446 | 0.61098 | T
b8 | 0.1017 | 0.073178 | 0.061853 | 0.028335 | -0.020821 | -0.004709 | -0.019784 | -0.084783 | -0.040232 | -0.073183 | -0.049587 | 0.00078138 | -0.025777 | -0.028089 | 0.074487 | 0.056518 | 0.00052334 | -0.046904 | -0.054793 | -0.022621 | -0.0056656 | b8 | -0.043046 | 0.014954 | 0.03446 | 0.096331 | 0.047679 | 0.12171 | 0.043422 | 0.076691 | 0.089027 | 0.042495 | 0.023514 | -0.021879 | 0.0019824 | 0.066062 | 0.02707 | -0.0089566 | 0.020565 | 0.030378 | 0.049972 | 0.061109 | 0.040068 | T
b9 | 0.26399 | 0.45946 | 0.41435 | 0.42281 | 0.41706 | 0.43771 | 0.41308 | 0.39427 | 0.39952 | 0.42294 | 0.4267 | 0.42901 | 0.41679 | 0.42736 | 0.44381 | 0.41034 | 0.42262 | 0.4237 | 0.43052 | 0.31693 | 0.415 | b9 | 0.30278 | 0.46083 | 0.41804 | 0.41324 | 0.40164 | 0.43266 | 0.42905 | 0.43384 | 0.42003 | 0.43321 | 0.44047 | 0.43498 | 0.42842 | 0.43702 | 0.44332 | 0.41758 | 0.42882 | 0.43161 | 0.43579 | 0.31935 | 0.42301 | T
b10 | 0.040769 | 0.039786 | -0.026512 | -0.013411 | -0.0549 | -0.064166 | -0.043155 | -0.025808 | 0.023616 | 0.0074197 | 0.024851 | -0.015416 | 0.018619 | 0.10106 | 0.11759 | 0.070864 | -0.00040275 | -0.068423 | -0.0723 | -0.073503 | 0 | b10 | 0.051509 | 0.049121 | 0.0004322 | 0.028776 | -0.0037336 | -0.050531 | -0.02685 | -0.026574 | 0.036505 | -0.0080167 | 0.024653 | -0.02417 | -0.037108 | 0.004336 | 0.025269 | 0.0043319 | 0.0067637 | -0.039379 | 0.028499 | -0.036684 | 0 | T
b11 | 0.5532 | 0.6498 | 0.6457 | 0.64556 | 0.64378 | 0.65149 | 0.64999 | 0.65577 | 0.65154 | 0.65059 | 0.64811 | 0.64638 | 0.65697 | 0.65997 | 0.66665 | 0.64837 | 0.6532 | 0.65163 | 0.63921 | 0.57421 | 0.64562 | b11 | 0.53166 | 0.64458 | 0.64725 | 0.65684 | 0.65693 | 0.66598 | 0.66289 | 0.66756 | 0.66107 | 0.66143 | 0.66238 | 0.65854 | 0.66354 | 0.65823 | 0.66425 | 0.6551 | 0.65462 | 0.65529 | 0.64644 | 0.56617 | 0.65152 | T
b12 | 0.022162 | 0.008094 | 0.0028254 | -0.00068881 | -0.0068104 | -0.0053141 | -0.0041131 | -0.0091545 | -0.0055206 | -0.010985 | -0.0092734 | -0.0036793 | -0.0045462 | -0.00097305 | 0.0079028 | 0.0067679 | 0.0025958 | -0.0050537 | -0.0057497 | -0.015568 | -0.0020955 | b12 | -7.4989e-05 | 0.0041599 | 0.0016067 | 0.0066282 | 0.00077355 | 0.015928 | 0.0031767 | 0.0073877 | 0.011643 | 0.0046575 | 0.00080042 | -0.002838 | -0.0010515 | 0.0091043 | 0.0036052 | 0.0008532 | 0.0016579 | 0.004242 | 0.0039383 | -0.0041093 | 0.0038623 | T
b13 | 0.069077 | 0.045335 | 0.049981 | 0.05151 | 0.050528 | 0.043889 | 0.049839 | 0.048802 | 0.045881 | 0.043517 | 0.043735 | 0.049314 | 0.042993 | 0.041428 | 0.036904 | 0.041945 | 0.042164 | 0.042936 | 0.04792 | 0.042531 | 0.046088 | b13 | 0.059952 | 0.04821 | 0.052809 | 0.044986 | 0.047024 | 0.035897 | 0.042856 | 0.035971 | 0.038385 | 0.038844 | 0.039042 | 0.043148 | 0.042252 | 0.03755 | 0.038455 | 0.042807 | 0.044522 | 0.041914 | 0.047327 | 0.033521 | 0.042583 | T
b14 | 0.0097107 | 0.0016445 | -0.0046326 | -0.0043697 | -0.008824 | -0.010839 | -0.0081798 | -0.0072988 | -0.0039705 | -0.0062359 | -0.0033585 | -0.0048441 | -0.0035541 | 0.0026242 | 0.0055895 | 0.0023535 | -0.0034294 | -0.0087906 | -0.0087575 | -0.013126 | -0.0040197 | b14 | 0.0031039 | -0.0045436 | -0.0066333 | -0.0052781 | -0.0099303 | -0.012142 | -0.0082397 | -0.0071538 | -0.0038074 | -0.0064738 | -0.0062211 | -0.0084578 | -0.0089619 | -0.0061783 | -0.0044618 | -0.0072303 | -0.0054602 | -0.0073247 | -0.005623 | -0.01016 | -0.0067021 | T
b15 | 0.0076986 | 0.037816 | 0.037798 | 0.036682 | 0.041174 | 0.044477 | 0.039109 | 0.033815 | 0.033119 | 0.037073 | 0.039401 | 0.038162 | 0.040537 | 0.044822 | 0.045742 | 0.046715 | 0.042456 | 0.038947 | 0.033507 | 0.045152 | 0.038748 | b15 | -0.0068221 | 0.026986 | 0.024469 | 0.030127 | 0.029851 | 0.035522 | 0.026477 | 0.030336 | 0.032282 | 0.029134 | 0.030172 | 0.024894 | 0.025662 | 0.03072 | 0.032856 | 0.031574 | 0.027579 | 0.026946 | 0.025506 | 0.01842 | 0.027593 | T
b16 | [21 empty] | b16 | [21 empty] | T
b17 | [21 empty] | b17 | [21 empty] | T
a1 | 38.6629 | 2.5887 | -1.5578 | -4.6243 | -1.963 | -4.1275 | -0.94205 | -1.9912 | -5.7856 | -4.4606 | -5.9487 | -5.2361 | -2.0725 | 0.78754 | 1.3968 | 3.2219 | 6.7313 | 7.8394 | 6.6978 | -21.5512 | 0 | a1 | 44.5718 | 7.4796 | -1.4853 | -4.8719 | 0.26097 | 3.6074 | 2.0841 | -1.3738 | -8.4003 | -8.2059 | -10.2561 | -8.9337 | -3.9779 | -0.054061 | 0.21157 | 1.5313 | 1.5292 | 6.5022 | 3.5934 | -7.1143 | 0 | T
a2 | 0.15687 | -1.6751 | -1.0851 | -1.3337 | -0.77637 | -0.5318 | -2.0387 | -1.6899 | -0.11083 | -1.0421 | -1.3234 | -0.8227 | -1.1257 | -1.5306 | -0.40316 | -0.9872 | -1.1864 | -0.01924 | -0.61459 | -7.0855 | -1.1648 | a2 | 3.514 | 0.64967 | 0.86257 | -0.20354 | -0.29612 | -0.65284 | 0.43439 | 0.48062 | 0.60886 | 0.67914 | 0.76515 | 0.81619 | 0.79964 | 0.39436 | 0.61471 | 1.0519 | 1.0668 | 1.9574 | 0.94854 | -5.4113 | 0.51438 | T
a3 | -1.3013 | 0.32999 | -0.34672 | -0.27425 | -0.52785 | 0.22621 | 0.16179 | -0.18659 | -0.35192 | -0.078678 | 0.62735 | 0.078713 | 0.21257 | 0.26697 | 0.44157 | -0.04251 | 0.16326 | 0.26318 | -0.043759 | -0.86807 | -0.016641 | a3 | -2.2618 | -0.4021 | -1.0125 | -0.91473 | -1.1184 | -0.92441 | -0.50943 | -0.014262 | -0.089836 | -0.63434 | 0.12109 | -0.1872 | -0.39893 | -0.065366 | 0.44912 | -0.22674 | 0.057843 | 0.46841 | -0.42618 | -0.87574 | -0.39674 | T
a4 | 0.33431 | 0.32404 | 0.21635 | 0.66176 | 0.93768 | 0.84274 | 0.85597 | 0.39785 | 0.18084 | 0.39646 | 0.40108 | 0.73191 | 0.41392 | 0.71373 | 0.018054 | 0.017992 | -0.028779 | 0.21927 | -0.20589 | 0.46875 | 0.39466 | a4 | 1.0353 | 0.69441 | 0.80543 | 0.74172 | 0.75795 | 1.7917 | 0.3693 | 0.7259 | 1.0476 | 1.0822 | 0.8765 | 1.3144 | 0.93249 | 1.2739 | 0.8087 | 0.29255 | 0.44366 | 0.080073 | 0.15386 | 1.0178 | 0.80253 | T
a5 | 1.9347 | 0.017111 | 0.11936 | 0.0376 | -0.11665 | -0.030629 | 0.11535 | 0.01042 | 0.033488 | -0.17229 | 0.12462 | 0.041061 | 0.01237 | -0.025974 | 0.32378 | 0.22893 | 0.10696 | 0.081369 | -0.09572 | 0.03087 | 0.10054 | a5 | 1.3512 | -0.10948 | -0.063732 | -0.13332 | -0.33281 | -0.28663 | -0.13633 | -0.12456 | -0.043062 | -0.30713 | 0.032678 | 0.017921 | -0.12327 | -0.01096 | 0.24979 | 0.24106 | 0.13526 | 0.2037 | -0.030364 | -0.11346 | -0.0069722 | T
a6 | -0.30617 | -0.064174 | -0.0034423 | -0.043777 | -0.092351 | -0.020815 | -0.081579 | -0.10631 | 0.15575 | 0.041198 | 0.085622 | -0.072084 | -0.14458 | -0.096009 | 0.1134 | 0.20971 | 0.20036 | 0.28068 | 0.13891 | 0.0051188 | 0.017239 | a6 | -0.13424 | -0.21689 | -0.1325 | -0.099362 | -0.15755 | -0.30094 | -0.067131 | -0.19088 | -0.14084 | -0.063014 | -0.078883 | -0.23444 | -0.13746 | -0.22581 | -0.055613 | 0.00028543 | 0.044503 | 0.085488 | 0.0010514 | -0.11684 | -0.11039 | T
a7 | 1.506 | -0.009932 | -0.05485 | 0.031513 | -0.024509 | 0.031675 | -0.028705 | -0.041133 | 0.14577 | 0.032234 | -0.050569 | -0.089357 | -0.11117 | -0.0086106 | 0.079473 | -0.027259 | 0.052186 | 0.066042 | -0.074563 | 0.063021 | 0.042175 | a7 | 1.6452 | 0.058398 | -0.022576 | 0.00020753 | 0.0010816 | 0.1546 | -0.037092 | 0.010081 | 0.069013 | -0.030656 | -0.084502 | -0.012914 | 0.0045997 | 0.079495 | 0.066716 | -0.021432 | -0.0091747 | -0.012349 | -0.065421 | -0.00019297 | 0.055652 | T
a8 | -0.02259 | -0.0067745 | -0.010114 | -0.012742 | -0.014441 | 0.0014073 | -0.012992 | 0.0025473 | 0.043112 | 0.042948 | 0.027669 | 0.031383 | 0.053911 | 0.0042138 | -0.0025691 | -0.036648 | 0.0015793 | -0.0072315 | -0.031761 | -0.099308 | 0.0001775 | a8 | 0.060035 | 0.015375 | 0.033235 | 0.040937 | 0.023756 | 0.043601 | 0.053429 | 0.059939 | 0.058143 | 0.0716 | 0.083632 | 0.11045 | 0.12364 | 0.086655 | 0.057367 | -0.0095777 | 0.03623 | 0.011656 | 0.03521 | -0.023825 | 0.049929 | T
a9 | -0.24144 | -0.0031186 | -0.015403 | 0.0052095 | -0.012889 | -0.0030153 | 0.024419 | 0.0080278 | -0.030466 | -0.037389 | 0.0085136 | 0.031372 | 0.019838 | 0.039839 | 0.065509 | -0.0054713 | 0.009958 | 0.033003 | 0.018713 | 0.020876 | 0.0016718 | a9 | -0.26913 | -0.026067 | -0.015165 | -0.02225 | -0.02384 | -0.042356 | 0.0063799 | 0.039788 | 0.045125 | 0.03383 | 0.057207 | 0.044613 | 0.034085 | 0.048193 | 0.081808 | 0.041147 | 0.055344 | 0.11524 | 0.052212 | 0.032456 | 0.020592 | T
a10 | -0.0023516 | -0.0003122 | -0.006929 | 0.0059884 | -0.037453 | -0.03262 | -0.0077519 | 0.041531 | 0.084081 | 0.026565 | 0.014949 | 0.012198 | 0.011538 | -0.023395 | -0.0079597 | -0.039984 | 0.0014102 | 0.022298 | 0.028202 | -0.15764 | 0 | a10 | 0.0060928 | -0.017236 | -0.011648 | -0.035638 | -0.045095 | -0.025114 | 0.020385 | 0.01556 | 0.01374 | 0.041543 | 0.02375 | 0.057282 | 0.049231 | 0.023045 | -0.012315 | -0.052773 | 0.0032258 | 0.0062139 | 0.012921 | -0.12316 | 0 | T
a11 | 0.14885 | -0.056694 | -0.048386 | -0.044188 | -0.050065 | -0.044698 | -0.047007 | -0.044129 | -0.022483 | -0.044413 | -0.050707 | -0.053248 | -0.053848 | -0.047508 | -0.034122 | -0.038626 | -0.028977 | -0.033149 | -0.047067 | -0.013862 | -0.037246 | a11 | 0.19295 | 0.0010726 | 0.0068839 | 0.011125 | 0.010834 | 0.016894 | -0.0079083 | -0.0043067 | -0.0021946 | -0.020023 | -0.025112 | -0.0084431 | -0.0055082 | 0.0060769 | 0.0061884 | 0.0047939 | 0.0067548 | 0.0096997 | 0.0048389 | 0.022252 | 0.0068953 | T
a12 | -0.0047652 | -0.0012035 | 0.0020488 | 0.00066834 | -0.0034653 | -0.00092547 | -0.0018986 | 0.0019494 | 0.015938 | 0.0074833 | 0.0068407 | -0.00050862 | -0.0008048 | -0.0026587 | 0.0091371 | 0.0078885 | 0.010009 | 0.017609 | 0.011008 | -0.012796 | 0.0036057 | a12 | -0.0029581 | -0.006319 | 0.00041802 | -0.0013965 | -0.0039668 | -0.0030844 | 0.0022428 | -0.0019324 | -0.00080473 | 0.0031138 | 0.00059845 | -0.0012011 | 0.0006891 | -0.00050357 | 0.0037002 | 0.0027105 | 0.0073668 | 0.0089937 | 0.0051594 | -0.010597 | 0.00042215 | T
a13 | -0.022753 | -0.0013575 | -0.0039036 | -0.0031096 | -0.0047036 | -0.0037271 | -0.003232 | -0.0037721 | -0.0042513 | -0.0039975 | -0.0016847 | -0.0026025 | -0.0048831 | -0.00061402 | 0.0018302 | -0.0032227 | -0.0012693 | 0.0024898 | -0.0018284 | -5.9628e-06 | -0.0029633 | a13 | -0.016007 | -0.00012434 | -0.0043252 | -0.0059404 | -0.0050384 | -0.0075515 | -0.0037979 | 0.0022414 | 0.002099 | -0.00084516 | 0.0019875 | -0.00065667 | -0.00012009 | 0.0032879 | 0.006839 | 0.0032068 | 0.0047533 | 0.0084564 | 0.0022504 | 0.0047919 | 2.8249e-05 | T
a14 | 0.013021 | -0.013465 | -0.016059 | -0.012943 | -0.013917 | -0.017148 | -0.013546 | -0.010748 | -0.011198 | -0.012643 | -0.014483 | -0.011725 | -0.012992 | -0.016297 | -0.0207 | -0.02388 | -0.020483 | -0.016876 | -0.014531 | -0.016804 | -0.014416 | a14 | 0.0070644 | -0.006063 | -0.0019851 | -0.0056374 | -0.0053962 | -0.0034057 | -0.0046853 | -0.0051452 | -0.0048382 | -0.0020864 | -0.0022632 | 0.0036001 | 0.00091179 | -0.0015551 | -0.0074284 | -0.011814 | -0.0085506 | -0.0076585 | -0.0063111 | 0.00011368 | -0.0039892 | T
a15 | 0.0028509 | 0.0004738 | -0.0069029 | -0.0086637 | -0.014526 | -0.014253 | -0.011822 | -0.011642 | -0.0076434 | -0.01124 | -0.0081659 | -0.0092667 | -0.007111 | -0.00080641 | 0.0047893 | 0.00092721 | -0.0070493 | -0.01442 | -0.01386 | -0.022258 | -0.0079617 | a15 | 0.00028238 | -0.0045343 | -0.0085892 | -0.0063254 | -0.0096019 | -0.0079914 | -0.0064719 | -0.0044782 | -0.0016543 | -0.0051907 | -0.0039598 | -0.0061967 | -0.007272 | -0.0038942 | -0.0047833 | -0.0088259 | -0.0057291 | -0.0072241 | -0.0058965 | -0.014057 | -0.0060904 | T
a16 | [21 empty] | a16 | [21 empty] | T
a17 | [21 empty] | a17 | [21 empty] | T
Dx (m) | -0.0001151 | -0.00010341 | 6.8047e-05 | 3.6227e-05 | 0.00013643 | 0.00016086 | 0.00011083 | 7.3792e-05 | -5.3981e-05 | -1.4602e-05 | -6.1714e-05 | 4.29e-05 | -4.5428e-05 | -0.00026316 | -0.00030034 | -0.0001916 | 0 | 0.00018091 | 0.00019669 | 0.00020396 | 7.1186e-05 | Dx (m) | -0.00015174 | -0.00012946 | 0 | -7.2881e-05 | 1.1586e-05 | 0.00013039 | 6.9462e-05 | 6.7921e-05 | -9.38e-05 | 2.3834e-05 | -6.0872e-05 | 6.4285e-05 | 9.6118e-05 | -1.1733e-05 | -6.4383e-05 | -1.0238e-05 | -1.7648e-05 | 0.00010186 | -7.519e-05 | 0.00012385 | 0.0003079 | T
Dy (m) | 3.8325e-05 | 0 | 2.3331e-05 | -1.329e-05 | 0.00010959 | 9.6212e-05 | 2.8308e-05 | -0.00010276 | -0.00022086 | -7.0384e-05 | -4.4018e-05 | -2.8549e-05 | -3.3571e-05 | 4.1356e-05 | 0 | 9.3386e-05 | 0 | -4.9257e-05 | -6.0745e-05 | 0.00047227 | 0.0008336 | Dy (m) | 3.56e-05 | 4.5733e-05 | 3.0603e-05 | 9.351e-05 | 0.00011649 | 6.0686e-05 | -5.1438e-05 | -3.9163e-05 | -3.5676e-05 | -0.00010603 | -6.3263e-05 | -0.00014709 | -0.00012534 | -5.9417e-05 | 3.2117e-05 | 0.00013706 | 0 | -1.7664e-05 | -3.3436e-05 | 0.00036462 | 0.00045779 | T
Magnetic length (m) |  | 14.3596 | [19 empty] | Magnetic length (m) |  | 14.3599 | [19 empty] | T
Transfer funct. (mT/KA) |  | 703.8547 | [19 empty] | Transfer funct. (mT/KA) |  | 703.4502 | [19 empty] | T
Coil positioning (mm) |  | 0.025281 | [19 empty] | Coil positioning (mm) |  | 0.028333 | [19 empty] | T
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10 | [19 empty] | T
T
T
Soft version | DIMM V2.0.8 | [20 empty] | Soft version | DIMM V2.0.8 | [20 empty] | T
Level Average | 0.05412 | 0.29395 | 0.46808 | 0.42856 | 0.67888 | 0.26325 | 0.69741 | 0.13776 | 0.16656 | 0.30392 | 0.29517 | 0.2073 | 0.1357 | -0.063811 | -0.037136 | -0.081846 | 0.4143 | 0.30601 | -0.049393 | -0.58946 |  | Level Average | -0.18976 | -0.14658 | -0.11557 | -0.085364 | 0.024985 | -0.087237 | -0.26144 | -0.30892 | 0.01939 | -0.24307 | -0.20747 | -0.56491 | 0.082503 | -0.6948 | -0.48953 | -0.10071 | -0.16706 | -0.31262 | -0.18443 | 0.23994 |  | T
Level Sigma2 | 0.076628 | 0.11462 | 0.01858 | 0.031981 | 0.029239 | 0.095505 | 0.00514 | 0.048504 | 0.03139 | 0.061772 | 0.032655 | 0.005104 | 0.049348 | 7.9e-05 | 0.003932 | 0.001171 | 0.001042 | 0.000708 | 0.98617 | 0.2195 |  | Level Sigma2 | 0.001653 | 0.031168 | 0.001193 | 0.002671 | 0.001546 | 0.016426 | 0.004826 | 0.006905 | 0.000978 | 0.002048 | 0.015382 | 0.007713 | 0.02104 | 0.00575 | 0.005431 | 0.003031 | 0.003528 | 0.018561 | 0.001773 | 0.006906 |  | T
Position (mm) | 0 | 751 | 1501 | 2251 | 3001 | 3752 | 4502 | 5251 | 6002 | 6752 | 7502 | 8252 | 9002 | 9752 | 10502 | 11252 | 12002 | 12752 | 13502 | 14252 |  | Position (mm) | 0 | 751 | 1502 | 2252 | 3003 | 3754 | 4504 | 5255 | 6006 | 6756 | 7507 | 8258 | 9009 | 9759 | 10511 | 11261 | 12012 | 12762 | 13513 | 14263 |  | T
Fist move (mm) | First move Ap1 | 940.8225 | First move Ap2 | 936.2705 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2 | [18 empty] | T
T
T
T
T
T
T
T
T
T
T
T

## Summary Data
File | HCMBALA001-02000097 |  |  | Aperture 1 - Cold mass | [16 empty] | Integrals | File | HCMBALA001-2000097 |  |  | Aperture 2 - Cold mass | [16 empty] | Integrals
C1/i (mT/KA) | 398.5513 | 703.5846 | 703.8522 | 703.9057 | 703.8248 | 703.94 | 703.9832 | 703.8475 | 703.8728 | 703.819 | 703.8771 | 703.883 | 703.9821 | 703.916 | 703.9616 | 703.8849 | 703.8289 | 703.8509 | 703.5698 | 406.3204 | 101.0436 | C1/i (mT/KA) | 391.9435 | 703.1351 | 703.5884 | 703.6321 | 703.562 | 703.5688 | 703.4837 | 703.3785 | 703.4148 | 703.4457 | 703.424 | 703.4666 | 703.443 | 703.4511 | 703.5287 | 703.4986 | 703.3847 | 703.5194 | 703.178 | 402.3598 | 100.9091
Angle (mrad) | 3.8663 | 0.25887 | -0.15578 | -0.46243 | -0.1963 | -0.41275 | -0.094205 | -0.19912 | -0.57856 | -0.44606 | -0.59487 | -0.52361 | -0.20725 | 0.078754 | 0.13968 | 0.32219 | 0.67313 | 0.78394 | 0.66978 | -2.1551 | 5.2077 | Angle (mrad) | 4.4572 | 0.74796 | -0.14853 | -0.48719 | 0.026097 | 0.36074 | 0.20841 | -0.13738 | -0.84003 | -0.82059 | -1.0256 | -0.89337 | -0.39779 | -0.005406 | 0.021157 | 0.15313 | 0.15292 | 0.65022 | 0.35934 | -0.71143 | 5.2543
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 26.3676 | 1.0905 | 1.2803 | 0.50981 | 0.35325 | 0.33209 | 0.67934 | 1.2559 | 0.39625 | -0.35371 | -0.14149 | 0.37119 | 1.2333 | 0.10943 | 0.96442 | 0.59567 | 1.2482 | 1.7404 | 1.3665 | 27.4403 |  | b2 | -29.9165 | -1.6009 | -0.43294 | 0.23973 | -0.24622 | 0.45217 | -0.53973 | -0.099193 | 0.25632 | -0.10051 | -0.89408 | -1.0373 | -1.0085 | -0.77679 | -0.93597 | -1.2964 | -1.163 | -0.18284 | -1.1099 | -26.7668
b3 | 32.2936 | 0.93379 | -0.56622 | -0.12519 | -1.423 | -1.2341 | -1.3961 | -1.9201 | -1.9516 | -1.6948 | -1.8877 | -1.2214 | -2.0007 | -1.6084 | -2.1609 | -1.7251 | -1.53 | -1.9373 | -0.34622 | 6.1493 |  | b3 | 31.3875 | 0.13952 | -1.2379 | -1.5297 | -2.458 | -3.4113 | -2.5415 | -2.9974 | -3.1878 | -2.4831 | -1.9461 | -1.7191 | -2.1414 | -2.5132 | -1.5791 | -1.7056 | -1.4879 | -2.2337 | -0.83701 | 4.5041
b4 | 1.1078 | -0.10347 | 0.17939 | 0.1611 | 0.044196 | 0.078086 | 0.023794 | 0.19689 | 0.22131 | 0.18507 | 0.19833 | 0.10645 | 0.30981 | 0.30418 | 0.30956 | 0.095704 | 0.075446 | 0.0035816 | 0.087799 | 0.76795 |  | b4 | -0.66653 | -0.13769 | -0.28539 | -0.20292 | -0.2151 | -0.42804 | -0.16683 | -0.22289 | -0.12991 | -0.10315 | -0.20724 | -0.10108 | -0.22013 | -0.30583 | 0.048394 | -0.067275 | -0.077024 | -0.35619 | -0.15584 | -1.1732
b5 | -5.5883 | -0.22757 | -0.29257 | -0.21282 | -0.1353 | 0.013799 | -0.17981 | -0.071482 | 0.038538 | 0.072523 | -0.047192 | -0.20302 | 0.22907 | 0.21305 | 0.43668 | 0.1246 | 0.07642 | 0.16511 | -0.1476 | -3.3162 |  | b5 | -5.9635 | -0.24783 | -0.075231 | 0.23952 | 0.38587 | 0.86509 | 0.243 | 0.62779 | 0.59048 | 0.46903 | 0.50308 | 0.39617 | 0.34331 | 0.63857 | 0.65703 | 0.32791 | 0.21276 | 0.33691 | 0.044201 | -3.7467
b6 | 0.34031 | 0.11584 | -0.029705 | 0.03173 | -0.0089573 | -0.050932 | 0.038863 | 0.072269 | 0.14473 | 0.063033 | 0.070458 | -0.047892 | -0.16245 | 0.022743 | 0.01616 | 0.020769 | -0.054205 | -0.065204 | -0.045572 | -0.00091002 |  | b6 | 0.08792 | -0.052631 | -0.14225 | -0.11791 | -0.094256 | -0.28538 | -0.16205 | -0.15766 | -0.20192 | -0.15024 | -0.022 | -0.0056955 | -0.11198 | -0.095224 | 0.006497 | -0.038764 | -0.032971 | -0.0086015 | -0.090102 | -0.14644
b7 | 1.9269 | 0.6159 | 0.74694 | 0.77206 | 0.73222 | 0.61454 | 0.70637 | 0.64233 | 0.65644 | 0.572 | 0.60321 | 0.63529 | 0.62692 | 0.57492 | 0.49883 | 0.66959 | 0.60204 | 0.65106 | 0.71146 | 0.43007 |  | b7 | 1.795 | 0.65982 | 0.74713 | 0.65398 | 0.63282 | 0.42886 | 0.5319 | 0.46315 | 0.5489 | 0.51682 | 0.51672 | 0.60127 | 0.54856 | 0.55476 | 0.50768 | 0.64876 | 0.64708 | 0.6426 | 0.69121 | 0.25446
b8 | 0.1017 | 0.073178 | 0.061853 | 0.028335 | -0.020821 | -0.004709 | -0.019784 | -0.084783 | -0.040232 | -0.073183 | -0.049587 | 0.00078138 | -0.025777 | -0.028089 | 0.074487 | 0.056518 | 0.00052334 | -0.046904 | -0.054793 | -0.022621 |  | b8 | -0.043046 | 0.014954 | 0.03446 | 0.096331 | 0.047679 | 0.12171 | 0.043422 | 0.076691 | 0.089027 | 0.042495 | 0.023514 | -0.021879 | 0.0019824 | 0.066062 | 0.02707 | -0.0089566 | 0.020565 | 0.030378 | 0.049972 | 0.061109
b9 | 0.26399 | 0.45946 | 0.41435 | 0.42281 | 0.41706 | 0.43771 | 0.41308 | 0.39427 | 0.39952 | 0.42294 | 0.4267 | 0.42901 | 0.41679 | 0.42736 | 0.44381 | 0.41034 | 0.42262 | 0.4237 | 0.43052 | 0.31693 |  | b9 | 0.30278 | 0.46083 | 0.41804 | 0.41324 | 0.40164 | 0.43266 | 0.42905 | 0.43384 | 0.42003 | 0.43321 | 0.44047 | 0.43498 | 0.42842 | 0.43702 | 0.44332 | 0.41758 | 0.42882 | 0.43161 | 0.43579 | 0.31935
b10 | 0.040769 | 0.039786 | -0.026512 | -0.013411 | -0.0549 | -0.064166 | -0.043155 | -0.025808 | 0.023616 | 0.0074197 | 0.024851 | -0.015416 | 0.018619 | 0.10106 | 0.11759 | 0.070864 | -0.00040275 | -0.068423 | -0.0723 | -0.073503 |  | b10 | 0.051509 | 0.049121 | 0.0004322 | 0.028776 | -0.0037336 | -0.050531 | -0.02685 | -0.026574 | 0.036505 | -0.0080167 | 0.024653 | -0.02417 | -0.037108 | 0.004336 | 0.025269 | 0.0043319 | 0.0067637 | -0.039379 | 0.028499 | -0.036684
b11 | 0.5532 | 0.6498 | 0.6457 | 0.64556 | 0.64378 | 0.65149 | 0.64999 | 0.65577 | 0.65154 | 0.65059 | 0.64811 | 0.64638 | 0.65697 | 0.65997 | 0.66665 | 0.64837 | 0.6532 | 0.65163 | 0.63921 | 0.57421 |  | b11 | 0.53166 | 0.64458 | 0.64725 | 0.65684 | 0.65693 | 0.66598 | 0.66289 | 0.66756 | 0.66107 | 0.66143 | 0.66238 | 0.65854 | 0.66354 | 0.65823 | 0.66425 | 0.6551 | 0.65462 | 0.65529 | 0.64644 | 0.56617
b12 | 0.022162 | 0.008094 | 0.0028254 | -0.00068881 | -0.0068104 | -0.0053141 | -0.0041131 | -0.0091545 | -0.0055206 | -0.010985 | -0.0092734 | -0.0036793 | -0.0045462 | -0.00097305 | 0.0079028 | 0.0067679 | 0.0025958 | -0.0050537 | -0.0057497 | -0.015568 |  | b12 | -7.4989e-05 | 0.0041599 | 0.0016067 | 0.0066282 | 0.00077355 | 0.015928 | 0.0031767 | 0.0073877 | 0.011643 | 0.0046575 | 0.00080042 | -0.002838 | -0.0010515 | 0.0091043 | 0.0036052 | 0.0008532 | 0.0016579 | 0.004242 | 0.0039383 | -0.0041093
b13 | 0.069077 | 0.045335 | 0.049981 | 0.05151 | 0.050528 | 0.043889 | 0.049839 | 0.048802 | 0.045881 | 0.043517 | 0.043735 | 0.049314 | 0.042993 | 0.041428 | 0.036904 | 0.041945 | 0.042164 | 0.042936 | 0.04792 | 0.042531 |  | b13 | 0.059952 | 0.04821 | 0.052809 | 0.044986 | 0.047024 | 0.035897 | 0.042856 | 0.035971 | 0.038385 | 0.038844 | 0.039042 | 0.043148 | 0.042252 | 0.03755 | 0.038455 | 0.042807 | 0.044522 | 0.041914 | 0.047327 | 0.033521
b14 | 0.0097107 | 0.0016445 | -0.0046326 | -0.0043697 | -0.008824 | -0.010839 | -0.0081798 | -0.0072988 | -0.0039705 | -0.0062359 | -0.0033585 | -0.0048441 | -0.0035541 | 0.0026242 | 0.0055895 | 0.0023535 | -0.0034294 | -0.0087906 | -0.0087575 | -0.013126 |  | b14 | 0.0031039 | -0.0045436 | -0.0066333 | -0.0052781 | -0.0099303 | -0.012142 | -0.0082397 | -0.0071538 | -0.0038074 | -0.0064738 | -0.0062211 | -0.0084578 | -0.0089619 | -0.0061783 | -0.0044618 | -0.0072303 | -0.0054602 | -0.0073247 | -0.005623 | -0.01016
b15 | 0.0076986 | 0.037816 | 0.037798 | 0.036682 | 0.041174 | 0.044477 | 0.039109 | 0.033815 | 0.033119 | 0.037073 | 0.039401 | 0.038162 | 0.040537 | 0.044822 | 0.045742 | 0.046715 | 0.042456 | 0.038947 | 0.033507 | 0.045152 |  | b15 | -0.0068221 | 0.026986 | 0.024469 | 0.030127 | 0.029851 | 0.035522 | 0.026477 | 0.030336 | 0.032282 | 0.029134 | 0.030172 | 0.024894 | 0.025662 | 0.03072 | 0.032856 | 0.031574 | 0.027579 | 0.026946 | 0.025506 | 0.01842
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 38.6629 | 2.5887 | -1.5578 | -4.6243 | -1.963 | -4.1275 | -0.94205 | -1.9912 | -5.7856 | -4.4606 | -5.9487 | -5.2361 | -2.0725 | 0.78754 | 1.3968 | 3.2219 | 6.7313 | 7.8394 | 6.6978 | -21.5512 |  | a1 | 44.5718 | 7.4796 | -1.4853 | -4.8719 | 0.26097 | 3.6074 | 2.0841 | -1.3738 | -8.4003 | -8.2059 | -10.2561 | -8.9337 | -3.9779 | -0.054061 | 0.21157 | 1.5313 | 1.5292 | 6.5022 | 3.5934 | -7.1143
a2 | 0.15687 | -1.6751 | -1.0851 | -1.3337 | -0.77637 | -0.5318 | -2.0387 | -1.6899 | -0.11083 | -1.0421 | -1.3234 | -0.8227 | -1.1257 | -1.5306 | -0.40316 | -0.9872 | -1.1864 | -0.01924 | -0.61459 | -7.0855 |  | a2 | 3.514 | 0.64967 | 0.86257 | -0.20354 | -0.29612 | -0.65284 | 0.43439 | 0.48062 | 0.60886 | 0.67914 | 0.76515 | 0.81619 | 0.79964 | 0.39436 | 0.61471 | 1.0519 | 1.0668 | 1.9574 | 0.94854 | -5.4113
a3 | -1.3013 | 0.32999 | -0.34672 | -0.27425 | -0.52785 | 0.22621 | 0.16179 | -0.18659 | -0.35192 | -0.078678 | 0.62735 | 0.078713 | 0.21257 | 0.26697 | 0.44157 | -0.04251 | 0.16326 | 0.26318 | -0.043759 | -0.86807 |  | a3 | -2.2618 | -0.4021 | -1.0125 | -0.91473 | -1.1184 | -0.92441 | -0.50943 | -0.014262 | -0.089836 | -0.63434 | 0.12109 | -0.1872 | -0.39893 | -0.065366 | 0.44912 | -0.22674 | 0.057843 | 0.46841 | -0.42618 | -0.87574
a4 | 0.33431 | 0.32404 | 0.21635 | 0.66176 | 0.93768 | 0.84274 | 0.85597 | 0.39785 | 0.18084 | 0.39646 | 0.40108 | 0.73191 | 0.41392 | 0.71373 | 0.018054 | 0.017992 | -0.028779 | 0.21927 | -0.20589 | 0.46875 |  | a4 | 1.0353 | 0.69441 | 0.80543 | 0.74172 | 0.75795 | 1.7917 | 0.3693 | 0.7259 | 1.0476 | 1.0822 | 0.8765 | 1.3144 | 0.93249 | 1.2739 | 0.8087 | 0.29255 | 0.44366 | 0.080073 | 0.15386 | 1.0178
a5 | 1.9347 | 0.017111 | 0.11936 | 0.0376 | -0.11665 | -0.030629 | 0.11535 | 0.01042 | 0.033488 | -0.17229 | 0.12462 | 0.041061 | 0.01237 | -0.025974 | 0.32378 | 0.22893 | 0.10696 | 0.081369 | -0.09572 | 0.03087 |  | a5 | 1.3512 | -0.10948 | -0.063732 | -0.13332 | -0.33281 | -0.28663 | -0.13633 | -0.12456 | -0.043062 | -0.30713 | 0.032678 | 0.017921 | -0.12327 | -0.01096 | 0.24979 | 0.24106 | 0.13526 | 0.2037 | -0.030364 | -0.11346
a6 | -0.30617 | -0.064174 | -0.0034423 | -0.043777 | -0.092351 | -0.020815 | -0.081579 | -0.10631 | 0.15575 | 0.041198 | 0.085622 | -0.072084 | -0.14458 | -0.096009 | 0.1134 | 0.20971 | 0.20036 | 0.28068 | 0.13891 | 0.0051188 |  | a6 | -0.13424 | -0.21689 | -0.1325 | -0.099362 | -0.15755 | -0.30094 | -0.067131 | -0.19088 | -0.14084 | -0.063014 | -0.078883 | -0.23444 | -0.13746 | -0.22581 | -0.055613 | 0.00028543 | 0.044503 | 0.085488 | 0.0010514 | -0.11684
a7 | 1.506 | -0.009932 | -0.05485 | 0.031513 | -0.024509 | 0.031675 | -0.028705 | -0.041133 | 0.14577 | 0.032234 | -0.050569 | -0.089357 | -0.11117 | -0.0086106 | 0.079473 | -0.027259 | 0.052186 | 0.066042 | -0.074563 | 0.063021 |  | a7 | 1.6452 | 0.058398 | -0.022576 | 0.00020753 | 0.0010816 | 0.1546 | -0.037092 | 0.010081 | 0.069013 | -0.030656 | -0.084502 | -0.012914 | 0.0045997 | 0.079495 | 0.066716 | -0.021432 | -0.0091747 | -0.012349 | -0.065421 | -0.00019297
a8 | -0.02259 | -0.0067745 | -0.010114 | -0.012742 | -0.014441 | 0.0014073 | -0.012992 | 0.0025473 | 0.043112 | 0.042948 | 0.027669 | 0.031383 | 0.053911 | 0.0042138 | -0.0025691 | -0.036648 | 0.0015793 | -0.0072315 | -0.031761 | -0.099308 |  | a8 | 0.060035 | 0.015375 | 0.033235 | 0.040937 | 0.023756 | 0.043601 | 0.053429 | 0.059939 | 0.058143 | 0.0716 | 0.083632 | 0.11045 | 0.12364 | 0.086655 | 0.057367 | -0.0095777 | 0.03623 | 0.011656 | 0.03521 | -0.023825
a9 | -0.24144 | -0.0031186 | -0.015403 | 0.0052095 | -0.012889 | -0.0030153 | 0.024419 | 0.0080278 | -0.030466 | -0.037389 | 0.0085136 | 0.031372 | 0.019838 | 0.039839 | 0.065509 | -0.0054713 | 0.009958 | 0.033003 | 0.018713 | 0.020876 |  | a9 | -0.26913 | -0.026067 | -0.015165 | -0.02225 | -0.02384 | -0.042356 | 0.0063799 | 0.039788 | 0.045125 | 0.03383 | 0.057207 | 0.044613 | 0.034085 | 0.048193 | 0.081808 | 0.041147 | 0.055344 | 0.11524 | 0.052212 | 0.032456
a10 | -0.0023516 | -0.0003122 | -0.006929 | 0.0059884 | -0.037453 | -0.03262 | -0.0077519 | 0.041531 | 0.084081 | 0.026565 | 0.014949 | 0.012198 | 0.011538 | -0.023395 | -0.0079597 | -0.039984 | 0.0014102 | 0.022298 | 0.028202 | -0.15764 |  | a10 | 0.0060928 | -0.017236 | -0.011648 | -0.035638 | -0.045095 | -0.025114 | 0.020385 | 0.01556 | 0.01374 | 0.041543 | 0.02375 | 0.057282 | 0.049231 | 0.023045 | -0.012315 | -0.052773 | 0.0032258 | 0.0062139 | 0.012921 | -0.12316
a11 | 0.14885 | -0.056694 | -0.048386 | -0.044188 | -0.050065 | -0.044698 | -0.047007 | -0.044129 | -0.022483 | -0.044413 | -0.050707 | -0.053248 | -0.053848 | -0.047508 | -0.034122 | -0.038626 | -0.028977 | -0.033149 | -0.047067 | -0.013862 |  | a11 | 0.19295 | 0.0010726 | 0.0068839 | 0.011125 | 0.010834 | 0.016894 | -0.0079083 | -0.0043067 | -0.0021946 | -0.020023 | -0.025112 | -0.0084431 | -0.0055082 | 0.0060769 | 0.0061884 | 0.0047939 | 0.0067548 | 0.0096997 | 0.0048389 | 0.022252
a12 | -0.0047652 | -0.0012035 | 0.0020488 | 0.00066834 | -0.0034653 | -0.00092547 | -0.0018986 | 0.0019494 | 0.015938 | 0.0074833 | 0.0068407 | -0.00050862 | -0.0008048 | -0.0026587 | 0.0091371 | 0.0078885 | 0.010009 | 0.017609 | 0.011008 | -0.012796 |  | a12 | -0.0029581 | -0.006319 | 0.00041802 | -0.0013965 | -0.0039668 | -0.0030844 | 0.0022428 | -0.0019324 | -0.00080473 | 0.0031138 | 0.00059845 | -0.0012011 | 0.0006891 | -0.00050357 | 0.0037002 | 0.0027105 | 0.0073668 | 0.0089937 | 0.0051594 | -0.010597
a13 | -0.022753 | -0.0013575 | -0.0039036 | -0.0031096 | -0.0047036 | -0.0037271 | -0.003232 | -0.0037721 | -0.0042513 | -0.0039975 | -0.0016847 | -0.0026025 | -0.0048831 | -0.00061402 | 0.0018302 | -0.0032227 | -0.0012693 | 0.0024898 | -0.0018284 | -5.9628e-06 |  | a13 | -0.016007 | -0.00012434 | -0.0043252 | -0.0059404 | -0.0050384 | -0.0075515 | -0.0037979 | 0.0022414 | 0.002099 | -0.00084516 | 0.0019875 | -0.00065667 | -0.00012009 | 0.0032879 | 0.006839 | 0.0032068 | 0.0047533 | 0.0084564 | 0.0022504 | 0.0047919
a14 | 0.013021 | -0.013465 | -0.016059 | -0.012943 | -0.013917 | -0.017148 | -0.013546 | -0.010748 | -0.011198 | -0.012643 | -0.014483 | -0.011725 | -0.012992 | -0.016297 | -0.0207 | -0.02388 | -0.020483 | -0.016876 | -0.014531 | -0.016804 |  | a14 | 0.0070644 | -0.006063 | -0.0019851 | -0.0056374 | -0.0053962 | -0.0034057 | -0.0046853 | -0.0051452 | -0.0048382 | -0.0020864 | -0.0022632 | 0.0036001 | 0.00091179 | -0.0015551 | -0.0074284 | -0.011814 | -0.0085506 | -0.0076585 | -0.0063111 | 0.00011368
a15 | 0.0028509 | 0.0004738 | -0.0069029 | -0.0086637 | -0.014526 | -0.014253 | -0.011822 | -0.011642 | -0.0076434 | -0.01124 | -0.0081659 | -0.0092667 | -0.007111 | -0.00080641 | 0.0047893 | 0.00092721 | -0.0070493 | -0.01442 | -0.01386 | -0.022258 |  | a15 | 0.00028238 | -0.0045343 | -0.0085892 | -0.0063254 | -0.0096019 | -0.0079914 | -0.0064719 | -0.0044782 | -0.0016543 | -0.0051907 | -0.0039598 | -0.0061967 | -0.007272 | -0.0038942 | -0.0047833 | -0.0088259 | -0.0057291 | -0.0072241 | -0.0058965 | -0.014057
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.11501 | 0.10337 | -0.06805 | -0.036227 | -0.13646 | -0.16084 | -0.11082 | -0.073815 | 0.053985 | 0.014581 | 0.061739 | -0.042897 | 0.045428 | 0.26329 | 0.30012 | 0.19137 | -0.0012086 | -0.181 | -0.19674 | -0.20622 |  | Dx (mm) | 0.15178 | 0.12947 | 0.00080969 | 0.072894 | -0.011583 | -0.13053 | -0.069472 | -0.067927 | 0.093757 | -0.023815 | 0.060872 | -0.064278 | -0.096112 | 0.011747 | 0.06437 | 0.010239 | 0.017649 | -0.1019 | 0.075195 | -0.12449
Dy (mm) | -0.038173 | 0.0082022 | -0.023342 | 0.01329 | -0.10951 | -0.096155 | -0.028289 | 0.1027 | 0.22125 | 0.07041 | 0.044041 | 0.028546 | 0.033579 | -0.04131 | -0.0049367 | -0.093436 | 0.0036164 | 0.048965 | 0.060517 | -0.47168 |  | Dy (mm) | -0.035603 | -0.045674 | -0.030602 | -0.09347 | -0.11651 | -0.060797 | 0.05145 | 0.039187 | 0.035645 | 0.10605 | 0.063262 | 0.14705 | 0.12533 | 0.059408 | -0.032117 | -0.13702 | 0.008195 | 0.017629 | 0.033417 | -0.36492
Mag. len. (m) | 14.3596 | [20 empty] | Mag. len. (m) | 14.3599
TF (mT/KA) | 703.8547 | [20 empty] | TF (mT/KA) | 703.4502
Coil pos. (mm) | 0.025281 | [20 empty] | Coil pos. (mm) | 0.028333
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.0021457 | [20 empty] | Twist (m2 rad) | 0.010115

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 703.8547 | 0.11308 | 0.022025 |  | -0.52486 | 4.4488 |  |  | 703.4502 | 0.12794 | -0.020427 |  | -1.1527 | 5.2891 |  | 0 | 4
2 | 2.2838 |  | 0.72666 | 0.56933 | -1.1491 |  | -1.0151 | 0.55358 | -2.2356 |  | -0.58557 | 0.5934 | 0.52776 |  | 0.61302 | 0.58834 |  | 0 | 5
3 | -0.1055 |  | -1.3241 | 0.81033 | -0.013289 |  | 0.052018 | 0.30288 | -0.82705 |  | -1.991 | 0.86913 | -0.39774 |  | -0.32215 | 0.46927 |  | 0 | 6
4 | 0.18084 |  | 0.1339 | 0.10829 | 0.37877 |  | 0.39224 | 0.33921 | -0.22431 |  | -0.18353 | 0.10643 | 0.78793 |  | 0.78631 | 0.42855 |  | 0 | 7
5 | -0.27345 |  | -0.0084051 | 0.1949 | 0.10023 |  | 0.044292 | 0.11662 | 0.056306 |  | 0.36332 | 0.27358 | -0.0075079 |  | -0.045774 | 0.17594 |  | 0 | 8
6 | 0.016691 |  | 0.0095101 | 0.0718 | 0.015303 |  | 0.024409 | 0.12624 | -0.09784 |  | -0.098784 | 0.07604 | -0.11235 |  | -0.11055 | 0.106 |  | 0 | 9
7 | 0.67728 |  | 0.64549 | 0.069552 | 0.042361 |  | -0.004017 | 0.065845 | 0.6123 |  | 0.5868 | 0.083429 | 0.056148 |  | 0.0086335 | 0.058486 |  | 0 | 10
8 | -0.005729 |  | -0.0090905 | 0.045611 | 0.00016065 |  | 0.0011835 | 0.021256 | 0.040915 |  | 0.042982 | 0.041062 | 0.050416 |  | 0.049922 | 0.022511 |  | 0 | 11
9 | 0.41321 |  | 0.42007 | 0.01312 | 0.000963 |  | 0.0090592 | 0.027287 | 0.42371 |  | 0.43022 | 0.012145 | 0.020752 |  | 0.029822 | 0.04177 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.64556 |  | 0.65074 | 0.0062393 | -0.037297 |  | -0.043725 | 0.0091735 | 0.65163 |  | 0.65799 | 0.0067829 | 0.0068923 |  | 0.00069343 | 0.01108 |  | 0 | 14
12 | -0.00018849 |  | -0.0022439 | 0.0052067 | 0.00036865 |  | 0.0039103 | 0.0056023 | 0.00037705 |  | 0.0042717 | 0.0051376 | 5.5284e-05 |  | 0.0006282 | 0.004191 |  | 0 | 15
13 | 0.0045704 |  | 0.044981 | 0.0040575 | -0.00030004 |  | -0.0022111 | 0.0033443 | 0.004255 |  | 0.042347 | 0.0048783 | 1.9111e-06 |  | 0.00047633 | 0.0045746 |  | 0 | 16
14 | -0.00039388 |  | -0.0045709 | 0.0016645 | -0.0014788 |  | -0.016038 | 0.0030268 | -0.00064582 |  | -0.0068569 | 0.00093461 | -0.000407 |  | -0.0046706 | 0.0025819 |  | 0 | 17
15 | 0.0038748 |  | 0.03952 | 0.0041123 | -0.00079617 |  | -0.0078434 | 0.0057756 | 0.0027593 |  | 0.02895 | 0.0030708 | -0.00060904 |  | -0.0060344 | 0.0020184 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.57732 | -0.52617
3 | [12 empty] | 3 | 9 | -0.70236 | -0.44842
4 | [12 empty] | 4 | 16 | -1.6521 | -1.5438
5 | [12 empty] | 5 | 25 | -1.8921 | -1.5169
6 | [12 empty] | 6 | 36 | -2.3517 | -2.4104
7 | [12 empty] | 7 | 49 | -2.6931 | -2.6614
8 | [12 empty] | 8 | 64 | -3.4694 | -3.4932
9 | [12 empty] | 9 | 81 | -3.9675 | -3.7932
10 | [12 empty] | sigma (mm) |  | 0.025281 | 0.028333
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05
17 | 0.00044619 | 0.00011155 | 2.6771e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 398.5513 | 703.5846 | 703.8522 | 703.9057 | 703.8248 | 703.94 | 703.9832 | 703.8475 | 703.8728 | 703.819 | 703.8771 | 703.883 | 703.9821 | 703.916 | 703.9616 | 703.8849 | 703.8289 | 703.8509 | 703.5698 | 406.3204
2 | 25.9365 | 1.0795 | 1.2767 | 0.50884 | 0.33721 | 0.30618 | 0.6606 | 1.237 | 0.39949 | -0.35146 | -0.12453 | 0.36529 | 1.2448 | 0.15796 | 1.0405 | 0.63498 | 1.248 | 1.7007 | 1.3582 | 27.6377 | 2.2838
3 | 32.2689 | 0.93615 | -0.56496 | -0.12261 | -1.4401 | -1.2462 | -1.4 | -1.9103 | -1.9466 | -1.6904 | -1.8867 | -1.2169 | -2.0007 | -1.6277 | -2.1773 | -1.7286 | -1.53 | -1.9353 | -0.34537 | 6.1383 | -0.1055
4 | 1.2417 | -0.097902 | 0.17405 | 0.1594 | 0.042857 | 0.079301 | 0.018337 | 0.1959 | 0.22256 | 0.18197 | 0.2003 | 0.10467 | 0.30746 | 0.29123 | 0.27835 | 0.08506 | 0.075559 | 0.011551 | 0.079603 | 0.60361 | 0.18084
5 | -5.5964 | -0.23125 | -0.29314 | -0.21266 | -0.13268 | 0.011978 | -0.17786 | -0.073125 | 0.046375 | 0.073106 | -0.047363 | -0.20423 | 0.22981 | 0.21245 | 0.43509 | 0.11767 | 0.076614 | 0.16568 | -0.14776 | -3.317 | -0.27345
6 | 0.2418 | 0.093341 | -0.011314 | 0.04175 | 0.027254 | -0.017121 | 0.066779 | 0.087512 | 0.1436 | 0.06089 | 0.056528 | -0.039174 | -0.17382 | -0.030557 | -0.036816 | -0.023558 | -0.053881 | -0.022472 | 0.0022368 | 0.019901 | 0.016691
7 | 1.9224 | 0.61276 | 0.74877 | 0.77241 | 0.7317 | 0.61417 | 0.70562 | 0.63986 | 0.66126 | 0.57369 | 0.60497 | 0.63567 | 0.62814 | 0.57789 | 0.48963 | 0.66654 | 0.60204 | 0.64742 | 0.70623 | 0.44743 | 0.67728
8 | 0.091746 | 0.050815 | 0.075292 | 0.035576 | 0.0066247 | 0.028557 | 0.0014341 | -0.070699 | -0.053554 | -0.077324 | -0.061808 | 0.0098631 | -0.034374 | -0.081815 | 0.011655 | 0.019804 | 0.00078066 | -0.010055 | -0.014402 | 0.0035019 | -0.005729
9 | 0.26156 | 0.45728 | 0.41348 | 0.4226 | 0.41527 | 0.43391 | 0.41066 | 0.39552 | 0.40869 | 0.42388 | 0.42623 | 0.42884 | 0.41655 | 0.41378 | 0.42514 | 0.40514 | 0.42262 | 0.41772 | 0.42389 | 0.34827 | 0.41321
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.55167 | 0.64925 | 0.6458 | 0.64555 | 0.64342 | 0.651 | 0.64973 | 0.65546 | 0.65401 | 0.65104 | 0.64868 | 0.64627 | 0.65709 | 0.66021 | 0.66509 | 0.64705 | 0.65323 | 0.65159 | 0.63891 | 0.57604 | 0.64556
12 | 0.17167 | 0.047781 | 0.052906 | 0.0059922 | -0.015798 | -0.00078209 | -0.0014971 | -0.068851 | -0.079329 | -0.11632 | -0.11232 | -0.022386 | -0.060406 | -0.086548 | 0.00078347 | 0.013143 | 0.026285 | 0.0051792 | 0.0082704 | -0.093747 | -0.0018849 |  | ten times!
13 | 0.67843 | 0.45121 | 0.50026 | 0.51257 | 0.50773 | 0.43817 | 0.49439 | 0.47546 | 0.44151 | 0.42906 | 0.43406 | 0.48899 | 0.42783 | 0.41415 | 0.357 | 0.43307 | 0.42104 | 0.41087 | 0.4593 | 0.46522 | 0.045704 |  | ten times!
14 | 0.088919 | -0.015715 | -0.023816 | -0.033701 | -0.02887 | -0.038196 | -0.04335 | -0.062278 | -0.068356 | -0.073328 | -0.056579 | -0.037138 | -0.052673 | -0.070671 | -0.057354 | -0.0508 | -0.034081 | -0.035668 | -0.040194 | 0.031882 | -0.0039388 |  | ten times!
15 | 0.076986 | 0.37816 | 0.37798 | 0.36682 | 0.41174 | 0.44477 | 0.39109 | 0.33815 | 0.33119 | 0.37073 | 0.39401 | 0.38162 | 0.40537 | 0.44822 | 0.45742 | 0.46715 | 0.42456 | 0.38947 | 0.33507 | 0.45152 | 0.038748 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 38.721 | 2.5984 | -1.5604 | -4.6275 | -1.9669 | -4.1307 | -0.95421 | -2.0061 | -5.7904 | -4.4582 | -5.9435 | -5.2388 | -2.0719 | 0.81151 | 1.4042 | 3.2363 | 6.7309 | 7.8341 | 6.6858 | -20.8758 | 0.022025
2 | 0.3195 | -1.68 | -1.0894 | -1.3347 | -0.80318 | -0.54148 | -2.0413 | -1.6683 | -0.057794 | -1.0279 | -1.3182 | -0.8182 | -1.1189 | -1.5467 | -0.42 | -1.0052 | -1.1857 | -0.002476 | -0.61314 | -6.7653 | -1.1491
3 | -1.3006 | 0.32423 | -0.34338 | -0.2704 | -0.50442 | 0.25145 | 0.17865 | -0.18498 | -0.36228 | -0.081998 | 0.62144 | 0.083717 | 0.20742 | 0.23602 | 0.44088 | -0.04154 | 0.16321 | 0.27016 | -0.051844 | -0.78708 | -0.013289
4 | 0.23176 | 0.32407 | 0.21665 | 0.66275 | 0.93044 | 0.84189 | 0.85778 | 0.39976 | 0.17841 | 0.39585 | 0.39975 | 0.73368 | 0.41198 | 0.71741 | -0.0043031 | 0.010423 | -0.028814 | 0.22083 | -0.20822 | 0.10219 | 0.37877
5 | 1.9489 | 0.018783 | 0.11909 | 0.037009 | -0.12064 | -0.033054 | 0.11301 | 0.0059286 | 0.021597 | -0.17377 | 0.12215 | 0.040553 | 0.015906 | -0.018263 | 0.3138 | 0.2177 | 0.10709 | 0.09725 | -0.086871 | 0.031054 | 0.10023
6 | -0.34134 | -0.065595 | 0.0013939 | -0.046995 | -0.06523 | 0.0018383 | -0.07565 | -0.13067 | 0.10171 | 0.026817 | 0.077348 | -0.079838 | -0.15023 | -0.086826 | 0.10585 | 0.23363 | 0.19961 | 0.27365 | 0.11854 | 0.081302 | 0.015303
7 | 1.5086 | -0.0098908 | -0.054539 | 0.031168 | -0.02626 | 0.031582 | -0.029529 | -0.037471 | 0.14847 | 0.034098 | -0.050373 | -0.088812 | -0.11182 | -0.0095453 | 0.079942 | -0.022196 | 0.052186 | 0.066449 | -0.075771 | 0.050194 | 0.042361
8 | -0.0047801 | -0.0083962 | -0.0060554 | -0.015297 | 0.0062245 | 0.020985 | -0.0062198 | -0.016228 | 0.0022897 | 0.029191 | 0.018578 | 0.026253 | 0.046901 | 0.0075855 | -0.01079 | -0.018112 | 0.00086576 | -0.014183 | -0.042289 | -0.026933 | 0.00016065
9 | -0.24048 | -0.0032743 | -0.015981 | 0.0054187 | -0.018778 | -0.0090593 | 0.023318 | 0.011054 | -0.035635 | -0.03787 | 0.0074456 | 0.031882 | 0.01923 | 0.04531 | 0.067081 | 0.002085 | 0.0099597 | 0.036914 | 0.023967 | -0.014689 | 0.000963
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14975 | -0.056657 | -0.048253 | -0.044166 | -0.050853 | -0.045125 | -0.047218 | -0.043428 | -0.022249 | -0.043983 | -0.050716 | -0.053194 | -0.053726 | -0.047081 | -0.035871 | -0.039194 | -0.028975 | -0.030926 | -0.045441 | -0.020321 | -0.037297
12 | -0.010566 | -0.013669 | 0.026848 | 0.001056 | -0.00012383 | 0.016303 | -0.011562 | -0.017848 | 0.089349 | 0.053616 | 0.055544 | -0.015811 | -0.016673 | -0.013366 | 0.088779 | 0.1109 | 0.099005 | 0.16443 | 0.087068 | 0.013635 | 0.0036865 |  | ten times!
13 | -0.23615 | -0.0030348 | -0.04822 | -0.034238 | -0.068948 | -0.066332 | -0.045569 | -0.038056 | -0.031172 | -0.035208 | -0.0088785 | -0.028813 | -0.043405 | 0.027501 | 0.066021 | 0.004401 | -0.012788 | 0.0048315 | -0.036094 | -0.073905 | -0.0030004 |  | ten times!
14 | 0.12993 | -0.1376 | -0.1572 | -0.13603 | -0.11836 | -0.15514 | -0.13714 | -0.14315 | -0.16892 | -0.14658 | -0.15497 | -0.12949 | -0.13847 | -0.14597 | -0.21698 | -0.20431 | -0.20617 | -0.20596 | -0.18447 | -0.030453 | -0.014788 |  | ten times!
15 | 0.028509 | 0.004738 | -0.069029 | -0.086637 | -0.14526 | -0.14253 | -0.11822 | -0.11642 | -0.076434 | -0.1124 | -0.081659 | -0.092667 | -0.07111 | -0.0080641 | 0.047893 | 0.0092721 | -0.070493 | -0.1442 | -0.1386 | -0.22258 | -0.0079617 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 391.9435 | 703.1351 | 703.5884 | 703.6321 | 703.562 | 703.5688 | 703.4837 | 703.3785 | 703.4148 | 703.4457 | 703.424 | 703.4666 | 703.443 | 703.4511 | 703.5287 | 703.4986 | 703.3847 | 703.5194 | 703.178 | 402.3598
2 | -30.4675 | -1.6008 | -0.42918 | 0.26291 | -0.23424 | 0.4064 | -0.56359 | -0.12321 | 0.29111 | -0.11538 | -0.87924 | -1.0535 | -1.0386 | -0.77378 | -0.9257 | -1.2907 | -1.1599 | -0.20865 | -1.1041 | -26.6632 | -2.2356
3 | 31.3989 | 0.13707 | -1.2422 | -1.5393 | -2.474 | -3.4404 | -2.5402 | -2.9951 | -3.1791 | -2.4633 | -1.9341 | -1.6861 | -2.1245 | -2.4992 | -1.5842 | -1.7125 | -1.4871 | -2.2399 | -0.83403 | 4.4128 | -0.82705
4 | -0.46488 | -0.12896 | -0.28492 | -0.2041 | -0.20493 | -0.39737 | -0.16451 | -0.214 | -0.1433 | -0.10819 | -0.21396 | -0.09447 | -0.216 | -0.30775 | 0.036555 | -0.075837 | -0.077647 | -0.34727 | -0.15686 | -1.2732 | -0.22431
5 | -5.966 | -0.24291 | -0.074004 | 0.24478 | 0.39095 | 0.85952 | 0.23867 | 0.62244 | 0.59458 | 0.46602 | 0.502 | 0.38592 | 0.33508 | 0.63495 | 0.65743 | 0.32802 | 0.21303 | 0.3371 | 0.046204 | -3.7396 | 0.056306
6 | -0.02891 | -0.083724 | -0.14222 | -0.13474 | -0.091713 | -0.26894 | -0.14968 | -0.14642 | -0.21922 | -0.14705 | -0.034989 | 0.0072751 | -0.093164 | -0.095857 | -0.0057932 | -0.040072 | -0.037029 | 0.014432 | -0.10922 | -0.13523 | -0.09784
7 | 1.7968 | 0.65873 | 0.7467 | 0.64952 | 0.6319 | 0.43431 | 0.53427 | 0.46626 | 0.54632 | 0.52036 | 0.51831 | 0.60738 | 0.55502 | 0.55656 | 0.5062 | 0.64934 | 0.64705 | 0.64396 | 0.69015 | 0.26117 | 0.6123
8 | -0.060164 | -0.012564 | 0.034519 | 0.083134 | 0.051175 | 0.14949 | 0.057603 | 0.091293 | 0.071252 | 0.049038 | 0.012599 | -0.0056346 | 0.02337 | 0.064994 | 0.012405 | -0.013622 | 0.017217 | 0.05203 | 0.035372 | 0.074244 | 0.040915
9 | 0.29853 | 0.45788 | 0.41823 | 0.41389 | 0.4044 | 0.42998 | 0.42862 | 0.4332 | 0.41848 | 0.43544 | 0.44047 | 0.43862 | 0.42979 | 0.43772 | 0.44267 | 0.42139 | 0.42877 | 0.42954 | 0.43488 | 0.34073 | 0.42371
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.53173 | 0.64442 | 0.64724 | 0.65661 | 0.65723 | 0.66745 | 0.66311 | 0.66783 | 0.66035 | 0.66172 | 0.66238 | 0.65831 | 0.66354 | 0.65814 | 0.66402 | 0.65485 | 0.65464 | 0.65567 | 0.64636 | 0.56835 | 0.65163
12 | -0.06096 | -0.0024217 | 0.0167 | 0.047054 | 0.015724 | 0.1956 | 0.051404 | 0.091745 | 0.09155 | 0.052473 | -0.0078841 | -0.0094846 | 0.018044 | 0.089308 | 0.017029 | 0.0023364 | 0.011307 | 0.073621 | 0.014794 | -0.023979 | 0.0037705 |  | ten times!
13 | 0.594 | 0.48872 | 0.5286 | 0.45683 | 0.47417 | 0.34844 | 0.42234 | 0.35445 | 0.38526 | 0.38557 | 0.39222 | 0.43137 | 0.41681 | 0.37535 | 0.38857 | 0.44102 | 0.44542 | 0.4124 | 0.47489 | 0.32522 | 0.04255 |  | ten times!
14 | 0.039484 | -0.072504 | -0.064332 | -0.065998 | -0.087243 | -0.079237 | -0.069992 | -0.056013 | -0.063484 | -0.063558 | -0.0794 | -0.078905 | -0.076812 | -0.06666 | -0.06077 | -0.065006 | -0.058997 | -0.051684 | -0.073647 | -0.040473 | -0.0064582 |  | ten times!
15 | -0.068221 | 0.26986 | 0.24469 | 0.30127 | 0.29851 | 0.35522 | 0.26477 | 0.30336 | 0.32282 | 0.29134 | 0.30172 | 0.24894 | 0.25662 | 0.3072 | 0.32856 | 0.31574 | 0.27579 | 0.26946 | 0.25506 | 0.1842 | 0.027593 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 44.4778 | 7.4704 | -1.4861 | -4.8697 | 0.25908 | 3.604 | 2.0875 | -1.3717 | -8.4042 | -8.2043 | -10.2556 | -8.9216 | -3.9659 | -0.051619 | 0.20747 | 1.5202 | 1.5287 | 6.5141 | 3.5913 | -7.7285 | -0.020427
2 | 3.6858 | 0.65655 | 0.85821 | -0.21252 | -0.33133 | -0.69144 | 0.44561 | 0.49433 | 0.62322 | 0.70834 | 0.77877 | 0.84451 | 0.82671 | 0.41202 | 0.60534 | 1.0247 | 1.0681 | 1.9677 | 0.9556 | -5.2308 | 0.52776
3 | -2.2937 | -0.41908 | -1.0142 | -0.92762 | -1.1213 | -0.88773 | -0.50339 | -0.0040192 | -0.10635 | -0.62786 | 0.11399 | -0.16967 | -0.37825 | -0.0648 | 0.44021 | -0.22889 | 0.056572 | 0.47096 | -0.4273 | -0.92894 | -0.39774
4 | 0.9371 | 0.69509 | 0.8049 | 0.74927 | 0.76762 | 1.7953 | 0.36413 | 0.71812 | 1.0436 | 1.0687 | 0.86854 | 1.3009 | 0.91957 | 1.2651 | 0.80988 | 0.30254 | 0.44269 | 0.083559 | 0.15405 | 0.69278 | 0.78793
5 | 1.3581 | -0.10192 | -0.064981 | -0.13443 | -0.33658 | -0.30328 | -0.13525 | -0.12655 | -0.037062 | -0.30289 | 0.034499 | 0.013735 | -0.12303 | -0.0085154 | 0.2509 | 0.2395 | 0.13511 | 0.2063 | -0.029502 | -0.13345 | -0.0075079
6 | -0.19981 | -0.20892 | -0.12443 | -0.077793 | -0.13152 | -0.28462 | -0.077699 | -0.19704 | -0.15003 | -0.082617 | -0.088605 | -0.26594 | -0.16157 | -0.23778 | -0.051374 | 0.031738 | 0.042688 | 0.081046 | -0.0053647 | -0.084071 | -0.11235
7 | 1.6408 | 0.05786 | -0.022153 | 0.0026863 | 0.0034822 | 0.15999 | -0.036484 | 0.01052 | 0.065462 | -0.031809 | -0.087211 | -0.0086661 | 0.0093907 | 0.07746 | 0.065553 | -0.021897 | -0.0095074 | -0.01208 | -0.067198 | 0.007768 | 0.056148
8 | 0.084331 | 0.026868 | 0.039261 | 0.059877 | 0.045646 | 0.053378 | 0.043249 | 0.053211 | 0.049106 | 0.050359 | 0.06888 | 0.081698 | 0.099916 | 0.074171 | 0.061589 | 0.017151 | 0.034116 | 0.013602 | 0.02651 | 0.032918 | 0.050416
9 | -0.26864 | -0.023698 | -0.015153 | -0.019451 | -0.024347 | -0.045718 | 0.007861 | 0.040899 | 0.043754 | 0.034804 | 0.055616 | 0.048444 | 0.039052 | 0.047913 | 0.082657 | 0.041747 | 0.055284 | 0.11595 | 0.051193 | 0.017251 | 0.020752
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.19324 | 0.0017249 | 0.0069155 | 0.011592 | 0.010863 | 0.01726 | -0.0079133 | -0.0045789 | -0.0024143 | -0.020295 | -0.025169 | -0.008223 | -0.0053801 | 0.0057307 | 0.0061092 | 0.0048516 | 0.0066619 | 0.010244 | 0.0045027 | 0.020428 | 0.0068923
12 | 0.0026351 | -0.047533 | 0.015613 | 0.018773 | -0.001408 | -0.022396 | 0.0050014 | -0.0282 | -0.019094 | 0.0019167 | -0.012304 | -0.057095 | -0.030571 | -0.021055 | 0.042613 | 0.068277 | 0.0705 | 0.090803 | 0.039235 | -0.015413 | 0.00055284 |  | ten times!
13 | -0.16742 | 0.0031726 | -0.044792 | -0.060034 | -0.059709 | -0.084559 | -0.037226 | 0.021885 | 0.025496 | -0.0035813 | 0.023938 | 0.0047136 | 0.0080586 | 0.035826 | 0.070951 | 0.025417 | 0.049029 | 0.079584 | 0.02757 | 0.019674 | 1.9111e-05 |  | ten times!
14 | 0.068291 | -0.045645 | -0.013627 | -0.029386 | -0.026237 | -0.024862 | -0.061774 | -0.063747 | -0.056581 | -0.047327 | -0.036366 | 0.0025757 | -0.023125 | -0.030204 | -0.063058 | -0.081766 | -0.086534 | -0.086559 | -0.066479 | 0.04208 | -0.00407 |  | ten times!
15 | 0.0028238 | -0.045343 | -0.085892 | -0.063254 | -0.096019 | -0.079914 | -0.064719 | -0.044782 | -0.016543 | -0.051907 | -0.039598 | -0.061967 | -0.07272 | -0.038942 | -0.047833 | -0.088259 | -0.057291 | -0.072241 | -0.058965 | -0.14057 | -0.0060904 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -0.0003837 | -3.5124e-06 | 7.2498e-05 | -4.2441e-05 | 0.00012123 | 0.00018261 | -1.019e-05 | 2.5755e-05 | -5.0681e-05 | 3.1864e-05 | 4.0247e-05 | 0.00018104 | 8.7131e-05 | 0.00015192 | 4.2946e-05 | -3.6616e-05 | -5.3594e-06 | -0.00040473 |  | 703.8547
|  | -0.0004479 | 0.00019649 | 0.00025861 | 0.00015896 | 0.00016863 | 4.7654e-05 | -0.00010189 | -5.0292e-05 | -6.3655e-06 | -3.7213e-05 | 2.3345e-05 | -1.0204e-05 | 1.311e-06 | 0.00011162 | 6.8835e-05 | -9.3081e-05 | 9.8404e-05 | -0.00038692 |  | 703.4502
Check Computation for Aperture1
Magnetic Length | -0.056573
dB/B Heads CS+NCS | -4282.4014
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.5103 | -4337.591 | -3.837 | -0.035124 | 0.72498 | -0.42441 | 1.2123 | 1.8261 | -0.1019 | 0.25755 | -0.50681 | 0.31864 | 0.40247 | 1.8104 | 0.87131 | 1.5192 | 0.42946 | -0.36616 | -0.053594 | -4.0473 | -4227.2117
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 3.9187 | 0.31135 | -0.1033 | -0.40995 | -0.14382 | -0.36027 | -0.041728 | -0.14664 | -0.52608 | -0.39358 | -0.54239 | -0.47113 | -0.15477 | 0.13123 | 0.19216 | 0.37467 | 0.72561 | 0.83641 | 0.72226 | -2.1026
db2 | 1.1483 | 26.4219 | 0.63803 | 0.15594 | 0.05956 | 0.0036512 | -0.21435 | -0.098574 | -0.077329 | -0.27659 | -0.12972 | -0.16832 | -0.19743 | -0.053952 | -0.064808 | 0.038719 | -0.01142 | -0.014683 | -0.033247 | 0.44453 | 28.2101
db3 | 4.6701 | 11.2495 | -0.12135 | -0.15308 | -0.031933 | 0.0078736 | 0.016515 | 0.041278 | 0.035485 | 0.12667 | 0.10781 | 0.084926 | 0.078077 | -0.0037463 | 0.047485 | -0.082956 | -0.030684 | 0.01845 | -0.013912 | -0.12691 | 9.7639
db4 | 0.014638 | 2.3399 | 0.015686 | 0.040794 | 0.011761 | 0.0038044 | -0.011705 | -0.022481 | -0.0098625 | -0.0027107 | -0.0040406 | -0.0017404 | -0.0095641 | -0.013545 | -0.011874 | -0.011281 | 0.009516 | -0.0094473 | 0.006154 | 0.020536 | 0.73207
db5 | 0.03515 | 0.48746 | 0.0061295 | -0.0036264 | 0.00013513 | 0.0070826 | -0.0024174 | 0.00083774 | 0.00041671 | 0.0097103 | 0.0093501 | 0.00075146 | 0.0025223 | 0.0010295 | 0.0089528 | -0.011117 | -0.016574 | -0.0025697 | -0.0065259 | -0.0040878 | 0.013491
db6 | -0.023351 | 0.30875 | 0.0042205 | 0.0040199 | 0.010121 | 0.0037403 | 0.0019834 | 0.0023469 | 0.0049222 | -0.0026606 | -0.0036739 | -0.0066881 | 0.0023373 | 0.0011327 | 0.00085402 | -0.0060826 | -0.011349 | -0.0022911 | -0.0041508 | 0.0012179 | -0.019646
db7 | 0.0019394 | 0.057896 | 0.0015308 | -0.0015385 | 0.0010161 | 0.0020166 | 0.0037936 | 0.0021376 | 0.0011551 | -0.0038768 | -0.0021057 | -0.0001951 | 0.00082358 | 0.0046195 | 0.00034934 | -0.0041471 | -0.0045426 | 0.0013201 | -0.0038125 | 0.0014559 | 0.03179
db8 | -0.0033699 | 0.085605 | 0.00051955 | 0.0024125 | 0.00027562 | 3.16e-05 | -0.00068467 | -0.00029969 | 0.00015735 | -0.00095378 | -0.0026597 | 0.0023895 | 0.0017818 | -0.0048145 | -0.0015715 | -0.0021624 | 0.00070517 | 0.0026372 | -0.00062023 | 0.0028563 | 0.0021029
db9 | 0.02237 | 0.033939 | -0.0013249 | -0.0043087 | 0.00016549 | -0.0021295 | 0.003597 | -0.00085295 | -0.0039866 | 0.0043551 | 0.0083482 | -7.0131e-05 | 0.0036675 | -0.00047229 | -0.0025986 | 0.0049952 | -0.00287 | 0.0021463 | -0.0015561 | -0.007105 | 0.047649
db10 | -2.5164e-18 | 0 | -9.2444e-18 | -9.5302e-19 | 7.8171e-19 | 8.3968e-18 | 2.5164e-18 | -1.5125e-17 | 1.7814e-18 | -9.5302e-19 | 3.3838e-18 | 2.5394e-18 | 7.8171e-19 | -9.5302e-19 | -3.4883e-17 | 1.216e-17 | -3.364e-18 | 1.4223e-17 | 2.5164e-18 | 1.6394e-17 | 2.6062e-17
db11 | 0.016372 | 0.0070515 | 0.0011895 | -0.00066067 | 0.00010991 | 0.00057096 | -0.0018142 | 0.00053831 | -0.00089797 | 0.00078482 | 1.6745e-05 | -0.00041587 | -0.0005579 | 0.00064271 | 0.001103 | -0.0010144 | -0.00089106 | 0.00029122 | 0.00044794 | 0.00055703 | 0.038218
db12 | -0.00067403 | 0.0041365 | 0.0011499 | 0.00074808 | 0.0006227 | -0.00086826 | -0.00091647 | -0.00031828 | -0.0014 | 0.00085657 | -0.0009642 | 0.00070164 | 0.00045122 | 0.00013185 | -0.00037136 | 0.00037237 | 0.0001096 | 0.0014687 | -0.0010049 | -0.00076924 | -0.0044863
db13 | 0.0014125 | 0.0046809 | -0.0016204 | -0.00039702 | 0.00024905 | 0.00028353 | 0.00015051 | -0.00032453 | 0.0010262 | 0.00071395 | 0.0007421 | 0.00064982 | 0.00064585 | 6.093e-05 | 0.00047338 | -2.5534e-05 | -0.00025287 | -0.00038069 | -0.0012336 | -0.00076068 | 0.0071434
db14 | -0.0058437 | 0.0021359 | 0.0026784 | 0.0012297 | 0.00013512 | 0.0011011 | 0.0006359 | 0.00021199 | 0.00067462 | 0.00026248 | -0.00027485 | -0.0011206 | -0.00077402 | -0.0014871 | -0.00068027 | -0.001383 | -0.00097477 | 0.00014109 | -0.00081989 | 0.00044399 | 0.00065187
db15 | 0.0062737 | 0.010531 | -0.0021402 | 0.0012023 | -0.0010306 | 0.0028162 | 0.0015863 | -0.0015351 | -0.0022375 | -0.0036123 | -0.003639 | -0.0013536 | -0.0028603 | -0.0011654 | 0.0030619 | 0.004122 | 0.0082582 | 0.00037489 | -0.0011066 | -0.00074098 | 0.017864
da2 | -0.20336 | -12.3733 | 0.073776 | 0.015344 | -0.074495 | 0.033894 | 0.28537 | 0.087064 | 0.16143 | 0.040144 | -0.14403 | -0.075791 | -0.020505 | -0.11 | 0.0069534 | -0.0046932 | -0.026733 | -0.13272 | 0.085717 | -0.20073 | 0.26485
da3 | 0.030818 | -1.2584 | 0.21343 | 0.00049806 | 0.018878 | 0.1134 | 0.27204 | -0.061391 | 0.083054 | 0.090152 | -0.06082 | -0.052253 | -0.028507 | -0.16814 | -0.054765 | -0.080356 | -0.02115 | -0.087904 | -0.07469 | -0.10147 | -0.29023
da4 | 0.011618 | -1.9467 | 0.035772 | 0.01138 | -0.0087053 | 0.0080202 | 0.045998 | -0.021356 | 0.05281 | 0.0073585 | -0.026739 | -0.022908 | -0.010001 | -0.047267 | 0.0056755 | -0.0031922 | 0.0095608 | -0.014169 | -0.0027237 | -0.019515 | 0.042776
da5 | 0.0089635 | -0.062766 | 0.012178 | -0.0055748 | 0.002143 | 0.0094834 | 0.025113 | -0.011904 | 0.010177 | 0.0098965 | -0.0034944 | -0.0054077 | -0.0063471 | -0.0063922 | -0.001594 | 0.0025638 | -0.0045827 | -0.0034017 | 0.0053583 | -0.028214 | -0.12936
da6 | 0.0044565 | -0.4196 | -0.0026644 | -0.0033887 | 0.0015648 | -0.0016396 | 0.0068075 | -0.00059276 | 0.0034148 | 0.0082663 | -0.0016984 | 0.013855 | 3.1367e-05 | -0.011058 | -0.0052772 | -0.0043379 | 0.0019119 | -0.00086715 | 0.00055647 | -0.0048834 | 0.005945
da7 | 0.01682 | 0.13165 | -0.0095197 | -0.0014209 | 0.00097939 | 0.00083957 | 0.0012991 | 0.0091673 | 0.0016047 | -0.0023028 | 0.00036846 | 0.00026041 | -0.0014071 | -0.006229 | -0.0031445 | -0.0020023 | 0.0014642 | 0.0013583 | 0.0032832 | 0.0054014 | 0.0012185
da8 | -0.0007481 | -0.016259 | -0.00092264 | 0.00045861 | 0.0029044 | -0.0011756 | 0.003766 | 0.00027581 | -0.0046508 | 0.0010075 | 0.0011402 | -0.00042786 | 0.0013398 | -0.00050373 | -0.0046677 | 0.0024748 | 0.0028291 | 0.0024066 | -0.0011059 | -0.0051485 | -0.0052083
da9 | 0.0098195 | 0.028654 | -0.010586 | 6.4281e-05 | -0.0013345 | -0.0025803 | -0.0031697 | 0.0081487 | 0.0028948 | -0.0073458 | 0.0012119 | -0.0048963 | 0.0010304 | 0.0027319 | 0.0013857 | -0.00039109 | 0.0038631 | -0.0003346 | 0.0051974 | 0.0041103 | -0.0035685
da10 | 1.467e-19 | 1.0211e-17 | 1.1397e-17 | 7.2066e-19 | -3.3515e-18 | 8.2623e-18 | 3.852e-18 | -2.2196e-18 | -4.9686e-18 | -1.9905e-17 | -1.0876e-18 | 2.9552e-18 | -1.3522e-18 | -2.4787e-18 | 6.7922e-18 | -1.467e-19 | -1.467e-19 | -1.4992e-18 | -2.5577e-18 | 5.7337e-18 | 6.7695e-17
da11 | 0.0024134 | 0.025973 | -0.007924 | -0.0031641 | -0.0044803 | -0.0024777 | 0.0010201 | 0.0069113 | 0.0049913 | 0.0022045 | -0.00048366 | -0.0027135 | -0.003949 | 0.00044356 | -0.00043906 | 0.00054013 | 0.0014912 | 0.00070497 | 0.0030922 | 0.0042321 | 0.023449
da12 | -0.00041521 | 0.005992 | 0.00027535 | -0.00034744 | 0.00021843 | -1.3766e-05 | 0.0016727 | -0.00013649 | 0.00042708 | -0.00044503 | 0.00021149 | 0.0006996 | 0.00052047 | -0.00040136 | -0.0015649 | -0.0021099 | -0.0014695 | 0.00043129 | 0.00085455 | 0.0011774 | 0.0010082
da13 | 0.0028897 | -0.0015267 | -0.0010573 | -0.0005994 | -0.0011852 | -0.00085248 | -0.00053317 | 0.0010261 | 0.00098883 | -0.0004524 | 0.00050945 | -0.00067596 | 0.00083956 | -0.0014545 | 0.00068795 | 0.00063957 | 0.0015555 | -0.00049409 | 0.00051565 | 0.00054185 | -0.0022286
da14 | -0.0049863 | 0.0067012 | -7.5993e-06 | 0.00012829 | 0.00044362 | -0.00041878 | 0.00049629 | -2.5432e-05 | 0.0002143 | 0.00083973 | 0.0014178 | -0.00033305 | 0.00026746 | -0.00016049 | 0.00016736 | 0.0009247 | -0.0011368 | -0.00047237 | -0.00093 | -0.001415 | 4.7073e-05
da15 | -0.0055662 | -0.012034 | -0.0011018 | -0.0031962 | -0.00065627 | -0.0051308 | -0.0019166 | -0.0013141 | -0.00079309 | 0.0058386 | -0.0044012 | 0.0023344 | 0.003602 | 0.00044747 | 0.00050371 | -0.0020347 | 0.0053815 | 0.0044114 | -0.0003 | -0.0016744 | -0.014338
Check Computation for Aperture2
Magnetic Length | -0.059653
dB/B Heads CS+NCS | -4354.232
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.1963 | -4428.2692 | -4.479 | 1.9649 | 2.5861 | 1.5896 | 1.6863 | 0.47654 | -1.0189 | -0.50292 | -0.063655 | -0.37213 | 0.23345 | -0.10204 | 0.01311 | 1.1162 | 0.68835 | -0.93081 | 0.98404 | -3.8692 | -4280.1948
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 4.5725 | 0.86329 | -0.033204 | -0.37186 | 0.14143 | 0.47607 | 0.32374 | -0.022052 | -0.7247 | -0.70526 | -0.91028 | -0.77804 | -0.28246 | 0.10992 | 0.13649 | 0.26846 | 0.26825 | 0.76555 | 0.47467 | -0.5961
db2 | 1.3847 | 25.7419 | 0.33078 | -0.014376 | -0.0064746 | -0.0046608 | -0.16526 | -0.11346 | 0.0048726 | -0.04036 | -0.24051 | -0.10066 | 0.022947 | -0.13241 | -0.045047 | -0.065856 | 0.14793 | 0.11806 | 0.051832 | 0.25266 | 26.6632
db3 | 4.5499 | -11.3741 | -0.073197 | -0.050385 | 0.004952 | 0.0064336 | 0.049994 | 0.062326 | 0.00965 | 0.039887 | 0.027941 | -0.0035407 | 0.003823 | 0.030671 | 0.0041847 | -0.011853 | -0.027662 | -0.017245 | -0.027118 | -0.028863 | 9.4607
db4 | 0.047661 | -0.39164 | 0.015242 | 0.01444 | 0.030381 | 0.018513 | 0.033434 | -0.015711 | -0.023347 | -0.015546 | -0.018793 | -0.00032954 | 0.0015991 | -0.013303 | -0.0097493 | -0.016975 | -0.001233 | -0.0065987 | 0.0017181 | 0.0062573 | 1.2732
db5 | 0.031063 | -0.38969 | 0.0096573 | 0.0052185 | 0.01172 | 0.0074005 | 0.016788 | -0.0015403 | -0.0023072 | -0.010931 | 0.0075419 | -0.010914 | 0.0014035 | 0.00017541 | -0.0020915 | -0.010737 | -0.012319 | -0.0024697 | -0.0060197 | -0.00057543 | -0.06702
db6 | -0.024204 | -0.33666 | 0.00019137 | 0.001956 | 0.0058346 | -0.0001041 | -0.010553 | 0.0009251 | 0.00039039 | -0.0018591 | -0.0016398 | 0.00041042 | -0.0027416 | -0.0045548 | -0.0063315 | 0.00035707 | 0.0055852 | 0.0072874 | 0.0020221 | 0.0028248 | 0.0028244
db7 | -0.0002895 | -0.076907 | 0.0018149 | 0.0051001 | -0.00088096 | 0.00031392 | -0.0062287 | 0.00090708 | 0.00051443 | -0.0043026 | 0.0026023 | 0.0032287 | 0.0029213 | 0.0012209 | -0.0058218 | -0.0040404 | 0.0016518 | -0.0010554 | -0.003763 | 0.0058173 | 0.011166
db8 | -0.0024946 | -0.069721 | 0.0032241 | 0.0022079 | 0.0029302 | -0.0015577 | -0.0031502 | -0.00049061 | -0.00015205 | -0.0014527 | -7.4247e-05 | -0.00093328 | 0.0016618 | 0.0014397 | 0.00039861 | -0.00049818 | 0.0015999 | -0.00087932 | -0.0018206 | -0.0024534 | -0.0035943
db9 | 0.023917 | -0.044044 | 0.00083415 | -0.0004086 | 0.00063795 | 0.0027137 | 0.0055582 | 0.0020255 | -0.00098224 | -0.0033476 | 0.0053917 | -0.0027716 | 0.0031726 | 0.0012373 | -0.0020801 | 0.0010221 | -0.00064923 | -0.0029635 | -0.0035843 | -0.0058059 | 0.032149
db10 | 1.7524e-18 | -1.9269e-17 | -2.5274e-17 | -4.801e-18 | 1.5333e-18 | -1.9361e-18 | 1.0008e-17 | 1.7171e-18 | -1.2231e-18 | -4.6926e-18 | -1.3513e-17 | 5.1865e-18 | 2.1769e-17 | 1.6947e-17 | -1.7524e-18 | -3.1049e-18 | -3.2225e-18 | -1.7662e-20 | -7.6328e-18 | 1.0008e-17 | 0
db11 | 0.011627 | -0.017558 | -0.00083503 | 0.00036743 | -0.00020354 | -0.00082646 | -0.001046 | -5.0504e-05 | 0.00054253 | -0.00043918 | 0.00065234 | -0.00068771 | 0.0016752 | 0.0012361 | -0.00037993 | -0.001313 | 0.00011634 | -0.00024446 | -0.00048731 | 0.0019231 | 0.022348
db12 | 0.00015883 | -0.0023759 | 0.00089088 | -2.2096e-05 | 0.00073219 | 7.6272e-05 | -0.0011194 | 0.00073598 | -0.0010222 | -0.00035341 | -0.00084653 | 5.1509e-05 | -0.00089927 | -0.0005674 | -2.5549e-05 | 0.00027627 | -0.000227 | 0.00075452 | 0.00042149 | 0.0011437 | 0.0023979
db13 | 0.0023466 | 0.004235 | 0.00059263 | 0.00058996 | -0.00027699 | 0.000182 | 0.0014411 | 0.00026436 | -0.0001787 | -0.00019435 | -0.00069135 | -0.0009851 | 0.00055102 | 9.117e-05 | -0.00048143 | -0.00047847 | -0.00021137 | -0.00046278 | -0.00058331 | 0.00083161 | -0.00086474
db14 | 0.0021178 | -0.0010819 | 4.5419e-05 | -0.00069004 | 0.00034663 | -0.00026147 | 0.001604 | 0.00017812 | 0.00037368 | 0.00068115 | -0.00020565 | 0.00025461 | -0.00029837 | -0.00013518 | -0.00010949 | -0.00047017 | 0.00010741 | -0.00020515 | -0.00072335 | -0.00049213 | 0.0040473
db15 | -0.0072631 | -0.0043244 | 0.0020906 | -0.0023984 | 0.0015425 | 0.00069815 | 0.0020055 | -0.0025685 | -0.00026498 | -0.00054169 | -0.0025761 | -0.0010839 | 0.0015199 | -0.0022609 | 0.0011352 | -0.00084259 | 0.0028923 | 0.0012227 | -0.00116 | 0.00059019 | -0.0066348
da2 | -0.40963 | -10.963 | 0.18545 | -0.023123 | -0.17862 | 0.016445 | 0.11065 | -0.036386 | -0.021029 | -0.13802 | -0.088003 | -0.091192 | 0.066913 | 0.042691 | 0.0050588 | 0.028856 | 0.03729 | -0.075662 | 0.04981 | 0.10886 | -0.99778
da3 | 0.20087 | -0.36675 | -0.012633 | 0.076267 | 0.11844 | 0.074757 | -0.19764 | -0.092837 | 0.092541 | 0.1109 | 0.031467 | 0.084044 | 0.02987 | -0.085295 | -0.029902 | -0.10065 | -0.05404 | 0.030344 | -0.045162 | -0.030476 | 0.39219
da4 | -0.068394 | -1.0156 | 0.020798 | 0.0011198 | -0.032557 | 0.0027076 | 0.01356 | 0.040269 | -0.030098 | -0.044575 | -0.024613 | -0.027285 | -0.0029683 | 0.010459 | -0.013945 | 0.007928 | 0.029401 | 0.00012168 | 0.014691 | 0.034986 | 0.050474
da5 | 0.019187 | -0.07419 | -0.01168 | -0.001887 | 0.010162 | 0.0034715 | 0.019472 | -0.011849 | 0.0056796 | 0.018137 | -0.010102 | 0.013415 | -0.0032835 | -0.0058457 | 5.3317e-05 | -0.0041443 | 0.0012546 | -0.0052311 | -0.0029775 | -0.014644 | -0.11377
da6 | -0.0025174 | -0.10205 | -0.0076117 | 0.00015094 | -0.0053015 | -0.0022471 | -0.016481 | 0.0096315 | -0.0027703 | -0.0057823 | 0.0078692 | 0.00075508 | 0.0042112 | 0.0071051 | -0.0012071 | 0.0065637 | -0.00079715 | 0.00060775 | 0.0052339 | 6.9544e-05 | 0.022151
da7 | 0.013485 | 0.08986 | -0.0059427 | 0.00042807 | 0.0030191 | -0.0023041 | 0.0040604 | 0.0025515 | 9.1646e-05 | -0.0026202 | 4.4854e-05 | -0.0070799 | -0.0069228 | 0.00025604 | -0.003405 | -0.0014183 | 0.0022427 | 0.0037235 | 0.0056174 | 0.007658 | 0.0012418
da8 | 0.0033651 | 0.021758 | -0.0017162 | -0.002026 | 0.0018509 | 0.00028666 | 0.0063695 | 0.0026357 | 0.00084619 | 0.0014554 | -0.00044968 | -0.00099216 | -0.0030239 | -0.0021264 | -0.00020846 | 0.00071447 | -0.0019581 | -0.0028588 | 0.00036858 | 0.00083235 | 0.0036403
da9 | 0.0092115 | 0.0078342 | -0.00035043 | -0.0039351 | 0.0021202 | -0.00049953 | 0.0091282 | 0.006986 | -0.00053586 | -0.0039823 | -0.0032516 | -0.0032314 | -0.0094232 | 0.0020707 | -0.0013171 | 0.0019467 | 0.0012923 | -0.0015449 | 0.0032405 | 0.0012868 | 0.017251
da10 | -2.0822e-18 | 5.697e-18 | 6.1214e-19 | -5.5329e-18 | 1.008e-17 | -4.8567e-18 | 2.8173e-18 | 9.0211e-18 | 3.4752e-19 | 1.8176e-18 | 9.0211e-18 | 1.553e-18 | -4.8567e-18 | -8.8553e-18 | -1.3872e-18 | 3.4752e-19 | -1.1796e-17 | -1.2917e-18 | 6.1214e-19 | 2.3469e-18 | 0
da11 | 0.025017 | 0.016531 | -0.0056627 | -0.0028278 | 0.0044935 | -0.0027619 | 0.018577 | 0.0062482 | 0.0035745 | -0.0022276 | -0.0026709 | -0.0060578 | -0.0074399 | -0.0015215 | -0.0010786 | -0.0026655 | -0.00010056 | -0.0025091 | 0.002451 | 0.002179 | 0.020428
da12 | 5.3039e-05 | 0.0046528 | -0.00044868 | -0.0001853 | -0.0010684 | -0.00044172 | 0.0018761 | -0.00025596 | 0.00025482 | -0.00023357 | -0.00071993 | -0.00016976 | -0.0015605 | 0.00062889 | 0.0010549 | 0.0011475 | -8.9016e-05 | -0.0010767 | 0.00039343 | 0.00089392 | -0.0015413
da13 | 0.0018158 | 0.00065548 | -0.00065387 | -0.00034736 | 0.0006959 | -0.0010845 | 0.00012224 | 0.00129 | 0.0011664 | 3.4098e-05 | -0.00097382 | -0.00032093 | -0.0014411 | 0.00082504 | -0.00033296 | -0.00019236 | -0.00022232 | 0.00026405 | 0.00048356 | 0.00068793 | 0.00049648
da14 | 0.0054879 | 0.0024121 | -0.00090793 | -0.00039689 | -0.0002496 | -1.8977e-06 | 0.00054868 | -0.00030269 | 0.00033441 | 0.00036625 | 0.00036445 | 9.4365e-05 | 0.0008494 | 0.0005018 | -0.00010433 | 0.00075874 | -0.00041269 | -0.0012502 | 0.00034873 | -0.00054061 | 0.004208
da15 | 0.00057533 | -0.0013958 | -0.004063 | -0.0018311 | 0.0013109 | -0.0019361 | 0.0019347 | 0.0029313 | 0.00086883 | 0.0033165 | -0.0048912 | -0.00051893 | 0.0051587 | 0.0027362 | 0.00090545 | -0.0047693 | -0.0007935 | 0.0018786 | 0.00018032 | -0.0024184 | -0.014057
Field Colinearity | -0.046612
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 3.8663 | 0.25887 | -0.15578 | -0.46243 | -0.1963 | -0.41275 | -0.094205 | -0.19912 | -0.57856 | -0.44606 | -0.59487 | -0.52361 | -0.20725 | 0.078754 | 0.13968 | 0.32219 | 0.67313 | 0.78394 | 0.66978 | -2.1551
Twist (m2 mrad) | 28.997 | 3.6889 | -2.1031 | -5.8959 | -2.3556 | -4.6434 | -0.98915 | -1.9414 | -5.2071 | -3.68 | -4.4615 | -3.5344 | -1.2435 | 0.41346 | 0.62855 | 1.2082 | 2.0194 | 1.7639 | 1.0047 | -0.80817
Twist Integral (I0) (m2 rad) | 0.0021457
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.4572 | 0.74796 | -0.14853 | -0.48719 | 0.026097 | 0.36074 | 0.20841 | -0.13738 | -0.84003 | -0.82059 | -1.0256 | -0.89337 | -0.39779 | -0.005406 | 0.021157 | 0.15313 | 0.15292 | 0.65022 | 0.35934 | -0.71143
Twist (m2 mrad) | 33.4286 | 10.6585 | -2.0052 | -6.2117 | 0.31316 | 4.0583 | 2.1883 | -1.3395 | -7.5602 | -6.7699 | -7.6921 | -6.0302 | -2.3867 | -0.028382 | 0.095206 | 0.57422 | 0.45877 | 1.463 | 0.539 | -0.26679
Twist Integral (I0) (m2 rad) | 0.010115
Note on B208: 3.8s

## Work sheet diff
| HCMBALA001-2000097 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 108.4094 | 0.30302 | 0 |  | 0.064626 | 0.50072 |  |  | 108.1046 | 0.27151 | 0 |  | 0.0075282 | 0.81783 |  | 0 | 4
2 | 0 |  | 1.1483 | 0.22556 | 0 |  | -0.20336 | 0.11735 | 0 |  | -1.3847 | 0.14272 | 0 |  | -0.40963 | 0.092299 |  | 0 | 5
3 | 0 |  | 4.6701 | 0.080509 | 0 |  | 0.030818 | 0.11391 | 0 |  | 4.5499 | 0.034783 | 0 |  | 0.20087 | 0.086837 |  | 0 | 6
4 | 0 |  | 0.014638 | 0.015362 | 0 |  | 0.011618 | 0.025737 | 0 |  | -0.047661 | 0.01691 | 0 |  | -0.068394 | 0.02453 |  | 0 | 7
5 | 0 |  | 0.03515 | 0.0070363 | 0 |  | 0.0089635 | 0.011471 | 0 |  | 0.031063 | 0.0084495 | 0 |  | 0.019187 | 0.010095 |  | 0 | 8
6 | 0 |  | -0.023351 | 0.0051255 | 0 |  | 0.0044565 | 0.0056504 | 0 |  | 0.024204 | 0.0043485 | 0 |  | -0.0025174 | 0.0064181 |  | 0 | 9
7 | 0 |  | 0.0019394 | 0.0027437 | 0 |  | 0.01682 | 0.0041069 | 0 |  | -0.0002895 | 0.0035798 | 0 |  | 0.013485 | 0.0042051 |  | 0 | 10
8 | 0 |  | -0.0033699 | 0.0020283 | 0 |  | -0.0007481 | 0.0026652 | 0 |  | 0.0024946 | 0.0018369 | 0 |  | 0.0033651 | 0.00229 |  | 0 | 11
9 | 0 |  | 0.02237 | 0.0038671 | 0 |  | 0.0098195 | 0.0045651 | 0 |  | 0.023917 | 0.0031181 | 0 |  | 0.0092115 | 0.004232 |  | 0 | 12
10 | 0 |  | -2.5164e-18 | 1.1554e-17 | 0 |  | 1.467e-19 | 6.67e-18 | 0 |  | -1.7524e-18 | 1.0701e-17 | 0 |  | -2.0822e-18 | 5.8583e-18 |  | 0 | 13
11 | 0 |  | 0.016372 | 0.00083051 | 0 |  | 0.024134 | 0.0036915 | 0 |  | 0.011627 | 0.00091558 | 0 |  | 0.025017 | 0.0059241 |  | 0 | 14
12 | 0 |  | -0.00067403 | 0.00085228 | 0 |  | -0.00041521 | 0.00096313 | 0 |  | -0.00015883 | 0.00070017 | 0 |  | 5.3039e-05 | 0.00088135 |  | 0 | 15
13 | 0 |  | 0.0014125 | 0.00070894 | 0 |  | 0.0028897 | 0.00090002 | 0 |  | 0.0023466 | 0.00062011 | 0 |  | 0.0018158 | 0.0007682 |  | 0 | 16
14 | 0 |  | -0.0058437 | 0.001064 | 0 |  | -0.0049863 | 0.00072236 | 0 |  | -0.0021178 | 0.00054994 | 0 |  | 0.0054879 | 0.00057042 |  | 0 | 17
15 | 0 |  | 0.0062737 | 0.0030586 | 0 |  | -0.0055662 | 0.0031809 | 0 |  | -0.0072631 | 0.0017617 | 0 |  | 0.00057533 | 0.0028832 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -1.8159 | -2.1648
3 | [12 empty] | 3 | 9 | -2.3459 | -2.9012
4 | [12 empty] | 4 | 16 | -3.9178 | -3.8938
5 | [12 empty] | 5 | 25 | -4.7123 | -4.6847
6 | [12 empty] | 6 | 36 | -5.2248 | -5.2433
7 | [12 empty] | 7 | 49 | -5.6968 | -5.5519
8 | [12 empty] | 8 | 64 | -6.064 | -6.1894
9 | [12 empty] | 9 | 81 | -5.4723 | -5.6178
10 | [12 empty] | sigma (mm) |  | 0.0026832 | 0.0023687
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05 | [5 empty] | Aperture 1 | Aperture 2
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | Magnetic length | -0.056573 | -0.059653
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 37.7574 | 107.4911 | 108.3124 | 108.428 | 108.4449 | 108.5393 | 108.5533 | 108.5647 | 108.5455 | 108.5495 | 108.5658 | 108.541 | 108.5296 | 108.5371 | 108.5497 | 108.5274 | 108.5316 | 108.4456 | 107.7134 | 36.5267
2 | 26.4219 | 1.7863 | 1.3042 | 1.2079 | 1.1519 | 0.93395 | 1.0497 | 1.071 | 0.8717 | 1.0186 | 0.97997 | 0.95087 | 1.0943 | 1.0835 | 1.187 | 1.1369 | 1.1336 | 1.115 | 1.5928 | 28.2101 | 0
3 | 11.2495 | 4.5487 | 4.517 | 4.6381 | 4.6779 | 4.6866 | 4.7113 | 4.7056 | 4.7967 | 4.7779 | 4.755 | 4.7481 | 4.6663 | 4.7176 | 4.5871 | 4.6394 | 4.6885 | 4.6562 | 4.5432 | 9.7639 | 0
4 | 2.3399 | 0.030324 | 0.055432 | 0.026398 | 0.018442 | 0.0029331 | -0.0078436 | 0.0047751 | 0.011927 | 0.010597 | 0.012897 | 0.0050735 | 0.0010922 | 0.0027635 | 0.0033567 | 0.024154 | 0.0051903 | 0.020792 | 0.035174 | 0.73207 | 0
5 | 0.48746 | 0.041279 | 0.031523 | 0.035285 | 0.042232 | 0.032732 | 0.035988 | 0.035567 | 0.04486 | 0.0445 | 0.035901 | 0.037672 | 0.036179 | 0.044103 | 0.024033 | 0.018576 | 0.03258 | 0.028624 | 0.031062 | 0.013491 | 0
6 | 0.30875 | -0.01913 | -0.019331 | -0.01323 | -0.01961 | -0.021367 | -0.021004 | -0.018429 | -0.026011 | -0.027025 | -0.030039 | -0.021013 | -0.022218 | -0.022497 | -0.029433 | -0.0347 | -0.025642 | -0.027502 | -0.022133 | -0.019646 | 0
7 | 0.057896 | 0.0034703 | 0.00040094 | 0.0029556 | 0.003956 | 0.005733 | 0.004077 | 0.0030945 | -0.0019373 | -0.00016622 | 0.0017443 | 0.002763 | 0.0065589 | 0.0022888 | -0.0022076 | -0.0026031 | 0.0032596 | -0.001873 | 0.0033954 | 0.03179 | 0
8 | 0.085605 | -0.0028504 | -0.00095738 | -0.0030943 | -0.0033383 | -0.0040546 | -0.0036696 | -0.0032126 | -0.0043237 | -0.0060296 | -0.00098044 | -0.0015881 | -0.0081844 | -0.0049415 | -0.0055324 | -0.0026648 | -0.00073269 | -0.0039902 | -0.00051367 | 0.0021029 | 0
9 | 0.033939 | 0.021045 | 0.018062 | 0.022536 | 0.020241 | 0.025967 | 0.021517 | 0.018384 | 0.026725 | 0.030719 | 0.0223 | 0.026038 | 0.021898 | 0.019772 | 0.027365 | 0.0195 | 0.024517 | 0.020814 | 0.015265 | 0.047649 | 0
10 | 0 | -1.1761e-17 | -3.4694e-18 | -1.7347e-18 | 5.8804e-18 | 0 | -1.7641e-17 | -7.3505e-19 | -3.4694e-18 | 8.6736e-19 | 2.297e-20 | -1.7347e-18 | -3.4694e-18 | -3.7399e-17 | 9.6439e-18 | -5.8804e-18 | 1.1707e-17 | 0 | 1.3878e-17 | 2.6062e-17 | 0
11 | 0.0070515 | 0.017561 | 0.015711 | 0.016482 | 0.016943 | 0.014558 | 0.01691 | 0.015474 | 0.017157 | 0.016389 | 0.015956 | 0.015814 | 0.017015 | 0.017475 | 0.015358 | 0.015481 | 0.016663 | 0.01682 | 0.016929 | 0.038218 | 0
12 | 0.041365 | 0.0047592 | 0.0007405 | -0.00051328 | -0.015423 | -0.015905 | -0.0099231 | -0.02074 | 0.0018254 | -0.016382 | 0.00027615 | -0.0022281 | -0.0054218 | -0.010454 | -0.0030165 | -0.0056443 | 0.0079465 | -0.016789 | -0.014433 | -0.044863 | 0 | ten times!
13 | 0.046809 | -0.0020793 | 0.010154 | 0.016615 | 0.01696 | 0.01563 | 0.010879 | 0.024386 | 0.021264 | 0.021546 | 0.020623 | 0.020583 | 0.014734 | 0.018858 | 0.013869 | 0.011596 | 0.010318 | 0.0017888 | 0.0065177 | 0.071434 | 0 | ten times!
14 | 0.021359 | -0.031653 | -0.04614 | -0.057086 | -0.047426 | -0.052078 | -0.056317 | -0.051691 | -0.055812 | -0.061186 | -0.069643 | -0.066177 | -0.073308 | -0.06524 | -0.072268 | -0.068185 | -0.057026 | -0.066636 | -0.053997 | 0.0065187 | 0 | ten times!
15 | 0.10531 | 0.041335 | 0.07476 | 0.05243 | 0.090899 | 0.078599 | 0.047386 | 0.040362 | 0.026614 | 0.026347 | 0.0492 | 0.034134 | 0.051083 | 0.093356 | 0.10396 | 0.14532 | 0.066486 | 0.051671 | 0.055327 | 0.17864 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -1.622 | -0.92502 | -0.40328 | -0.60476 | 0.028494 | 0.18828 | 1.0246 | 0.53758 | 0.10094 | 0.12689 | -0.38541 | -0.44567 | 0.26311 | -0.026151 | -0.13256 | 0.27291 | 0.15461 | 0.52876 | 0.8599 | -0.19212
2 | -12.3733 | -0.12958 | -0.18801 | -0.27785 | -0.16946 | 0.082014 | -0.11629 | -0.041925 | -0.16321 | -0.34739 | -0.27915 | -0.22386 | -0.31335 | -0.1964 | -0.20805 | -0.23009 | -0.33608 | -0.11764 | -0.40409 | 0.26485 | 0
3 | -1.2584 | 0.24424 | 0.031316 | 0.049696 | 0.14422 | 0.30286 | -0.030572 | 0.11387 | 0.12097 | -0.030002 | -0.021435 | 0.0023107 | -0.13733 | -0.023946 | -0.049538 | 0.0096681 | -0.057086 | -0.043872 | -0.070652 | -0.29023 | 0
4 | -1.9467 | 0.04739 | 0.022998 | 0.0029125 | 0.019638 | 0.057616 | -0.009738 | 0.064428 | 0.018976 | -0.015121 | -0.01129 | 0.0016171 | -0.035649 | 0.017293 | 0.0084256 | 0.021179 | -0.0025516 | 0.0088942 | -0.0078968 | 0.042776 | 0
5 | -0.062766 | 0.021141 | 0.0033886 | 0.011106 | 0.018447 | 0.034077 | -0.0029408 | 0.019141 | 0.01886 | 0.005469 | 0.0035557 | 0.0026163 | 0.0025713 | 0.0073695 | 0.011527 | 0.0043808 | 0.0055618 | 0.014322 | -0.019251 | -0.12936 | 0
6 | -0.4196 | 0.0017922 | 0.0010679 | 0.0060213 | 0.0028169 | 0.011264 | 0.0038638 | 0.0078713 | 0.012723 | 0.0027582 | 0.018311 | 0.0044879 | -0.0066018 | -0.00082063 | 0.00011865 | 0.0063685 | 0.0035894 | 0.005013 | -0.00042685 | 0.005945 | 0
7 | 0.13165 | 0.0073001 | 0.015399 | 0.017799 | 0.017659 | 0.018119 | 0.025987 | 0.018425 | 0.014517 | 0.017188 | 0.01708 | 0.015413 | 0.010591 | 0.013675 | 0.014817 | 0.018284 | 0.018178 | 0.020103 | 0.022221 | 0.0012185 | 0
8 | -0.016259 | -0.0016707 | -0.00028948 | 0.0021563 | -0.0019237 | 0.0030179 | -0.00047229 | -0.0053989 | 0.00025945 | 0.0003921 | -0.001176 | 0.00059168 | -0.0012518 | -0.0054158 | 0.0017267 | 0.002081 | 0.0016585 | -0.001854 | -0.0058966 | -0.0052083 | 0
9 | 0.028654 | -0.00076672 | 0.0098838 | 0.008485 | 0.0072392 | 0.0066498 | 0.017968 | 0.012714 | 0.0024737 | 0.011031 | 0.0049231 | 0.01085 | 0.012551 | 0.011205 | 0.0094284 | 0.013683 | 0.0094849 | 0.015017 | 0.01393 | -0.0035685 | 0
10 | 1.0211e-17 | 1.1544e-17 | 8.6736e-19 | -3.2048e-18 | 8.409e-18 | 3.9987e-18 | -2.0728e-18 | -4.8219e-18 | -1.9758e-17 | -9.4087e-19 | 3.1019e-18 | -1.2055e-18 | -2.332e-18 | 6.9389e-18 | 0 | 0 | -1.3525e-18 | -2.411e-18 | 5.8804e-18 | 6.7695e-17 | 0
11 | 0.025973 | 0.01621 | 0.02097 | 0.019653 | 0.021656 | 0.025154 | 0.031045 | 0.029125 | 0.026338 | 0.02365 | 0.02142 | 0.020185 | 0.024577 | 0.023695 | 0.024674 | 0.025625 | 0.024839 | 0.027226 | 0.028366 | 0.023449 | 0 | ten times!
12 | 0.05992 | -0.0013986 | -0.0076265 | -0.0019678 | -0.0042897 | 0.012575 | -0.005517 | 0.00011869 | -0.0086024 | -0.0020372 | 0.0028439 | 0.0010526 | -0.0081656 | -0.019801 | -0.025251 | -0.018847 | 0.00016085 | 0.0043935 | 0.0076214 | 0.010082 | 0 | ten times!
13 | -0.015267 | 0.018324 | 0.022903 | 0.017045 | 0.020372 | 0.023565 | 0.039158 | 0.038785 | 0.024373 | 0.033991 | 0.022137 | 0.037293 | 0.014352 | 0.035777 | 0.035293 | 0.044452 | 0.023956 | 0.034054 | 0.034315 | -0.022286 | 0 | ten times!
14 | 0.067012 | -0.049939 | -0.04858 | -0.045427 | -0.054051 | -0.0449 | -0.050117 | -0.04772 | -0.041466 | -0.035685 | -0.053194 | -0.047188 | -0.051468 | -0.048189 | -0.040616 | -0.061231 | -0.054587 | -0.059163 | -0.064013 | 0.00047073 | 0 | ten times!
15 | -0.12034 | -0.06668 | -0.087624 | -0.062225 | -0.10697 | -0.074828 | -0.068802 | -0.063593 | 0.0027244 | -0.099674 | -0.032318 | -0.019642 | -0.051187 | -0.050625 | -0.076009 | -0.0018466 | -0.011548 | -0.058662 | -0.072406 | -0.14338 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 37.1062 | 107.3116 | 108.1205 | 108.2047 | 108.2054 | 108.2054 | 108.2279 | 108.1871 | 108.2187 | 108.2357 | 108.2312 | 108.2214 | 108.2474 | 108.2286 | 108.2347 | 108.1692 | 108.1214 | 108.0813 | 107.4307 | 27.2815
2 | -25.7419 | -1.7155 | -1.3704 | -1.3783 | -1.3801 | -1.2195 | -1.2713 | -1.3896 | -1.3444 | -1.1442 | -1.2841 | -1.4077 | -1.2523 | -1.3397 | -1.3189 | -1.5327 | -1.5028 | -1.4366 | -1.6374 | -26.6946 | 0
3 | 11.3741 | 4.4767 | 4.4996 | 4.5549 | 4.5564 | 4.5999 | 4.6123 | 4.5596 | 4.5898 | 4.5779 | 4.5464 | 4.5538 | 4.5806 | 4.5541 | 4.5381 | 4.5223 | 4.5327 | 4.5228 | 4.5211 | 9.4607 | 0
4 | 0.39164 | -0.062903 | -0.062101 | -0.078042 | -0.066174 | -0.081095 | -0.03195 | -0.024314 | -0.032114 | -0.028868 | -0.047331 | -0.04926 | -0.034358 | -0.037911 | -0.030686 | -0.046428 | -0.041062 | -0.049379 | -0.053918 | -1.0373 | 0
5 | 0.38969 | 0.04072 | 0.036281 | 0.042783 | 0.038463 | 0.047851 | 0.029523 | 0.028756 | 0.020132 | 0.038605 | 0.020149 | 0.032466 | 0.031238 | 0.028971 | 0.020326 | 0.018744 | 0.028593 | 0.025043 | 0.030487 | -0.06702 | 0
6 | 0.33666 | 0.024012 | 0.022248 | 0.018369 | 0.024308 | 0.034757 | 0.023279 | 0.023813 | 0.026063 | 0.025844 | 0.023793 | 0.026945 | 0.028759 | 0.030535 | 0.023847 | 0.018619 | 0.016916 | 0.022182 | 0.021379 | -0.0028244 | 0
7 | 0.076907 | 0.0015254 | 0.0048106 | -0.0011705 | 2.4415e-05 | -0.0065182 | 0.00061758 | 0.00022493 | -0.0045921 | 0.0023128 | 0.0029392 | 0.0026318 | 0.00093137 | -0.0061113 | -0.0043299 | 0.0013623 | -0.0013449 | -0.0040525 | 0.0055278 | 0.011166 | 0
8 | 0.069721 | -0.00072949 | 0.00028671 | -0.00043552 | 0.0040523 | 0.0056449 | 0.0029853 | 0.0026467 | 0.0039473 | 0.0025689 | 0.0034279 | 0.00083282 | 0.001055 | 0.002096 | 0.0029928 | 0.00089478 | 0.003374 | 0.0043152 | 0.004948 | 0.0035943 | 0
9 | 0.044044 | 0.024751 | 0.023509 | 0.024555 | 0.026631 | 0.029476 | 0.025943 | 0.022935 | 0.02057 | 0.029309 | 0.021146 | 0.02709 | 0.025155 | 0.021837 | 0.024939 | 0.023268 | 0.020954 | 0.020333 | 0.018111 | 0.032149 | 0
10 | 1.9269e-17 | 2.3522e-17 | 3.0486e-18 | -3.2857e-18 | 1.8376e-19 | -1.1761e-17 | -3.4694e-18 | -5.2924e-19 | 2.9402e-18 | 1.1761e-17 | -6.9389e-18 | -2.3522e-17 | -1.87e-17 | 0 | 1.3525e-18 | 1.4701e-18 | -1.7347e-18 | 5.8804e-18 | -1.1761e-17 | 1.3031e-17 | 0
11 | 0.017558 | 0.010792 | 0.011995 | 0.011424 | 0.010801 | 0.010581 | 0.011577 | 0.01217 | 0.011188 | 0.01228 | 0.01094 | 0.013303 | 0.012863 | 0.011247 | 0.010314 | 0.011744 | 0.011383 | 0.01114 | 0.01355 | 0.022348 | 0
12 | -0.023759 | -0.010497 | -0.0013673 | -0.0089101 | -0.002351 | 0.009606 | -0.008948 | 0.0086332 | 0.0019459 | 0.0068771 | -0.0021034 | 0.0074045 | 0.0040858 | -0.0013328 | -0.004351 | 0.00068169 | -0.0091335 | -0.0058032 | -0.013026 | -0.028483 | 0 | ten times!
13 | 0.04235 | 0.029393 | 0.029366 | 0.020696 | 0.025286 | 0.037877 | 0.02611 | 0.021679 | 0.021523 | 0.016553 | 0.013615 | 0.028977 | 0.024378 | 0.018652 | 0.018682 | 0.021353 | 0.018839 | 0.017633 | 0.031782 | -0.0086474 | 0 | ten times!
14 | -0.010819 | -0.021632 | -0.014277 | -0.024644 | -0.018563 | -0.037217 | -0.022959 | -0.024915 | -0.027989 | -0.019121 | -0.023724 | -0.018194 | -0.019826 | -0.020083 | -0.016476 | -0.022252 | -0.019126 | -0.013944 | -0.016256 | 0.0052848 | 0 | ten times!
15 | -0.043244 | -0.051725 | -0.096615 | -0.057206 | -0.065649 | -0.052576 | -0.098316 | -0.075281 | -0.078048 | -0.098392 | -0.083469 | -0.057432 | -0.09524 | -0.061279 | -0.081057 | -0.043707 | -0.060404 | -0.084231 | -0.066729 | -0.066348 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -0.4756 | -0.82249 | -0.47267 | 0.76448 | -0.41387 | 2.5563 | 0.82451 | 0.31142 | -0.4119 | -0.3176 | -0.82706 | -0.82132 | -0.23301 | -0.13326 | -0.35878 | -0.007034 | -0.39176 | 0.33254 | 0.55704 | 0.4469
2 | -10.963 | -0.22418 | -0.43275 | -0.58825 | -0.39319 | -0.29898 | -0.44602 | -0.43066 | -0.54765 | -0.49763 | -0.50082 | -0.34272 | -0.36694 | -0.40457 | -0.38078 | -0.37234 | -0.48529 | -0.35982 | -0.30077 | -0.99778 | 0
3 | -0.36675 | 0.18824 | 0.27714 | 0.31931 | 0.27563 | 0.0032345 | 0.10804 | 0.29341 | 0.31177 | 0.23234 | 0.28492 | 0.23074 | 0.11558 | 0.17097 | 0.10023 | 0.14683 | 0.23122 | 0.15571 | 0.1704 | 0.39219 | 0
4 | -1.0156 | -0.047595 | -0.067274 | -0.10095 | -0.065686 | -0.054833 | -0.028125 | -0.098492 | -0.11297 | -0.093007 | -0.095679 | -0.071362 | -0.057935 | -0.082339 | -0.060466 | -0.038993 | -0.068272 | -0.053703 | -0.033408 | 0.050474 | 0
5 | -0.07419 | 0.0075071 | 0.0173 | 0.029349 | 0.022658 | 0.038658 | 0.0073381 | 0.024866 | 0.037323 | 0.0090847 | 0.032602 | 0.015903 | 0.013341 | 0.01924 | 0.015043 | 0.020441 | 0.013956 | 0.016209 | 0.0045427 | -0.11377 | 0
6 | -0.10205 | -0.010129 | -0.0023665 | -0.007819 | -0.0047645 | -0.018998 | 0.007114 | -0.0052878 | -0.0082997 | 0.0053518 | -0.0017624 | 0.0016937 | 0.0045877 | -0.0037245 | 0.0040463 | -0.0033146 | -0.0019097 | 0.0027164 | -0.0024479 | 0.022151 | 0
7 | 0.08986 | 0.0075426 | 0.013913 | 0.016504 | 0.011181 | 0.017546 | 0.016037 | 0.013577 | 0.010865 | 0.01353 | 0.0064053 | 0.0065624 | 0.013741 | 0.01008 | 0.012067 | 0.015728 | 0.017209 | 0.019103 | 0.021143 | 0.0012418 | 0
8 | 0.021758 | 0.0016488 | 0.0013391 | 0.005216 | 0.0036517 | 0.0097345 | 0.0060008 | 0.0042113 | 0.0048205 | 0.0029154 | 0.0023729 | 0.00034113 | 0.0012387 | 0.0031566 | 0.0040795 | 0.001407 | 0.00050629 | 0.0037337 | 0.0041974 | 0.0036403 | 0
9 | 0.0078342 | 0.0088611 | 0.0052764 | 0.011332 | 0.0087119 | 0.01834 | 0.016197 | 0.0086756 | 0.0052291 | 0.0059599 | 0.0059801 | -0.00021173 | 0.011282 | 0.0078944 | 0.011158 | 0.010504 | 0.0076666 | 0.012452 | 0.010498 | -0.0030979 | 0
10 | 5.697e-18 | -1.4701e-18 | -7.6151e-18 | 7.9974e-18 | -6.9389e-18 | 7.3505e-19 | 6.9389e-18 | -1.7347e-18 | -2.6462e-19 | 6.9389e-18 | -5.2924e-19 | -6.9389e-18 | -1.0938e-17 | -3.4694e-18 | -1.7347e-18 | -1.3878e-17 | -3.3739e-18 | -1.4701e-18 | 2.6462e-19 | 1.2184e-17 | 0
11 | 0.016531 | 0.019354 | 0.022189 | 0.02951 | 0.022255 | 0.043594 | 0.031265 | 0.028592 | 0.022789 | 0.022346 | 0.018959 | 0.017577 | 0.023496 | 0.023938 | 0.022352 | 0.024916 | 0.022508 | 0.027468 | 0.027196 | 0.034081 | 0
12 | 0.046528 | -0.0039564 | -0.0013226 | -0.010153 | -0.0038868 | 0.019291 | -0.0020292 | 0.0030786 | -0.0018053 | -0.0066689 | -0.0011672 | -0.015075 | 0.0068193 | 0.011079 | 0.012005 | -0.00035977 | -0.010236 | 0.0044647 | 0.0094696 | 0.0054373 | 0 | ten times!
13 | 0.0065548 | 0.01162 | 0.014685 | 0.025117 | 0.0073132 | 0.019381 | 0.031058 | 0.029823 | 0.018499 | 0.0084202 | 0.014949 | 0.0037475 | 0.026409 | 0.014829 | 0.016235 | 0.015935 | 0.020799 | 0.022994 | 0.025038 | 0.0049648 | 0 | ten times!
14 | 0.024121 | 0.045799 | 0.05091 | 0.052383 | 0.05486 | 0.060366 | 0.051852 | 0.058223 | 0.058541 | 0.058523 | 0.055822 | 0.063373 | 0.059897 | 0.053835 | 0.062466 | 0.050752 | 0.042377 | 0.058366 | 0.049473 | 0.049657 | 0 | ten times!
15 | -0.013958 | -0.034877 | -0.012557 | 0.018862 | -0.013608 | 0.0251 | 0.035066 | 0.014442 | 0.038918 | -0.043159 | 0.00056406 | 0.057341 | 0.033115 | 0.014808 | -0.04194 | -0.0021817 | 0.02454 | 0.0075565 | -0.018431 | -0.090835 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.72666 | -1.3241 | 0.1339 | -0.0084051 | 0.64549 | -0.52486 | -1.0151 | 0.052018 | 0.39224 | 25.2806 | 0.70385 | -0.58557 | -1.991 | -0.18353 | 0.36332 | 0.5868 | -1.1527 | 0.61302 | -0.32215 | 0.78631 | 28.3331 | 0.70345 | 2.2838 | -0.1055 | 0.18084 | -0.27345 | 0.67728 | 0.022025 | -1.1491 | -0.013289 | 0.37877 | 14359.625 | -2.2356 | -0.82705 | -0.22431 | 0.056306 | 0.6123 | -0.020427 | 0.52776 | -0.39774 | 0.78793 | 14359.852
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | English | 0 | 0