## Alarm sheet
File name | HCMBARA001-3000024_cm.xls
Component ID | HCMBARA001 |  |  |  | Serial Number | 3000024
Date of test Ap 1 | 2003-05-21 |  |  |  | Date of test Ap 2 | 2003-05-21
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | red alarm | yellow alarm | yellow alarm
Angle |  | status ok | status ok | status ok | Angle |  | status ok | yellow alarm | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 |  |  |  |  | b10
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 |  |  |  |  | a10
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | red alarm

## Assembly Data
Magnet name
| HCMBARA001 | 3000024 |  | Additional lamin. (mm) |  | 0
| COMMENTS: | Red alarms in position 2 and 19 of aperture 2 due to wrong positioning of measuring mole along magnet axis
| HOLDING POINT | OK | Data at CERN |  | 2003-05-21 | Analysis by | E.wildner | Answer to MMS-LD |  | 2003-05-21 | Answer to manufacturer |  | 2003-05-21
1 | Fill in the yellow fields in Original Data | [10 empty] | English
2 | CTRL-g to load the measurement files .txt
3 | Save as … with the "official" name in Original Data C1
4 | Open the file of the collared coil measurement of the same magnet
5 | Go back to cold_mass file. Tools - Macro - Macros - Make difference - Run (to make the difference between collared coil and cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (to perform the analysis - results in Alarm sheet)
7 | Save the file
| Cold mass template - version 11.10.2002 | [7 empty] | Version : | cern
1 | Remplir les champs jaunes dans Original Data | [10 empty] | Francais
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec le nom "officiel" dans Original Data C1
4 | Ouvrir le fichier avec les mesures de la collared coil
5 | Revenir au ficher cold_mass. Tools - Macro - Macros - Make difference - Run (pour faire la difference entre collared coil et cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (pour faire l'analyse - resultats dans Alarm sheet)
7 | Sauver le fichier
1 | Riempire I campi gialli in Original Data | [10 empty] | Italiano
2 | CTRL-g per caricare I file di misura .txt
3 | Salvare il file con il nome "ufficiale" in Original Data C1
4 | Aprire il file delle misure della collared coil corrispondente
5 | Tornare sul file cold_mass. Tools - Macro - Macros - Make difference - Run (per fare la differenza tra collared coil e cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (per fare l'analisi - risultati in Alarm sheet)
7 | Salvare il file
1 | Die gelben Felder in Original Data ausfüllen | [10 empty] | Deutsch
2 | Ctrl-g um die Messdatei ( Datei .txt) zu laden.
3 | Die Datei mit der ”offiziellen” bezeichnung im Original Data C1 speichern mit “Save as…”.
4 | Öffnen “Collared Coils” Messdatei (gleiche Magnete versteht sich)
5 | Zurück zum cold mass Datei. Tools - Macro - Macros - Make difference - Run (Zeigt allen unterschieden in Messdatei zw. cc und cm)
6 | Tools - Macro - Macros - Collared coil analysis - Run (Um Maßdaten zu analysieren - Resultaten ins Alarm Sheet)
7 | Datei speichern
| Last modif on | by | Modifications:
| 2002-10-20 | E. Todesco | Cell for iron additional iron laminations - check on magnetic length includes additional iron laminations
| 2002-11-13 | E. Wildner | German version of instructions added

## Original data
File name |  | HCMBBRA001-03000024_cm.xls
Component ID |  | HCMBARA001 |  |  | Serial Number |  |  | 3000024 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3024_cm1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 34 |  |  |  | File name for aperture 2 |  |  | 3024_cm2 |  |  |  | a even | 1
Test Operator |  | Schuetze |  |  | Test Controller |  |  | E. Wildner
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-05-21 |  |  | Time of test |  |  | 13:04 |  |  |  | Date of test |  | 2003-05-21 |  |  | Time of test |  |  | 13:04
Mole name |  | Taupe_dipole#4 |  |  | Calibration date |  |  | 1970-01-01 |  |  |  | Mole name |  | Taupe_Dipole#5 |  |  | Calibration date |  |  | 1970-01-01
Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 21 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 21
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.5 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.5 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00018745 |  |  | Dy average |  |  | 0.00019448 |  |  |  | Dx average |  | 5.688e-05 |  |  | Dy average |  |  | 0.00020692
Magnetic length |  | 14.359 |  |  | Tranfs. Function (mT/KA) |  |  | 704.7339 |  |  |  | Magnetic length |  | 14.3593 |  |  | Tranfs. Function (mT/KA) |  |  | 704.585
Phase |  | -1.3832 |  |  | Integral field (T.m) |  |  | 101.1371 |  |  |  | Phase |  | -19.0388 |  |  | Integral field (T.m) |  |  | 100.9132
File | Aperture 1 - Cold mass | [19 empty] | Integrals | File | Aperture 2 - Cold mass | [19 empty] | Integrals
C1 (mT) | 3.8919 | 7.042 | 7.0473 | 7.048 | 7.0476 | 7.0477 | 7.0477 | 7.0478 | 7.0479 | 7.048 | 7.0474 | 7.048 | 7.0475 | 7.0481 | 7.0476 | 7.0474 | 7.0481 | 7.048 | 7.0459 | 4.1236 | 101.1371 | C1 (mT) | 3.3886 | 7.0273 | 7.046 | 7.0463 | 7.0462 | 7.0465 | 7.047 | 7.0468 | 7.0477 | 7.0478 | 7.0472 | 7.0478 | 7.0471 | 7.0472 | 7.0465 | 7.0468 | 7.0474 | 7.0477 | 7.0458 | 4.355 | 100.9132
Angle (mrad) | 5.519 | -0.36626 | -0.17473 | 0.54611 | 0.18211 | -0.12869 | -0.30781 | -0.43454 | -0.75957 | -0.35462 | -0.30999 | -0.053296 | 0.31351 | 0.30426 | 0.079158 | -0.11315 | 0.046418 | -0.59318 | -0.60434 | -0.4671 | -1.3832 | Angle (mrad) | 6.4481 | -0.34176 | -0.89955 | 0.48048 | -0.039006 | -0.13517 | -0.11058 | 0.065165 | -0.18315 | -0.07698 | -0.22087 | -0.11083 | 0.34552 | -0.01205 | -0.10532 | -0.34207 | 0.32755 | -0.033236 | -0.5106 | -1.9907 | -19.0388
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 25.1972 | -0.1341 | -0.086412 | 0.12956 | -0.13682 | 0.077111 | -0.4719 | -0.29072 | -0.2645 | 0.13395 | -0.069256 | 0.21904 | 0.93862 | 0.46906 | 0.11103 | 0.31322 | 0.44084 | 0.10899 | 0.28325 | 21.0454 | 1.4631 | b2 | -27.8884 | -0.6511 | -0.093736 | -0.051472 | -0.19495 | -0.23475 | -1.1812 | -0.6771 | 0.18201 | -0.37024 | -0.16109 | -0.46049 | -0.32595 | -0.25452 | -0.30736 | -0.1247 | -0.1194 | -0.027144 | 0.36821 | -20.9256 | -1.6247
b3 | 38.6014 | 3.3788 | 4.1922 | 4.6919 | 4.8216 | 4.6253 | 4.582 | 5.6586 | 5.5983 | 4.9897 | 4.2292 | 4.4569 | 5.1897 | 4.5935 | 4.5019 | 4.6572 | 5.2409 | 4.8733 | 4.9587 | 9.3059 | 5.8526 | b3 | 44.8241 | 4.4216 | 4.6486 | 4.1424 | 4.538 | 3.7875 | 4.4393 | 4.5893 | 3.9457 | 5.1368 | 4.4717 | 3.9246 | 4.3228 | 3.8433 | 4.4532 | 4.4914 | 5.0591 | 4.9459 | 4.4349 | 7.9492 | 5.5535
b4 | 0.61372 | -0.027364 | -0.044419 | -0.074159 | -0.17234 | -0.096559 | -0.079515 | 0.10164 | 0.012027 | 0.1696 | -0.043694 | -0.014527 | 0.083486 | -0.062547 | -0.024676 | 0.018667 | -0.090715 | -0.14066 | -0.23076 | 1.1189 | 0.014481 | b4 | -1.1609 | -0.033123 | 0.13918 | 0.21533 | 0.35692 | 0.061745 | -0.15139 | 0.05577 | 0.024192 | -0.086377 | -0.12833 | -0.20818 | -0.061743 | 0.22156 | 0.2009 | 0.29831 | 0.15465 | 0.11737 | -0.041954 | -0.47623 | 0.014782
b5 | -5.1639 | -0.24005 | -0.61128 | -0.23137 | -0.17013 | -0.062998 | -0.21894 | -0.11997 | -0.093566 | 0.08282 | 0.14656 | 0.08799 | 0.079415 | 0.098782 | 0.023994 | 0.030717 | 0.099924 | 0.22406 | 0.084946 | -2.7478 | -0.27425 | b5 | -6.1714 | -0.50108 | -0.44693 | -0.14532 | -0.011193 | -0.027486 | 0.023192 | -0.10096 | 0.036837 | -0.088288 | 0.088495 | 0.10366 | -0.091928 | -0.20315 | -0.15082 | -0.23256 | -0.18432 | 0.0091012 | -0.094499 | -2.7602 | -0.35028
b6 | 0.13651 | 0.044521 | 0.0062766 | 0.056016 | 0.076264 | 0.044402 | 0.023597 | -0.013773 | -0.068715 | -0.044871 | -0.14376 | 0.013672 | -0.024756 | -0.026872 | -0.0081476 | 0.013393 | 0.035509 | 0.00036307 | -0.12283 | -0.12688 | -0.0072409 | b6 | 0.10879 | 0.084076 | 0.062543 | 0.090833 | 0.083908 | 0.09675 | 0.083567 | 0.073134 | -0.029268 | 0.050706 | -0.032592 | -0.069867 | 0.058245 | 0.061903 | -0.011176 | 0.052769 | 0.028865 | 0.10262 | -0.0073423 | -0.1574 | 0.03846
b7 | 2.4162 | 0.75189 | 0.79025 | 0.95283 | 0.9618 | 1.0098 | 1.0002 | 0.94951 | 0.98211 | 0.98972 | 0.96953 | 0.93297 | 0.96761 | 0.98782 | 0.96979 | 0.97469 | 0.9758 | 1.0219 | 0.98098 | 0.72591 | 0.98908 | b7 | 2.7326 | 0.68548 | 0.83589 | 0.97413 | 0.96883 | 1.0175 | 1.0045 | 0.92746 | 0.95509 | 0.95226 | 0.96618 | 0.95252 | 0.96337 | 0.91653 | 0.92961 | 0.93134 | 0.93935 | 1.0056 | 0.96445 | 0.65245 | 0.9743
b8 | 0.024779 | -0.018783 | -0.021715 | -0.0033964 | -0.030327 | 0.0074601 | -0.04551 | -0.0024166 | 3.2814e-05 | -0.0021741 | -0.04292 | -0.0063777 | -9.6074e-05 | -0.0083777 | -0.020692 | 0.0091329 | 0.03266 | 0.025396 | 0.039706 | -0.08927 | -0.0066331 | b8 | -0.03135 | 0.010515 | 0.004904 | 0.032211 | 0.052847 | 0.051392 | -0.033917 | -0.027132 | -0.0087085 | -0.0014524 | -0.0058015 | 0.015536 | 0.00079673 | -0.026065 | -0.027685 | 0.00029084 | 0.012833 | 0.036254 | -0.0090472 | -0.027818 | 0.002381
b9 | 0.29963 | 0.39058 | 0.37813 | 0.4148 | 0.41841 | 0.41208 | 0.41944 | 0.43292 | 0.4355 | 0.4114 | 0.41 | 0.40858 | 0.42405 | 0.41009 | 0.39751 | 0.40015 | 0.40214 | 0.41152 | 0.41027 | 0.40117 | 0.40694 | b9 | 0.33361 | 0.41383 | 0.40165 | 0.39919 | 0.40233 | 0.38861 | 0.40034 | 0.39953 | 0.39554 | 0.41522 | 0.40007 | 0.39857 | 0.4055 | 0.41163 | 0.4094 | 0.39637 | 0.3863 | 0.40285 | 0.39828 | 0.4194 | 0.40027
b10 | 0.043131 | 0.028405 | -0.015077 | 0.025576 | -0.011446 | 0.032484 | -0.056487 | -0.0021981 | -0.014796 | -0.023631 | -0.057858 | 0.0092271 | 0.0079331 | 0.0043634 | 0.013975 | 0.028337 | 0.046087 | 0.044514 | 0.0088818 | -0.15739 | 0 | b10 | 0.047523 | 0.0056115 | -0.013403 | 0.040452 | 0.025717 | 0.065566 | -0.029107 | -0.010172 | -0.024428 | -0.033982 | -0.038155 | 0.025678 | 0.016471 | -0.048319 | -0.054459 | -0.01109 | 0.031932 | 0.091578 | 0.047255 | -0.17795 | 0
b11 | 0.57371 | 0.64828 | 0.64227 | 0.64414 | 0.64953 | 0.65094 | 0.64227 | 0.64896 | 0.6452 | 0.64456 | 0.65361 | 0.65414 | 0.65163 | 0.64369 | 0.64532 | 0.64565 | 0.64354 | 0.64063 | 0.64098 | 0.59874 | 0.64285 | b11 | 0.57964 | 0.64625 | 0.63336 | 0.64081 | 0.64198 | 0.64123 | 0.64353 | 0.6453 | 0.64963 | 0.64218 | 0.64955 | 0.64989 | 0.6493 | 0.6515 | 0.64815 | 0.64332 | 0.63595 | 0.6377 | 0.63985 | 0.59709 | 0.64073
b12 | 0.011285 | 0.0038805 | -0.0023626 | 0.0010811 | -0.0042571 | 0.00062981 | -0.0072512 | -0.002726 | -0.0028932 | -0.0010375 | -0.0062509 | 0.0008262 | 0.002423 | 0.001744 | 0.00086968 | 0.004633 | 0.0088901 | 0.0073326 | 0.0048264 | -0.016027 | 0.00037667 | b12 | 0.009428 | 0.0039379 | -0.00251 | 0.0039287 | -0.0016293 | 0.0067805 | -0.0042838 | -0.0032926 | -0.003815 | -0.0021697 | -0.003884 | 0.0021659 | 0.0026308 | -0.0081159 | -0.0072753 | -0.0046797 | 0.0030764 | 0.0094764 | 0.00501 | -0.0092103 | -0.00030468
b13 | 0.082853 | 0.063374 | 0.0687 | 0.064169 | 0.06627 | 0.065524 | 0.066952 | 0.0673 | 0.064633 | 0.062757 | 0.062496 | 0.065084 | 0.065355 | 0.062049 | 0.063698 | 0.063968 | 0.063699 | 0.06133 | 0.061816 | 0.045999 | 0.064368 | b13 | 0.093291 | 0.063449 | 0.063518 | 0.063304 | 0.063472 | 0.060933 | 0.062017 | 0.064253 | 0.062314 | 0.062638 | 0.064795 | 0.062865 | 0.06685 | 0.065365 | 0.06491 | 0.065485 | 0.064435 | 0.062763 | 0.06189 | 0.045608 | 0.063789
b14 | 0.002887 | -0.0017543 | -0.0031583 | -0.0010998 | -0.0034373 | -0.0015931 | -0.0047707 | -0.0027781 | -0.0034029 | -0.0034046 | -0.0045407 | -0.0014137 | -0.0010388 | -0.001454 | -0.001263 | -0.0010223 | 0.00036562 | 0.00049541 | -0.0013006 | 0.00015146 | -0.0018232 | b14 | 0.0043663 | -0.0023128 | -0.0024771 | -0.00014876 | -0.001591 | 0.0014844 | -0.0029035 | -0.0025345 | -0.0017074 | -0.0028053 | -0.003419 | -0.00020376 | -0.0016912 | -0.0045045 | -0.0042056 | -0.0037005 | -7.7075e-05 | 0.0010257 | 0.00027677 | -0.0022005 | -0.0016102
b15 | -0.01212 | 0.021877 | 0.019803 | 0.019278 | 0.018214 | 0.018504 | 0.016047 | 0.019164 | 0.01846 | 0.018825 | 0.017618 | 0.020015 | 0.019569 | 0.01791 | 0.017163 | 0.018836 | 0.019652 | 0.020472 | 0.019154 | -0.0053507 | 0.017282 | b15 | -0.013788 | 0.023128 | 0.020775 | 0.018997 | 0.020707 | 0.018315 | 0.018069 | 0.018943 | 0.019302 | 0.018457 | 0.018696 | 0.020252 | 0.020009 | 0.01971 | 0.018528 | 0.019869 | 0.017952 | 0.019468 | 0.020429 | -0.00037122 | 0.01805
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 55.1908 | -3.6626 | -1.7473 | 5.4611 | 1.8211 | -1.2869 | -3.0781 | -4.3454 | -7.5957 | -3.5462 | -3.0999 | -0.53296 | 3.1351 | 3.0426 | 0.79158 | -1.1315 | 0.46418 | -5.9318 | -6.0434 | -4.671 | 0 | a1 | 64.4822 | -3.4176 | -8.9955 | 4.8048 | -0.39006 | -1.3517 | -1.1058 | 0.65165 | -1.8315 | -0.7698 | -2.2086 | -1.1083 | 3.4552 | -0.1205 | -1.0532 | -3.4207 | 3.2755 | -0.33236 | -5.106 | -19.9075 | 0
a2 | 0.44884 | 0.092078 | 0.40073 | 0.71258 | 1.5312 | 1.9257 | 1.687 | 0.98203 | 0.99767 | 0.84226 | 0.19251 | 1.3958 | 1.0703 | 1.5033 | 1.2764 | 1.0807 | 1.3039 | 1.6466 | 0.63711 | 1.9626 | 1.0803 | a2 | -4.5991 | -0.3103 | -0.64926 | -0.23987 | -0.10693 | 0.40652 | 0.22323 | 0.44007 | 0.73272 | 0.32157 | 0.355 | 0.23979 | 0.98192 | -0.10263 | 0.62269 | 0.35346 | 0.77681 | 0.10386 | 0.25387 | -1.7874 | 0.056946
a3 | -1.2087 | 0.41924 | -0.094515 | -0.30976 | -0.38226 | -0.60088 | -0.27008 | -0.61541 | 0.13345 | 0.21985 | -0.472 | -0.54458 | -0.70152 | -0.44897 | -0.021204 | -0.25667 | -0.59622 | -0.29138 | -0.10697 | -0.31853 | -0.30282 | a3 | -2.694 | -0.024508 | -0.37513 | -0.24649 | -0.24577 | -0.33256 | -0.21013 | -0.11823 | -0.15924 | -0.1022 | -0.4327 | -0.42444 | -0.54605 | -0.23344 | -0.21337 | -0.38287 | -0.033356 | 0.1173 | 0.30868 | 0.078067 | -0.25665
a4 | -0.15598 | 0.42561 | -0.39824 | -0.26169 | -0.091126 | -0.017281 | -0.20106 | -0.058873 | -0.088101 | 0.11828 | 0.038917 | -0.14447 | 0.14549 | -0.20778 | -0.18202 | -0.021917 | -0.077347 | -0.070687 | 0.1938 | 0.067448 | -0.049412 | a4 | -0.065292 | 0.2005 | 0.30009 | 0.28452 | 0.10871 | 0.48981 | 0.38962 | 0.40295 | 0.22517 | 0.16898 | 0.19919 | 0.14702 | 0.23105 | 0.28301 | 0.17244 | 0.39403 | 0.24794 | 0.29708 | 0.51963 | 0.14173 | 0.26798
a5 | 1.8939 | 0.13303 | -0.12108 | -0.087476 | -0.045937 | -0.13456 | -0.067156 | 0.078313 | 0.041454 | -0.11748 | -0.0013283 | 0.080275 | 0.015022 | -0.16364 | -0.22558 | -0.23114 | -0.12511 | 0.033954 | -0.021758 | -0.40461 | -0.0078968 | a5 | 2.2214 | 0.11464 | -0.026482 | -0.032829 | -0.1134 | -0.089182 | 0.0041081 | 0.10657 | -0.038085 | -0.11953 | -0.076438 | 0.009303 | 0.031347 | -0.0074123 | 0.051826 | -0.048558 | 0.0074494 | -0.1298 | 0.074505 | -0.1924 | 0.034921
a6 | -0.013542 | -0.044409 | -0.10229 | -0.044776 | 0.012476 | 0.014845 | 0.060284 | 0.043136 | 0.0051632 | -0.026847 | -0.016935 | 0.041871 | 0.10039 | 0.13319 | 0.12605 | 0.07368 | -0.038909 | -0.096073 | -0.057348 | -0.025676 | 0.0084159 | a6 | -0.13738 | 0.047072 | 0.04241 | 0.11016 | 0.069885 | 0.11897 | 0.15078 | 0.12887 | 0.10058 | 0.027665 | 0.0017691 | 0.064369 | 0.093448 | 0.059481 | 0.10728 | 0.040853 | 0.015778 | -0.036989 | 0.074475 | 0.21468 | 0.067203
a7 | 1.4251 | -0.055549 | -0.049573 | -0.10873 | -0.11013 | -0.066009 | -0.10402 | -0.10628 | -0.070242 | -0.028186 | -0.078243 | -0.064079 | -0.077793 | -0.073536 | -0.026365 | -0.03992 | -0.085155 | -0.10138 | -0.043491 | -0.0070354 | -0.026446 | a7 | 1.634 | -0.044679 | -0.065689 | -0.026042 | -0.063147 | -0.032017 | -0.015449 | 0.01133 | -0.1069 | -0.0768 | -0.077611 | -0.046045 | -0.036308 | -0.030047 | -0.042231 | -0.038366 | 0.0059785 | -0.02316 | -0.026783 | -0.059903 | 0.00078749
a8 | -0.041803 | 0.0085266 | 0.028435 | 0.0053099 | 0.014046 | 0.024576 | 0.027674 | 0.00087246 | -0.0057546 | -0.014179 | -0.024069 | 0.0068192 | 0.016318 | 0.049274 | 0.054898 | 0.033443 | -0.012572 | -0.024172 | -0.03378 | 0.0059026 | 0.0071087 | a8 | -0.054256 | -0.016427 | -0.0076877 | 0.043036 | 0.049588 | 0.069499 | 0.055859 | 0.045465 | 0.0068377 | -0.019747 | 0.0010809 | 0.010188 | 0.025376 | 0.012581 | 0.012889 | -0.00023908 | 0.017212 | -0.017416 | 0.010923 | 0.023584 | 0.015057
a9 | -0.22661 | 0.018031 | -0.024429 | -0.022902 | -0.014807 | -0.026067 | -0.01955 | -0.035304 | -0.030795 | -0.010754 | -0.030534 | -0.042671 | -0.038825 | -0.020423 | -0.031118 | -0.030521 | -0.036885 | -0.035486 | -0.037699 | -0.072482 | -0.033364 | a9 | -0.28767 | -0.010566 | -0.0078771 | -0.01581 | -0.014959 | -0.019176 | -0.028144 | -0.020213 | -0.0052786 | -0.010451 | -0.030896 | -0.021956 | -0.024127 | -0.017679 | -0.020088 | -0.0127 | -0.030056 | -0.046636 | -0.024023 | -0.054632 | -0.02789
a10 | 0.00067424 | 0.012036 | -0.023697 | -0.0139 | -0.0067629 | 0.014198 | -0.0014332 | 0.0055307 | 0.0099066 | -0.025809 | -0.0251 | -0.016768 | 0.010809 | 0.037259 | 0.059902 | 0.029421 | -0.0089298 | -0.041029 | -6.2547e-05 | -0.027225 | 0 | a10 | -0.036187 | -0.022356 | -0.018704 | 0.011127 | -0.0073972 | 0.017132 | 0.020716 | 0.039329 | -0.0065814 | -0.0043465 | -0.011045 | -0.0067508 | 0.024654 | 0.0051049 | 0.033351 | -0.0090173 | 0.023804 | -0.047863 | -0.023337 | -0.00077373 | 0
a11 | 0.16625 | -0.031447 | -0.033313 | -0.031877 | -0.031254 | -0.031746 | -0.0299 | -0.033347 | -0.034751 | -0.033921 | -0.03589 | -0.032706 | -0.028912 | -0.028457 | -0.027232 | -0.029151 | -0.031385 | -0.032971 | -0.030795 | -0.0098189 | -0.025255 | a11 | 0.20818 | -0.040698 | -0.052689 | -0.039946 | -0.039727 | -0.038237 | -0.037546 | -0.037516 | -0.044741 | -0.041708 | -0.04087 | -0.038858 | -0.033231 | -0.038885 | -0.034788 | -0.037466 | -0.034233 | -0.040168 | -0.04168 | -0.054267 | -0.033834
a12 | -0.0089248 | -0.001797 | -0.0060157 | -0.0030879 | -0.0012316 | 0.0031275 | 0.0031032 | 0.0025457 | 0.0017278 | -0.0032204 | -0.0059969 | -0.00081936 | 0.00059147 | 0.0059642 | 0.0070501 | 0.0026146 | -0.0019804 | -0.0052248 | -0.0035367 | -0.0013986 | -0.00062346 | a12 | -0.0069619 | -0.0016131 | -0.0037708 | 0.0031981 | -0.0019877 | 0.0016114 | 0.00045747 | 0.0026154 | 0.0011921 | 0.00019258 | -0.0035626 | -6.6129e-05 | 0.0055919 | 0.00090502 | 0.0021967 | -0.00264 | 0.00097104 | -0.0085992 | -0.0013411 | 0.0044649 | -0.00027405
a13 | -0.018624 | -0.0013983 | -0.0033505 | -0.0040212 | -0.0023291 | -0.0045322 | -0.0044305 | -0.0038878 | -0.0030542 | -0.0029052 | -0.0038237 | -0.0029824 | -0.0044211 | -0.0025114 | -0.0029423 | -0.0050235 | -0.0051716 | -0.0039356 | -0.0032575 | -0.0043674 | -0.0040162 | a13 | -0.021554 | -0.00091634 | -0.0037939 | -0.0023078 | -0.0053633 | -0.0027311 | -0.0047278 | -0.0035165 | -0.0050367 | -0.0018659 | -0.0031601 | -0.002019 | -0.0022153 | -0.0028856 | -0.0032607 | -0.0039734 | -0.0028761 | -0.0031801 | -0.0027173 | -0.0037383 | -0.0036234
a14 | -0.00022275 | -0.0024597 | -0.0046677 | -0.0051677 | -0.0047872 | -0.0037061 | -0.0019323 | -0.0037981 | -0.0038997 | -0.0047753 | -0.0043329 | -0.0048226 | -0.0037079 | -0.0038311 | -0.0034256 | -0.0035817 | -0.0042748 | -0.0053347 | -0.0042304 | -0.0016131 | -0.0038565 | a14 | 0.0075644 | -0.0053303 | -0.0044191 | -0.0017581 | -0.0033311 | -0.00068131 | -0.0011837 | -0.0010452 | -0.0012635 | -0.0012561 | -0.001925 | -0.0025053 | -0.0027209 | -0.0011743 | -0.00069602 | -0.0025128 | -0.00038754 | -0.002656 | -0.002247 | -9.2457e-05 | -0.0017542
a15 | 0.005158 | -0.0016851 | -0.0050099 | -0.0029169 | -0.0035926 | -0.0031995 | -0.005227 | -0.0032905 | -0.002782 | -0.0042486 | -0.0052909 | -0.0037097 | -0.0045022 | -0.0023069 | -0.0014223 | -0.00174 | -0.0015861 | -0.0017026 | -0.0015312 | -0.0073043 | -0.002988 | a15 | 0.0054963 | -0.0025336 | -0.0037368 | -0.0010388 | -0.0022123 | -0.00106 | -0.003067 | -0.0020927 | -0.003156 | -0.0033632 | -0.0030637 | -0.0011459 | -0.0024987 | -0.0029235 | -0.0024996 | -0.0030393 | -0.0014168 | 0.00048552 | -0.0010651 | -0.005 | -0.0020873
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00011844 | -7.2796e-05 | 3.6541e-05 | -6.9142e-05 | 2.9038e-05 | -8.2821e-05 | 0.00014908 | 0 | 4.0274e-05 | 5.8573e-05 | 0.00014643 | -2.6093e-05 | -1.9406e-05 | 0 | -3.0078e-05 | -7.0957e-05 | -0.00012265 | -0.00012337 | -2.3511e-05 | 0.0004466 | 0.00018745 | Dx (m) | -8.9576e-05 | -1.8386e-05 | 3.1573e-05 | -0.00010507 | -6.904e-05 | -0.00017055 | 7.982e-05 | 3.2693e-05 | 6.245e-05 | 8.884e-05 | 9.7651e-05 | -6.7983e-05 | -3.9703e-05 | 0.00012648 | 0.00014714 | 2.7822e-05 | -8.1698e-05 | -0.00025115 | -0.00012906 | 0.00050297 | 5.688e-05
Dy (m) | 3.2421e-05 | -3.5095e-05 | 6.4615e-05 | 3.3268e-05 | 1.91e-05 | -4.1123e-05 | 1.0707e-05 | -1.4155e-05 | -2.3936e-05 | 7.1152e-05 | 7.3408e-05 | 4.2273e-05 | -2.9061e-05 | -9.8697e-05 | -0.00015902 | -8.0688e-05 | 1.7634e-05 | 0.00010261 | 0 | 8.5179e-05 | 0.00019448 | Dy (m) | 0.00013839 | 5.7653e-05 | 5.2832e-05 | -3.6092e-05 | 1.5317e-05 | -5.5643e-05 | -5.008e-05 | -0.00010171 | 2.1525e-05 | 1.7276e-05 | 3.5073e-05 | 1.3595e-05 | -6.658e-05 | 0 | -7.9645e-05 | 2.5452e-05 | -6.804e-05 | 0.00011212 | 5.3621e-05 | 4.7398e-05 | 0.00020692
Magnetic length (m) |  | 14.359 | [19 empty] | Magnetic length (m) |  | 14.3593
Transfer funct. (mT/KA) |  | 704.7339 | [19 empty] | Transfer funct. (mT/KA) |  | 704.585
Coil positioning (mm) |  | 0.017709 | [19 empty] | Coil positioning (mm) |  | 0.015152
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMBARA001-03000024 |  |  | Aperture 1 - Cold mass | [16 empty] | Integrals | File | HCMBARA001-3000024 |  |  | Aperture 2 - Cold mass | [16 empty] | Integrals
C1/i (mT/KA) | 389.1852 | 704.1982 | 704.7288 | 704.8044 | 704.7638 | 704.7724 | 704.7723 | 704.7774 | 704.7934 | 704.8032 | 704.7396 | 704.8022 | 704.7458 | 704.8058 | 704.7586 | 704.7411 | 704.8115 | 704.8001 | 704.5915 | 412.3583 | 101.1371 | C1/i (mT/KA) | 338.858 | 702.7316 | 704.6029 | 704.632 | 704.6238 | 704.6475 | 704.7042 | 704.6793 | 704.7694 | 704.7815 | 704.7185 | 704.7799 | 704.7126 | 704.7216 | 704.6501 | 704.683 | 704.7404 | 704.7695 | 704.5819 | 435.5028 | 100.9132
Angle (mrad) | 5.519 | -0.36626 | -0.17473 | 0.54611 | 0.18211 | -0.12869 | -0.30781 | -0.43454 | -0.75957 | -0.35462 | -0.30999 | -0.053296 | 0.31351 | 0.30426 | 0.079158 | -0.11315 | 0.046418 | -0.59318 | -0.60434 | -0.4671 | -1.3832 | Angle (mrad) | 6.4481 | -0.34176 | -0.89955 | 0.48048 | -0.039006 | -0.13517 | -0.11058 | 0.065165 | -0.18315 | -0.07698 | -0.22087 | -0.11083 | 0.34552 | -0.01205 | -0.10532 | -0.34207 | 0.32755 | -0.033236 | -0.5106 | -1.9907 | -19.0388
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 25.1972 | -0.1341 | -0.086412 | 0.12956 | -0.13682 | 0.077111 | -0.4719 | -0.29072 | -0.2645 | 0.13395 | -0.069256 | 0.21904 | 0.93862 | 0.46906 | 0.11103 | 0.31322 | 0.44084 | 0.10899 | 0.28325 | 21.0454 | 1.4631 | b2 | -27.8884 | -0.6511 | -0.093736 | -0.051472 | -0.19495 | -0.23475 | -1.1812 | -0.6771 | 0.18201 | -0.37024 | -0.16109 | -0.46049 | -0.32595 | -0.25452 | -0.30736 | -0.1247 | -0.1194 | -0.027144 | 0.36821 | -20.9256 | -1.6247
b3 | 38.6014 | 3.3788 | 4.1922 | 4.6919 | 4.8216 | 4.6253 | 4.582 | 5.6586 | 5.5983 | 4.9897 | 4.2292 | 4.4569 | 5.1897 | 4.5935 | 4.5019 | 4.6572 | 5.2409 | 4.8733 | 4.9587 | 9.3059 | 5.8526 | b3 | 44.8241 | 4.4216 | 4.6486 | 4.1424 | 4.538 | 3.7875 | 4.4393 | 4.5893 | 3.9457 | 5.1368 | 4.4717 | 3.9246 | 4.3228 | 3.8433 | 4.4532 | 4.4914 | 5.0591 | 4.9459 | 4.4349 | 7.9492 | 5.5535
b4 | 0.61372 | -0.027364 | -0.044419 | -0.074159 | -0.17234 | -0.096559 | -0.079515 | 0.10164 | 0.012027 | 0.1696 | -0.043694 | -0.014527 | 0.083486 | -0.062547 | -0.024676 | 0.018667 | -0.090715 | -0.14066 | -0.23076 | 1.1189 | 0.014481 | b4 | -1.1609 | -0.033123 | 0.13918 | 0.21533 | 0.35692 | 0.061745 | -0.15139 | 0.05577 | 0.024192 | -0.086377 | -0.12833 | -0.20818 | -0.061743 | 0.22156 | 0.2009 | 0.29831 | 0.15465 | 0.11737 | -0.041954 | -0.47623 | 0.014782
b5 | -5.1639 | -0.24005 | -0.61128 | -0.23137 | -0.17013 | -0.062998 | -0.21894 | -0.11997 | -0.093566 | 0.08282 | 0.14656 | 0.08799 | 0.079415 | 0.098782 | 0.023994 | 0.030717 | 0.099924 | 0.22406 | 0.084946 | -2.7478 | -0.27425 | b5 | -6.1714 | -0.50108 | -0.44693 | -0.14532 | -0.011193 | -0.027486 | 0.023192 | -0.10096 | 0.036837 | -0.088288 | 0.088495 | 0.10366 | -0.091928 | -0.20315 | -0.15082 | -0.23256 | -0.18432 | 0.0091012 | -0.094499 | -2.7602 | -0.35028
b6 | 0.13651 | 0.044521 | 0.0062766 | 0.056016 | 0.076264 | 0.044402 | 0.023597 | -0.013773 | -0.068715 | -0.044871 | -0.14376 | 0.013672 | -0.024756 | -0.026872 | -0.0081476 | 0.013393 | 0.035509 | 0.00036307 | -0.12283 | -0.12688 | -0.0072409 | b6 | 0.10879 | 0.084076 | 0.062543 | 0.090833 | 0.083908 | 0.09675 | 0.083567 | 0.073134 | -0.029268 | 0.050706 | -0.032592 | -0.069867 | 0.058245 | 0.061903 | -0.011176 | 0.052769 | 0.028865 | 0.10262 | -0.0073423 | -0.1574 | 0.03846
b7 | 2.4162 | 0.75189 | 0.79025 | 0.95283 | 0.9618 | 1.0098 | 1.0002 | 0.94951 | 0.98211 | 0.98972 | 0.96953 | 0.93297 | 0.96761 | 0.98782 | 0.96979 | 0.97469 | 0.9758 | 1.0219 | 0.98098 | 0.72591 | 0.98908 | b7 | 2.7326 | 0.68548 | 0.83589 | 0.97413 | 0.96883 | 1.0175 | 1.0045 | 0.92746 | 0.95509 | 0.95226 | 0.96618 | 0.95252 | 0.96337 | 0.91653 | 0.92961 | 0.93134 | 0.93935 | 1.0056 | 0.96445 | 0.65245 | 0.9743
b8 | 0.024779 | -0.018783 | -0.021715 | -0.0033964 | -0.030327 | 0.0074601 | -0.04551 | -0.0024166 | 3.2814e-05 | -0.0021741 | -0.04292 | -0.0063777 | -9.6074e-05 | -0.0083777 | -0.020692 | 0.0091329 | 0.03266 | 0.025396 | 0.039706 | -0.08927 | -0.0066331 | b8 | -0.03135 | 0.010515 | 0.004904 | 0.032211 | 0.052847 | 0.051392 | -0.033917 | -0.027132 | -0.0087085 | -0.0014524 | -0.0058015 | 0.015536 | 0.00079673 | -0.026065 | -0.027685 | 0.00029084 | 0.012833 | 0.036254 | -0.0090472 | -0.027818 | 0.002381
b9 | 0.29963 | 0.39058 | 0.37813 | 0.4148 | 0.41841 | 0.41208 | 0.41944 | 0.43292 | 0.4355 | 0.4114 | 0.41 | 0.40858 | 0.42405 | 0.41009 | 0.39751 | 0.40015 | 0.40214 | 0.41152 | 0.41027 | 0.40117 | 0.40694 | b9 | 0.33361 | 0.41383 | 0.40165 | 0.39919 | 0.40233 | 0.38861 | 0.40034 | 0.39953 | 0.39554 | 0.41522 | 0.40007 | 0.39857 | 0.4055 | 0.41163 | 0.4094 | 0.39637 | 0.3863 | 0.40285 | 0.39828 | 0.4194 | 0.40027
b10 | 0.043131 | 0.028405 | -0.015077 | 0.025576 | -0.011446 | 0.032484 | -0.056487 | -0.0021981 | -0.014796 | -0.023631 | -0.057858 | 0.0092271 | 0.0079331 | 0.0043634 | 0.013975 | 0.028337 | 0.046087 | 0.044514 | 0.0088818 | -0.15739 | 0 | b10 | 0.047523 | 0.0056115 | -0.013403 | 0.040452 | 0.025717 | 0.065566 | -0.029107 | -0.010172 | -0.024428 | -0.033982 | -0.038155 | 0.025678 | 0.016471 | -0.048319 | -0.054459 | -0.01109 | 0.031932 | 0.091578 | 0.047255 | -0.17795 | 0
b11 | 0.57371 | 0.64828 | 0.64227 | 0.64414 | 0.64953 | 0.65094 | 0.64227 | 0.64896 | 0.6452 | 0.64456 | 0.65361 | 0.65414 | 0.65163 | 0.64369 | 0.64532 | 0.64565 | 0.64354 | 0.64063 | 0.64098 | 0.59874 | 0.64285 | b11 | 0.57964 | 0.64625 | 0.63336 | 0.64081 | 0.64198 | 0.64123 | 0.64353 | 0.6453 | 0.64963 | 0.64218 | 0.64955 | 0.64989 | 0.6493 | 0.6515 | 0.64815 | 0.64332 | 0.63595 | 0.6377 | 0.63985 | 0.59709 | 0.64073
b12 | 0.011285 | 0.0038805 | -0.0023626 | 0.0010811 | -0.0042571 | 0.00062981 | -0.0072512 | -0.002726 | -0.0028932 | -0.0010375 | -0.0062509 | 0.0008262 | 0.002423 | 0.001744 | 0.00086968 | 0.004633 | 0.0088901 | 0.0073326 | 0.0048264 | -0.016027 | 0.00037667 | b12 | 0.009428 | 0.0039379 | -0.00251 | 0.0039287 | -0.0016293 | 0.0067805 | -0.0042838 | -0.0032926 | -0.003815 | -0.0021697 | -0.003884 | 0.0021659 | 0.0026308 | -0.0081159 | -0.0072753 | -0.0046797 | 0.0030764 | 0.0094764 | 0.00501 | -0.0092103 | -0.00030468
b13 | 0.082853 | 0.063374 | 0.0687 | 0.064169 | 0.06627 | 0.065524 | 0.066952 | 0.0673 | 0.064633 | 0.062757 | 0.062496 | 0.065084 | 0.065355 | 0.062049 | 0.063698 | 0.063968 | 0.063699 | 0.06133 | 0.061816 | 0.045999 | 0.064368 | b13 | 0.093291 | 0.063449 | 0.063518 | 0.063304 | 0.063472 | 0.060933 | 0.062017 | 0.064253 | 0.062314 | 0.062638 | 0.064795 | 0.062865 | 0.06685 | 0.065365 | 0.06491 | 0.065485 | 0.064435 | 0.062763 | 0.06189 | 0.045608 | 0.063789
b14 | 0.002887 | -0.0017543 | -0.0031583 | -0.0010998 | -0.0034373 | -0.0015931 | -0.0047707 | -0.0027781 | -0.0034029 | -0.0034046 | -0.0045407 | -0.0014137 | -0.0010388 | -0.001454 | -0.001263 | -0.0010223 | 0.00036562 | 0.00049541 | -0.0013006 | 0.00015146 | -0.0018232 | b14 | 0.0043663 | -0.0023128 | -0.0024771 | -0.00014876 | -0.001591 | 0.0014844 | -0.0029035 | -0.0025345 | -0.0017074 | -0.0028053 | -0.003419 | -0.00020376 | -0.0016912 | -0.0045045 | -0.0042056 | -0.0037005 | -7.7075e-05 | 0.0010257 | 0.00027677 | -0.0022005 | -0.0016102
b15 | -0.01212 | 0.021877 | 0.019803 | 0.019278 | 0.018214 | 0.018504 | 0.016047 | 0.019164 | 0.01846 | 0.018825 | 0.017618 | 0.020015 | 0.019569 | 0.01791 | 0.017163 | 0.018836 | 0.019652 | 0.020472 | 0.019154 | -0.0053507 | 0.017282 | b15 | -0.013788 | 0.023128 | 0.020775 | 0.018997 | 0.020707 | 0.018315 | 0.018069 | 0.018943 | 0.019302 | 0.018457 | 0.018696 | 0.020252 | 0.020009 | 0.01971 | 0.018528 | 0.019869 | 0.017952 | 0.019468 | 0.020429 | -0.00037122 | 0.01805
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 55.1908 | -3.6626 | -1.7473 | 5.4611 | 1.8211 | -1.2869 | -3.0781 | -4.3454 | -7.5957 | -3.5462 | -3.0999 | -0.53296 | 3.1351 | 3.0426 | 0.79158 | -1.1315 | 0.46418 | -5.9318 | -6.0434 | -4.671 | 0 | a1 | 64.4822 | -3.4176 | -8.9955 | 4.8048 | -0.39006 | -1.3517 | -1.1058 | 0.65165 | -1.8315 | -0.7698 | -2.2086 | -1.1083 | 3.4552 | -0.1205 | -1.0532 | -3.4207 | 3.2755 | -0.33236 | -5.106 | -19.9075 | 0
a2 | 0.44884 | 0.092078 | 0.40073 | 0.71258 | 1.5312 | 1.9257 | 1.687 | 0.98203 | 0.99767 | 0.84226 | 0.19251 | 1.3958 | 1.0703 | 1.5033 | 1.2764 | 1.0807 | 1.3039 | 1.6466 | 0.63711 | 1.9626 | 1.0803 | a2 | -4.5991 | -0.3103 | -0.64926 | -0.23987 | -0.10693 | 0.40652 | 0.22323 | 0.44007 | 0.73272 | 0.32157 | 0.355 | 0.23979 | 0.98192 | -0.10263 | 0.62269 | 0.35346 | 0.77681 | 0.10386 | 0.25387 | -1.7874 | 0.056946
a3 | -1.2087 | 0.41924 | -0.094515 | -0.30976 | -0.38226 | -0.60088 | -0.27008 | -0.61541 | 0.13345 | 0.21985 | -0.472 | -0.54458 | -0.70152 | -0.44897 | -0.021204 | -0.25667 | -0.59622 | -0.29138 | -0.10697 | -0.31853 | -0.30282 | a3 | -2.694 | -0.024508 | -0.37513 | -0.24649 | -0.24577 | -0.33256 | -0.21013 | -0.11823 | -0.15924 | -0.1022 | -0.4327 | -0.42444 | -0.54605 | -0.23344 | -0.21337 | -0.38287 | -0.033356 | 0.1173 | 0.30868 | 0.078067 | -0.25665
a4 | -0.15598 | 0.42561 | -0.39824 | -0.26169 | -0.091126 | -0.017281 | -0.20106 | -0.058873 | -0.088101 | 0.11828 | 0.038917 | -0.14447 | 0.14549 | -0.20778 | -0.18202 | -0.021917 | -0.077347 | -0.070687 | 0.1938 | 0.067448 | -0.049412 | a4 | -0.065292 | 0.2005 | 0.30009 | 0.28452 | 0.10871 | 0.48981 | 0.38962 | 0.40295 | 0.22517 | 0.16898 | 0.19919 | 0.14702 | 0.23105 | 0.28301 | 0.17244 | 0.39403 | 0.24794 | 0.29708 | 0.51963 | 0.14173 | 0.26798
a5 | 1.8939 | 0.13303 | -0.12108 | -0.087476 | -0.045937 | -0.13456 | -0.067156 | 0.078313 | 0.041454 | -0.11748 | -0.0013283 | 0.080275 | 0.015022 | -0.16364 | -0.22558 | -0.23114 | -0.12511 | 0.033954 | -0.021758 | -0.40461 | -0.0078968 | a5 | 2.2214 | 0.11464 | -0.026482 | -0.032829 | -0.1134 | -0.089182 | 0.0041081 | 0.10657 | -0.038085 | -0.11953 | -0.076438 | 0.009303 | 0.031347 | -0.0074123 | 0.051826 | -0.048558 | 0.0074494 | -0.1298 | 0.074505 | -0.1924 | 0.034921
a6 | -0.013542 | -0.044409 | -0.10229 | -0.044776 | 0.012476 | 0.014845 | 0.060284 | 0.043136 | 0.0051632 | -0.026847 | -0.016935 | 0.041871 | 0.10039 | 0.13319 | 0.12605 | 0.07368 | -0.038909 | -0.096073 | -0.057348 | -0.025676 | 0.0084159 | a6 | -0.13738 | 0.047072 | 0.04241 | 0.11016 | 0.069885 | 0.11897 | 0.15078 | 0.12887 | 0.10058 | 0.027665 | 0.0017691 | 0.064369 | 0.093448 | 0.059481 | 0.10728 | 0.040853 | 0.015778 | -0.036989 | 0.074475 | 0.21468 | 0.067203
a7 | 1.4251 | -0.055549 | -0.049573 | -0.10873 | -0.11013 | -0.066009 | -0.10402 | -0.10628 | -0.070242 | -0.028186 | -0.078243 | -0.064079 | -0.077793 | -0.073536 | -0.026365 | -0.03992 | -0.085155 | -0.10138 | -0.043491 | -0.0070354 | -0.026446 | a7 | 1.634 | -0.044679 | -0.065689 | -0.026042 | -0.063147 | -0.032017 | -0.015449 | 0.01133 | -0.1069 | -0.0768 | -0.077611 | -0.046045 | -0.036308 | -0.030047 | -0.042231 | -0.038366 | 0.0059785 | -0.02316 | -0.026783 | -0.059903 | 0.00078749
a8 | -0.041803 | 0.0085266 | 0.028435 | 0.0053099 | 0.014046 | 0.024576 | 0.027674 | 0.00087246 | -0.0057546 | -0.014179 | -0.024069 | 0.0068192 | 0.016318 | 0.049274 | 0.054898 | 0.033443 | -0.012572 | -0.024172 | -0.03378 | 0.0059026 | 0.0071087 | a8 | -0.054256 | -0.016427 | -0.0076877 | 0.043036 | 0.049588 | 0.069499 | 0.055859 | 0.045465 | 0.0068377 | -0.019747 | 0.0010809 | 0.010188 | 0.025376 | 0.012581 | 0.012889 | -0.00023908 | 0.017212 | -0.017416 | 0.010923 | 0.023584 | 0.015057
a9 | -0.22661 | 0.018031 | -0.024429 | -0.022902 | -0.014807 | -0.026067 | -0.01955 | -0.035304 | -0.030795 | -0.010754 | -0.030534 | -0.042671 | -0.038825 | -0.020423 | -0.031118 | -0.030521 | -0.036885 | -0.035486 | -0.037699 | -0.072482 | -0.033364 | a9 | -0.28767 | -0.010566 | -0.0078771 | -0.01581 | -0.014959 | -0.019176 | -0.028144 | -0.020213 | -0.0052786 | -0.010451 | -0.030896 | -0.021956 | -0.024127 | -0.017679 | -0.020088 | -0.0127 | -0.030056 | -0.046636 | -0.024023 | -0.054632 | -0.02789
a10 | 0.00067424 | 0.012036 | -0.023697 | -0.0139 | -0.0067629 | 0.014198 | -0.0014332 | 0.0055307 | 0.0099066 | -0.025809 | -0.0251 | -0.016768 | 0.010809 | 0.037259 | 0.059902 | 0.029421 | -0.0089298 | -0.041029 | -6.2547e-05 | -0.027225 | 0 | a10 | -0.036187 | -0.022356 | -0.018704 | 0.011127 | -0.0073972 | 0.017132 | 0.020716 | 0.039329 | -0.0065814 | -0.0043465 | -0.011045 | -0.0067508 | 0.024654 | 0.0051049 | 0.033351 | -0.0090173 | 0.023804 | -0.047863 | -0.023337 | -0.00077373 | 0
a11 | 0.16625 | -0.031447 | -0.033313 | -0.031877 | -0.031254 | -0.031746 | -0.0299 | -0.033347 | -0.034751 | -0.033921 | -0.03589 | -0.032706 | -0.028912 | -0.028457 | -0.027232 | -0.029151 | -0.031385 | -0.032971 | -0.030795 | -0.0098189 | -0.025255 | a11 | 0.20818 | -0.040698 | -0.052689 | -0.039946 | -0.039727 | -0.038237 | -0.037546 | -0.037516 | -0.044741 | -0.041708 | -0.04087 | -0.038858 | -0.033231 | -0.038885 | -0.034788 | -0.037466 | -0.034233 | -0.040168 | -0.04168 | -0.054267 | -0.033834
a12 | -0.0089248 | -0.001797 | -0.0060157 | -0.0030879 | -0.0012316 | 0.0031275 | 0.0031032 | 0.0025457 | 0.0017278 | -0.0032204 | -0.0059969 | -0.00081936 | 0.00059147 | 0.0059642 | 0.0070501 | 0.0026146 | -0.0019804 | -0.0052248 | -0.0035367 | -0.0013986 | -0.00062346 | a12 | -0.0069619 | -0.0016131 | -0.0037708 | 0.0031981 | -0.0019877 | 0.0016114 | 0.00045747 | 0.0026154 | 0.0011921 | 0.00019258 | -0.0035626 | -6.6129e-05 | 0.0055919 | 0.00090502 | 0.0021967 | -0.00264 | 0.00097104 | -0.0085992 | -0.0013411 | 0.0044649 | -0.00027405
a13 | -0.018624 | -0.0013983 | -0.0033505 | -0.0040212 | -0.0023291 | -0.0045322 | -0.0044305 | -0.0038878 | -0.0030542 | -0.0029052 | -0.0038237 | -0.0029824 | -0.0044211 | -0.0025114 | -0.0029423 | -0.0050235 | -0.0051716 | -0.0039356 | -0.0032575 | -0.0043674 | -0.0040162 | a13 | -0.021554 | -0.00091634 | -0.0037939 | -0.0023078 | -0.0053633 | -0.0027311 | -0.0047278 | -0.0035165 | -0.0050367 | -0.0018659 | -0.0031601 | -0.002019 | -0.0022153 | -0.0028856 | -0.0032607 | -0.0039734 | -0.0028761 | -0.0031801 | -0.0027173 | -0.0037383 | -0.0036234
a14 | -0.00022275 | -0.0024597 | -0.0046677 | -0.0051677 | -0.0047872 | -0.0037061 | -0.0019323 | -0.0037981 | -0.0038997 | -0.0047753 | -0.0043329 | -0.0048226 | -0.0037079 | -0.0038311 | -0.0034256 | -0.0035817 | -0.0042748 | -0.0053347 | -0.0042304 | -0.0016131 | -0.0038565 | a14 | 0.0075644 | -0.0053303 | -0.0044191 | -0.0017581 | -0.0033311 | -0.00068131 | -0.0011837 | -0.0010452 | -0.0012635 | -0.0012561 | -0.001925 | -0.0025053 | -0.0027209 | -0.0011743 | -0.00069602 | -0.0025128 | -0.00038754 | -0.002656 | -0.002247 | -9.2457e-05 | -0.0017542
a15 | 0.005158 | -0.0016851 | -0.0050099 | -0.0029169 | -0.0035926 | -0.0031995 | -0.005227 | -0.0032905 | -0.002782 | -0.0042486 | -0.0052909 | -0.0037097 | -0.0045022 | -0.0023069 | -0.0014223 | -0.00174 | -0.0015861 | -0.0017026 | -0.0015312 | -0.0073043 | -0.002988 | a15 | 0.0054963 | -0.0025336 | -0.0037368 | -0.0010388 | -0.0022123 | -0.00106 | -0.003067 | -0.0020927 | -0.003156 | -0.0033632 | -0.0030637 | -0.0011459 | -0.0024987 | -0.0029235 | -0.0024996 | -0.0030393 | -0.0014168 | 0.00048552 | -0.0010651 | -0.005 | -0.0020873
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.11844 | 0.072796 | -0.036541 | 0.069142 | -0.029038 | 0.082821 | -0.14908 | 0 | -0.040274 | -0.058573 | -0.14643 | 0.026093 | 0.019406 | 0 | 0.030078 | 0.070958 | 0.12265 | 0.12337 | 0.023511 | -0.4466 | -0.18745 | Dx (mm) | 0.089576 | 0.018386 | -0.031573 | 0.10507 | 0.06904 | 0.17055 | -0.07982 | -0.032693 | -0.06245 | -0.08884 | -0.097651 | 0.067983 | 0.039703 | -0.12648 | -0.14714 | -0.027822 | 0.081698 | 0.25115 | 0.12906 | -0.50297 | -0.05688
Dy (mm) | -0.032421 | 0.035095 | -0.064615 | -0.033268 | -0.0191 | 0.041123 | -0.010707 | 0.014155 | 0.023936 | -0.071152 | -0.073408 | -0.042273 | 0.029061 | 0.098697 | 0.15902 | 0.080688 | -0.017634 | -0.10261 | 0 | -0.085179 | -0.19448 | Dy (mm) | -0.13839 | -0.057653 | -0.052832 | 0.036092 | -0.015317 | 0.055643 | 0.05008 | 0.10171 | -0.021525 | -0.017276 | -0.035073 | -0.013595 | 0.06658 | 0 | 0.079645 | -0.025452 | 0.06804 | -0.11212 | -0.053621 | -0.047398 | -0.20692
Mag. len. (m) | 14.359 | [20 empty] | Mag. len. (m) | 14.3593
TF (mT/KA) | 704.7339 | [20 empty] | TF (mT/KA) | 704.585
Coil pos. (mm) | 0.017709 | [20 empty] | Coil pos. (mm) | 0.015152
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.01705 | [20 empty] | Twist (m2 rad) | 0.022439

## Harmonics
Normal:
Skew:
*10

## Work sheet
| HCMBARA001-3000024 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 1.4631 | 7.125 | 0.092642 | 0.30933 | 1.0803 | 0.55165 | 1.0703 | 0.51644 | -1.6247 | 7.5246 | -0.26785 | 0.32756 | 0.056946 | 1.2169 | 0.24291 | 0.39854 |  | 0
3 | 5.8526 | 7.6044 | 4.7359 | 0.53041 | -0.30282 | 0.36333 | -0.27479 | 0.31387 | 5.5535 | 9.0347 | 4.4219 | 0.38879 | -0.25665 | 0.59767 | -0.20442 | 0.20716 |  | 0
4 | 0.014481 | 0.30478 | -0.041116 | 0.097888 | -0.049412 | 0.18231 | -0.049815 | 0.19025 | 0.014782 | 0.33051 | 0.063178 | 0.1608 | 0.26798 | 0.13768 | 0.28148 | 0.11711 |  | 0
5 | -0.27425 | 1.2802 | -0.043388 | 0.19926 | -0.0078968 | 0.45757 | -0.053242 | 0.10614 | -0.35028 | 1.458 | -0.11212 | 0.16343 | 0.034921 | 0.50867 | -0.01569 | 0.07389 |  | 0
6 | -0.0072409 | 0.069999 | -0.010928 | 0.052617 | 0.0084159 | 0.067304 | 0.0094681 | 0.054357 | 0.03846 | 0.068354 | 0.038477 | 0.053254 | 0.067203 | 0.073766 | 0.066106 | 0.037594 |  | 0
7 | 0.98908 | 0.34011 | 0.95382 | 0.070107 | -0.026446 | 0.33524 | -0.071603 | 0.027215 | 0.9743 | 0.41566 | 0.93789 | 0.07447 | 0.00078749 | 0.37577 | -0.040888 | 0.029695 |  | 0
8 | -0.0066331 | 0.029965 | -0.0067502 | 0.016037 | 0.0071087 | 0.02627 | 0.0083285 | 0.018217 | 0.002381 | 0.026747 | 0.0019718 | 0.016776 | 0.015057 | 0.029723 | 0.016437 | 0.020819 |  | 0
9 | 0.40694 | 0.028014 | 0.41006 | 0.01364 | -0.033364 | 0.047367 | -0.026531 | 0.013728 | 0.40027 | 0.017486 | 0.39978 | 0.0082408 | -0.02789 | 0.060671 | -0.019328 | 0.0084714 |  | 0
10 | 0 | 0.047056 | 1.4888e-05 | 0.00089041 | 0 | 0.024401 | 2.5132e-05 | 9.4552e-05 | 0 | 0.057487 | 8.8312e-06 | 3.6146e-05 | 0 | 0.022997 | 0.00012763 | 0.00052115 |  | 0
11 | 0.64285 | 0.019381 | 0.64635 | 0.0042575 | -0.025255 | 0.044309 | -0.031643 | 0.0023375 | 0.64073 | 0.018007 | 0.64364 | 0.0052555 | -0.033834 | 0.055826 | -0.039503 | 0.0043767 |  | 0
12 | 0.00037667 | 0.0061534 | 0.00014611 | 0.0020766 | -0.00062346 | 0.0041708 | -0.00042598 | 0.0019158 | -0.00030468 | 0.005614 | -0.00081788 | 0.0017514 | -0.00027405 | 0.0035631 | -0.00039216 | 0.0017106 |  | 0
13 | 0.064368 | 0.0062784 | 0.06435 | 0.0020157 | -0.0040162 | 0.003492 | -0.0035212 | 0.00096847 | 0.063789 | 0.0080652 | 0.063542 | 0.0014628 | -0.0036234 | 0.0042519 | -0.0030886 | 0.001169 |  | 0
14 | -0.0018232 | 0.0018485 | -0.002213 | 0.00057029 | -0.0038565 | 0.0012901 | -0.0040996 | 0.00074217 | -0.0016102 | 0.0021785 | -0.0019875 | 0.00067466 | -0.0017542 | 0.0025072 | -0.0021348 | 0.00097145 |  | 0
15 | 0.017282 | 0.008672 | 0.01892 | 0.0013125 | -0.002988 | 0.0024717 | -0.0030969 | 0.0013515 | 0.01805 | 0.0085599 | 0.019534 | 0.0012655 | -0.0020873 | 0.0021314 | -0.0021904 | 0.0010989 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.91706 | -1.018
3 | [12 empty] | 3 | 9 | -0.89644 | -1.2595
4 | [12 empty] | 4 | 16 | -1.9917 | -1.9861
5 | [12 empty] | 5 | 25 | -1.9281 | -2.2083
6 | [12 empty] | 6 | 36 | -2.9284 | -3.1068
7 | [12 empty] | 7 | 49 | -3.1309 | -3.0571
8 | [12 empty] | 8 | 64 | -4.0691 | -3.9799
9 | [12 empty] | 9 | 81 | -4.2915 | -4.7849
10 | [12 empty] | sigma (mm) |  | 0.017709 | 0.015152
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05
17 | 0.00044619 | 0.00011155 | 2.6771e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 24.6639 | -0.1613 | -0.067671 | 0.092603 | -0.11949 | 0.029137 | -0.3912 | -0.29175 | -0.2376 | 0.16649 | 0.0076779 | 0.20807 | 0.92437 | 0.46384 | 0.094703 | 0.27191 | 0.36646 | 0.041772 | 0.26953 | 21.5376 | 1.4631
3 | 38.5895 | 3.3818 | 4.1964 | 4.6944 | 4.821 | 4.6265 | 4.5803 | 5.6584 | 5.598 | 4.99 | 4.2276 | 4.458 | 5.1902 | 4.5899 | 4.4969 | 4.6567 | 5.2431 | 4.8777 | 4.9597 | 9.393 | 5.8526
4 | 0.74318 | -0.022153 | -0.047834 | -0.06971 | -0.1733 | -0.096633 | -0.087025 | 0.1019 | 0.011373 | 0.17271 | -0.038621 | -0.015866 | 0.083226 | -0.066347 | -0.033286 | 0.013766 | -0.09308 | -0.14799 | -0.23123 | 0.83823 | 0.014481
5 | -5.1685 | -0.24146 | -0.60927 | -0.23207 | -0.16955 | -0.0639 | -0.2181 | -0.11979 | -0.094344 | 0.082609 | 0.14073 | 0.087365 | 0.080414 | 0.10265 | 0.029961 | 0.032186 | 0.098845 | 0.22694 | 0.085795 | -2.7638 | -0.27425
6 | 0.019194 | 0.024515 | 0.017599 | 0.034041 | 0.086863 | 0.013926 | 0.076615 | -0.014304 | -0.055349 | -0.023702 | -0.091621 | 0.0060358 | -0.032181 | -0.029434 | -0.019922 | -0.012154 | -0.0062029 | -0.040461 | -0.13097 | -0.01225 | -0.0072409
7 | 2.4156 | 0.75258 | 0.78917 | 0.95285 | 0.96133 | 1.01 | 0.99727 | 0.94952 | 0.98205 | 0.99008 | 0.96767 | 0.93292 | 0.96781 | 0.98983 | 0.97364 | 0.97554 | 0.97424 | 1.0216 | 0.9806 | 0.70929 | 0.98908
8 | 0.011536 | -0.031866 | -0.01447 | -0.016534 | -0.024477 | -0.0091049 | -0.015986 | -0.0026517 | 0.0079398 | 0.0095258 | -0.013612 | -0.010546 | -0.0044996 | -0.0093263 | -0.028648 | -0.0053877 | 0.0097548 | 0.0032177 | 0.035167 | -0.0020528 | -0.0066331
9 | 0.29691 | 0.38971 | 0.37865 | 0.41411 | 0.4183 | 0.41097 | 0.41499 | 0.43297 | 0.43531 | 0.41164 | 0.40649 | 0.40883 | 0.42413 | 0.41204 | 0.40233 | 0.40034 | 0.39923 | 0.41084 | 0.41016 | 0.36519 | 0.40694
10 | -1.1158e-05 | -4.384e-06 | -5.4214e-06 | 2.1528e-06 | 1.4711e-07 | 4.0275e-06 | 2.2474e-05 | -0.0024758 | -2.4271e-08 | -3.5433e-06 | -8.3325e-06 | 1.601e-07 | 2.2352e-07 | 0.0027113 | 9.975e-06 | 4.2096e-06 | -1.8549e-05 | 1.2451e-05 | 1.694e-05 | 0.0003925 | 0
11 | 0.57303 | 0.64806 | 0.64247 | 0.64416 | 0.64947 | 0.65098 | 0.64155 | 0.64898 | 0.64515 | 0.64467 | 0.6533 | 0.65415 | 0.65161 | 0.64407 | 0.64603 | 0.64558 | 0.64286 | 0.6404 | 0.64091 | 0.59418 | 0.64285
12 | 0.047843 | 0.0058937 | -0.0043781 | -0.019563 | -0.028674 | -0.033324 | -0.0017234 | -0.027649 | -0.011074 | 0.017031 | 0.0040699 | -0.0028356 | 0.01437 | 0.01569 | -0.00813 | 0.011428 | 0.034395 | 0.022766 | 0.038005 | -0.012636 | 0.0037667 | ten times!
13 | 0.82597 | 0.63405 | 0.68842 | 0.64358 | 0.66264 | 0.65509 | 0.66424 | 0.67259 | 0.64457 | 0.62864 | 0.6223 | 0.65268 | 0.65288 | 0.6176 | 0.6331 | 0.63802 | 0.63723 | 0.61702 | 0.61839 | 0.46156 | 0.64368 | ten times!
14 | 0.039314 | -0.031145 | -0.022958 | -0.021176 | -0.029453 | -0.029636 | -0.027545 | -0.028165 | -0.028455 | -0.022476 | -0.020962 | -0.017146 | -0.014593 | -0.016415 | -0.018744 | -0.022386 | -0.015963 | -0.014407 | -0.016715 | -0.013041 | -0.018232 | ten times!
15 | -0.1212 | 0.21877 | 0.19803 | 0.19278 | 0.18214 | 0.18504 | 0.16047 | 0.19164 | 0.1846 | 0.18825 | 0.17618 | 0.20015 | 0.19569 | 0.1791 | 0.17163 | 0.18836 | 0.19652 | 0.20472 | 0.19154 | -0.053507 | 0.17282 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 0.61292 | 0.074537 | 0.43219 | 0.73346 | 1.5408 | 1.9092 | 1.688 | 0.9726 | 0.98254 | 0.88554 | 0.2209 | 1.4196 | 1.0542 | 1.4499 | 1.1923 | 1.0386 | 1.3234 | 1.7096 | 0.6374 | 2.0391 | 1.0803
3 | -1.202 | 0.41394 | -0.097589 | -0.307 | -0.3833 | -0.59993 | -0.27552 | -0.61567 | 0.13278 | 0.22321 | -0.47156 | -0.54403 | -0.70245 | -0.44788 | -0.019546 | -0.25666 | -0.59483 | -0.29239 | -0.10778 | -0.29639 | -0.30282
4 | -0.24816 | 0.42531 | -0.40857 | -0.26208 | -0.092205 | -0.014049 | -0.20397 | -0.058474 | -0.087181 | 0.11805 | 0.041403 | -0.14409 | 0.14488 | -0.21007 | -0.18132 | -0.018641 | -0.073321 | -0.066263 | 0.19392 | -0.030141 | -0.049412
5 | 1.8957 | 0.13352 | -0.12206 | -0.086017 | -0.045402 | -0.13546 | -0.064439 | 0.07837 | 0.041999 | -0.11888 | -0.0051614 | 0.080124 | 0.014661 | -0.16286 | -0.22632 | -0.233 | -0.12352 | 0.037451 | -0.021361 | -0.41116 | -0.0078968
6 | -0.045469 | -0.052295 | -0.084904 | -0.030935 | 0.017831 | 0.0021184 | 0.05859 | 0.038392 | -0.0041322 | -0.0025755 | 0.0041405 | 0.056381 | 0.090997 | 0.098781 | 0.071901 | 0.046922 | -0.02915 | -0.054651 | -0.056987 | -0.0049614 | 0.0084159
7 | 1.4275 | -0.055533 | -0.049723 | -0.10893 | -0.1102 | -0.066973 | -0.10252 | -0.10626 | -0.070338 | -0.028592 | -0.080992 | -0.064263 | -0.077923 | -0.073196 | -0.02569 | -0.0412 | -0.084283 | -0.099082 | -0.043164 | -0.009081 | -0.026446
8 | -0.024601 | 0.0014583 | 0.039512 | 0.012549 | 0.017605 | 0.017617 | 0.028416 | -0.0020113 | -0.011244 | -0.0007002 | -0.01201 | 0.015471 | 0.010873 | 0.030227 | 0.025593 | 0.019268 | -0.0071058 | -0.0022409 | -0.033363 | 0.0067501 | 0.0071087
9 | -0.22591 | 0.017039 | -0.025403 | -0.021943 | -0.015027 | -0.027396 | -0.019984 | -0.035287 | -0.030396 | -0.012445 | -0.034728 | -0.042233 | -0.039059 | -0.020651 | -0.033249 | -0.032837 | -0.035874 | -0.030388 | -0.037698 | -0.086017 | -0.033364
10 | 3.2529e-05 | -9.284e-07 | -1.1536e-06 | 1.8204e-06 | 8.8643e-07 | -1.6455e-06 | -1.0137e-05 | 0.00012719 | -1.1106e-06 | -4.9698e-07 | 3.1922e-05 | 1.5428e-07 | -4.3659e-07 | -0.00011175 | 2.1208e-05 | -7.1485e-06 | 1.022e-05 | 3.0439e-05 | 0.00036335 | 0.00019515 | 0
11 | 0.16717 | -0.03145 | -0.033554 | -0.031716 | -0.03133 | -0.03193 | -0.02965 | -0.033322 | -0.034661 | -0.034091 | -0.036755 | -0.03267 | -0.028965 | -0.028569 | -0.027458 | -0.029513 | -0.031127 | -0.032067 | -0.030741 | -0.011106 | -0.025255 | ten times!
12 | -0.054715 | -0.032951 | -0.029687 | -0.013847 | -0.0038586 | 0.014904 | 0.031429 | 0.018732 | 0.0054891 | -0.0018853 | -0.031538 | 0.011777 | -0.0068865 | 0.016413 | -0.00037264 | -0.0077715 | -0.0073972 | -0.0044002 | -0.034826 | -9.6247e-05 | -0.0062346 | ten times!
13 | -0.18533 | -0.012143 | -0.03637 | -0.03776 | -0.024856 | -0.042473 | -0.046898 | -0.038578 | -0.03112 | -0.033043 | -0.045637 | -0.029319 | -0.04343 | -0.024016 | -0.027099 | -0.047661 | -0.047658 | -0.033934 | -0.031814 | -0.049084 | -0.040162 | ten times!
14 | -0.010494 | -0.02991 | -0.037647 | -0.044735 | -0.045866 | -0.041145 | -0.024325 | -0.040215 | -0.043559 | -0.038772 | -0.039058 | -0.040461 | -0.041043 | -0.052868 | -0.056379 | -0.047317 | -0.038292 | -0.034319 | -0.042008 | -0.046748 | -0.038565 | ten times!
15 | 0.05158 | -0.016851 | -0.050099 | -0.029169 | -0.035926 | -0.031995 | -0.05227 | -0.032905 | -0.02782 | -0.042486 | -0.052909 | -0.037097 | -0.045022 | -0.023069 | -0.014223 | -0.017399 | -0.015861 | -0.017026 | -0.015312 | -0.073043 | -0.02988 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -28.3169 | -0.6605 | -0.074137 | -0.10372 | -0.23136 | -0.31292 | -1.1407 | -0.66086 | 0.2114 | -0.31635 | -0.10793 | -0.4912 | -0.35042 | -0.19734 | -0.23227 | -0.10885 | -0.1683 | -0.17483 | 0.29893 | -20.4557 | -1.6247
3 | 44.8441 | 4.4197 | 4.6466 | 4.1403 | 4.5333 | 3.7904 | 4.4407 | 4.5968 | 3.9451 | 5.135 | 4.4683 | 3.9267 | 4.3259 | 3.8482 | 4.4608 | 4.4911 | 5.0598 | 4.9348 | 4.4309 | 7.9058 | 5.5535
4 | -1.1032 | -0.03251 | 0.13619 | 0.21864 | 0.35751 | 0.061681 | -0.15091 | 0.057543 | 0.024926 | -0.087737 | -0.12566 | -0.20986 | -0.060393 | 0.21552 | 0.19665 | 0.29708 | 0.15832 | 0.12026 | -0.040025 | -0.80075 | 0.014782
5 | -6.1686 | -0.50233 | -0.44701 | -0.14696 | -0.013212 | -0.030393 | 0.027375 | -0.096405 | 0.035663 | -0.087103 | 0.087541 | 0.1048 | -0.090778 | -0.20085 | -0.14879 | -0.23243 | -0.1847 | 0.0027407 | -0.095395 | -2.7865 | -0.35028
6 | -0.057417 | 0.080537 | 0.073083 | 0.054379 | 0.060642 | 0.034874 | 0.11159 | 0.084242 | -0.0074047 | 0.081032 | 0.0016677 | -0.092501 | 0.043892 | 0.10282 | 0.035914 | 0.062259 | 0.0019224 | 0.014399 | -0.050768 | -0.040573 | 0.03846
7 | 2.7369 | 0.68579 | 0.83612 | 0.97337 | 0.96701 | 1.0154 | 1.0045 | 0.929 | 0.95481 | 0.95235 | 0.96593 | 0.95203 | 0.96405 | 0.91517 | 0.92836 | 0.93135 | 0.9394 | 1.0026 | 0.96469 | 0.64623 | 0.9743
8 | -0.026678 | 0.0072215 | 0.011068 | 0.012206 | 0.039883 | 0.019699 | -0.019542 | -0.021953 | 0.0029691 | 0.015991 | 0.013093 | 0.0029251 | -0.0075355 | -0.0015656 | -8.9872e-05 | 0.0056325 | -0.0029814 | -0.0088981 | -0.03263 | 0.07267 | 0.002381
9 | 0.33401 | 0.41446 | 0.40195 | 0.39715 | 0.40145 | 0.3832 | 0.39966 | 0.40147 | 0.3948 | 0.41366 | 0.3983 | 0.39769 | 0.40602 | 0.40839 | 0.40657 | 0.39633 | 0.38578 | 0.39351 | 0.39571 | 0.37204 | 0.40027
10 | 3.3687e-05 | 2.3922e-06 | -2.5891e-06 | -4.2158e-07 | 2.9531e-06 | -1.7939e-05 | 2.8778e-06 | -6.419e-06 | 2.6078e-06 | 1.469e-06 | -2.9874e-07 | -1.0407e-06 | 5.6604e-06 | 0.00015083 | 1.0933e-05 | -5.1133e-07 | -7.0751e-07 | 1.6197e-05 | -7.0267e-06 | 0.00022139 | 0
11 | 0.57972 | 0.64627 | 0.63344 | 0.64061 | 0.64207 | 0.64054 | 0.64333 | 0.64541 | 0.64946 | 0.64205 | 0.64939 | 0.6498 | 0.64947 | 0.65083 | 0.64757 | 0.64327 | 0.63583 | 0.63678 | 0.63948 | 0.59396 | 0.64073
12 | 0.056347 | 0.031517 | -0.0095284 | -0.0082497 | -0.046646 | -0.0066255 | -0.0095673 | -0.020623 | -0.0099151 | 0.017811 | 0.0066061 | -0.0083155 | 0.0065313 | -0.022804 | -0.0071679 | -0.033222 | -0.0077767 | -0.013988 | -0.005255 | 0.071073 | -0.0030468 | ten times!
13 | 0.92192 | 0.63716 | 0.63637 | 0.63268 | 0.63595 | 0.60711 | 0.61794 | 0.64109 | 0.62253 | 0.62464 | 0.64591 | 0.62902 | 0.66763 | 0.64929 | 0.64394 | 0.65455 | 0.6442 | 0.62793 | 0.61955 | 0.44765 | 0.63789 | ten times!
14 | 0.047571 | -0.025427 | -0.017743 | -0.018234 | -0.027404 | -0.011367 | -0.018423 | -0.021998 | -0.0065874 | -0.014071 | -0.01827 | -0.013248 | -0.024824 | -0.024516 | -0.021244 | -0.031816 | -0.013643 | -0.030458 | -0.018476 | -0.021591 | -0.016102 | ten times!
15 | -0.13788 | 0.23128 | 0.20775 | 0.18997 | 0.20707 | 0.18316 | 0.18069 | 0.18943 | 0.19302 | 0.18457 | 0.18696 | 0.20252 | 0.20009 | 0.1971 | 0.18528 | 0.19869 | 0.17952 | 0.19468 | 0.20429 | -0.0037122 | 0.1805 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -3.8409 | -0.28026 | -0.62176 | -0.25441 | -0.096757 | 0.3884 | 0.1951 | 0.3847 | 0.74154 | 0.33094 | 0.36848 | 0.24946 | 0.95061 | -0.1061 | 0.57727 | 0.36566 | 0.73664 | 0.16563 | 0.27716 | -1.7385 | 0.056946
3 | -2.7213 | -0.025495 | -0.37216 | -0.25314 | -0.24613 | -0.3479 | -0.20331 | -0.11691 | -0.15666 | -0.099816 | -0.43006 | -0.42671 | -0.54694 | -0.22712 | -0.21171 | -0.37959 | -0.038788 | 0.10646 | 0.29645 | 0.086664 | -0.25665
4 | -0.31307 | 0.19321 | 0.29434 | 0.28657 | 0.11051 | 0.49375 | 0.38942 | 0.40619 | 0.2248 | 0.16612 | 0.19817 | 0.14721 | 0.2322 | 0.28279 | 0.17706 | 0.39231 | 0.25074 | 0.30499 | 0.51618 | 0.088182 | 0.26798
5 | 2.2294 | 0.11581 | -0.025116 | -0.037197 | -0.11444 | -0.096733 | 0.006417 | 0.10562 | -0.036423 | -0.11855 | -0.076724 | 0.0077365 | 0.029115 | -0.0051997 | 0.056731 | -0.047829 | 0.0064927 | -0.12368 | 0.071562 | -0.16284 | 0.034921
6 | -0.055564 | 0.06131 | 0.057264 | 0.098717 | 0.076661 | 0.10091 | 0.13259 | 0.095709 | 0.10548 | 0.031063 | 0.011054 | 0.070045 | 0.071319 | 0.05814 | 0.078951 | 0.048843 | -0.0069515 | 0.0048551 | 0.093947 | 0.21496 | 0.067203
7 | 1.6342 | -0.044305 | -0.065682 | -0.028382 | -0.064223 | -0.038075 | -0.012913 | 0.013078 | -0.1068 | -0.077533 | -0.077651 | -0.046243 | -0.036745 | -0.029392 | -0.040543 | -0.038366 | 0.00504 | -0.019686 | -0.027563 | -0.055561 | 0.00078749
8 | -0.020403 | -0.0051076 | 0.0021812 | 0.037038 | 0.052974 | 0.060862 | 0.045368 | 0.02603 | 0.010689 | -0.016808 | 0.0062642 | 0.01344 | 0.013122 | 0.011529 | -0.0038467 | 0.0043423 | 0.0059987 | 0.0093508 | 0.022432 | 0.020008 | 0.015057
9 | -0.28247 | -0.010178 | -0.0085647 | -0.017202 | -0.01448 | -0.022654 | -0.026497 | -0.018985 | -0.0057746 | -0.010966 | -0.032176 | -0.021528 | -0.025225 | -0.017337 | -0.015194 | -0.012982 | -0.032235 | -0.034836 | -0.021087 | -0.059303 | -0.02789
10 | 3.1167e-05 | 1.0597e-06 | 2.416e-07 | -8.4148e-06 | 3.7813e-07 | -2.0094e-05 | -4.5628e-06 | -1.7851e-06 | 5.3389e-07 | -9.2327e-08 | 7.9907e-06 | 2.6954e-07 | 4.9327e-07 | 0.002212 | -2.5237e-05 | 1.1959e-06 | -3.7819e-06 | 0.00012766 | 9.5881e-06 | -0.00018183 | 0
11 | 0.20942 | -0.040532 | -0.052851 | -0.040255 | -0.039654 | -0.038659 | -0.037384 | -0.037243 | -0.044746 | -0.041721 | -0.041183 | -0.038836 | -0.033488 | -0.038811 | -0.034204 | -0.037591 | -0.03442 | -0.038083 | -0.041394 | -0.053096 | -0.033834
12 | 0.035145 | 0.0098087 | -0.014866 | 0.017565 | -0.010401 | -0.0045311 | -0.020012 | -0.02079 | 0.019169 | 0.0083944 | -0.021763 | 0.0063403 | 0.025122 | 0.006474 | -0.017912 | -0.015415 | -0.019578 | -0.030683 | 0.01249 | 0.046635 | -0.0027405 | ten times!
13 | -0.2161 | -0.0094336 | -0.040006 | -0.021625 | -0.052061 | -0.027054 | -0.046889 | -0.033455 | -0.051251 | -0.019883 | -0.033955 | -0.018909 | -0.020466 | -0.029992 | -0.030829 | -0.040989 | -0.028479 | -0.025821 | -0.024842 | -0.038536 | -0.036234 | ten times!
14 | 0.055875 | -0.041939 | -0.036124 | -0.022328 | -0.029441 | -0.013717 | -0.021305 | -0.026883 | -0.010837 | -0.012396 | -0.016313 | -0.022145 | -0.037363 | -0.014788 | -0.022142 | -0.02166 | -0.012981 | -0.0095886 | -0.012316 | -0.02178 | -0.017542 | ten times!
15 | 0.054963 | -0.025336 | -0.037368 | -0.010388 | -0.022123 | -0.0106 | -0.03067 | -0.020927 | -0.03156 | -0.033632 | -0.030637 | -0.011459 | -0.024987 | -0.029235 | -0.024996 | -0.030393 | -0.014168 | 0.0048552 | -0.010651 | -0.05 | -0.020873 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -0.00076014 | -7.2289e-06 | 0.00010005 | 4.2435e-05 | 5.4638e-05 | 5.4497e-05 | 6.1733e-05 | 8.4437e-05 | 9.8343e-05 | 8.096e-06 | 9.6924e-05 | 1.6894e-05 | 0.00010203 | 3.5057e-05 | 1.0225e-05 | 0.00011012 | 9.3944e-05 | -0.00020205 |  | 704.7339
|  | -0.0026305 | 2.5429e-05 | 6.673e-05 | 5.5092e-05 | 8.8728e-05 | 0.0001692 | 0.00013386 | 0.00026174 | 0.00027891 | 0.0001895 | 0.00027664 | 0.00018112 | 0.0001939 | 9.2418e-05 | 0.00013911 | 0.00022058 | 0.00026188 | -4.3761e-06 |  | 704.585
Magnetic Length | -0.076459
mbh dB/B | -4313.1478
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.1292 | -4477.5581 | -7.6014 | -0.072289 | 1.0005 | 0.42435 | 0.54638 | 0.54497 | 0.61733 | 0.84437 | 0.98343 | 0.08096 | 0.96924 | 0.16894 | 1.0203 | 0.35057 | 0.10225 | 1.1012 | 0.93944 | -2.0205 | -4148.7375
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 5.6706 | -0.21467 | -0.023144 | 0.6977 | 0.3337 | 0.022903 | -0.15622 | -0.28295 | -0.60798 | -0.20303 | -0.1584 | 0.098293 | 0.4651 | 0.45585 | 0.23075 | 0.038441 | 0.19801 | -0.44159 | -0.45275 | -0.31551
db2 | 1.1256 | 27.0124 | 0.058765 | -0.067974 | 0.0079605 | -0.042443 | 0.0022335 | 0.11968 | 0.03848 | 0.0058369 | 0.0068488 | -0.023118 | 0.029155 | -0.028537 | 0.012237 | -0.050671 | -0.06258 | 0.064786 | -0.06308 | -0.0075828 | 21.5376
db3 | 4.6389 | 6.6859 | 0.052538 | -0.05143 | -0.032155 | 0.075498 | -0.0065155 | -0.017134 | -0.045911 | 0.0076484 | 0.022122 | -0.0088456 | -0.028115 | 0.003421 | 0.045909 | 0.021008 | -0.0040611 | 0.00060186 | -0.016816 | -0.017763 | 9.0924
db4 | -0.037551 | 0.74318 | -0.016477 | 0.015402 | -0.00093174 | -0.0029778 | -0.00089013 | 0.0029002 | 0.0049278 | -0.0011635 | 0.0011811 | 0.0031182 | -0.00093359 | 0.020888 | -0.0075488 | -0.0067481 | -0.0041763 | -0.0047306 | -0.0015898 | -0.00024975 | 0.83823
db5 | 0.048532 | -1.1418 | -0.0082308 | 0.010653 | -0.029243 | 0.010901 | 0.00024738 | -0.0027304 | -0.0044648 | 0.0051525 | 0.001088 | 0.0014079 | -0.0010145 | 0.00055183 | 0.0032605 | 0.0014393 | 0.0035271 | 0.0056756 | 0.003001 | -0.0012212 | -0.0018385
db6 | -0.028635 | 0.019194 | 0.0063649 | 0.0096066 | -0.0060007 | -0.0062259 | -0.001214 | 0.0010762 | 0.0018764 | 0.0046269 | 0.0024905 | -0.0021369 | -0.00074516 | 0.0009063 | -0.0022929 | -0.0047811 | -0.0017302 | -0.0020953 | -0.0023545 | 0.0026288 | 0.011127
db7 | 0.0092242 | 0.074115 | -0.00092152 | 0.0045477 | -0.004752 | -0.0019167 | -0.0047069 | 0.00070536 | -0.0038195 | 0.0023757 | 0.0028318 | -0.00093571 | -0.0015758 | 0.00057608 | 0.0011737 | 0.001664 | 0.00044973 | 0.0047931 | 0.0001521 | -0.00064093 | -0.0014466
db8 | -0.0075899 | 0.011536 | -0.0071195 | -0.0050157 | -0.001248 | -0.0010817 | 0.0022135 | 0.002475 | -0.0039295 | 0.00072167 | 0.0011754 | 0.0011852 | 0.0020676 | 0.0024922 | 0.0014021 | -0.0019444 | 0.000541 | 0.0020249 | 0.002378 | 0.0016623 | -0.0018493
db9 | 0.028656 | 0.0069245 | -0.003254 | -0.0071079 | 0.0015928 | 0.0033329 | 0.0013696 | 0.0048615 | -0.0026296 | -0.00070693 | -0.0014762 | -0.0022014 | 0.00012293 | 6.6439e-05 | 0.0028957 | 0.00046312 | 4.8426e-05 | 0.0034069 | -0.00029643 | -0.00048785 | 0.0061727
db10 | -0.00022453 | -1.1158e-05 | 0.00022496 | 0.00022304 | 0.00013032 | -0.0031323 | 0.00022268 | 0.00030288 | -0.0023327 | 0.00022552 | 0.00014876 | 0.0001994 | 0.00022533 | 0.0002193 | 0.0022598 | 6.0847e-05 | 0.00023112 | 0.00028105 | 0.00026648 | 0.00024347 | 0.0003925
db11 | 0.013403 | 0.013817 | 0.0006534 | 0.00032591 | 0.001984 | -0.00060604 | -2.1876e-05 | -0.0018364 | -0.0014268 | -0.00072958 | -9.877e-05 | 0.00022258 | -0.001219 | 0.0013744 | -0.00088239 | 0.0011273 | 0.00023909 | 0.00069317 | 6.2068e-06 | 0.00019477 | 0.011888
db12 | -0.00048812 | 0.0047843 | 0.0018202 | 0.00062676 | -0.00058448 | -0.00015423 | 0.0003783 | -0.00044126 | -0.00040585 | 0.00054436 | -0.0001488 | -0.00083207 | 0.0010069 | -1.7559e-05 | 0.0010549 | -0.0022231 | 0.00087875 | -0.00026997 | -0.0005894 | -0.00064352 | -0.0012636
db13 | 0.0025506 | 0.0065634 | -0.00095661 | -0.00075961 | 2.6152e-06 | 0.00025064 | 0.00050049 | -0.00019731 | 0.00073381 | 0.00036145 | 0.00019413 | -0.00032347 | 0.0010321 | 0.00056437 | 0.00035472 | -0.0010876 | -0.00070625 | 0.00041751 | 0.00027367 | -0.00065459 | 0.00055007
db14 | 0.0016657 | 0.0039314 | -4.266e-05 | -0.00047111 | -0.00020747 | -0.00091094 | -0.00083457 | 0.00035317 | -0.00056315 | -0.00063988 | 1.2095e-05 | 0.00015331 | -1.5581e-05 | 0.00014773 | -0.00039489 | 0.00063581 | 0.0009977 | 0.00090325 | 0.00067847 | 0.00019873 | -0.0013041
db15 | -0.0032088 | -0.01212 | 0.00038251 | -0.0009856 | 0.00046787 | 0.0004996 | 0.00014862 | -0.0033972 | 0.00010419 | -3.3437e-05 | 0.0012696 | 0.0023553 | 0.0018169 | 0.00087738 | -0.0020277 | -0.001221 | -0.00042828 | 4.8841e-05 | 0.00014772 | -2.5429e-05 | -0.013859
da2 | -0.24648 | 0.61292 | -0.099109 | -0.0976 | -0.11094 | -0.022934 | 0.037374 | 0.088692 | 0.12253 | 0.15707 | 0.086841 | -0.066411 | 0.03564 | -0.038797 | 0.0154 | -0.023288 | -0.056196 | -0.032253 | 0.0038396 | 0.00014316 | -0.15599
da3 | -0.12785 | -1.202 | 0.055921 | 0.026011 | 0.021981 | -0.055799 | -0.082375 | -0.045762 | -0.046546 | 0.072323 | 0.029123 | -0.05956 | -0.033383 | -0.11557 | -0.016483 | -0.033828 | 0.050137 | 0.021216 | 0.071336 | 0.14126 | -0.1215
da4 | -0.059655 | -0.24816 | 0.024095 | 0.002632 | -0.044126 | -0.013582 | -0.0012326 | -0.0056707 | 0.0204 | 0.039781 | 0.01732 | -0.018284 | -0.002185 | -0.015668 | 0.0039483 | -0.013588 | -0.0062973 | -0.0014951 | 0.0053468 | 0.0086045 | -0.09055
da5 | -0.0028114 | 1.8957 | 0.0087281 | -0.00056467 | -0.015154 | 0.0010737 | -0.0050218 | -0.0019755 | -0.0074861 | 0.011917 | 0.003849 | -0.006858 | 0.00080593 | 3.1595e-05 | 0.0005868 | -0.006052 | -0.0017085 | 0.0032797 | 0.006825 | 0.0077242 | -0.08564
da6 | -0.003521 | -0.045469 | 0.0034317 | -0.0037685 | -0.0091005 | -0.00044326 | 0.00091923 | 0.0036689 | 0.00068399 | 0.0029186 | -0.0019192 | -0.0018721 | -0.0039606 | 0.0030975 | 0.0054643 | -0.00082528 | 0.002347 | -0.0035979 | 0.0017141 | 0.001242 | -0.0032422
da7 | 0.012355 | 1.4275 | -0.0075584 | 0.0036178 | 0.019582 | 0.0071625 | 0.00094105 | -0.0029884 | -0.0071693 | -0.0047968 | -0.0038846 | -0.007269 | -0.0080095 | -0.004356 | -0.0059268 | -0.0030391 | 0.0038363 | 0.012544 | 0.0066199 | 0.00069434 | 0.0072557
da8 | 0.00081233 | -0.024601 | -0.0044483 | -0.00059399 | 0.0051607 | -0.00075035 | -0.00098311 | 0.00085291 | -0.0019586 | -4.8506e-05 | -0.0036948 | 0.0030478 | 0.00068933 | 0.001438 | 0.0017311 | -0.00083689 | -4.3988e-05 | 0.00053086 | -0.00024311 | 0.00015101 | 0.0071799
da9 | 0.0074242 | -0.22591 | 0.0039651 | -0.0035115 | 0.0084846 | 0.0050955 | 0.004082 | -0.0002407 | -0.0012109 | -0.0046876 | -0.0016646 | -0.0061798 | -0.002978 | -0.0016795 | -0.0046674 | -0.0013794 | -0.00093204 | 0.0080041 | 0.0035101 | -0.0040099 | -0.086017
da10 | -0.00023876 | 3.2529e-05 | 0.0003011 | 0.00024024 | -0.00080151 | 0.00053563 | 0.00023987 | 0.0011077 | -0.00085396 | 0.00020069 | -0.00075627 | 0.0003057 | 0.00024247 | 0.00022738 | 0.00016836 | -0.002577 | 0.00024564 | 0.00032376 | 0.00024753 | 0.00060264 | 0.00019515
da11 | 0.0014609 | 0.16717 | 0.0020516 | 0.0048346 | 0.011808 | 0.010188 | 0.0031741 | -0.0063913 | -0.0054564 | -0.0073192 | -0.0052975 | -0.0065163 | -0.0063687 | -0.0067219 | -0.0050428 | -0.0034006 | 0.0039976 | 0.011076 | 0.0051236 | 0.00026046 | -0.011106
da12 | 0.00021858 | -0.0054715 | -0.00069196 | -0.0010428 | -0.00056434 | 0.00052153 | -3.305e-05 | -4.8608e-05 | 0.0013753 | 0.00038888 | -0.00061242 | -0.001088 | 0.00061527 | 0.00086645 | 0.0018076 | -0.0012163 | -0.0019075 | 0.0008234 | 0.00079894 | 7.5846e-06 | -9.6247e-06
da13 | -0.0035212 | -0.00064747 | -0.00025686 | 8.9472e-05 | 0.0011472 | 0.0014362 | 0.00051303 | 0.0004635 | -4.5902e-05 | -0.0012941 | -0.001736 | -0.0010031 | 6.7384e-05 | 9.4574e-05 | -0.00035803 | 0.000638 | -0.00080379 | 0.00082556 | 0.00023984 | -1.6988e-05 | 0.0018759
da14 | -0.0040996 | -0.0010494 | 0.00042161 | -0.00010144 | -0.00057628 | -5.7709e-05 | 0.00016556 | -0.00014419 | 0.00023071 | 0.00030442 | 0.0003203 | -0.00057069 | 1.7137e-05 | 0.00040383 | -0.00040288 | -0.00046889 | 0.00052647 | 0.00020618 | -0.00037388 | 9.9738e-05 | -0.0046748
da15 | 0.0018179 | 0.005158 | 0.0014465 | 0.00073375 | 2.1993e-05 | -0.00064287 | -0.00094784 | -0.0029733 | -3.9573e-05 | 0.0038201 | -0.00038497 | -0.0022465 | -0.00018962 | 0.0022287 | 0.00083143 | 0.00029023 | -0.00039543 | -0.0011938 | -0.00076907 | 0.00041027 | -0.0073043
Check Computation for Aperture2
Magnetic Length | -0.076576
mbh dB/B | -4504.8446
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.0613 | -5190.6724 | -26.3046 | 0.25429 | 0.6673 | 0.55092 | 0.88728 | 1.692 | 1.3386 | 2.6174 | 2.7891 | 1.895 | 2.7664 | 1.8112 | 1.939 | 0.92418 | 1.3911 | 2.2058 | 2.6188 | -0.043761 | -3819.0167
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 6.5538 | -0.23607 | -0.79386 | 0.58617 | 0.066686 | -0.029482 | -0.0048851 | 0.17086 | -0.077459 | 0.028712 | -0.11517 | -0.0051341 | 0.45121 | 0.093642 | 0.00037089 | -0.23638 | 0.43324 | 0.072456 | -0.40491 | -1.8851
db2 | 1.1 | 27.1895 | 0.6427 | -0.03075 | -0.046117 | 0.0034786 | -0.11298 | 0.011886 | 0.045255 | 0.033728 | -0.055234 | -0.079046 | -0.0082613 | -0.11466 | 0.0098627 | -0.076728 | -0.060073 | -0.07083 | -0.051607 | -0.040621 | 20.4557
db3 | 4.7106 | 11.0317 | 0.15984 | -0.12359 | 0.015539 | -0.058899 | -0.019289 | -0.05985 | -0.039333 | 0.040234 | -0.025834 | 0.019059 | 0.09519 | -0.001504 | 0.032796 | -0.0039161 | -0.016457 | 0.0073823 | -0.011881 | -0.0094844 | 9.0293
db4 | -0.04424 | 1.1032 | 0.016281 | 0.032015 | 0.0069708 | 0.0057718 | -0.059143 | 0.035083 | 0.010231 | -0.028231 | 0.0056289 | 0.00092701 | -0.0032573 | 0.022282 | -0.027282 | 0.0046925 | 0.014588 | -0.028979 | -0.010205 | 0.0026255 | 0.80075
db5 | 0.055224 | -2.2719 | 0.017508 | -0.014068 | -0.096497 | 0.055659 | -0.01651 | 0.031884 | -0.019958 | -0.014405 | 0.025977 | -0.012484 | 0.0085667 | 0.025431 | -0.0066674 | 0.0042546 | 0.0052808 | -0.005372 | -0.0051013 | 0.016502 | 0.13536
db6 | -0.032177 | 0.057417 | -0.0031871 | -0.0020237 | 0.0012829 | -0.0037026 | 0.010136 | -0.0075543 | -0.003741 | 0.0036575 | -0.0044081 | 0.0038872 | -0.011571 | 0.021256 | -0.011321 | 0.0017568 | 0.0056071 | -0.0024698 | 0.0022327 | 0.0001622 | -0.011859
db7 | 0.0065589 | 0.30859 | -0.0040799 | -0.01558 | -0.0038833 | 0.011632 | 0.010838 | 0.0098026 | -0.0090597 | -0.001865 | 0.0020893 | 0.0017986 | 0.002453 | 0.0034205 | -0.0029601 | -0.0013601 | -0.0019584 | -0.00021442 | -0.0009006 | -0.00017232 | 0.0070712
db8 | -0.0045854 | 0.026678 | -0.006857 | 0.010383 | 0.005383 | 0.0085198 | -0.00028987 | -0.0014734 | -0.0016681 | -0.0015865 | -0.0058667 | 0.00080531 | 0.0033566 | -0.0054453 | -0.0004522 | -0.00016911 | -0.0016111 | -0.0016942 | -0.0014541 | 0.00012003 | -0.01408
db9 | 0.026773 | -0.0041977 | 0.0038963 | -0.0048208 | 0.00069736 | -0.013602 | -0.0058177 | 0.002948 | 1.381e-05 | 0.0051959 | 0.0031734 | 0.00061508 | 0.0014218 | -0.0011573 | -0.00019328 | 0.0060002 | -0.00071861 | 0.0017324 | 0.0041057 | -0.0034905 | -0.0031855
db10 | -6.938e-06 | -3.3687e-05 | 6.8479e-07 | 9.6277e-06 | 9.2816e-07 | 1.1002e-05 | 1.7257e-05 | 2.162e-05 | -2.1701e-05 | 4.0139e-05 | 1.4103e-05 | 8.1952e-06 | 2.0248e-05 | -6.0999e-06 | -0.00014925 | 2.7868e-05 | 0.00011999 | -2.0319e-05 | -8.6475e-05 | -7.8165e-06 | -0.00022139
db11 | 0.013046 | 0.017292 | 0.0019421 | -5.3456e-05 | 0.0007268 | 0.0034801 | -0.0020709 | 0.00053192 | -0.0017492 | -0.00023744 | -2.936e-05 | -0.0016814 | -0.0031998 | 0.00073542 | 0.00022413 | 0.00036187 | 0.00028155 | -0.00081761 | 0.001242 | 0.0003133 | 0.0044397
db12 | 0.0003087 | -0.0056347 | -0.0018035 | -0.00014178 | 0.0012949 | -0.00043938 | 0.00022662 | 0.00045613 | -0.00073022 | 0.0012134 | -0.0012247 | 0.00073956 | 0.00018419 | -0.0018276 | 0.00069513 | 0.00065628 | 0.00020138 | -0.00033815 | 0.00031403 | 0.00052372 | -0.0071073
db13 | 0.0023887 | 0.011885 | -0.00011929 | 0.00043919 | 0.0006238 | 0.00062903 | -0.0010198 | -0.0017281 | 0.00062985 | 0.00094531 | -8.3682e-06 | 0.00043875 | 0.00081666 | -0.00037915 | 0.0005384 | -0.00048534 | -0.0010514 | -4.8507e-05 | 0.00017459 | -0.00039566 | 0.00044833
db14 | -0.00058062 | -0.0047571 | 0.00091588 | -0.00035043 | -0.00016588 | 9.4574e-05 | -0.00072038 | -4.0416e-05 | 0.0004416 | 1.3257e-05 | 0.00037672 | -0.00056439 | 9.6576e-05 | 0.0011268 | -0.00057464 | -0.0003551 | -0.0003207 | -0.00041068 | 0.00060063 | -0.00016339 | 0.0021591
db15 | -0.0031964 | -0.013788 | -0.00017228 | 0.00043698 | -0.0014362 | 0.00051315 | -0.001118 | -0.00042318 | 0.0010937 | 0.00039403 | 0.00034391 | 0.00098359 | 0.0010637 | -0.0018516 | -0.0014939 | 0.00058304 | 0.00068079 | -0.00069293 | 0.00057843 | 0.00051683 | -0.012205
da2 | -0.24469 | -3.8409 | -0.057105 | -0.19872 | 0.044192 | -0.26779 | -0.051665 | -0.080368 | 0.15944 | 0.056224 | 0.12753 | 0.0092686 | -0.056754 | 0.085586 | 0.01539 | -0.0014846 | -0.042131 | 0.053691 | 0.034289 | 0.17041 | -0.37928
da3 | 0.038201 | -2.7213 | 0.04189 | -0.013418 | -0.010339 | -0.10097 | -0.036372 | -0.087225 | 0.075667 | 0.029397 | 0.067707 | 0.032702 | 0.060396 | 0.005262 | 0.065314 | 0.033947 | 0.018485 | -0.10339 | -0.025518 | -0.053537 | -0.12778
da4 | 0.016144 | -0.31307 | -0.054383 | 0.0049036 | -0.048858 | -0.011022 | 0.051533 | 0.0060593 | -0.020562 | 0.048528 | -0.027584 | -0.0056102 | 0.026375 | -0.01432 | 0.00605 | 0.0057141 | 0.0044489 | 0.014501 | 0.016176 | -0.0019504 | -0.051835
da5 | 0.0058829 | 2.2294 | 0.0037187 | -0.0001506 | 0.012645 | -0.01626 | -0.019139 | -0.0047505 | 0.014365 | 0.00017721 | 0.0028202 | 0.0050722 | -0.015387 | 0.0066127 | 0.009405 | 0.013862 | 0.0020163 | -0.0052711 | 0.001942 | -0.011678 | -0.0078868
da6 | 0.0055738 | -0.055564 | -0.0035606 | -0.00024675 | 0.00032603 | -0.0007956 | -0.010894 | 0.015111 | -0.0043427 | 0.002858 | 0.0005888 | -0.0088724 | -0.0054853 | 0.0073083 | 0.0098771 | -0.0015818 | 0.0013893 | -0.0023801 | 0.00038425 | 0.00031638 | -0.00017705
da7 | 0.0006342 | 1.6342 | -0.0029031 | 0.00019548 | 0.0021982 | 0.015775 | -0.0012891 | 0.0011257 | -0.0045619 | 0.0045253 | -0.017008 | -0.005253 | -0.0015636 | -0.011979 | -0.0028294 | 0.0014507 | 0.00030565 | 0.0088682 | 0.0085961 | 0.0043468 | -0.010748
da8 | 0.00082749 | -0.020403 | 0.0054741 | -0.0034761 | -0.0028004 | 0.0029033 | -0.00023375 | -0.0025185 | 0.0033125 | -0.0012655 | -0.0016709 | -0.0018181 | 0.0024401 | 0.0018531 | 0.0016498 | -0.00072003 | -0.0025766 | 2.0607e-05 | -0.002261 | 0.0016873 | -5.4435e-05
da9 | 0.0014243 | -0.28247 | 0.0034397 | 0.0038983 | 0.010749 | 0.0063392 | -0.0027658 | -0.0046391 | -0.00315 | -0.0032177 | -0.0014952 | -0.0035295 | -0.002355 | -0.0019958 | -0.0057261 | -0.0015307 | -0.0032732 | 0.0037977 | 0.0060013 | -0.00054763 | -0.059303
da10 | 1.8609e-05 | 3.1167e-05 | -1.3928e-05 | -1.8539e-05 | -3.0951e-05 | -1.7946e-05 | -3.1886e-05 | -1.2462e-05 | -1.3584e-05 | -5.8397e-05 | -1.3105e-05 | -8.9855e-06 | -2.1952e-05 | -2.3971e-05 | 0.0021638 | -9.6537e-05 | -0.0017931 | -2.0184e-05 | 6.9469e-06 | 4.8045e-06 | -0.00018183
da11 | -0.039503 | 0.20942 | 0.0027446 | -0.0031721 | 0.010248 | 0.0076427 | 0.0034172 | -0.0029046 | -0.0064752 | -0.0084498 | -0.0042715 | -1.8854e-06 | -0.0045739 | -0.0027418 | -0.0069642 | -0.0033748 | 0.00051413 | 0.0097686 | 0.0082407 | 0.00035452 | -0.053096
da12 | 0.00013683 | 0.0035145 | 0.0004731 | -0.0014669 | 0.0017356 | -0.0015787 | -0.00012628 | -0.0014493 | -0.0019626 | -0.00047522 | 0.0018196 | -0.0014474 | 0.00053274 | 0.0013238 | 0.002427 | -7.9104e-05 | -0.0011676 | 0.00090859 | 0.00049713 | 3.5653e-05 | 0.0046635
da13 | -0.0030886 | -0.0042986 | -2.3715e-05 | -0.00050546 | 0.00099976 | -0.0003308 | 0.00064427 | -0.0011707 | -0.0019102 | -0.00086323 | -0.00053367 | 0.00039734 | -0.00034831 | 0.0010668 | 0.0011337 | 0.00047247 | -0.00082704 | 0.0017084 | 0.000299 | -0.00020868 | -4.2079e-05
da14 | -0.0021348 | 0.0055875 | -0.00045892 | -9.909e-05 | -0.00033024 | -0.0002516 | -1.6195e-05 | 0.00074413 | -0.00020737 | 0.00018252 | 0.0001697 | 0.00060357 | 0.00091246 | -0.00059953 | 0.00029526 | -9.4422e-05 | -0.00095316 | -0.000195 | 5.1096e-05 | 0.00024678 | -0.002178
da15 | 0.0013623 | 0.0054963 | 0.0012456 | 3.88e-05 | -0.0003141 | -0.0015238 | -0.000526 | -0.0025593 | -0.00028916 | 0.0024581 | -0.00029787 | 7.9865e-06 | -0.00053027 | 0.0011376 | 0.001971 | 0.0022905 | 0.00042812 | -0.0012686 | -0.0011679 | -0.0011007 | -0.005
Field Colinearity | 17.6556
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.519 | -0.36626 | -0.17473 | 0.54611 | 0.18211 | -0.12869 | -0.30781 | -0.43454 | -0.75957 | -0.35462 | -0.30999 | -0.053296 | 0.31351 | 0.30426 | 0.079158 | -0.11315 | 0.046418 | -0.59318 | -0.60434 | -0.4671
Twist (m2 mrad) | 41.3927 | -5.2192 | -2.3589 | 6.9629 | 2.1854 | -1.4477 | -3.232 | -4.2367 | -6.8362 | -2.9256 | -2.3249 | -0.35975 | 1.881 | 1.5974 | 0.35621 | -0.4243 | 0.13925 | -1.3347 | -0.90651 | -0.17516
Twist Integral (I0) (m2 rad) | 0.01705
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 6.4481 | -0.34176 | -0.89955 | 0.48048 | -0.039006 | -0.13517 | -0.11058 | 0.065165 | -0.18315 | -0.07698 | -0.22087 | -0.11083 | 0.34552 | -0.01205 | -0.10532 | -0.34207 | 0.32755 | -0.033236 | -0.5106 | -1.9907
Twist (m2 mrad) | 48.361 | -4.8701 | -12.1439 | 6.1261 | -0.46807 | -1.5207 | -1.1611 | 0.63536 | -1.6484 | -0.63509 | -1.6565 | -0.74808 | 2.0731 | -0.063262 | -0.47394 | -1.2828 | 0.98264 | -0.074781 | -0.7659 | -0.74653
Twist Integral (I0) (m2 rad) | 0.022439

## Work sheet diff
| HCMBARA001-3000024 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 1.4631 | 7.125 | 1.1256 | 0.050232 | 1.0803 | 0.55165 | -0.24648 | 0.077021 | -1.6247 | 7.5246 | -1.1 | 0.16692 | 0.056946 | 1.2169 | -0.24469 | 0.11276 |  | 0
3 | 5.8526 | 7.6044 | 4.6389 | 0.033585 | -0.30282 | 0.36333 | -0.12785 | 0.065064 | 5.5535 | 9.0347 | 4.7106 | 0.061066 | -0.25665 | 0.59767 | 0.038201 | 0.057932 |  | 0
4 | 0.014481 | 0.30478 | -0.037551 | 0.0082122 | -0.049412 | 0.18231 | -0.059655 | 0.018756 | 0.014782 | 0.33051 | 0.04424 | 0.023604 | 0.26798 | 0.13768 | 0.016144 | 0.027884 |  | 0
5 | -0.27425 | 1.2802 | 0.048532 | 0.0086826 | -0.0078968 | 0.45757 | -0.0028114 | 0.0066592 | -0.35028 | 1.458 | 0.055224 | 0.031415 | 0.034921 | 0.50867 | 0.0058829 | 0.01029 |  | 0
6 | -0.0072409 | 0.069999 | -0.028635 | 0.0041673 | 0.0084159 | 0.067304 | -0.003521 | 0.0035829 | 0.03846 | 0.068354 | 0.032177 | 0.0077111 | 0.067203 | 0.073766 | 0.0055738 | 0.0062214 |  | 0
7 | 0.98908 | 0.34011 | 0.0092242 | 0.0027668 | -0.026446 | 0.33524 | 0.012355 | 0.0076735 | 0.9743 | 0.41566 | 0.0065589 | 0.0066401 | 0.00078749 | 0.37577 | 0.0006342 | 0.0074787 |  | 0
8 | -0.0066331 | 0.029965 | -0.0075899 | 0.0028428 | 0.0071087 | 0.02627 | 0.00081233 | 0.0022018 | 0.002381 | 0.026747 | 0.0045854 | 0.0045229 | 0.015057 | 0.029723 | 0.00082749 | 0.0025508 |  | 0
9 | 0.40694 | 0.028014 | 0.028656 | 0.0028207 | -0.033364 | 0.047367 | 0.0074242 | 0.0044447 | 0.40027 | 0.017486 | 0.026773 | 0.004702 | -0.02789 | 0.060671 | 0.0014243 | 0.0045569 |  | 0
10 | 0 | 0.047056 | -0.00022453 | 0.001114 | 0 | 0.024401 | -0.00023876 | 0.00081047 | 0 | 0.057487 | 6.938e-06 | 5.3786e-05 | 0 | 0.022997 | 1.8609e-05 | 0.00068234 |  | 0
11 | 0.64285 | 0.019381 | 0.013403 | 0.0009952 | -0.025255 | 0.044309 | 0.014609 | 0.0067009 | 0.64073 | 0.018007 | 0.013046 | 0.0015393 | -0.033834 | 0.055826 | 0.0016877 | 0.0058649 |  | 0
12 | 0.00037667 | 0.0061534 | -0.00048812 | 0.00091588 | -0.00062346 | 0.0041708 | 0.00021858 | 0.00098795 | -0.00030468 | 0.005614 | -0.0003087 | 0.00091619 | -0.00027405 | 0.0035631 | 0.00013683 | 0.0013251 |  | 0
13 | 0.064368 | 0.0062784 | 0.0025506 | 0.00061952 | -0.0040162 | 0.003492 | 0.0015612 | 0.00082468 | 0.063789 | 0.0080652 | 0.0023887 | 0.00073292 | -0.0036234 | 0.0042519 | 0.00045803 | 0.00092326 |  | 0
14 | -0.0018232 | 0.0018485 | 0.0016657 | 0.00057196 | -0.0038565 | 0.0012901 | 0.0020516 | 0.00035652 | -0.0016102 | 0.0021785 | 0.00058062 | 0.0005184 | -0.0017542 | 0.0025072 | 0.0020485 | 0.00046743 |  | 0
15 | 0.017282 | 0.008672 | -0.0032088 | 0.0013329 | -0.002988 | 0.0024717 | 0.0018179 | 0.0015471 | 0.01805 | 0.0085599 | -0.0031964 | 0.00094214 | -0.0020873 | 0.0021314 | 0.0013623 | 0.0013739 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -2.7774 | -1.9864
3 | [12 empty] | 3 | 9 | -3.063 | -2.8221
4 | [12 empty] | 4 | 16 | -4.3892 | -3.663
5 | [12 empty] | 5 | 25 | -4.8791 | -4.0185
6 | [12 empty] | 6 | 36 | -5.556 | -4.9724
7 | [12 empty] | 7 | 49 | -5.38 | -4.9552
8 | [12 empty] | 8 | 64 | -5.9907 | -5.685
9 | [12 empty] | 9 | 81 | -5.6434 | -5.3754
10 | [12 empty] | sigma (mm) |  | 0.0019805 | 0.0033483
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05 | [5 empty] | Aperture 1 | Aperture 2
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | Magnetic length | -0.076459 | -0.076576
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 6.7146 | 107.257 | 108.0229 | 108.0985 | 108.0579 | 108.1842 | 108.1841 | 108.1892 | 108.0875 | 108.215 | 108.1514 | 108.214 | 108.1576 | 108.0999 | 108.1704 | 108.0352 | 108.1056 | 108.0942 | 107.415 | 33.0642
2 | 27.0124 | 1.1844 | 1.0577 | 1.1336 | 1.0832 | 1.1279 | 1.2453 | 1.1641 | 1.1315 | 1.1325 | 1.1025 | 1.1548 | 1.0971 | 1.1379 | 1.075 | 1.0631 | 1.1904 | 1.0626 | 1.1181 | 23.2493 | 1.4631
3 | 6.6859 | 4.6914 | 4.5875 | 4.6067 | 4.7144 | 4.6324 | 4.6218 | 4.593 | 4.6465 | 4.661 | 4.6301 | 4.6108 | 4.6423 | 4.6848 | 4.6599 | 4.6348 | 4.6395 | 4.6221 | 4.6211 | 9.0924 | 5.8526
4 | 1.6996 | -0.054028 | -0.02215 | -0.038483 | -0.040529 | -0.038442 | -0.034651 | -0.032624 | -0.038715 | -0.03637 | -0.034433 | -0.038485 | -0.016664 | -0.0451 | -0.044299 | -0.041728 | -0.042282 | -0.039141 | -0.037801 | 0.99876 | 0.014481
5 | -1.1418 | 0.040301 | 0.059185 | 0.019289 | 0.059433 | 0.048779 | 0.045802 | 0.044067 | 0.053685 | 0.04962 | 0.04994 | 0.047517 | 0.049084 | 0.051793 | 0.049971 | 0.052059 | 0.054208 | 0.051533 | 0.047311 | -0.0018385 | -0.27425
6 | 0.24012 | -0.02227 | -0.019028 | -0.034635 | -0.03486 | -0.029849 | -0.027558 | -0.026758 | -0.024008 | -0.026144 | -0.030772 | -0.02938 | -0.027728 | -0.030928 | -0.033416 | -0.030365 | -0.03073 | -0.030989 | -0.026006 | 0.011127 | -0.0072409
7 | 0.074115 | 0.0083027 | 0.013772 | 0.0044722 | 0.0073075 | 0.0045173 | 0.0099296 | 0.0054047 | 0.0116 | 0.012056 | 0.0082885 | 0.0076484 | 0.0098003 | 0.010398 | 0.010888 | 0.0096739 | 0.014017 | 0.0093763 | 0.0085833 | -0.0014466 | 0.98908
8 | 0.071001 | -0.014709 | -0.012606 | -0.0088378 | -0.0086716 | -0.0053763 | -0.0051149 | -0.011519 | -0.0068682 | -0.0064145 | -0.0064047 | -0.0055223 | -0.0050977 | -0.0061878 | -0.0095343 | -0.0070489 | -0.0055649 | -0.0052119 | -0.0059275 | -0.0018493 | -0.0066331
9 | 0.0069245 | 0.025402 | 0.021548 | 0.030248 | 0.031988 | 0.030025 | 0.033517 | 0.026026 | 0.027949 | 0.027179 | 0.026454 | 0.028779 | 0.028722 | 0.031551 | 0.029119 | 0.028704 | 0.032063 | 0.028359 | 0.028168 | 0.0061727 | 0.40694
10 | -1.5838e-06 | 4.3164e-07 | -1.49e-06 | -9.421e-05 | -0.0033568 | -1.8551e-06 | 7.8346e-05 | -0.0025572 | 9.9083e-07 | -7.5773e-05 | -2.5135e-05 | 7.9856e-07 | -5.2295e-06 | 0.0020353 | -0.00016368 | 6.5927e-06 | 5.652e-05 | 4.1945e-05 | 1.894e-05 | 0.0004101 | 0
11 | 0.013817 | 0.014056 | 0.013729 | 0.015387 | 0.012797 | 0.013381 | 0.011566 | 0.011976 | 0.012673 | 0.013304 | 0.013625 | 0.012184 | 0.014777 | 0.01252 | 0.01453 | 0.013642 | 0.014096 | 0.013409 | 0.013597 | 0.011888 | 0.64285
12 | 0.023881 | 0.01332 | 0.0013863 | -0.010726 | -0.0064235 | -0.0010983 | -0.0092939 | -0.0089397 | 0.00056237 | -0.0063692 | -0.013202 | 0.0051883 | -0.0050568 | 0.0056683 | -0.027112 | 0.0039063 | -0.0075809 | -0.010775 | -0.011316 | 0.0015338 | 0.0037667 | ten times!
13 | 0.065634 | 0.01594 | 0.017909 | 0.025532 | 0.028012 | 0.03051 | 0.023532 | 0.032844 | 0.02912 | 0.027447 | 0.022271 | 0.035826 | 0.031149 | 0.029053 | 0.014629 | 0.018443 | 0.029681 | 0.028242 | 0.01896 | 0.0055007 | 0.64368 | ten times!
14 | 0.015449 | 0.01623 | 0.011946 | 0.014582 | 0.0075475 | 0.0083113 | 0.020189 | 0.011025 | 0.010258 | 0.016778 | 0.01819 | 0.016501 | 0.018134 | 0.012708 | 0.023015 | 0.026634 | 0.025689 | 0.023442 | 0.018644 | 0.0023855 | -0.018232 | ten times!
15 | -0.081182 | -0.028263 | -0.041944 | -0.02741 | -0.027092 | -0.030602 | -0.06606 | -0.031046 | -0.032423 | -0.019392 | -0.0085352 | -0.013919 | -0.023314 | -0.052365 | -0.044298 | -0.036371 | -0.0316 | -0.030611 | -0.032342 | -0.13859 | 0.17282 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 0.27531 | 0.009815 | 0.70656 | 1.7763 | 1.4368 | 0.41798 | -0.62954 | -0.86741 | -1.2308 | -0.78256 | -0.48456 | -0.94638 | -0.82145 | -0.81006 | -0.23497 | 0.46983 | 1.6296 | 0.85879 | 0.13515 | -0.48324
2 | -7.1828 | -0.34559 | -0.34408 | -0.35742 | -0.26941 | -0.2091 | -0.15779 | -0.12395 | -0.089411 | -0.15964 | -0.31289 | -0.21084 | -0.28528 | -0.23108 | -0.26977 | -0.30267 | -0.27873 | -0.24264 | -0.24634 | -0.15599 | 1.0803
3 | 0.22151 | -0.071932 | -0.10184 | -0.10587 | -0.18365 | -0.21023 | -0.17362 | -0.1744 | -0.05553 | -0.098731 | -0.18741 | -0.16124 | -0.24342 | -0.14434 | -0.16168 | -0.077717 | -0.10664 | -0.056518 | 0.013408 | -0.1215 | -0.30282
4 | -2.5018 | -0.03556 | -0.057023 | -0.10378 | -0.073237 | -0.060887 | -0.065325 | -0.039254 | -0.019873 | -0.042335 | -0.077939 | -0.06184 | -0.075322 | -0.055706 | -0.073242 | -0.065952 | -0.06115 | -0.054308 | -0.05105 | -0.09055 | -0.049412
5 | 0.14355 | 0.0059168 | -0.0033761 | -0.017966 | -0.0017377 | -0.0078332 | -0.0047869 | -0.010297 | 0.0091055 | 0.0010376 | -0.0096694 | -0.0020055 | -0.0027798 | -0.0022246 | -0.0088634 | -0.0045199 | 0.0004683 | 0.0040136 | 0.0049128 | -0.08564 | -0.0078968
6 | -0.45652 | -8.9302e-05 | -0.0072895 | -0.012621 | -0.0039643 | -0.0026018 | 0.00014792 | -0.002837 | -0.00060237 | -0.0054402 | -0.0053931 | -0.0074816 | -0.00042351 | 0.0019433 | -0.0043463 | -0.001174 | -0.0071189 | -0.0018069 | -0.002279 | -0.0032422 | 0.0084159
7 | 0.23812 | 0.0047962 | 0.015972 | 0.031937 | 0.019517 | 0.013296 | 0.0093662 | 0.0051853 | 0.0075578 | 0.0084701 | 0.0050856 | 0.0043452 | 0.0079986 | 0.0064278 | 0.0093156 | 0.016191 | 0.024898 | 0.018974 | 0.013049 | 0.0072557 | -0.026446
8 | -0.079851 | -0.0036359 | 0.00021833 | 0.005973 | 6.1973e-05 | -0.00017078 | 0.0016652 | -0.0011463 | 0.00076382 | -0.0028825 | 0.0038601 | 0.0015017 | 0.0022503 | 0.0025434 | -2.4563e-05 | 0.00076834 | 0.0013432 | 0.00056922 | 0.00096333 | 0.0071799 | 0.0071087
9 | -0.01623 | 0.011389 | 0.0039127 | 0.015909 | 0.01252 | 0.011506 | 0.0071835 | 0.0062133 | 0.0027366 | 0.0057597 | 0.0012444 | 0.0044462 | 0.0057447 | 0.0027569 | 0.0060448 | 0.0064922 | 0.015428 | 0.010934 | 0.0034143 | 0.0030295 | -0.033364
10 | 5.2574e-05 | 6.2341e-05 | 1.48e-06 | -0.0010403 | 0.00029687 | 1.11e-06 | 0.00086897 | -0.0010927 | -3.8077e-05 | -0.00099503 | 6.6937e-05 | 3.7014e-06 | -1.1387e-05 | -7.0406e-05 | -0.0028158 | 6.8734e-06 | 8.5e-05 | 8.7647e-06 | 0.00036387 | 0.00019403 | 0
11 | 0.033678 | 0.01666 | 0.019443 | 0.026417 | 0.024797 | 0.017783 | 0.0082173 | 0.0091523 | 0.0072895 | 0.0093111 | 0.0080923 | 0.0082399 | 0.0078867 | 0.0095659 | 0.011208 | 0.018606 | 0.025685 | 0.019732 | 0.014869 | 0.019819 | -0.025255 | ten times!
12 | -0.0017253 | -0.0047338 | -0.0082427 | -0.0034576 | 0.0074011 | 0.0018553 | 0.0016997 | 0.015939 | 0.0060746 | -0.0039384 | -0.0086944 | 0.0083384 | 0.01085 | 0.020262 | -0.0099769 | -0.016889 | 0.01042 | 0.010175 | 0.0022616 | 0.002067 | -0.0062346 | ten times!
13 | -0.0064747 | 0.013043 | 0.016507 | 0.027084 | 0.029974 | 0.020742 | 0.020247 | 0.015153 | 0.0026706 | -0.001748 | 0.0055816 | 0.016286 | 0.016558 | 0.012032 | 0.021992 | 0.0075741 | 0.023868 | 0.01801 | 0.015442 | 0.018759 | -0.040162 | ten times!
14 | -0.023921 | 0.024732 | 0.019502 | 0.014753 | 0.019939 | 0.022172 | 0.019074 | 0.022823 | 0.02356 | 0.023719 | 0.014809 | 0.020688 | 0.024554 | 0.016487 | 0.015827 | 0.025781 | 0.022578 | 0.016777 | 0.021514 | -0.021142 | -0.038565 | ten times!
15 | 0.017958 | 0.032643 | 0.025516 | 0.018399 | 0.01175 | 0.0087002 | -0.011554 | 0.017783 | 0.056379 | 0.014329 | -0.004286 | 0.016282 | 0.040466 | 0.026493 | 0.021081 | 0.014224 | 0.0062407 | 0.010488 | 0.022281 | -0.0090294 | -0.02988 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -43.9655 | 105.7904 | 108.0147 | 108.0438 | 108.0356 | 108.0593 | 107.9983 | 107.9734 | 108.0635 | 108.0756 | 108.1303 | 108.1917 | 108.1244 | 108.1334 | 108.0619 | 108.0948 | 108.0345 | 107.946 | 107.5231 | 56.3263
2 | -27.1895 | -1.7427 | -1.0693 | -1.0539 | -1.1035 | -0.98704 | -1.1119 | -1.1453 | -1.1338 | -1.0448 | -1.021 | -1.0918 | -0.98537 | -1.1099 | -1.0233 | -1.04 | -1.0292 | -1.0484 | -1.0594 | -22.1318 | -1.6247
3 | 11.0317 | 4.8704 | 4.587 | 4.7261 | 4.6517 | 4.6913 | 4.6507 | 4.6713 | 4.7508 | 4.6848 | 4.7296 | 4.8058 | 4.7091 | 4.7434 | 4.7067 | 4.6941 | 4.718 | 4.6987 | 4.7011 | 9.0293 | 5.5535
4 | -0.29903 | 0.027959 | 0.012224 | 0.037269 | 0.038468 | 0.10338 | 0.0091564 | 0.034008 | 0.072471 | 0.038611 | 0.043313 | 0.047497 | 0.021957 | 0.071521 | 0.039547 | 0.029652 | 0.073219 | 0.054445 | 0.041614 | -0.69415 | 0.014782
5 | -2.2719 | 0.072732 | 0.041156 | -0.041273 | 0.11088 | 0.038714 | 0.087108 | 0.035266 | 0.040819 | 0.081201 | 0.04274 | 0.063791 | 0.080655 | 0.048557 | 0.059479 | 0.060505 | 0.049852 | 0.050123 | 0.071726 | 0.13536 | -0.35028
6 | 0.14989 | 0.035364 | 0.034201 | 0.030895 | 0.03588 | 0.022041 | 0.039732 | 0.035918 | 0.02852 | 0.036586 | 0.02829 | 0.043748 | 0.010921 | 0.043498 | 0.030421 | 0.02657 | 0.034647 | 0.029945 | 0.032015 | 0.011859 | 0.03846
7 | 0.30859 | 0.002479 | -0.0090214 | 0.0026756 | 0.018191 | 0.017397 | 0.016361 | -0.0025008 | 0.0046939 | 0.0086481 | 0.0083575 | 0.0090119 | 0.0099794 | 0.0035988 | 0.0051988 | 0.0046005 | 0.0063445 | 0.0056583 | 0.0063866 | 0.0070712 | 0.9743
8 | 0.089533 | 0.011442 | -0.0057976 | -0.00079752 | -0.0039343 | 0.0048753 | 0.0060588 | 0.0062535 | 0.0061719 | 0.010452 | 0.0037801 | 0.0012288 | 0.010031 | 0.0050376 | 0.0047545 | 0.0061965 | 0.0062796 | 0.0060396 | 0.0044654 | 0.01408 | 0.002381
9 | -0.0041977 | 0.030669 | 0.021952 | 0.02747 | 0.013171 | 0.020955 | 0.029721 | 0.026786 | 0.031968 | 0.029946 | 0.027388 | 0.028194 | 0.025615 | 0.026579 | 0.032773 | 0.026054 | 0.028505 | 0.030878 | 0.023282 | -0.0031855 | 0.40027
10 | 0.0025152 | 6.2533e-06 | -2.6897e-06 | 6.0099e-06 | -4.0642e-06 | -1.0319e-05 | -1.4682e-05 | 2.8639e-05 | -3.32e-05 | -7.1646e-06 | -1.2571e-06 | -1.331e-05 | 1.3038e-05 | 0.00015619 | -2.093e-05 | -0.00011305 | 2.7257e-05 | 9.3413e-05 | 1.4755e-05 | 0.00019194 | 0
11 | 0.017292 | 0.014988 | 0.012993 | 0.013773 | 0.016526 | 0.010975 | 0.013578 | 0.011297 | 0.012809 | 0.013017 | 0.011365 | 0.0098466 | 0.013782 | 0.01327 | 0.013408 | 0.013328 | 0.012229 | 0.014288 | 0.01336 | 0.0044397 | 0.64073
12 | 0.085968 | 0.014948 | -0.0016693 | -0.016036 | 0.0013067 | -0.0053532 | -0.0076484 | 0.0042151 | -0.015221 | 0.0091595 | -0.010483 | -0.0049289 | 0.015189 | -0.010038 | -0.0096498 | -0.0051009 | 0.00029449 | -0.0062274 | -0.0083242 | 0.0066324 | -0.0030468 | ten times!
13 | 0.11885 | 0.022694 | 0.028279 | 0.030125 | 0.030177 | 0.013689 | 0.0066055 | 0.030185 | 0.03334 | 0.023803 | 0.028274 | 0.032053 | 0.020095 | 0.029271 | 0.019033 | 0.013373 | 0.023402 | 0.025633 | 0.01993 | 0.0044833 | 0.63789 | ten times!
14 | 0.043546 | -0.0033526 | 0.0093105 | 0.007465 | 0.0048605 | 0.01301 | 0.0062104 | 0.0013902 | 0.0056737 | 0.002039 | 0.01145 | 0.0048405 | -0.0054614 | 0.011553 | 0.0093572 | 0.0090132 | 0.009913 | -0.00020009 | 0.0074401 | -0.017736 | -0.016102 | ten times!
15 | -0.1165 | -0.033687 | -0.027594 | -0.046326 | -0.026832 | -0.043143 | -0.036196 | -0.021027 | -0.028024 | -0.028525 | -0.022128 | -0.021327 | -0.05048 | -0.046903 | -0.026134 | -0.025156 | -0.038893 | -0.02618 | -0.026796 | -0.12205 | 0.1805 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1.8435 | 0.055706 | 0.083387 | 1.6397 | 1.1683 | 0.58997 | -0.35846 | -0.77038 | -1.0139 | -0.44894 | -0.18919 | -0.73714 | -0.45611 | -0.77054 | -0.50461 | 0.26882 | 1.5338 | 0.90348 | 0.010571 | -2.0058
2 | -6.0361 | -0.30179 | -0.44341 | -0.2005 | -0.51248 | -0.29635 | -0.32506 | -0.085246 | -0.18846 | -0.11716 | -0.23542 | -0.30144 | -0.1591 | -0.2293 | -0.24617 | -0.28682 | -0.191 | -0.2104 | -0.074282 | -0.37928 | 0.056946
3 | -0.96727 | 0.080092 | 0.024783 | 0.027863 | -0.062764 | 0.0018294 | -0.049024 | 0.11387 | 0.067599 | 0.10591 | 0.070904 | 0.098598 | 0.043463 | 0.10352 | 0.072149 | 0.056686 | -0.065193 | 0.012684 | -0.015336 | -0.12778 | -0.25665
4 | -2.1269 | -0.038239 | 0.021048 | -0.032714 | 0.0051216 | 0.067677 | 0.022203 | -0.0044177 | 0.064672 | -0.01144 | 0.010534 | 0.042519 | 0.0018237 | 0.022194 | 0.021858 | 0.020593 | 0.030645 | 0.03232 | 0.014194 | -0.051835 | 0.26798
5 | 0.11816 | 0.0096016 | 0.0057323 | 0.018528 | -0.010377 | -0.013256 | 0.0011324 | 0.020248 | 0.0060601 | 0.0087031 | 0.010955 | -0.0095041 | 0.012496 | 0.015288 | 0.019745 | 0.0078992 | 0.00061181 | 0.007825 | -0.0057948 | -0.0078868 | 0.034921
6 | -0.44122 | 0.0020132 | 0.0053271 | 0.0058998 | 0.0047782 | -0.00532 | 0.020685 | 0.0012311 | 0.0084318 | 0.0061626 | -0.0032986 | 8.8552e-05 | 0.012882 | 0.015451 | 0.003992 | 0.0069631 | 0.0031937 | 0.0059581 | 0.0058902 | -0.00017705 | 0.067203
7 | 0.30449 | -0.0022689 | 0.00082968 | 0.0028324 | 0.016409 | -0.00065491 | 0.0017599 | -0.0039277 | 0.0051595 | -0.016374 | -0.0046188 | -0.00092938 | -0.011345 | -0.0021952 | 0.0020849 | 0.00093985 | 0.0095024 | 0.0092303 | 0.004981 | -0.010748 | 0.00078749
8 | -0.096866 | 0.0063016 | -0.0026486 | -0.0019729 | 0.0037308 | 0.00059374 | -0.001691 | 0.00414 | -0.000438 | -0.00084338 | -0.00099061 | 0.0032676 | 0.0026805 | 0.0024773 | 0.00010747 | -0.0017491 | 0.0008481 | -0.0014335 | 0.0025148 | -5.4435e-05 | 0.015057
9 | -0.078838 | 0.004864 | 0.0053226 | 0.012174 | 0.0077635 | -0.0013415 | -0.0032149 | -0.0017257 | -0.0017934 | -7.0919e-05 | -0.0021052 | -0.00093067 | -0.00057151 | -0.0043018 | -0.00010646 | -0.0018489 | 0.005222 | 0.0074256 | 0.00087665 | 0.0022646 | -0.02789
10 | 0.00068742 | 4.6815e-06 | 7.0253e-08 | -1.2342e-05 | 6.6359e-07 | -1.3277e-05 | 6.1471e-06 | 5.0249e-06 | -3.9788e-05 | 5.5039e-06 | 9.6236e-06 | -3.3425e-06 | -5.3618e-06 | 0.0021824 | -7.7928e-05 | -0.0017745 | -1.5753e-06 | 2.5556e-05 | 2.3414e-05 | -0.00015216 | 0
11 | 0.060605 | 0.0044323 | -0.0014845 | 0.011935 | 0.0093304 | 0.0051048 | -0.001217 | -0.0047876 | -0.0067621 | -0.0025838 | 0.0016858 | -0.0028863 | -0.0010541 | -0.0052765 | -0.0016872 | 0.0022018 | 0.011456 | 0.0099284 | 0.0020422 | -0.0056066 | -0.033834
12 | 0.013221 | 0.0060992 | -0.0133 | 0.018724 | -0.014419 | 0.0001055 | -0.013125 | -0.018258 | -0.0033839 | 0.019564 | -0.013106 | 0.0066957 | 0.014606 | 0.025638 | 0.00057721 | -0.010308 | 0.010454 | 0.0063395 | 0.0017248 | -0.007327 | -0.0027405 | ten times!
13 | -0.042986 | 0.0043432 | -0.00047427 | 0.014578 | 0.0012723 | 0.011023 | -0.0071262 | -0.014522 | -0.004052 | -0.00075637 | 0.0085537 | 0.0010972 | 0.015248 | 0.015918 | 0.009305 | -0.0036901 | 0.021665 | 0.0075703 | 0.0024935 | -0.00042079 | -0.036234 | ten times!
14 | 0.022177 | 0.015896 | 0.019494 | 0.017183 | 0.017969 | 0.020323 | 0.027926 | 0.018411 | 0.02231 | 0.022182 | 0.026521 | 0.02961 | 0.01449 | 0.023438 | 0.019541 | 0.010953 | 0.018535 | 0.020996 | 0.022953 | -0.014736 | -0.017542 | ten times!
15 | 0.018208 | 0.026079 | 0.014011 | 0.010482 | -0.001615 | 0.0083631 | -0.01197 | 0.010731 | 0.038204 | 0.010644 | 0.013703 | 0.0083204 | 0.024999 | 0.033333 | 0.036528 | 0.017904 | 0.00093718 | 0.0019439 | 0.0026164 | 0.0013213 | -0.020873 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.098386 | 4.7355 | -0.039807 | -0.043839 | 0.95384 | -1.5159 | 1.071 | -0.27444 | -0.049916 | 17.7087 | 0.70473 | -0.26027 | 4.422 | 0.063047 | -0.11207 | 0.93833 | -1.0569 | 0.24458 | -0.20303 | 0.28121 | 15.1524 | 0.70458 | 1.4631 | 5.8526 | 0.014481 | -0.27425 | 0.98908 | 0 | 1.0803 | -0.30282 | -0.049412 | 14359.014 | -1.6247 | 5.5535 | 0.014782 | -0.35028 | 0.9743 | 0 | 0.056946 | -0.25665 | 0.26798 | 14359.263
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | English | 0 | 0