## Alarm sheet
File name | HCMBBLA001-3000076_cm.xls
Component ID | HCMBBLA001 |  |  |  | Serial Number | 3000076
Date of test Ap 1 | 2003-10-30 |  |  |  | Date of test Ap 2 | 2003-10-30
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | yellow alarm | yellow alarm | yellow alarm
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 |  |  |  |  | b10
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | yellow alarm | status ok | status ok | status ok | b15 | yellow alarm | status ok | status ok | status ok
a2 | status ok | status ok | status ok | yellow alarm | a2 | status ok | status ok | status ok | status ok
a3 | yellow alarm | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | yellow alarm | status ok | status ok | status ok | a4 | yellow alarm | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | yellow alarm | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 |  |  |  |  | a10
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | yellow alarm | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity

## Assembly Data
Magnet name
| HCMBBLA001 | 3000076 |  | Additional lamin. (mm) |  | 0
| COMMENTS: | Yellow alarms in position 1, 2 and 20 in aperture 2 due to wrong positioning of measuring coil along the magnet axis (17 mm difference at p 1) + yellow alarms in ap 1 on a3 (4.6 sigma), a4 (5.9 sigma), a6 (3.5 sigma) average straight, a2 p20 (3.6 sigma) and in ap 2 on a4 average straight (6 sigma)
| HOLDING POINT | OK-W | Data at CERN |  | 2003-10-30 | Analysis by | V. Remondino | Answer to MMS-LD |  | 2003-10-30 | Answer to manufacturer |  | 2003-10-31
1 | Fill in the yellow fields in Original Data | [10 empty] | English
2 | CTRL-g to load the measurement files .txt
3 | Save as … with the "official" name in Original Data C1
4 | Open the file of the collared coil measurement of the same magnet
5 | Go back to cold_mass file. Tools - Macro - Macros - Make difference - Run (to make the difference between collared coil and cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (to perform the analysis - results in Alarm sheet)
7 | Save the file
| Cold mass template - version 11.10.2002 | [7 empty] | Version : | cern
1 | Remplir les champs jaunes dans Original Data | [10 empty] | Francais
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec le nom "officiel" dans Original Data C1
4 | Ouvrir le fichier avec les mesures de la collared coil
5 | Revenir au ficher cold_mass. Tools - Macro - Macros - Make difference - Run (pour faire la difference entre collared coil et cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (pour faire l'analyse - resultats dans Alarm sheet)
7 | Sauver le fichier
1 | Riempire I campi gialli in Original Data | [10 empty] | Italiano
2 | CTRL-g per caricare I file di misura .txt
3 | Salvare il file con il nome "ufficiale" in Original Data C1
4 | Aprire il file delle misure della collared coil corrispondente
5 | Tornare sul file cold_mass. Tools - Macro - Macros - Make difference - Run (per fare la differenza tra collared coil e cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (per fare l'analisi - risultati in Alarm sheet)
7 | Salvare il file
1 | Die gelben Felder in Original Data ausfüllen | [10 empty] | Deutsch
2 | Ctrl-g um die Messdatei ( Datei .txt) zu laden.
3 | Die Datei mit der ”offiziellen” bezeichnung im Original Data C1 speichern mit “Save as…”.
4 | Öffnen “Collared Coils” Messdatei (gleiche Magnete versteht sich)
5 | Zurück zum cold mass Datei. Tools - Macro - Macros - Make difference - Run (Zeigt allen unterschieden in Messdatei zw. cc und cm)
6 | Tools - Macro - Macros - Collared coil analysis - Run (Um Maßdaten zu analysieren - Resultaten ins Alarm Sheet)
7 | Datei speichern
| Last modif on | by | Modifications:
| 2002-10-20 | E. Todesco | Cell for iron additional iron laminations - check on magnetic length includes additional iron laminations
| 2002-11-13 | E. Wildner | German version of instructions added

## Original data
File name |  | HCMBBLA001-03000076_cm.xls
Component ID |  | HCMBBLA001 |  |  | Serial Number |  |  | 3000076 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 3076BL_cm1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 34 |  |  |  | File name for aperture 2 |  |  | 3076BL_cm2.txt |  |  |  | a even | 1
Test Operator |  | A.Bomba |  |  | Test Controller |  |  | V. Remondino
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2003-10-30 |  |  | Time of test |  |  | 10:57 |  |  |  | Date of test |  | 2003-10-30 |  |  | Time of test |  |  | 10:57
Mole name |  | Dipole 4 | [9 empty] | Mole name |  | Dipole 5
Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 20 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 20
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.51 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.49 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00012655 |  |  | Dy average |  |  | 0.00046097 |  |  |  | Dx average |  | 0 |  |  | Dy average |  |  | 0.00030534
Magnetic length |  | 14.3569 |  |  | Tranfs. Function (mT/KA) |  |  | 704.298 |  |  |  | Magnetic length |  | 14.3546 |  |  | Tranfs. Function (mT/KA) |  |  | 704.263
Phase |  | -1.864 |  |  | Integral field (T.m) |  |  | 101.053 |  |  |  | Phase |  | -19.8469 |  |  | Integral field (T.m) |  |  | 100.9117
File | Aperture 1 - Cold mass | [19 empty] | Integrals | File | Aperture 2 - Cold mass | [19 empty] | Integrals
C1 (mT) | 3.8781 | 7.0373 | 7.0421 | 7.0428 | 7.0427 | 7.0435 | 7.0438 | 7.0442 | 7.0441 | 7.0435 | 7.0435 | 7.0448 | 7.0437 | 7.0432 | 7.0434 | 7.0438 | 7.044 | 7.0435 | 7.0399 | 4.1034 | 101.053 | C1 (mT) | 3.6697 | 7.0328 | 7.0424 | 7.0428 | 7.0436 | 7.0445 | 7.0445 | 7.0436 | 7.0437 | 7.0432 | 7.0432 | 7.0442 | 7.0428 | 7.043 | 7.0433 | 7.0433 | 7.0439 | 7.0424 | 7.0401 | 4.1299 | 100.9117
Angle (mrad) | 4.65 | 0.35995 | 0.71975 | 0.89856 | 0.40201 | 0.7347 | 0.59896 | 0.46114 | -0.07979 | 0.11271 | -0.10943 | -0.57716 | -0.51499 | -0.73752 | -0.75328 | -0.95461 | -0.76457 | -0.54817 | -0.68062 | -1.9356 | -1.864 | Angle (mrad) | 5.45 | 0.36308 | 0.71559 | 1.3403 | 0.93818 | 0.8176 | 0.69938 | 0.40529 | 0.44398 | 0.13427 | -0.043608 | -0.5274 | -0.77367 | -0.24917 | -0.64062 | -0.99175 | -1.5003 | -1.3347 | -1.4243 | -2.067 | -19.8469
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 26.333 | 0.89268 | 0.78182 | 0.86659 | 1.217 | 1.1526 | 0.23826 | 1.3729 | 0.5094 | 0.69132 | 0.49283 | 1.1559 | 1.2269 | 0.5075 | 0.70068 | 0.68767 | 0.56529 | 1.0111 | 0.70299 | 22.8765 | 2.2266 | b2 | -25.8844 | -1.6516 | -1.8065 | -1.2312 | -1.449 | -1.3406 | -1.5105 | -1.4818 | -1.4266 | -1.0569 | -1.0709 | -1.2703 | -1.1507 | -0.87558 | -1.6229 | -1.4351 | -1.6087 | -1.8988 | -1.5091 | -22.2604 | -2.7182
b3 | 38.0976 | 3.287 | 3.9677 | 3.7909 | 3.9273 | 3.3784 | 4.0046 | 3.6071 | 4.1301 | 3.6524 | 3.0328 | 4.3614 | 4.3793 | 1.9317 | 2.4113 | 4.0916 | 3.974 | 3.6524 | 4.101 | 7.9313 | 4.7708 | b3 | 39.5352 | 3.4551 | 3.9465 | 3.718 | 4.3005 | 3.476 | 3.3395 | 4.3126 | 4.3454 | 4.8686 | 3.3998 | 3.9552 | 3.9887 | 2.4229 | 3.0354 | 3.3434 | 3.7813 | 4.1955 | 4.9491 | 9.1331 | 4.9609
b4 | 1.0112 | -0.073974 | 0.051262 | 0.10196 | 0.056802 | 0.13903 | -0.29319 | -0.055441 | -0.039342 | 0.036774 | -0.041946 | 0.09994 | 0.023779 | -0.0086531 | 0.047913 | 0.062027 | 0.019147 | 0.049814 | -0.075557 | 0.96038 | 0.063595 | b4 | -1.3822 | -0.27526 | -0.18965 | -0.23115 | -0.18666 | -0.4242 | -0.35581 | -0.25799 | -0.37565 | -0.13092 | -0.087673 | -0.25289 | -0.022199 | -0.14144 | -0.28334 | -0.13601 | -0.1925 | -0.21661 | -0.22294 | -1.2699 | -0.28511
b5 | -5.341 | -0.12479 | -0.17963 | 0.099318 | 0.10546 | 0.1667 | 0.20575 | 0.41524 | 0.37032 | 0.21313 | 0.38815 | 0.22937 | 0.080033 | 0.17098 | 0.20235 | -0.093838 | -0.14156 | -0.1533 | -0.12583 | -2.7315 | -0.14133 | b5 | -5.348 | 0.15044 | 0.009369 | 0.32506 | 0.43098 | 0.56366 | 0.60926 | 0.48017 | 0.62603 | 0.51594 | 0.63841 | 0.42349 | 0.42847 | 0.17357 | 0.31233 | 0.25875 | 0.085732 | 0.12528 | 0.089877 | -2.3457 | 0.10913
b6 | 0.36668 | 0.011793 | -0.07655 | 0.010634 | -0.017076 | 0.1016 | 0.23938 | 0.05506 | 0.18025 | 0.056781 | 0.088619 | -0.0032499 | -0.022112 | 0.15617 | -0.024442 | 0.034083 | 0.098595 | -0.066927 | -0.029631 | -0.013956 | 0.051578 | b6 | 0.40458 | -0.059977 | -0.062139 | 0.017503 | -0.021185 | -0.066751 | -0.050054 | -0.078037 | 0.0030579 | 0.082788 | 0.023024 | 0.062688 | 0.061018 | -0.10307 | -0.074064 | 0.028976 | 0.052383 | -0.036658 | -0.016429 | -0.19668 | -0.0073988
b7 | 2.4645 | 0.81461 | 0.97271 | 1.1699 | 1.2028 | 1.19 | 1.1811 | 1.1881 | 1.1252 | 1.1858 | 1.1819 | 1.1476 | 1.1258 | 1.1754 | 1.1612 | 1.1428 | 1.1571 | 1.1907 | 1.1144 | 1.022 | 1.1697 | b7 | 2.6651 | 0.86505 | 0.86219 | 1.0943 | 1.0167 | 1.0836 | 1.0696 | 1.0946 | 1.073 | 1.0973 | 1.1696 | 1.1339 | 1.0844 | 1.0454 | 1.0195 | 1.074 | 1.0586 | 1.1275 | 1.0997 | 0.93785 | 1.0995
b8 | -0.0007265 | -0.042509 | -0.019056 | 0.0028045 | -0.017295 | -0.0048116 | -0.046338 | -0.018633 | -0.059905 | -0.033216 | 0.0061778 | 0.03144 | -0.0010692 | 0.016557 | -0.039508 | 0.00206 | -0.014191 | -0.038554 | -0.048194 | -0.061196 | -0.018829 | b8 | 0.017076 | -0.036703 | -0.040222 | 0.0031232 | -0.02215 | 0.024579 | 0.00066096 | 0.0099592 | 0.013451 | -0.017401 | -0.043953 | -0.035017 | -0.050972 | -0.035212 | -0.031529 | 0.012651 | 0.023868 | -0.016986 | -0.0036791 | -0.041098 | -0.013643
b9 | 0.29912 | 0.47412 | 0.46577 | 0.4698 | 0.48299 | 0.47198 | 0.48477 | 0.46038 | 0.47174 | 0.4831 | 0.46365 | 0.47315 | 0.47167 | 0.46817 | 0.46142 | 0.47638 | 0.47011 | 0.47784 | 0.47206 | 0.43657 | 0.46611 | b9 | 0.34052 | 0.49172 | 0.49543 | 0.48507 | 0.50139 | 0.48798 | 0.4836 | 0.47764 | 0.50658 | 0.50682 | 0.48593 | 0.47343 | 0.48685 | 0.4638 | 0.47532 | 0.48564 | 0.47826 | 0.51585 | 0.50124 | 0.45421 | 0.48391
b10 | 0.10254 | -0.041191 | -0.026477 | 0.033314 | -0.024092 | 0.0266 | 0.004864 | -0.0095011 | 0.0059732 | -0.0019099 | 0.0026733 | 0.047216 | -0.022187 | 0.062854 | -0.051663 | 0.022718 | 0.031704 | -0.019305 | -0.043664 | -0.093486 | 0 | b10 | 0.10266 | -0.043062 | -0.035477 | 0.033539 | 0.020529 | 0.046088 | -0.00091957 | 0.0081423 | 0.0045701 | 0.0085471 | -0.012166 | 0.026431 | -0.031863 | 0.024037 | -0.049549 | 0.0171 | 0.040362 | -0.037854 | -0.020474 | -0.098871 | -0.00033548
b11 | 0.56173 | 0.6535 | 0.64729 | 0.65708 | 0.65343 | 0.65523 | 0.65184 | 0.65995 | 0.66119 | 0.65323 | 0.66151 | 0.65552 | 0.65616 | 0.65814 | 0.66245 | 0.65231 | 0.65274 | 0.65074 | 0.64717 | 0.60259 | 0.65069 | b11 | 0.58517 | 0.65895 | 0.66058 | 0.66426 | 0.66619 | 0.66894 | 0.6698 | 0.66485 | 0.66754 | 0.65647 | 0.66634 | 0.65979 | 0.66345 | 0.66759 | 0.671 | 0.66383 | 0.66263 | 0.663 | 0.65057 | 0.60408 | 0.65969
b12 | 0.010924 | -0.0031806 | -0.0024639 | 0.0034136 | -0.003641 | 0.00023186 | 0.00037771 | -0.00045437 | 0.0012244 | -0.0010117 | -0.0011769 | 0.0025413 | -0.0035791 | 0.0037291 | -0.0098426 | 0.0006297 | 0.0032593 | -0.0044448 | -0.0079604 | -0.013654 | -0.0012691 | b12 | 0.011034 | -0.0061743 | -0.0045299 | 0.0058696 | 0.0017371 | 0.009311 | 0.0038727 | 0.0048627 | 0.0068717 | 0.001078 | -0.0024139 | 0.001524 | -0.0062613 | 0.0036762 | -0.0016564 | 0.00051518 | 0.0050739 | -0.0041924 | -0.00026177 | -0.0061394 | 0.0011026
b13 | 0.077535 | 0.068829 | 0.06899 | 0.068884 | 0.069172 | 0.065884 | 0.067414 | 0.065667 | 0.067369 | 0.067796 | 0.067767 | 0.067186 | 0.06906 | 0.06595 | 0.068138 | 0.069235 | 0.071233 | 0.073641 | 0.0721 | 0.055482 | 0.068432 | b13 | 0.079892 | 0.065986 | 0.065459 | 0.066154 | 0.064985 | 0.063425 | 0.0618 | 0.066818 | 0.060876 | 0.064132 | 0.065998 | 0.06773 | 0.067445 | 0.067113 | 0.067205 | 0.064764 | 0.06674 | 0.068556 | 0.06805 | 0.04985 | 0.065634
b14 | -0.002213 | -0.0032169 | -0.0024608 | -9.5792e-05 | -0.0037339 | -0.0012222 | -3.8982e-05 | -0.0014578 | -0.00099054 | -0.0021045 | -0.0024859 | -7.4831e-05 | -0.0034344 | -0.00088086 | -0.0041594 | -0.0021303 | -0.0006686 | -0.003363 | -0.0038724 | -0.0049422 | -0.002116 | b14 | -0.0015886 | -0.0028788 | -0.0012211 | 0.0010149 | -0.00089894 | 0.001128 | -0.0003856 | -0.00022799 | 0.00062901 | -0.00040017 | -0.00084658 | 0.00041289 | -0.0019954 | 0.00073644 | -0.0026029 | -0.00060906 | -7.9286e-06 | -0.0033017 | -0.00099326 | -0.0022466 | -0.00076352
b15 | -0.011652 | 0.025207 | 0.021928 | 0.017971 | 0.017983 | 0.019543 | 0.017894 | 0.019621 | 0.020076 | 0.018897 | 0.021327 | 0.019081 | 0.019248 | 0.020108 | 0.016922 | 0.019222 | 0.018653 | 0.019782 | 0.017643 | 0.011939 | 0.018379 | b15 | -0.012783 | 0.021324 | 0.020284 | 0.016901 | 0.019832 | 0.019122 | 0.020079 | 0.018921 | 0.020313 | 0.017907 | 0.01867 | 0.017991 | 0.017657 | 0.018512 | 0.017959 | 0.018698 | 0.018309 | 0.017668 | 0.017849 | 0.0065677 | 0.017542
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 46.5007 | 3.5995 | 7.1975 | 8.9856 | 4.0201 | 7.347 | 5.9896 | 4.6114 | -0.7979 | 1.1271 | -1.0943 | -5.7716 | -5.1499 | -7.3752 | -7.5328 | -9.5461 | -7.6457 | -5.4817 | -6.8062 | -19.3562 | 0 | a1 | 54.5007 | 3.6308 | 7.1559 | 13.4029 | 9.3818 | 8.176 | 6.9938 | 4.0529 | 4.4398 | 1.3427 | -0.43608 | -5.274 | -7.7367 | -2.4917 | -6.4062 | -9.9175 | -15.0029 | -13.3474 | -14.243 | -20.6701 | 0
a2 | -3.2344 | -0.71641 | -0.94025 | -0.16779 | -0.35037 | -0.41358 | -1.083 | -0.17223 | -0.65703 | -0.93791 | -0.46434 | -0.94223 | -0.43327 | -0.63883 | -0.53223 | -0.48766 | -0.46894 | -0.70923 | -0.3053 | 14.3891 | -0.19957 | a2 | 0.4051 | -0.068983 | -0.21402 | 0.39789 | 1.1027 | 0.57 | 0.35901 | 1.0894 | 0.45193 | 0.37037 | 0.56829 | 1.18 | 1.0105 | 1.8139 | 1.372 | 2.0657 | 2.0108 | 0.97112 | 0.19466 | 6.6561 | 1.0133
a3 | -1.8338 | 1.1935 | 1.2556 | 1.115 | 0.61598 | 0.56641 | 0.49029 | 0.62832 | 0.93915 | 0.90571 | 0.79651 | 0.60467 | 0.4983 | 0.30973 | 0.54849 | 0.13762 | 0.2449 | 0.5977 | 0.81356 | -0.44383 | 0.57452 | a3 | -2.0997 | 1.5565 | 1.1453 | 0.66359 | 0.77772 | 0.61058 | 0.33905 | 0.5824 | 0.72327 | 0.67763 | 0.64071 | 0.39579 | 0.37769 | 0.92411 | 0.64018 | 0.63256 | 0.78113 | 0.69358 | 0.92603 | -0.013813 | 0.6272
a4 | -0.12177 | 0.16174 | -0.00063252 | -0.051878 | -0.021398 | 0.12668 | -0.12914 | -0.017677 | 0.36552 | 0.54199 | 0.25091 | 0.17167 | 0.056712 | 0.13852 | 0.085414 | 0.27593 | 0.0062051 | 0.1481 | -0.047646 | 1.0588 | 0.13646 | a4 | -0.025094 | 0.34176 | -0.0091064 | -0.073491 | 0.16321 | -0.2368 | -0.39637 | 0.022112 | -0.21141 | -0.14784 | -0.12424 | -0.56029 | -0.25092 | 0.19452 | -0.080935 | -0.50427 | -0.11587 | 0.10395 | 0.059071 | 1.1728 | -0.060351
a5 | 1.8751 | 0.10424 | 0.15412 | 0.4032 | 0.37274 | 0.34881 | 0.54593 | 0.36976 | 0.30692 | 0.31042 | 0.32933 | 0.28323 | 0.20973 | 0.39675 | 0.23596 | 0.18391 | 0.077535 | 0.22204 | 0.34993 | 0.11952 | 0.32963 | a5 | 2.0223 | 0.23601 | 0.054895 | 0.072078 | 0.065871 | 0.09986 | -0.017326 | 0.075205 | 0.03441 | 0.12026 | 0.15806 | 0.0025911 | 0.020923 | 0.25465 | 0.21478 | 0.20956 | 0.12905 | 0.11356 | 0.21777 | -0.063356 | 0.16112
a6 | -0.012925 | -0.032392 | -0.065583 | -0.05539 | 0.02436 | -0.044928 | -0.0039747 | -0.086166 | -0.095489 | 0.016892 | -0.074334 | -0.073096 | -0.11699 | -0.041344 | -0.04576 | -0.0080835 | 0.028122 | -0.019596 | -0.034202 | -0.082738 | -0.040939 | a6 | 0.056897 | 0.03588 | 0.0081171 | 0.099877 | 0.11778 | 0.059936 | 0.03653 | 0.070602 | 0.071678 | -0.020962 | -0.012705 | -0.076786 | 0.055022 | 0.13005 | 0.1437 | 0.045345 | 0.025781 | 0.052782 | 0.09544 | 0.095752 | 0.053584
a7 | 1.3683 | -0.073428 | 0.034331 | 0.023632 | 0.0097696 | -0.051916 | -0.084241 | -0.0085557 | -0.027168 | -0.041022 | 0.048135 | -0.015087 | 0.060878 | 0.0053266 | 0.033624 | -0.037443 | -0.014983 | 0.032875 | -0.0015881 | 0.027861 | 0.034643 | a7 | 1.3688 | -0.088449 | -0.045997 | -0.076034 | -0.075064 | -0.051653 | -0.083842 | 0.00023089 | 0.034339 | -0.072919 | -0.011345 | -0.053255 | -0.053309 | -0.0044957 | -0.090432 | -0.0031173 | -0.00057821 | -0.00094069 | -0.032298 | 0.020321 | 0.00084166
a8 | -0.030357 | 0.0059248 | -0.016452 | 0.0013733 | -0.019123 | -0.011432 | -0.0015836 | -0.006486 | -0.0059075 | -0.02658 | -0.039264 | -0.03256 | -0.054453 | -0.01293 | -0.012803 | -0.006718 | -0.0011944 | -0.047509 | 0.021069 | -0.015451 | -0.015281 | a8 | -0.070884 | -0.0059599 | -0.037104 | -0.0019742 | -0.039134 | -0.044189 | -0.046379 | -0.02561 | -0.045441 | -0.022338 | -0.066002 | -0.055319 | -0.013667 | -0.027143 | -0.045714 | -0.018023 | -0.039129 | -0.052138 | -0.013064 | -0.011666 | -0.033608
a9 | -0.2808 | 0.018501 | 0.022755 | 0.020161 | -0.014909 | -0.012311 | 0.012541 | 0.0216 | 0.01473 | -0.0088948 | -0.00051485 | -0.0068765 | 0.012847 | 0.0079706 | -0.0065881 | -0.02182 | -0.0049768 | 0.014198 | -0.00080112 | -0.069015 | -0.0066498 | a9 | -0.28926 | 0.03524 | 0.017929 | 0.0047521 | -0.013673 | 0.011205 | -0.0096328 | -0.00632 | -0.0023212 | -0.020832 | -0.024698 | -0.050437 | -0.029986 | -0.038657 | -0.027647 | -0.023039 | -0.0040392 | -0.0062531 | -0.005382 | -0.10334 | -0.021206
a10 | 0.017139 | -0.010111 | -0.011675 | 0.0029064 | -0.0018316 | -0.0028219 | 0.0199 | 0.0034996 | 0.013027 | 0.020701 | -0.034996 | -0.0014139 | -0.017087 | -0.0076814 | 0.020406 | 0.019611 | 0.021709 | -0.014242 | 0.037395 | -0.11454 | 0 | a10 | -0.0086307 | 0.0058913 | -0.0025281 | 0.020951 | -0.00067201 | 0.0060024 | 0.00090346 | 0.0078115 | -0.017507 | 0.007881 | -0.025712 | -0.014115 | 0.011396 | -0.0053021 | 0.034283 | -0.0089792 | -0.001015 | -0.0093785 | 0.024163 | -0.049543 | 2.6446e-05
a11 | 0.13642 | -0.056547 | -0.037033 | -0.024221 | -0.032169 | -0.027908 | -0.022117 | -0.02552 | -0.033554 | -0.03533 | -0.028242 | -0.033868 | -0.02886 | -0.03332 | -0.034493 | -0.03241 | -0.036327 | -0.028657 | -0.032333 | -0.028086 | -0.027394 | a11 | 0.15854 | -0.055277 | -0.047493 | -0.034517 | -0.040256 | -0.040323 | -0.046968 | -0.036815 | -0.036252 | -0.049886 | -0.040563 | -0.051009 | -0.045301 | -0.046258 | -0.054645 | -0.039586 | -0.047829 | -0.044378 | -0.046863 | -0.03534 | -0.038849
a12 | -0.0094597 | -0.0024604 | -0.007765 | -0.0065653 | -0.0062359 | -0.0072227 | -0.0028226 | -0.0057151 | -0.0083893 | -0.0032602 | -0.011937 | -0.0069276 | -0.0087089 | -0.0053604 | -0.0016263 | 0.00055221 | 0.0013734 | -0.003428 | 0.0023536 | -0.0093165 | -0.004954 | a12 | -0.0077956 | -0.00098289 | -0.0021014 | 0.0044967 | 0.0031423 | 0.00092209 | -0.0026628 | 0.0004697 | -0.0020945 | -0.0015972 | -0.0036274 | -0.0018663 | 0.0014583 | -0.00027846 | 0.0073136 | 0.0040129 | 0.0030587 | -0.00086678 | 0.0039741 | -0.0058969 | 0.00027484
a13 | -0.024172 | 0.0018496 | 0.0023574 | 0.00040754 | -0.00082778 | -0.002621 | -7.713e-05 | -2.7835e-05 | -0.00078924 | -0.00052338 | 0.0019519 | -0.00096247 | -7.389e-05 | -0.0010553 | 0.00036203 | -0.0026576 | -0.0024553 | 6.4284e-05 | 0.00068103 | -0.0013631 | -0.00096714 | a13 | -0.025226 | 0.0030294 | 0.0013544 | -0.00025975 | 0.00092182 | 0.00010245 | 0.00025801 | -0.0019602 | -0.00029864 | -0.0024658 | -0.0021755 | -0.0062194 | -0.0036622 | -0.0037819 | -0.0043377 | -0.002768 | -0.002343 | -0.0007591 | -0.00035826 | -0.0041928 | -0.0021632
a14 | 0.0085194 | -0.0075067 | -0.0085321 | -0.0053979 | -0.0070084 | -0.0076674 | -0.0067646 | -0.0071915 | -0.00557 | -0.0066879 | -0.0085684 | -0.0072893 | -0.0074561 | -0.006781 | -0.0057513 | -0.0045643 | -0.0056193 | -0.0075396 | -0.0034842 | -0.00062454 | -0.0060133 | a14 | 0.0088477 | -0.004952 | -0.0064833 | -0.0045943 | -0.0058618 | -0.0082629 | -0.0079717 | -0.0059881 | -0.0096045 | -0.0053341 | -0.0073134 | -0.0072961 | -0.0065512 | -0.0063385 | -0.0067713 | -0.0073948 | -0.0052316 | -0.0053951 | -0.0044617 | 0.0018912 | -0.0057617
a15 | 0.0058449 | -0.006534 | -0.0043186 | 0.00071483 | -0.0043947 | -0.0023315 | -0.0027519 | -0.0024892 | -0.0012977 | -0.0032415 | -0.0041795 | -0.0017625 | -0.0043938 | 0.00055919 | -0.0061184 | -0.0028449 | -0.0024302 | -0.0056524 | -0.0037267 | -0.0091881 | -0.0031006 | a15 | 0.010535 | -0.0017903 | -0.0019069 | 0.000801 | 0.00026412 | 0.0019641 | 0.00010773 | 0.00079446 | 0.00063081 | -0.00064944 | -0.0017348 | -0.0015201 | -0.0020753 | -0.00075943 | -0.0024346 | -0.00034348 | -0.00018609 | -0.0023849 | -0.00017891 | -0.0024575 | -0.00038474
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00030507 | 0.0001041 | 6.7546e-05 | -8.581e-05 | 6.2295e-05 | -6.9197e-05 | -1.0913e-05 | 2.4787e-05 | -1.3626e-05 | 0 | -1.0684e-05 | -0.00012232 | 5.5408e-05 | -0.00016295 | 0.00013502 | -5.6527e-05 | -7.9178e-05 | 4.8696e-05 | 0.00011932 | 0.00024737 | 0.00012655 | Dx (m) | -0.00027152 | 0.00011161 | 9.0375e-05 | -8.2814e-05 | -5.2299e-05 | -0.00011583 | 0 | -1.9653e-05 | -1.4012e-05 | -2.0465e-05 | 2.6937e-05 | -7.0491e-05 | 8.3265e-05 | -6.1853e-05 | 0.00013168 | -4.5006e-05 | -0.00010322 | 9.5036e-05 | 5.7738e-05 | 0.00026903 | 0
Dy (m) | 2.2654e-05 | 3.5326e-05 | 3.4543e-05 | -1.066e-05 | 0 | 0 | -5.2275e-05 | 0 | -3.419e-05 | -5.3455e-05 | 8.9439e-05 | 0 | 4.6716e-05 | 1.166e-05 | -4.5364e-05 | -5.3917e-05 | -6.0957e-05 | 3.9358e-05 | -9.2336e-05 | 0.00033545 | 0.00046097 | Dy (m) | 9.8983e-05 | 0 | 1.3015e-05 | -5.7949e-05 | 0 | -2.2259e-05 | 0 | -2.1064e-05 | 4.3827e-05 | -2.1964e-05 | 6.7236e-05 | 3.0922e-05 | -2.3528e-05 | 0 | -7.6163e-05 | 2.0307e-05 | 0 | 3.0418e-05 | -5.8978e-05 | 0.0001554 | 0.00030534
Magnetic length (m) |  | 14.3569 | [19 empty] | Magnetic length (m) |  | 14.3546
Transfer funct. (mT/KA) |  | 704.298 | [19 empty] | Transfer funct. (mT/KA) |  | 704.263
Coil positioning (mm) |  | 0.017216 | [19 empty] | Coil positioning (mm) |  | 0.019032
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10

## Summary Data
File | HCMBBLA001-03000076 |  |  | Aperture 1 - Cold mass | [16 empty] | Integrals | File | HCMBBLA001-3000076 |  |  | Aperture 2 - Cold mass | [16 empty] | Integrals
C1/i (mT/KA) | 387.8078 | 703.7273 | 704.2111 | 704.2847 | 704.2686 | 704.355 | 704.3799 | 704.4201 | 704.4111 | 704.3456 | 704.3456 | 704.4807 | 704.3658 | 704.32 | 704.3361 | 704.379 | 704.4026 | 704.3455 | 703.9862 | 410.3395 | 101.053 | C1/i (mT/KA) | 366.9721 | 703.2843 | 704.2379 | 704.2808 | 704.3601 | 704.4471 | 704.4491 | 704.3646 | 704.3694 | 704.324 | 704.3235 | 704.4183 | 704.2792 | 704.2997 | 704.3271 | 704.3294 | 704.3928 | 704.2389 | 704.0082 | 412.9943 | 100.9117
Angle (mrad) | 4.65 | 0.35995 | 0.71975 | 0.89856 | 0.40201 | 0.7347 | 0.59896 | 0.46114 | -0.07979 | 0.11271 | -0.10943 | -0.57716 | -0.51499 | -0.73752 | -0.75328 | -0.95461 | -0.76457 | -0.54817 | -0.68062 | -1.9356 | -1.864 | Angle (mrad) | 5.45 | 0.36308 | 0.71559 | 1.3403 | 0.93818 | 0.8176 | 0.69938 | 0.40529 | 0.44398 | 0.13427 | -0.043608 | -0.5274 | -0.77367 | -0.24917 | -0.64062 | -0.99175 | -1.5003 | -1.3347 | -1.4243 | -2.067 | -19.8469
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 26.333 | 0.89268 | 0.78182 | 0.86659 | 1.217 | 1.1526 | 0.23826 | 1.3729 | 0.5094 | 0.69132 | 0.49283 | 1.1559 | 1.2269 | 0.5075 | 0.70068 | 0.68767 | 0.56529 | 1.0111 | 0.70299 | 22.8765 | 2.2266 | b2 | -25.8844 | -1.6516 | -1.8065 | -1.2312 | -1.449 | -1.3406 | -1.5105 | -1.4818 | -1.4266 | -1.0569 | -1.0709 | -1.2703 | -1.1507 | -0.87558 | -1.6229 | -1.4351 | -1.6087 | -1.8988 | -1.5091 | -22.2604 | -2.7182
b3 | 38.0976 | 3.287 | 3.9677 | 3.7909 | 3.9273 | 3.3784 | 4.0046 | 3.6071 | 4.1301 | 3.6524 | 3.0328 | 4.3614 | 4.3793 | 1.9317 | 2.4113 | 4.0916 | 3.974 | 3.6524 | 4.101 | 7.9313 | 4.7708 | b3 | 39.5352 | 3.4551 | 3.9465 | 3.718 | 4.3005 | 3.476 | 3.3395 | 4.3126 | 4.3454 | 4.8686 | 3.3998 | 3.9552 | 3.9887 | 2.4229 | 3.0354 | 3.3434 | 3.7813 | 4.1955 | 4.9491 | 9.1331 | 4.9609
b4 | 1.0112 | -0.073974 | 0.051262 | 0.10196 | 0.056802 | 0.13903 | -0.29319 | -0.055441 | -0.039342 | 0.036774 | -0.041946 | 0.09994 | 0.023779 | -0.0086531 | 0.047913 | 0.062027 | 0.019147 | 0.049814 | -0.075557 | 0.96038 | 0.063595 | b4 | -1.3822 | -0.27526 | -0.18965 | -0.23115 | -0.18666 | -0.4242 | -0.35581 | -0.25799 | -0.37565 | -0.13092 | -0.087673 | -0.25289 | -0.022199 | -0.14144 | -0.28334 | -0.13601 | -0.1925 | -0.21661 | -0.22294 | -1.2699 | -0.28511
b5 | -5.341 | -0.12479 | -0.17963 | 0.099318 | 0.10546 | 0.1667 | 0.20575 | 0.41524 | 0.37032 | 0.21313 | 0.38815 | 0.22937 | 0.080033 | 0.17098 | 0.20235 | -0.093838 | -0.14156 | -0.1533 | -0.12583 | -2.7315 | -0.14133 | b5 | -5.348 | 0.15044 | 0.009369 | 0.32506 | 0.43098 | 0.56366 | 0.60926 | 0.48017 | 0.62603 | 0.51594 | 0.63841 | 0.42349 | 0.42847 | 0.17357 | 0.31233 | 0.25875 | 0.085732 | 0.12528 | 0.089877 | -2.3457 | 0.10913
b6 | 0.36668 | 0.011793 | -0.07655 | 0.010634 | -0.017076 | 0.1016 | 0.23938 | 0.05506 | 0.18025 | 0.056781 | 0.088619 | -0.0032499 | -0.022112 | 0.15617 | -0.024442 | 0.034083 | 0.098595 | -0.066927 | -0.029631 | -0.013956 | 0.051578 | b6 | 0.40458 | -0.059977 | -0.062139 | 0.017503 | -0.021185 | -0.066751 | -0.050054 | -0.078037 | 0.0030579 | 0.082788 | 0.023024 | 0.062688 | 0.061018 | -0.10307 | -0.074064 | 0.028976 | 0.052383 | -0.036658 | -0.016429 | -0.19668 | -0.0073988
b7 | 2.4645 | 0.81461 | 0.97271 | 1.1699 | 1.2028 | 1.19 | 1.1811 | 1.1881 | 1.1252 | 1.1858 | 1.1819 | 1.1476 | 1.1258 | 1.1754 | 1.1612 | 1.1428 | 1.1571 | 1.1907 | 1.1144 | 1.022 | 1.1697 | b7 | 2.6651 | 0.86505 | 0.86219 | 1.0943 | 1.0167 | 1.0836 | 1.0696 | 1.0946 | 1.073 | 1.0973 | 1.1696 | 1.1339 | 1.0844 | 1.0454 | 1.0195 | 1.074 | 1.0586 | 1.1275 | 1.0997 | 0.93785 | 1.0995
b8 | -0.0007265 | -0.042509 | -0.019056 | 0.0028045 | -0.017295 | -0.0048116 | -0.046338 | -0.018633 | -0.059905 | -0.033216 | 0.0061778 | 0.03144 | -0.0010692 | 0.016557 | -0.039508 | 0.00206 | -0.014191 | -0.038554 | -0.048194 | -0.061196 | -0.018829 | b8 | 0.017076 | -0.036703 | -0.040222 | 0.0031232 | -0.02215 | 0.024579 | 0.00066096 | 0.0099592 | 0.013451 | -0.017401 | -0.043953 | -0.035017 | -0.050972 | -0.035212 | -0.031529 | 0.012651 | 0.023868 | -0.016986 | -0.0036791 | -0.041098 | -0.013643
b9 | 0.29912 | 0.47412 | 0.46577 | 0.4698 | 0.48299 | 0.47198 | 0.48477 | 0.46038 | 0.47174 | 0.4831 | 0.46365 | 0.47315 | 0.47167 | 0.46817 | 0.46142 | 0.47638 | 0.47011 | 0.47784 | 0.47206 | 0.43657 | 0.46611 | b9 | 0.34052 | 0.49172 | 0.49543 | 0.48507 | 0.50139 | 0.48798 | 0.4836 | 0.47764 | 0.50658 | 0.50682 | 0.48593 | 0.47343 | 0.48685 | 0.4638 | 0.47532 | 0.48564 | 0.47826 | 0.51585 | 0.50124 | 0.45421 | 0.48391
b10 | 0.10254 | -0.041191 | -0.026477 | 0.033314 | -0.024092 | 0.0266 | 0.004864 | -0.0095011 | 0.0059732 | -0.0019099 | 0.0026733 | 0.047216 | -0.022187 | 0.062854 | -0.051663 | 0.022718 | 0.031704 | -0.019305 | -0.043664 | -0.093486 | 0 | b10 | 0.10266 | -0.043062 | -0.035477 | 0.033539 | 0.020529 | 0.046088 | -0.00091957 | 0.0081423 | 0.0045701 | 0.0085471 | -0.012166 | 0.026431 | -0.031863 | 0.024037 | -0.049549 | 0.0171 | 0.040362 | -0.037854 | -0.020474 | -0.098871 | -0.00033548
b11 | 0.56173 | 0.6535 | 0.64729 | 0.65708 | 0.65343 | 0.65523 | 0.65184 | 0.65995 | 0.66119 | 0.65323 | 0.66151 | 0.65552 | 0.65616 | 0.65814 | 0.66245 | 0.65231 | 0.65274 | 0.65074 | 0.64717 | 0.60259 | 0.65069 | b11 | 0.58517 | 0.65895 | 0.66058 | 0.66426 | 0.66619 | 0.66894 | 0.6698 | 0.66485 | 0.66754 | 0.65647 | 0.66634 | 0.65979 | 0.66345 | 0.66759 | 0.671 | 0.66383 | 0.66263 | 0.663 | 0.65057 | 0.60408 | 0.65969
b12 | 0.010924 | -0.0031806 | -0.0024639 | 0.0034136 | -0.003641 | 0.00023186 | 0.00037771 | -0.00045437 | 0.0012244 | -0.0010117 | -0.0011769 | 0.0025413 | -0.0035791 | 0.0037291 | -0.0098426 | 0.0006297 | 0.0032593 | -0.0044448 | -0.0079604 | -0.013654 | -0.0012691 | b12 | 0.011034 | -0.0061743 | -0.0045299 | 0.0058696 | 0.0017371 | 0.009311 | 0.0038727 | 0.0048627 | 0.0068717 | 0.001078 | -0.0024139 | 0.001524 | -0.0062613 | 0.0036762 | -0.0016564 | 0.00051518 | 0.0050739 | -0.0041924 | -0.00026177 | -0.0061394 | 0.0011026
b13 | 0.077535 | 0.068829 | 0.06899 | 0.068884 | 0.069172 | 0.065884 | 0.067414 | 0.065667 | 0.067369 | 0.067796 | 0.067767 | 0.067186 | 0.06906 | 0.06595 | 0.068138 | 0.069235 | 0.071233 | 0.073641 | 0.0721 | 0.055482 | 0.068432 | b13 | 0.079892 | 0.065986 | 0.065459 | 0.066154 | 0.064985 | 0.063425 | 0.0618 | 0.066818 | 0.060876 | 0.064132 | 0.065998 | 0.06773 | 0.067445 | 0.067113 | 0.067205 | 0.064764 | 0.06674 | 0.068556 | 0.06805 | 0.04985 | 0.065634
b14 | -0.002213 | -0.0032169 | -0.0024608 | -9.5792e-05 | -0.0037339 | -0.0012222 | -3.8982e-05 | -0.0014578 | -0.00099054 | -0.0021045 | -0.0024859 | -7.4831e-05 | -0.0034344 | -0.00088086 | -0.0041594 | -0.0021303 | -0.0006686 | -0.003363 | -0.0038724 | -0.0049422 | -0.002116 | b14 | -0.0015886 | -0.0028788 | -0.0012211 | 0.0010149 | -0.00089894 | 0.001128 | -0.0003856 | -0.00022799 | 0.00062901 | -0.00040017 | -0.00084658 | 0.00041289 | -0.0019954 | 0.00073644 | -0.0026029 | -0.00060906 | -7.9286e-06 | -0.0033017 | -0.00099326 | -0.0022466 | -0.00076352
b15 | -0.011652 | 0.025207 | 0.021928 | 0.017971 | 0.017983 | 0.019543 | 0.017894 | 0.019621 | 0.020076 | 0.018897 | 0.021327 | 0.019081 | 0.019248 | 0.020108 | 0.016922 | 0.019222 | 0.018653 | 0.019782 | 0.017643 | 0.011939 | 0.018379 | b15 | -0.012783 | 0.021324 | 0.020284 | 0.016901 | 0.019832 | 0.019122 | 0.020079 | 0.018921 | 0.020313 | 0.017907 | 0.01867 | 0.017991 | 0.017657 | 0.018512 | 0.017959 | 0.018698 | 0.018309 | 0.017668 | 0.017849 | 0.0065677 | 0.017542
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 46.5007 | 3.5995 | 7.1975 | 8.9856 | 4.0201 | 7.347 | 5.9896 | 4.6114 | -0.7979 | 1.1271 | -1.0943 | -5.7716 | -5.1499 | -7.3752 | -7.5328 | -9.5461 | -7.6457 | -5.4817 | -6.8062 | -19.3562 | 0 | a1 | 54.5007 | 3.6308 | 7.1559 | 13.4029 | 9.3818 | 8.176 | 6.9938 | 4.0529 | 4.4398 | 1.3427 | -0.43608 | -5.274 | -7.7367 | -2.4917 | -6.4062 | -9.9175 | -15.0029 | -13.3474 | -14.243 | -20.6701 | 0
a2 | -3.2344 | -0.71641 | -0.94025 | -0.16779 | -0.35037 | -0.41358 | -1.083 | -0.17223 | -0.65703 | -0.93791 | -0.46434 | -0.94223 | -0.43327 | -0.63883 | -0.53223 | -0.48766 | -0.46894 | -0.70923 | -0.3053 | 14.3891 | -0.19957 | a2 | 0.4051 | -0.068983 | -0.21402 | 0.39789 | 1.1027 | 0.57 | 0.35901 | 1.0894 | 0.45193 | 0.37037 | 0.56829 | 1.18 | 1.0105 | 1.8139 | 1.372 | 2.0657 | 2.0108 | 0.97112 | 0.19466 | 6.6561 | 1.0133
a3 | -1.8338 | 1.1935 | 1.2556 | 1.115 | 0.61598 | 0.56641 | 0.49029 | 0.62832 | 0.93915 | 0.90571 | 0.79651 | 0.60467 | 0.4983 | 0.30973 | 0.54849 | 0.13762 | 0.2449 | 0.5977 | 0.81356 | -0.44383 | 0.57452 | a3 | -2.0997 | 1.5565 | 1.1453 | 0.66359 | 0.77772 | 0.61058 | 0.33905 | 0.5824 | 0.72327 | 0.67763 | 0.64071 | 0.39579 | 0.37769 | 0.92411 | 0.64018 | 0.63256 | 0.78113 | 0.69358 | 0.92603 | -0.013813 | 0.6272
a4 | -0.12177 | 0.16174 | -0.00063252 | -0.051878 | -0.021398 | 0.12668 | -0.12914 | -0.017677 | 0.36552 | 0.54199 | 0.25091 | 0.17167 | 0.056712 | 0.13852 | 0.085414 | 0.27593 | 0.0062051 | 0.1481 | -0.047646 | 1.0588 | 0.13646 | a4 | -0.025094 | 0.34176 | -0.0091064 | -0.073491 | 0.16321 | -0.2368 | -0.39637 | 0.022112 | -0.21141 | -0.14784 | -0.12424 | -0.56029 | -0.25092 | 0.19452 | -0.080935 | -0.50427 | -0.11587 | 0.10395 | 0.059071 | 1.1728 | -0.060351
a5 | 1.8751 | 0.10424 | 0.15412 | 0.4032 | 0.37274 | 0.34881 | 0.54593 | 0.36976 | 0.30692 | 0.31042 | 0.32933 | 0.28323 | 0.20973 | 0.39675 | 0.23596 | 0.18391 | 0.077535 | 0.22204 | 0.34993 | 0.11952 | 0.32963 | a5 | 2.0223 | 0.23601 | 0.054895 | 0.072078 | 0.065871 | 0.09986 | -0.017326 | 0.075205 | 0.03441 | 0.12026 | 0.15806 | 0.0025911 | 0.020923 | 0.25465 | 0.21478 | 0.20956 | 0.12905 | 0.11356 | 0.21777 | -0.063356 | 0.16112
a6 | -0.012925 | -0.032392 | -0.065583 | -0.05539 | 0.02436 | -0.044928 | -0.0039747 | -0.086166 | -0.095489 | 0.016892 | -0.074334 | -0.073096 | -0.11699 | -0.041344 | -0.04576 | -0.0080835 | 0.028122 | -0.019596 | -0.034202 | -0.082738 | -0.040939 | a6 | 0.056897 | 0.03588 | 0.0081171 | 0.099877 | 0.11778 | 0.059936 | 0.03653 | 0.070602 | 0.071678 | -0.020962 | -0.012705 | -0.076786 | 0.055022 | 0.13005 | 0.1437 | 0.045345 | 0.025781 | 0.052782 | 0.09544 | 0.095752 | 0.053584
a7 | 1.3683 | -0.073428 | 0.034331 | 0.023632 | 0.0097696 | -0.051916 | -0.084241 | -0.0085557 | -0.027168 | -0.041022 | 0.048135 | -0.015087 | 0.060878 | 0.0053266 | 0.033624 | -0.037443 | -0.014983 | 0.032875 | -0.0015881 | 0.027861 | 0.034643 | a7 | 1.3688 | -0.088449 | -0.045997 | -0.076034 | -0.075064 | -0.051653 | -0.083842 | 0.00023089 | 0.034339 | -0.072919 | -0.011345 | -0.053255 | -0.053309 | -0.0044957 | -0.090432 | -0.0031173 | -0.00057821 | -0.00094069 | -0.032298 | 0.020321 | 0.00084166
a8 | -0.030357 | 0.0059248 | -0.016452 | 0.0013733 | -0.019123 | -0.011432 | -0.0015836 | -0.006486 | -0.0059075 | -0.02658 | -0.039264 | -0.03256 | -0.054453 | -0.01293 | -0.012803 | -0.006718 | -0.0011944 | -0.047509 | 0.021069 | -0.015451 | -0.015281 | a8 | -0.070884 | -0.0059599 | -0.037104 | -0.0019742 | -0.039134 | -0.044189 | -0.046379 | -0.02561 | -0.045441 | -0.022338 | -0.066002 | -0.055319 | -0.013667 | -0.027143 | -0.045714 | -0.018023 | -0.039129 | -0.052138 | -0.013064 | -0.011666 | -0.033608
a9 | -0.2808 | 0.018501 | 0.022755 | 0.020161 | -0.014909 | -0.012311 | 0.012541 | 0.0216 | 0.01473 | -0.0088948 | -0.00051485 | -0.0068765 | 0.012847 | 0.0079706 | -0.0065881 | -0.02182 | -0.0049768 | 0.014198 | -0.00080112 | -0.069015 | -0.0066498 | a9 | -0.28926 | 0.03524 | 0.017929 | 0.0047521 | -0.013673 | 0.011205 | -0.0096328 | -0.00632 | -0.0023212 | -0.020832 | -0.024698 | -0.050437 | -0.029986 | -0.038657 | -0.027647 | -0.023039 | -0.0040392 | -0.0062531 | -0.005382 | -0.10334 | -0.021206
a10 | 0.017139 | -0.010111 | -0.011675 | 0.0029064 | -0.0018316 | -0.0028219 | 0.0199 | 0.0034996 | 0.013027 | 0.020701 | -0.034996 | -0.0014139 | -0.017087 | -0.0076814 | 0.020406 | 0.019611 | 0.021709 | -0.014242 | 0.037395 | -0.11454 | 0 | a10 | -0.0086307 | 0.0058913 | -0.0025281 | 0.020951 | -0.00067201 | 0.0060024 | 0.00090346 | 0.0078115 | -0.017507 | 0.007881 | -0.025712 | -0.014115 | 0.011396 | -0.0053021 | 0.034283 | -0.0089792 | -0.001015 | -0.0093785 | 0.024163 | -0.049543 | 2.6446e-05
a11 | 0.13642 | -0.056547 | -0.037033 | -0.024221 | -0.032169 | -0.027908 | -0.022117 | -0.02552 | -0.033554 | -0.03533 | -0.028242 | -0.033868 | -0.02886 | -0.03332 | -0.034493 | -0.03241 | -0.036327 | -0.028657 | -0.032333 | -0.028086 | -0.027394 | a11 | 0.15854 | -0.055277 | -0.047493 | -0.034517 | -0.040256 | -0.040323 | -0.046968 | -0.036815 | -0.036252 | -0.049886 | -0.040563 | -0.051009 | -0.045301 | -0.046258 | -0.054645 | -0.039586 | -0.047829 | -0.044378 | -0.046863 | -0.03534 | -0.038849
a12 | -0.0094597 | -0.0024604 | -0.007765 | -0.0065653 | -0.0062359 | -0.0072227 | -0.0028226 | -0.0057151 | -0.0083893 | -0.0032602 | -0.011937 | -0.0069276 | -0.0087089 | -0.0053604 | -0.0016263 | 0.00055221 | 0.0013734 | -0.003428 | 0.0023536 | -0.0093165 | -0.004954 | a12 | -0.0077956 | -0.00098289 | -0.0021014 | 0.0044967 | 0.0031423 | 0.00092209 | -0.0026628 | 0.0004697 | -0.0020945 | -0.0015972 | -0.0036274 | -0.0018663 | 0.0014583 | -0.00027846 | 0.0073136 | 0.0040129 | 0.0030587 | -0.00086678 | 0.0039741 | -0.0058969 | 0.00027484
a13 | -0.024172 | 0.0018496 | 0.0023574 | 0.00040754 | -0.00082778 | -0.002621 | -7.713e-05 | -2.7835e-05 | -0.00078924 | -0.00052338 | 0.0019519 | -0.00096247 | -7.389e-05 | -0.0010553 | 0.00036203 | -0.0026576 | -0.0024553 | 6.4284e-05 | 0.00068103 | -0.0013631 | -0.00096714 | a13 | -0.025226 | 0.0030294 | 0.0013544 | -0.00025975 | 0.00092182 | 0.00010245 | 0.00025801 | -0.0019602 | -0.00029864 | -0.0024658 | -0.0021755 | -0.0062194 | -0.0036622 | -0.0037819 | -0.0043377 | -0.002768 | -0.002343 | -0.0007591 | -0.00035826 | -0.0041928 | -0.0021632
a14 | 0.0085194 | -0.0075067 | -0.0085321 | -0.0053979 | -0.0070084 | -0.0076674 | -0.0067646 | -0.0071915 | -0.00557 | -0.0066879 | -0.0085684 | -0.0072893 | -0.0074561 | -0.006781 | -0.0057513 | -0.0045643 | -0.0056193 | -0.0075396 | -0.0034842 | -0.00062454 | -0.0060133 | a14 | 0.0088477 | -0.004952 | -0.0064833 | -0.0045943 | -0.0058618 | -0.0082629 | -0.0079717 | -0.0059881 | -0.0096045 | -0.0053341 | -0.0073134 | -0.0072961 | -0.0065512 | -0.0063385 | -0.0067713 | -0.0073948 | -0.0052316 | -0.0053951 | -0.0044617 | 0.0018912 | -0.0057617
a15 | 0.0058449 | -0.006534 | -0.0043186 | 0.00071483 | -0.0043947 | -0.0023315 | -0.0027519 | -0.0024892 | -0.0012977 | -0.0032415 | -0.0041795 | -0.0017625 | -0.0043938 | 0.00055919 | -0.0061184 | -0.0028449 | -0.0024302 | -0.0056524 | -0.0037267 | -0.0091881 | -0.0031006 | a15 | 0.010535 | -0.0017903 | -0.0019069 | 0.000801 | 0.00026412 | 0.0019641 | 0.00010773 | 0.00079446 | 0.00063081 | -0.00064944 | -0.0017348 | -0.0015201 | -0.0020753 | -0.00075943 | -0.0024346 | -0.00034348 | -0.00018609 | -0.0023849 | -0.00017891 | -0.0024575 | -0.00038474
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.30507 | -0.1041 | -0.067546 | 0.08581 | -0.062295 | 0.069197 | 0.010913 | -0.024787 | 0.013626 | 0 | 0.010684 | 0.12232 | -0.055408 | 0.16295 | -0.13502 | 0.056527 | 0.079178 | -0.048696 | -0.11932 | -0.24737 | -0.12655 | Dx (mm) | 0.27152 | -0.11161 | -0.090375 | 0.082814 | 0.052299 | 0.11583 | 0 | 0.019653 | 0.014012 | 0.020465 | -0.026937 | 0.070491 | -0.083265 | 0.061853 | -0.13168 | 0.045006 | 0.10322 | -0.095036 | -0.057738 | -0.26903 | 0
Dy (mm) | -0.022654 | -0.035326 | -0.034543 | 0.01066 | 0 | 0 | 0.052275 | 0 | 0.03419 | 0.053455 | -0.089439 | 0 | -0.046716 | -0.01166 | 0.045364 | 0.053917 | 0.060957 | -0.039358 | 0.092336 | -0.33545 | -0.46097 | Dy (mm) | -0.098983 | 0 | -0.013015 | 0.057949 | 0 | 0.022259 | 0 | 0.021064 | -0.043827 | 0.021964 | -0.067236 | -0.030922 | 0.023528 | 0 | 0.076163 | -0.020307 | 0 | -0.030418 | 0.058978 | -0.1554 | -0.30534
Mag. len. (m) | 14.3569 | [20 empty] | Mag. len. (m) | 14.3546
TF (mT/KA) | 704.298 | [20 empty] | TF (mT/KA) | 704.263
Coil pos. (mm) | 0.017216 | [20 empty] | Coil pos. (mm) | 0.019032
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.046007 | [20 empty] | Twist (m2 rad) | 0.062099

## Harmonics
Normal:
Skew:
*10

## Work sheet
| HCMBBLA001-3000076 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 2.2266 | 7.348 | 0.82158 | 0.318 | -0.19957 | 3.4393 | -0.58196 | 0.26989 | -2.7182 | 7.0044 | -1.4099 | 0.25462 | 1.0133 | 1.4545 | 0.84575 | 0.67042 |  | 0
3 | 4.7708 | 7.7365 | 3.6487 | 0.64952 | 0.57452 | 0.67282 | 0.68093 | 0.31448 | 4.9609 | 8.0332 | 3.8245 | 0.63216 | 0.6272 | 0.69956 | 0.7284 | 0.2872 |  | 0
4 | 0.063595 | 0.31542 | 0.0053923 | 0.095637 | 0.13646 | 0.27132 | 0.11405 | 0.16631 | -0.28511 | 0.35377 | -0.222 | 0.10364 | -0.060351 | 0.3628 | -0.10116 | 0.23832 |  | 0
5 | -0.14133 | 1.3549 | 0.10066 | 0.19667 | 0.32963 | 0.37552 | 0.28875 | 0.11817 | 0.10913 | 1.3933 | 0.34713 | 0.20627 | 0.16112 | 0.43785 | 0.11482 | 0.084769 |  | 0
6 | 0.051578 | 0.11035 | 0.042826 | 0.074893 | -0.040939 | 0.040871 | -0.043806 | 0.036187 | -0.0073988 | 0.11697 | -0.014044 | 0.057377 | 0.053584 | 0.053742 | 0.050432 | 0.049218 |  | 0
7 | 1.1697 | 0.31297 | 1.1345 | 0.09562 | 0.034643 | 0.30946 | -0.0056716 | 0.04092 | 1.0995 | 0.36941 | 1.059 | 0.081007 | 0.00084166 | 0.31653 | -0.039404 | 0.038146 |  | 0
8 | -0.018829 | 0.02601 | -0.018045 | 0.018503 | -0.015281 | 0.018517 | -0.016567 | 0.014874 | -0.013643 | 0.025155 | -0.013344 | 0.019989 | -0.033608 | 0.019771 | -0.034051 | 0.015336 |  | 0
9 | 0.46611 | 0.039657 | 0.47123 | 0.0076244 | -0.0066498 | 0.066179 | 0.0038016 | 0.013812 | 0.48391 | 0.036031 | 0.48813 | 0.013622 | -0.021206 | 0.067495 | -0.010578 | 0.020302 |  | 0
10 | 0 | 0.044337 | -0.00016993 | 0.00071334 | 0 | 0.031664 | -2.1711e-05 | 0.0011569 | -0.00033548 | 0.043484 | -4.6574e-05 | 0.00023386 | 2.6446e-05 | 0.018345 | 0.00010967 | 0.0011557 |  | 0
11 | 0.65069 | 0.023757 | 0.65485 | 0.0045155 | -0.027394 | 0.038345 | -0.032336 | 0.0074322 | 0.65969 | 0.021996 | 0.66352 | 0.0050643 | -0.038849 | 0.045742 | -0.044696 | 0.0059565 |  | 0
12 | -0.0012691 | 0.0053505 | -0.001201 | 0.0015916 | -0.004954 | 0.0038796 | -0.0050292 | 0.0025747 | 0.0011026 | 0.0051719 | 0.0011052 | 0.0025215 | 0.00027484 | 0.0037022 | 0.00052119 | 0.0020722 |  | 0
13 | 0.068432 | 0.0041431 | 0.068498 | 0.0020568 | -0.00096714 | 0.0055178 | -0.00023934 | 0.0012275 | 0.065634 | 0.0052714 | 0.065666 | 0.002028 | -0.0021632 | 0.0057478 | -0.0014151 | 0.0022525 |  | 0
14 | -0.002116 | 0.0014733 | -0.0020096 | 0.00085009 | -0.0060133 | 0.0037944 | -0.0068314 | 0.00081497 | -0.00076352 | 0.0013059 | -0.00066733 | 0.00063975 | -0.0057617 | 0.0040199 | -0.0065448 | 0.0012273 |  | 0
15 | 0.018379 | 0.0073068 | 0.019506 | 0.0018991 | -0.0031006 | 0.0031236 | -0.0031774 | 0.0020044 | 0.017542 | 0.0075136 | 0.018778 | 0.0011703 | -0.00038474 | 0.0028223 | -0.00063345 | 0.0012754 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -1.2277 | -0.88391
3 | [12 empty] | 3 | 9 | -0.79417 | -0.8531
4 | [12 empty] | 4 | 16 | -2.0706 | -1.8505
5 | [12 empty] | 5 | 25 | -1.8809 | -2.0232
6 | [12 empty] | 6 | 36 | -2.9554 | -2.9348
7 | [12 empty] | 7 | 49 | -2.7718 | -2.8898
8 | [12 empty] | 8 | 64 | -4.099 | -4.0451
9 | [12 empty] | 9 | 81 | -4.5793 | -4.0966
10 | [12 empty] | sigma (mm) |  | 0.017216 | 0.019032
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05
17 | 0.00044619 | 0.00011155 | 2.6771e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 24.9705 | 0.92797 | 0.80825 | 0.82972 | 1.2458 | 1.1251 | 0.23613 | 1.3834 | 0.50656 | 0.69702 | 0.48064 | 1.0931 | 1.2527 | 0.47004 | 0.74191 | 0.66134 | 0.53002 | 1.0293 | 0.76939 | 23.1249 | 2.2266
3 | 38.0437 | 3.2846 | 3.9683 | 3.7893 | 3.9279 | 3.3767 | 4.0039 | 3.6068 | 4.1324 | 3.6575 | 3.0289 | 4.3593 | 4.3791 | 1.9316 | 2.4131 | 4.0936 | 3.9738 | 3.6518 | 4.0986 | 7.9105 | 4.7708
4 | 1.3846 | -0.077897 | 0.047154 | 0.10096 | 0.058347 | 0.13631 | -0.287 | -0.053019 | -0.03806 | 0.040678 | -0.049853 | 0.093339 | 0.022518 | -0.016297 | 0.05686 | 0.065608 | 0.022896 | 0.046001 | -0.071487 | 0.79196 | 0.063595
5 | -5.3738 | -0.12409 | -0.18048 | 0.098876 | 0.10514 | 0.16463 | 0.20492 | 0.41564 | 0.36864 | 0.21339 | 0.38983 | 0.22949 | 0.081281 | 0.16364 | 0.20077 | -0.094533 | -0.14335 | -0.15403 | -0.1278 | -2.7244 | -0.14133
6 | 0.090374 | 0.042638 | -0.05378 | -0.024707 | 0.0093695 | 0.072535 | 0.23327 | 0.065455 | 0.17451 | 0.056007 | 0.082643 | -0.052796 | -0.0010992 | 0.088546 | 0.03143 | 0.010571 | 0.065936 | -0.04692 | 0.01725 | 0.071976 | 0.051578
7 | 2.4649 | 0.81271 | 0.97242 | 1.1698 | 1.2024 | 1.1902 | 1.1812 | 1.1879 | 1.1254 | 1.1852 | 1.1833 | 1.1461 | 1.1269 | 1.1743 | 1.1587 | 1.1426 | 1.1576 | 1.1906 | 1.1129 | 1.0179 | 1.1697
8 | -0.040676 | -0.019591 | -0.0046206 | -0.016066 | -0.0031365 | -0.020181 | -0.048519 | -0.013263 | -0.062693 | -0.03344 | 0.0038683 | 0.0042044 | 0.010947 | -0.019387 | -0.010332 | -0.011166 | -0.03185 | -0.027867 | -0.021723 | 0.00052097 | -0.018829
9 | 0.28235 | 0.47203 | 0.46504 | 0.4683 | 0.48219 | 0.471 | 0.48529 | 0.46025 | 0.47194 | 0.48368 | 0.46529 | 0.47009 | 0.47144 | 0.46279 | 0.45821 | 0.47626 | 0.46948 | 0.47764 | 0.47113 | 0.44467 | 0.46611
10 | -7.741e-05 | 1.6539e-07 | -5.4735e-06 | -5.1417e-06 | -0.00014768 | -7.0494e-05 | -6.9092e-07 | 0.00012137 | -1.3965e-06 | -0.0030208 | 1.7426e-06 | 4.9274e-05 | -7.3865e-06 | -2.3346e-06 | 2.9075e-05 | -3.1663e-07 | -1.9021e-07 | -1.4369e-06 | 3.0471e-06 | -0.00025954 | 0
11 | 0.55971 | 0.65334 | 0.64735 | 0.65685 | 0.65328 | 0.65522 | 0.65174 | 0.65995 | 0.66099 | 0.65311 | 0.66221 | 0.65532 | 0.6563 | 0.65779 | 0.66154 | 0.65231 | 0.65262 | 0.65069 | 0.64669 | 0.60242 | 0.65069
12 | -0.053864 | 0.018309 | 0.0076806 | -0.0075576 | -0.005993 | -0.029862 | -0.0014446 | 0.0069458 | 0.0055736 | -0.010314 | -0.018113 | -0.032598 | -0.0087557 | -0.03848 | -0.033371 | -0.02234 | -0.0082758 | -0.019152 | -0.018434 | -0.036436 | -0.012691 | ten times!
13 | 0.77904 | 0.68776 | 0.69088 | 0.68846 | 0.68995 | 0.65948 | 0.67144 | 0.65639 | 0.67234 | 0.67523 | 0.68373 | 0.67193 | 0.69181 | 0.6612 | 0.67509 | 0.69139 | 0.71012 | 0.73743 | 0.71501 | 0.54707 | 0.68432 | ten times!
14 | 0.0060532 | -0.0086582 | -0.011182 | -0.013594 | -0.028113 | -0.023359 | -0.0031827 | -0.010573 | -0.012524 | -0.022472 | -0.023657 | -0.019969 | -0.02387 | -0.035846 | -0.025064 | -0.031515 | -0.020069 | -0.023865 | -0.024221 | 0.00028384 | -0.02116 | ten times!
15 | -0.11652 | 0.25207 | 0.21928 | 0.17971 | 0.17983 | 0.19543 | 0.17894 | 0.19621 | 0.20076 | 0.18897 | 0.21327 | 0.19081 | 0.19248 | 0.20108 | 0.16922 | 0.19222 | 0.18653 | 0.19782 | 0.17643 | 0.11939 | 0.18379 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -3.0671 | -0.68813 | -0.91415 | -0.1838 | -0.34586 | -0.41819 | -1.1083 | -0.1704 | -0.67515 | -0.96088 | -0.43343 | -0.95094 | -0.40595 | -0.64212 | -0.53639 | -0.51453 | -0.49972 | -0.6889 | -0.33843 | 14.6892 | -0.19957
3 | -1.8232 | 1.196 | 1.2559 | 1.1156 | 0.61575 | 0.56486 | 0.49324 | 0.62824 | 0.93851 | 0.90536 | 0.79538 | 0.60096 | 0.49905 | 0.30573 | 0.55014 | 0.13428 | 0.24461 | 0.59932 | 0.81379 | -0.34076 | 0.57452
4 | -0.28483 | 0.16325 | 0.00035699 | -0.060268 | -0.015934 | 0.121 | -0.13307 | -0.015521 | 0.36156 | 0.53931 | 0.25825 | 0.16352 | 0.060326 | 0.12378 | 0.09075 | 0.27467 | 0.006791 | 0.14923 | -0.035088 | 0.85018 | 0.13646
5 | 1.8787 | 0.10337 | 0.15204 | 0.40456 | 0.37318 | 0.34973 | 0.54226 | 0.36914 | 0.30549 | 0.30952 | 0.33189 | 0.28586 | 0.20752 | 0.39927 | 0.23446 | 0.1835 | 0.075112 | 0.22098 | 0.34953 | 0.11212 | 0.32963
6 | -0.14055 | -0.024934 | -0.052906 | -0.060507 | 0.024574 | -0.04366 | -0.025441 | -0.086241 | -0.10894 | -0.0054795 | -0.037206 | -0.072445 | -0.09724 | -0.036813 | -0.06275 | -0.029083 | 0.0036456 | -0.0024917 | -0.070587 | 0.040697 | -0.040939
7 | 1.3721 | -0.073792 | 0.033603 | 0.023571 | 0.0092791 | -0.051591 | -0.083237 | -0.0086219 | -0.026291 | -0.040291 | 0.048535 | -0.013447 | 0.059615 | 0.0062736 | 0.03365 | -0.037333 | -0.014588 | 0.031297 | 0.0012794 | 0.017834 | 0.034643
8 | 0.013145 | 0.014713 | -0.0081574 | -0.0017975 | -0.01956 | -0.011032 | -0.013573 | -0.0062341 | -0.013592 | -0.038733 | -0.019747 | -0.032164 | -0.043748 | -0.010972 | -0.023072 | -0.018225 | -0.014494 | -0.038333 | 0.00051178 | 0.045432 | -0.015281
9 | -0.28234 | 0.017174 | 0.021853 | 0.019841 | -0.014969 | -0.012208 | 0.012292 | 0.021646 | 0.014528 | -0.0088408 | -0.00019032 | -0.006785 | 0.011797 | 0.0090213 | -0.0038887 | -0.023055 | -0.00691 | 0.013429 | 0.0036955 | -0.10062 | -0.0066498
10 | 0.00014302 | 5.4832e-06 | 5.8246e-06 | 8.6034e-06 | -0.0030104 | -0.0016859 | -1.8622e-06 | 0.0031275 | -1.8436e-06 | 0.00016106 | -1.629e-05 | 0.001023 | 3.9799e-06 | 2.6721e-05 | -1.0681e-05 | 3.9543e-07 | -3.6863e-06 | 3.0578e-06 | -2.5714e-05 | 0.00027662 | 0
11 | 0.13845 | -0.056785 | -0.037427 | -0.02388 | -0.03242 | -0.027584 | -0.02211 | -0.025612 | -0.033507 | -0.035295 | -0.028227 | -0.033319 | -0.029281 | -0.032727 | -0.034346 | -0.032453 | -0.036526 | -0.028878 | -0.031676 | -0.032541 | -0.027394 | ten times!
12 | -0.030146 | -0.0060822 | -0.059704 | -0.071084 | -0.062724 | -0.070946 | -0.053095 | -0.057156 | -0.10008 | -0.058184 | -0.076731 | -0.068445 | -0.064345 | -0.046962 | -0.037737 | -0.019768 | -0.015544 | -0.013799 | -0.022884 | 0.03583 | -0.04954 | ten times!
13 | -0.26197 | 0.011651 | 0.018517 | 0.0076253 | -0.011616 | -0.022153 | -0.00019118 | -0.0016415 | -0.0070531 | -0.0043736 | 0.018518 | -0.0028063 | -0.005125 | -0.0021821 | -0.0008749 | -0.023725 | -0.020839 | -0.0031769 | 0.0063655 | -0.027491 | -0.0096714 | ten times!
14 | 0.068336 | -0.073335 | -0.081486 | -0.056061 | -0.072338 | -0.075345 | -0.075102 | -0.072423 | -0.061207 | -0.075198 | -0.069608 | -0.071118 | -0.069161 | -0.066629 | -0.070638 | -0.052853 | -0.063972 | -0.071251 | -0.05192 | 0.00802 | -0.060133 | ten times!
15 | 0.058449 | -0.06534 | -0.043186 | 0.0071483 | -0.043947 | -0.023315 | -0.027519 | -0.024892 | -0.012977 | -0.032415 | -0.041795 | -0.017625 | -0.043938 | 0.0055919 | -0.061184 | -0.028449 | -0.024302 | -0.056524 | -0.037267 | -0.091881 | -0.031006 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -27.1229 | -1.6063 | -1.7663 | -1.2629 | -1.4755 | -1.3864 | -1.5105 | -1.4904 | -1.4374 | -1.0669 | -1.0652 | -1.3046 | -1.1106 | -0.89321 | -1.5702 | -1.4543 | -1.6546 | -1.8544 | -1.4691 | -21.971 | -2.7182
3 | 39.6018 | 3.4497 | 3.9435 | 3.7206 | 4.3023 | 3.4837 | 3.3395 | 4.3136 | 4.348 | 4.8685 | 3.4008 | 3.9614 | 3.9874 | 2.4244 | 3.0277 | 3.3463 | 3.7848 | 4.1913 | 4.9474 | 9.0407 | 4.9609
4 | -1.0876 | -0.27131 | -0.18962 | -0.2365 | -0.19196 | -0.43904 | -0.35581 | -0.25984 | -0.37807 | -0.13278 | -0.086128 | -0.25993 | -0.013688 | -0.14396 | -0.26981 | -0.13975 | -0.19458 | -0.21462 | -0.2187 | -1.4161 | -0.28511
5 | -5.382 | 0.14847 | 0.0076862 | 0.32634 | 0.43131 | 0.56632 | 0.60926 | 0.48105 | 0.62509 | 0.51531 | 0.63884 | 0.42289 | 0.43035 | 0.17545 | 0.31268 | 0.25809 | 0.084142 | 0.12378 | 0.091253 | -2.3656 | 0.10913
6 | 0.10136 | -0.025901 | -0.034426 | -0.016037 | -0.039952 | -0.11146 | -0.050054 | -0.085628 | -0.0027799 | 0.074297 | 0.034413 | 0.035058 | 0.092445 | -0.12589 | -0.029112 | 0.011938 | 0.013821 | 0.0011703 | 0.0053085 | -0.10875 | -0.0073988
7 | 2.6661 | 0.86336 | 0.86089 | 1.0941 | 1.0172 | 1.0821 | 1.0696 | 1.0943 | 1.0738 | 1.0972 | 1.171 | 1.1357 | 1.0826 | 1.0463 | 1.0164 | 1.0739 | 1.0575 | 1.1275 | 1.0993 | 0.93404 | 1.0995
8 | -0.012959 | -0.010877 | -0.019261 | -0.015651 | -0.03449 | -0.0019028 | 0.00066096 | 0.0054791 | 0.010158 | -0.022497 | -0.037012 | -0.049988 | -0.032227 | -0.048712 | -0.0030665 | 0.0025859 | 0.00063787 | 0.0061738 | 0.0097907 | 0.023963 | -0.013643
9 | 0.32621 | 0.48918 | 0.49375 | 0.48424 | 0.50082 | 0.48522 | 0.4836 | 0.47764 | 0.50695 | 0.50682 | 0.48667 | 0.47267 | 0.48558 | 0.46301 | 0.47325 | 0.48533 | 0.47605 | 0.5141 | 0.50137 | 0.44421 | 0.48391
10 | -3.5178e-05 | 0.00020039 | 3.9472e-06 | 3.5222e-06 | 3.393e-05 | -1.8062e-05 | -0.00091957 | 7.311e-08 | 2.5329e-06 | -3.1796e-07 | -3.2065e-06 | 7.1381e-07 | 5.2683e-06 | -0.00025272 | -2.3559e-05 | -1.4688e-06 | 0.00013043 | 3.6642e-06 | -3.9039e-06 | -4.3063e-05 | -0.00033548
11 | 0.58373 | 0.6585 | 0.66033 | 0.66412 | 0.66613 | 0.66826 | 0.6698 | 0.6648 | 0.66753 | 0.65643 | 0.66645 | 0.65975 | 0.66313 | 0.66744 | 0.67121 | 0.66377 | 0.6623 | 0.66276 | 0.65071 | 0.60361 | 0.65969
12 | -0.025157 | -0.0097562 | -0.0036648 | 0.019918 | -0.0066193 | 0.041269 | 0.038727 | 0.039067 | 0.062788 | 0.0011335 | -0.010557 | -0.017103 | -0.02358 | 0.0074601 | 0.04357 | -0.015027 | 0.0021134 | 0.0042293 | 0.024968 | 0.037871 | 0.011026 | ten times!
13 | 0.79552 | 0.6574 | 0.65439 | 0.65886 | 0.65021 | 0.63184 | 0.618 | 0.66725 | 0.61191 | 0.64049 | 0.66356 | 0.6788 | 0.672 | 0.67078 | 0.66548 | 0.649 | 0.66741 | 0.68442 | 0.67805 | 0.49163 | 0.65634 | ten times!
14 | 0.0041092 | -0.009188 | 0.0030901 | -0.00099525 | -0.017531 | -0.0065997 | -0.003856 | -0.0052043 | 0.0037184 | -0.0071371 | -0.0033636 | -0.0059279 | -0.008248 | -0.0020654 | -0.0080817 | -0.012964 | -0.015642 | -0.018592 | -0.0015323 | -0.0047696 | -0.0076352 | ten times!
15 | -0.12783 | 0.21324 | 0.20284 | 0.16901 | 0.19832 | 0.19122 | 0.20079 | 0.18921 | 0.20313 | 0.17907 | 0.1867 | 0.17991 | 0.17657 | 0.18512 | 0.17959 | 0.18698 | 0.18309 | 0.17668 | 0.17849 | 0.065677 | 0.17542 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | 0.93256 | -0.048545 | -0.1958 | 0.36608 | 1.0979 | 0.55258 | 0.35901 | 1.0774 | 0.47315 | 0.35615 | 0.59721 | 1.1911 | 1.0031 | 1.8072 | 1.3547 | 2.0704 | 2.0014 | 0.99389 | 0.16662 | 6.8226 | 1.0133
3 | -2.1226 | 1.5632 | 1.1448 | 0.66703 | 0.77622 | 0.61709 | 0.33905 | 0.58328 | 0.72089 | 0.67867 | 0.63908 | 0.40138 | 0.3741 | 0.92198 | 0.64211 | 0.63608 | 0.78324 | 0.69416 | 0.92895 | 0.0070378 | 0.6272
4 | -0.27885 | 0.34796 | -0.0079104 | -0.079328 | 0.1624 | -0.24247 | -0.39637 | 0.019385 | -0.20507 | -0.15108 | -0.11314 | -0.55726 | -0.25288 | 0.19082 | -0.079877 | -0.50526 | -0.119 | 0.10739 | 0.060783 | 1.083 | -0.060351
5 | 2.0296 | 0.23719 | 0.054873 | 0.069347 | 0.06406 | 0.098255 | -0.017326 | 0.07528 | 0.034154 | 0.11985 | 0.15841 | 0.0047532 | 0.021848 | 0.25229 | 0.222 | 0.20913 | 0.12827 | 0.1147 | 0.21968 | -0.064769 | 0.16112
6 | 0.018827 | 0.032396 | 0.01061 | 0.079718 | 0.11916 | 0.053534 | 0.03653 | 0.062463 | 0.088106 | -0.028942 | 0.014943 | -0.063086 | 0.04445 | 0.13015 | 0.11209 | 0.053092 | 0.025802 | 0.064854 | 0.071891 | 0.14912 | 0.053584
7 | 1.3775 | -0.088723 | -0.047593 | -0.076041 | -0.074221 | -0.049771 | -0.083842 | 0.00035175 | 0.034844 | -0.072573 | -0.013293 | -0.052095 | -0.053284 | -0.0038044 | -0.091922 | -0.0026775 | 0.0010848 | -0.0031937 | -0.032519 | 0.016399 | 0.00084166
8 | -0.018062 | -0.004109 | -0.033307 | -0.015387 | -0.038798 | -0.049911 | -0.046379 | -0.030286 | -0.034977 | -0.027376 | -0.050941 | -0.046757 | -0.020232 | -0.026018 | -0.064463 | -0.012894 | -0.038933 | -0.045034 | -0.027122 | 0.0084673 | -0.033608
9 | -0.28264 | 0.035588 | 0.017563 | 0.0028046 | -0.013655 | 0.010294 | -0.0096328 | -0.0064921 | -0.0020853 | -0.021016 | -0.025498 | -0.049477 | -0.029087 | -0.038483 | -0.023259 | -0.022641 | -0.0039837 | -0.0073345 | -0.0040041 | -0.11853 | -0.021206
10 | 0.00011921 | 0.0022622 | 4.2559e-06 | -1.076e-05 | 0.00056643 | -8.8833e-06 | 0.00090346 | -7.1231e-07 | 1.0803e-06 | -1.5226e-07 | -5.1922e-07 | 1.3403e-06 | -4.9294e-06 | -0.0036191 | -1.1806e-05 | -1.3918e-06 | 0.001889 | 3.5269e-06 | 9.3251e-07 | 8.4937e-05 | 2.6446e-05
11 | 0.16062 | -0.055348 | -0.047654 | -0.034978 | -0.040362 | -0.040527 | -0.046968 | -0.036888 | -0.036038 | -0.049881 | -0.040731 | -0.050893 | -0.045127 | -0.046247 | -0.05394 | -0.039696 | -0.048033 | -0.044514 | -0.046705 | -0.036984 | -0.038849
12 | 0.026213 | -0.0074423 | -0.014136 | 0.018058 | 0.031083 | -0.00082819 | -0.026628 | -0.0049659 | -0.0020819 | -0.025559 | -0.0053646 | -0.00078438 | 0.0012291 | -0.0011334 | 0.032973 | 0.050292 | 0.032295 | 0.0055429 | 0.011265 | -0.012248 | 0.0027484 | ten times!
13 | -0.27183 | 0.026067 | 0.008942 | -0.00013781 | 0.011562 | 0.0081513 | 0.0025801 | -0.018666 | -0.0017464 | -0.023756 | -0.023696 | -0.058164 | -0.040435 | -0.034821 | -0.048679 | -0.025229 | -0.019301 | -0.01228 | -0.0051046 | -0.040707 | -0.021632 | ten times!
14 | 0.054501 | -0.051166 | -0.064078 | -0.054555 | -0.058731 | -0.088008 | -0.079717 | -0.063292 | -0.088786 | -0.05647 | -0.063182 | -0.067497 | -0.070356 | -0.062999 | -0.081618 | -0.070693 | -0.052158 | -0.051391 | -0.053372 | 0.021873 | -0.057617 | ten times!
15 | 0.10535 | -0.017903 | -0.019069 | 0.00801 | 0.0026412 | 0.019641 | 0.0010773 | 0.0079446 | 0.0063081 | -0.0064944 | -0.017348 | -0.015201 | -0.020753 | -0.0075943 | -0.024346 | -0.0034348 | -0.0018609 | -0.023849 | -0.0017891 | -0.024575 | -0.0038474 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -0.00081038 | -0.00012346 | -1.8955e-05 | -4.1815e-05 | 8.0861e-05 | 0.00011622 | 0.00017329 | 0.00016051 | 6.7514e-05 | 6.7514e-05 | 0.00025934 | 9.6195e-05 | 3.1166e-05 | 5.4025e-05 | 0.00011494 | 0.00014845 | 6.7372e-05 | -0.00044278 |  | 704.298
|  | -0.0013897 | -3.5672e-05 | 2.5243e-05 | 0.00013784 | 0.00026138 | 0.00026422 | 0.00014423 | 0.00015105 | 8.6584e-05 | 8.5874e-05 | 0.00022048 | 2.2971e-05 | 5.208e-05 | 9.0986e-05 | 9.4251e-05 | 0.00018427 | -3.4252e-05 | -0.00036183 |  | 704.263
Magnetic Length | -0.080914
mbh dB/B | -4333.7391
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.0297 | -4493.6977 | -8.1038 | -1.2346 | -0.18955 | -0.41815 | 0.80861 | 1.1622 | 1.7329 | 1.6051 | 0.67514 | 0.67514 | 2.5934 | 0.96195 | 0.31166 | 0.54025 | 1.1494 | 1.4845 | 0.67372 | -4.4278 | -4173.7805
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 4.7296 | 0.43952 | 0.79932 | 0.97814 | 0.48159 | 0.81427 | 0.67853 | 0.54072 | -0.00021378 | 0.19228 | -0.029853 | -0.49759 | -0.43541 | -0.65794 | -0.6737 | -0.87504 | -0.68499 | -0.4686 | -0.60104 | -1.856
db2 | 1.0641 | 28.4661 | 0.23352 | -0.020985 | 0.11711 | 0.085945 | -0.089562 | -0.13332 | 0.11198 | 0.057667 | 0.052053 | 0.025 | -0.020169 | 0.074402 | -0.049492 | 0.020605 | -0.18781 | -0.22493 | -0.091966 | 0.039942 | 23.1249
db3 | 4.6988 | 7.6199 | -0.014773 | 0.043366 | 0.022164 | -0.41237 | 0.036836 | -0.0053975 | -0.053295 | 0.032705 | 0.11674 | 0.026778 | 0.14454 | 0.073092 | 0.042162 | 0.011275 | 0.063746 | 0.077022 | -0.0629 | -0.14168 | 9.4586
db4 | 0.0084333 | 1.3846 | 0.021508 | -0.010645 | 0.01004 | 0.0048539 | -0.0027044 | -0.0096276 | 0.016896 | 0.023645 | 0.017958 | 0.01746 | -0.0089683 | -0.013672 | -0.014426 | 0.0038006 | -0.018986 | -0.028404 | -0.020197 | 0.011469 | 0.79196
db5 | 0.040204 | -0.95626 | 0.031759 | 0.0069556 | 0.0085095 | -0.38571 | 0.021344 | 0.027107 | 0.015622 | 0.021603 | 0.03062 | 0.023116 | 0.029603 | 0.033005 | 0.031477 | 0.019355 | 0.025905 | 0.024866 | 0.012699 | 0.022164 | 0.045529
db6 | -0.021568 | 0.090374 | -0.0014855 | 0.0030496 | 0.0078522 | 0.0022875 | 0.0033214 | -0.002187 | 0.0017738 | 0.0023987 | 0.00092304 | 0.0011436 | -0.0042355 | -0.004656 | -0.0012932 | -0.0028577 | -0.0022471 | -0.0015013 | -0.0011216 | -0.0011651 | -0.0030335
db7 | 0.005659 | 0.15726 | -0.014366 | -0.014194 | -0.0023595 | 0.061876 | -0.00019707 | 0.00067012 | -0.0070911 | -0.0037952 | -0.0017936 | 0.00095031 | -0.0041089 | -0.00048749 | -0.00023414 | 0.0022729 | -0.0059128 | -0.004221 | -0.0030179 | -0.00399 | 0.039637
db8 | -0.0033393 | -0.040676 | -0.0023469 | -0.0013106 | -0.0012183 | 0.027483 | 0.0010047 | -0.0031754 | 0.0011277 | -0.0010915 | -0.0037832 | -0.0011699 | -0.0035467 | -0.0038381 | -0.00029955 | -0.0023472 | -0.0037468 | 0.0015795 | -0.0022279 | -0.0010933 | -0.0022288
db9 | 0.02303 | -0.0084501 | -0.0023837 | 0.0028331 | 0.0015057 | -0.017591 | -0.0015395 | 0.0022052 | -0.0016379 | -0.00030125 | 0.0030056 | 0.0031663 | 0.0039493 | 0.0017338 | 0.0031286 | 0.0029247 | 0.0024296 | 0.0013878 | -0.0016955 | -0.003121 | 0.039915
db10 | -0.0002396 | -7.741e-05 | 0.00025822 | 0.00024102 | 0.00022543 | 0.0001013 | 0.00016199 | -0.0028886 | 0.00036792 | -0.00051944 | -0.0027817 | 0.00024524 | 0.0002967 | 0.00021163 | 0.00021623 | 0.00026984 | 0.00024386 | 0.0029368 | 0.00017628 | 0.00023726 | -0.00025954
db11 | 0.0068804 | 0.0070283 | 0.0022819 | 0.0012068 | 0.0019463 | -0.0084354 | -0.00038681 | -0.00029975 | -0.00029378 | 0.00063587 | -0.00060702 | 0.001349 | 0.0006484 | 0.0015721 | 0.00045428 | 0.0002363 | -0.00036414 | 0.0014561 | -0.0015032 | 0.000103 | 0.024252
db12 | 3.9582e-05 | -0.0053864 | 0.0010809 | -7.0301e-05 | -2.061e-05 | 0.00091275 | 0.00041394 | 1.9414e-05 | -0.00021001 | 0.00010722 | 0.0012528 | 0.00011924 | -0.00035765 | -0.0011342 | -0.0016393 | 0.00035105 | -9.6894e-05 | 0.0001885 | -0.00073746 | -0.00017927 | -0.0036436
db13 | 0.0036801 | 0.00097819 | -0.0003761 | -0.0001329 | -0.0002802 | 0.0041767 | -0.00046853 | -0.00043921 | -0.00095865 | 8.914e-05 | -6.3364e-05 | 0.00042199 | -0.00014103 | -0.00063814 | -0.00050213 | 0.00062285 | -7.3255e-05 | -0.00045031 | -0.00032687 | -0.00045997 | 0.0040747
db14 | -0.001391 | 0.00060532 | 0.00032025 | -0.00019916 | 0.00041744 | 0.00020698 | 0.00095171 | 0.00030446 | -0.00076448 | 0.00021286 | 3.8522e-05 | -0.00047323 | 7.428e-05 | -0.00070089 | -0.00014606 | -4.6747e-05 | 0.00020908 | 0.0006708 | -0.00035755 | -0.00071827 | 2.8384e-05
db15 | -0.015042 | -0.011652 | -3.7766e-05 | -1.4073e-05 | 0.0002438 | 0.00061656 | -0.0010787 | -0.00037161 | 0.0011763 | 0.00010965 | -0.00097115 | 0.00081737 | 0.00095678 | -0.0014956 | -0.0017724 | -0.0038286 | 0.00414 | 0.00062625 | 0.00016133 | 0.00072193 | -0.0035379
da2 | 0.070074 | -3.0671 | 0.35554 | 0.1465 | 0.06578 | -2.1704 | 0.056464 | 0.091734 | 0.057918 | 0.12184 | 0.19881 | 0.23157 | 0.161 | 0.08951 | 0.23328 | 0.13865 | -0.003076 | -0.007599 | 0.14861 | 0.083878 | 1.8444
da3 | -0.38244 | -1.8232 | 0.25244 | 0.11951 | 0.14789 | 0.16621 | -0.054453 | -0.036561 | 0.080255 | 0.11078 | 0.08304 | 0.070043 | -0.088115 | -0.21217 | -0.083007 | -0.0077454 | -0.19131 | -0.29836 | -0.13155 | 0.073101 | -0.37032
da4 | 0.10591 | -0.28483 | 0.071977 | 0.011817 | 0.070669 | -0.051781 | 0.011743 | -0.019612 | 0.0074436 | 0.012151 | 0.026032 | 0.0019517 | -0.010597 | -0.044776 | -0.0027295 | -0.027888 | -0.036652 | -0.042913 | -0.012133 | 0.035298 | 0.32973
da5 | -0.023732 | 1.8787 | 0.0055939 | -0.0094227 | -0.015261 | 0.18985 | -0.012797 | -0.0094518 | -0.010808 | -0.0070942 | -0.008887 | 0.001231 | -0.023042 | -0.024927 | -0.016749 | -0.0099201 | -0.016776 | -0.020747 | -0.0088762 | -0.0019141 | 0.035401
da6 | 0.017476 | -0.14055 | -0.0040133 | -0.011023 | 0.00083789 | 0.15668 | -0.00451 | -0.012303 | -0.0040625 | -0.0068766 | -0.012277 | -0.0092852 | -0.0078993 | -0.013219 | -0.011853 | -0.0097137 | -0.01312 | -0.017655 | -0.013751 | -0.0059575 | 0.028696
da7 | -0.012571 | 1.3721 | -0.0054219 | 0.0093222 | 0.0024992 | 0.094378 | -0.0001028 | -0.00239 | -0.013653 | -0.0088591 | -0.0078766 | 0.00041594 | -0.008703 | -0.011223 | -0.011019 | -0.0082009 | -0.012262 | -0.0052073 | -0.005104 | -0.0065932 | -0.040583
da8 | -0.0038529 | 0.013145 | 0.0026951 | 0.002336 | -1.8205e-06 | -0.059946 | -0.00035199 | 0.0058342 | 0.006741 | 0.0045907 | 0.0007903 | 0.0014093 | 0.0066542 | 0.0041179 | 0.0026252 | 0.0038085 | 0.0087477 | 0.0048508 | 0.0012408 | 0.0038581 | 0.00045292
da9 | -0.0050604 | -0.28234 | 0.001146 | 0.0040255 | 0.0046387 | 0.010311 | -0.00051136 | 0.0030353 | -0.0024014 | -0.006584 | -0.0039911 | 0.004168 | 0.00034776 | -0.0029293 | -0.0065097 | 0.0013975 | -0.0053851 | -0.00052915 | 0.0024929 | -0.002722 | -0.10062
da10 | -8.2657e-05 | 0.00014302 | 0.00034093 | 8.1574e-05 | 5.8387e-05 | -0.0029227 | -0.0016008 | 0.0025822 | 0.0032069 | 9.8463e-05 | 0.00024377 | 6.5091e-05 | 0.0010807 | 7.0459e-05 | 4.2691e-05 | 6.288e-05 | 0.00011213 | 8.128e-06 | -0.0035948 | 6.415e-05 | 0.00027662
da11 | -0.001488 | 0.13845 | 0.0088427 | 0.010511 | 0.011349 | 0.018727 | 0.0029882 | 0.0018303 | -0.0044837 | -0.0036464 | -0.00081285 | 0.0025376 | -0.0044532 | -0.0056474 | -0.0093369 | -0.0075154 | -0.0076042 | -0.0046089 | -0.0031384 | -0.0055387 | -0.032541
da12 | 0.0025153 | -0.0030146 | 0.00023845 | 0.00029468 | -0.00036193 | 0.0024641 | 0.00028487 | -0.00052236 | -0.00027963 | -0.00075199 | -0.00012239 | -0.000112 | 0.0005963 | -0.00016537 | 0.00028049 | 0.00082363 | -0.00057646 | -0.0012567 | -0.0010063 | 0.00017275 | 0.003583
da13 | -0.00023934 | -0.0046896 | 0.0012131 | 0.0020576 | 0.00092831 | 0.0011866 | 0.00017305 | 0.0012949 | -0.00084496 | -0.00054673 | -0.00094335 | 0.0018491 | -0.0010986 | -0.0010935 | -0.0016656 | -0.00018137 | -0.00049729 | -0.00030678 | -0.00097841 | -0.00054628 | -0.0018642
da14 | -0.0068314 | 0.0068336 | 0.0024282 | 0.00047365 | -0.00083597 | -0.0021319 | -0.0010847 | -0.00034314 | -0.00058497 | -0.00043635 | -0.00033038 | 0.00066487 | -0.00021745 | -0.0002249 | 0.00046629 | 0.00085258 | 0.000605 | 0.00051566 | -0.00026088 | 0.00044438 | 0.000802
da15 | -0.0025612 | 0.0058449 | 0.0020359 | 0.0022993 | -0.0014201 | -0.00084458 | 0.0009935 | 0.001066 | -0.0034147 | 0.0015132 | 0.0017952 | 0.00078597 | -0.0021101 | -0.0055855 | -0.0035379 | 0.00078787 | 0.0049573 | 0.0017841 | -0.0028488 | 0.0017433 | -0.0091881
Check Computation for Aperture2
Magnetic Length | -0.077853
mbh dB/B | -4462.5348
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 107.9725 | -4789.2749 | -13.8971 | -0.35672 | 0.25243 | 1.3784 | 2.6138 | 2.6422 | 1.4423 | 1.5105 | 0.86584 | 0.85874 | 2.2048 | 0.22971 | 0.5208 | 0.90986 | 0.94251 | 1.8427 | -0.34252 | -3.6183 | -4135.7946
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 5.5405 | 0.45352 | 0.80603 | 1.4307 | 1.0286 | 0.90804 | 0.78981 | 0.49573 | 0.53442 | 0.22471 | 0.04683 | -0.43696 | -0.68323 | -0.15874 | -0.55019 | -0.90131 | -1.4098 | -1.2443 | -1.3339 | -1.9766
db2 | 1.231 | 24.8054 | 0.33687 | 0.10623 | 0.0058502 | -0.123 | -0.045722 | -0.069964 | -0.065116 | 0.15222 | 0.033957 | -0.043842 | -0.049511 | -0.20785 | -0.0012297 | 0.046092 | -0.13177 | 0.047072 | -0.02339 | 0.033114 | 21.971
db3 | 4.7723 | 10.492 | -0.033483 | -0.0056871 | -0.082628 | 0.1338 | 0.056905 | -0.10854 | 0.0126 | -0.037178 | 0.0072492 | -0.077911 | 0.036761 | 0.039666 | 0.013105 | -0.031562 | 0.061886 | 0.067905 | 0.031559 | -0.08444 | 9.6435
db4 | 0.023033 | 1.0876 | 0.011672 | -0.014122 | 0.016635 | -0.032263 | -0.010837 | 0.054488 | -0.0035121 | 0.022087 | 0.005765 | 0.0055913 | 0.009913 | -0.026222 | -0.011341 | -0.00016948 | -0.015875 | -0.012255 | -0.01569 | 0.016135 | 1.4161
db5 | 0.064686 | -0.68244 | 0.0034824 | -0.003043 | -0.035024 | 0.014896 | -0.014298 | 0.0015267 | -0.014447 | 0.0040542 | 0.0017993 | 0.0047952 | -0.0042074 | 0.0029175 | 0.013254 | 0.0018604 | 0.012658 | 0.0036676 | 0.0078145 | -0.0017055 | 0.049362
db6 | -0.019342 | -0.10136 | -0.0021001 | -0.012426 | 0.013051 | -0.0097004 | 0.009728 | 0.0030591 | 0.0014426 | 0.0019298 | -0.00066868 | -0.00055882 | 0.0051114 | -0.0028005 | -0.002603 | 0.0035843 | -0.006995 | -0.0031263 | -0.00016931 | 0.0032431 | 0.0034518
db7 | 0.0026407 | 0.42866 | -0.0054027 | 0.00062885 | -0.0046666 | 0.0017937 | -0.0011596 | 0.0034527 | -0.0060887 | 0.0073055 | -0.0037377 | 0.0032383 | 0.0027291 | -0.00037359 | 0.003015 | -0.00065754 | 0.0022328 | -0.0016111 | 0.0014348 | -0.0021334 | 0.024901
db8 | -0.005149 | 0.012959 | 0.0030846 | 0.0014919 | -0.0014017 | 0.00093851 | -0.0020602 | 0.0038914 | -0.0022866 | -6.0931e-05 | -0.0042154 | 0.00014634 | 0.00033621 | -0.0016817 | 0.00074497 | 0.00065114 | -8.2853e-05 | 0.00013828 | 0.00011816 | 0.00024773 | -0.0073132
db9 | 0.026698 | 0.043334 | -0.0082505 | 0.0053866 | -0.0042076 | 0.013752 | -0.0020615 | 0.0035402 | -0.0054777 | 0.0041849 | 0.0014636 | 0.0035685 | -0.0023972 | -0.0012796 | -0.0020175 | 0.0027868 | -0.0012901 | -0.0026068 | -0.00086463 | -0.0042296 | 0.0099338
db10 | -4.0724e-05 | 3.5178e-05 | -0.00015832 | -5.6285e-05 | 3.1367e-05 | 0.00010673 | 6.451e-05 | 0.00095969 | 4.2012e-05 | 4.4007e-05 | -0.0034737 | 5.6604e-05 | 3.7311e-05 | 0.00013072 | 0.00029314 | 6.0658e-05 | 4.1733e-05 | 0.0017383 | 3.731e-05 | 4.4216e-05 | 4.3063e-05
db11 | 0.0080419 | 0.0243 | 0.00019019 | 0.001011 | -2.3196e-05 | 0.0038606 | -0.00064845 | 0.00066162 | -0.00055218 | 0.00073135 | -0.0013392 | -0.00060879 | -0.00025428 | 0.00012206 | -0.00045446 | 0.00063808 | -0.00058001 | -0.00071626 | -0.00048352 | -0.0015546 | 0.012691
db12 | -0.0013433 | 0.0025157 | 0.0005918 | 0.00098514 | 0.00026377 | 0.00032107 | 0.0008975 | 2.0167e-05 | 3.6252e-05 | -0.0023036 | 4.2612e-06 | -0.0009324 | 0.0014149 | -1.1075e-05 | -0.00015663 | 0.00063976 | -0.00021283 | -0.00093139 | -0.00071069 | 8.4e-05 | -0.0037871
db13 | 0.0041984 | 0.0057622 | -0.00063158 | 0.00048056 | 4.7254e-05 | 0.001335 | 0.00030377 | -0.00041003 | -0.00039148 | 0.0014959 | -0.00051906 | 1.3105e-05 | -0.00023085 | -0.00022239 | 0.00054434 | 0.00077788 | 0.00028152 | -0.0011008 | -0.0011466 | -0.00062651 | -0.000713
db14 | 0.00057445 | -0.00041092 | -0.00028706 | -0.00062449 | -0.00056365 | 0.0028137 | 0.00087847 | 7.5594e-06 | 0.00084477 | 0.00025222 | -0.00027959 | -0.00050081 | -0.0013124 | -0.00038223 | -0.00063923 | 0.00022465 | -0.0008268 | 6.7893e-05 | 0.00046678 | -0.00013977 | 0.00047696
db15 | -0.017888 | -0.012783 | 0.00071004 | -9.1667e-05 | -0.0015481 | 0.0089561 | -0.0037065 | -0.0028973 | -0.0013825 | 0.0058048 | -0.0022333 | -0.00076461 | 0.002131 | 3.1902e-05 | 0.00077575 | -0.0037884 | 0.00046567 | -0.0031577 | 0.0001925 | 0.00050231 | -0.027282
da2 | -0.28992 | 0.93256 | 0.17751 | -0.035872 | -0.04248 | 0.034515 | 0.14783 | -0.13342 | 0.0051726 | -0.026489 | -0.023798 | -0.029365 | 0.012699 | -0.025135 | 0.070129 | 0.12153 | -0.087849 | -0.056849 | -0.024899 | -0.083229 | 0.4498
da3 | -0.27443 | -2.1226 | 0.22404 | -0.034802 | 0.13833 | -0.17265 | 0.013392 | -0.015847 | 0.073679 | 0.13766 | 0.054957 | 0.061726 | -0.046957 | -0.14083 | -0.075508 | -0.052949 | -0.0723 | -0.13211 | -0.084285 | 0.12446 | -0.33216
da4 | -0.10752 | -0.27885 | -0.024211 | -0.0012203 | -0.034136 | 0.079425 | -0.039498 | 0.0025085 | -0.028588 | -0.015509 | -0.026086 | -0.0095208 | 0.016054 | 0.01431 | 0.033569 | 0.015923 | 0.0083683 | 0.014595 | 0.015609 | -0.021593 | 0.12873
da5 | -0.033109 | 2.0296 | 0.01758 | 0.0012983 | 0.0068324 | -0.0066947 | -9.7664e-05 | -0.0004657 | -0.0079305 | -0.01167 | -0.0095743 | 0.0077628 | -0.0037624 | -0.008772 | 0.004409 | -0.00018875 | 0.0077815 | -0.0022256 | -0.0040219 | 0.0097396 | 0.028882
da6 | -0.0050657 | 0.018827 | 0.014101 | -0.016193 | 0.0011321 | 0.0021245 | 0.01403 | -0.0040671 | 0.0026241 | -0.0070866 | -0.0031288 | -0.0047669 | -0.0016273 | -0.004222 | 0.00071833 | 0.0019544 | 0.0026228 | 0.0033948 | 0.00082543 | -0.0024354 | 0.0058384
da7 | -0.022341 | 1.3775 | -0.013997 | 0.017761 | 0.0074099 | 0.015146 | 0.0098462 | 0.0012003 | -0.0089576 | -0.0060975 | 3.3178e-05 | 0.00404 | 0.003953 | -0.0086669 | 0.0042672 | -0.0050902 | -0.0011796 | -0.0075733 | -0.0095177 | -0.0025779 | -0.019044
da8 | 0.0026423 | -0.018062 | -0.0045104 | 0.0049854 | -0.00136 | -0.0018129 | 2.5411e-05 | 0.0022964 | -0.0025172 | 0.0022096 | 0.0030393 | 0.0019938 | -8.2488e-05 | 0.00033849 | -0.0019237 | -0.0011595 | 0.0018626 | -0.0001363 | -0.0022342 | -0.0010144 | -0.0020004
da9 | -0.0098182 | -0.28264 | -0.0093024 | 0.0076532 | 0.0025484 | 0.021934 | -0.00016312 | 0.0028563 | -0.0043012 | -0.011282 | -0.0021401 | 0.0029353 | 0.0017272 | -0.0035405 | 0.0022804 | 0.001578 | -1.3107e-05 | -0.0031757 | -0.006462 | -0.0031323 | -0.11853
da10 | 3.8249e-05 | 0.00011921 | 0.0022308 | 0.001888 | -4.4988e-05 | 0.0010241 | -5.0474e-05 | 0.00087381 | -4.7691e-05 | 1.7402e-05 | -0.00021478 | -7.3861e-05 | -3.7884e-05 | -0.0036544 | -0.0036568 | -4.9815e-05 | -3.9355e-05 | 0.0019085 | -3.4932e-05 | -3.7678e-05 | 8.4937e-05
da11 | -0.044696 | 0.16062 | -0.0018545 | 0.0090819 | 0.0094097 | 0.01149 | 0.00044858 | 0.0016062 | -0.0037207 | -0.0064638 | 0.0032391 | 0.0035421 | 0.0023059 | -0.0076899 | 0.0016333 | -0.005605 | -0.0015342 | -0.0078473 | -0.0085999 | 0.00055833 | -0.036984
da12 | 0.0014037 | 0.0026213 | 0.00022342 | 0.00010865 | 0.00034558 | -0.0019911 | 0.0010329 | 0.00015219 | 0.00013814 | -0.0014893 | 2.3182e-05 | 0.00045093 | -0.00029284 | -0.00050351 | -0.00053073 | 0.0016112 | 0.00068335 | 0.00072447 | -0.00036767 | -0.00031878 | -0.0012248
da13 | -0.0014151 | -0.0025044 | -6.2837e-05 | 0.0014902 | 0.0013685 | 0.00088869 | 0.0004405 | 0.0010434 | -0.0010848 | 0.00012206 | -0.00020377 | 0.002235 | -8.931e-05 | -0.0020502 | 0.00037249 | 0.00038518 | -0.00017992 | -0.001635 | -0.0024585 | -0.00058172 | -0.00021632
da14 | -0.0065448 | 0.0054501 | 0.0016881 | 0.0018224 | -0.00022548 | 0.0019721 | -8.8678e-05 | 3.4822e-06 | -0.00036234 | 0.00094454 | -0.0011697 | 0.00036634 | -0.0010592 | -0.001237 | -0.00024553 | -0.00031709 | -0.00028749 | -0.0011984 | 0.00020296 | -0.00080897 | 0.0021873
da15 | 0.00075541 | 0.010535 | 0.0012707 | 0.0041162 | -0.0021993 | -0.012028 | -0.00089215 | 0.00013299 | -0.00325 | 0.0065659 | 0.00095599 | 0.0062871 | -0.0015642 | -0.0025878 | 0.0014138 | 0.0019231 | 0.0013936 | 0.0013081 | -0.0033104 | 0.00046403 | -0.0024575
Field Colinearity | 17.9828
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 4.65 | 0.35995 | 0.71975 | 0.89856 | 0.40201 | 0.7347 | 0.59896 | 0.46114 | -0.07979 | 0.11271 | -0.10943 | -0.57716 | -0.51499 | -0.73752 | -0.75328 | -0.95461 | -0.76457 | -0.54817 | -0.68062 | -1.9356
Twist (m2 mrad) | 34.8753 | 5.1292 | 9.7166 | 11.4566 | 4.8241 | 8.2653 | 6.289 | 4.4961 | -0.71811 | 0.92982 | -0.82072 | -3.8959 | -3.0899 | -3.872 | -3.3897 | -3.5798 | -2.2937 | -1.2334 | -1.0209 | -0.72586
Twist Integral (I0) (m2 rad) | 0.046007
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 5.45 | 0.36308 | 0.71559 | 1.3403 | 0.93818 | 0.8176 | 0.69938 | 0.40529 | 0.44398 | 0.13427 | -0.043608 | -0.5274 | -0.77367 | -0.24917 | -0.64062 | -0.99175 | -1.5003 | -1.3347 | -1.4243 | -2.067
Twist (m2 mrad) | 40.8751 | 5.1739 | 9.6605 | 17.0887 | 11.2582 | 9.198 | 7.3434 | 3.9516 | 3.9958 | 1.1078 | -0.32706 | -3.5599 | -4.642 | -1.3082 | -2.8828 | -3.7191 | -4.5009 | -3.0032 | -2.1364 | -0.77513
Twist Integral (I0) (m2 rad) | 0.062099

## Work sheet diff
| HCMBBLA001-3000076 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 2.2266 | 7.348 | 1.0641 | 0.11525 | -0.19957 | 3.4393 | 0.070074 | 0.54894 | -2.7182 | 7.0044 | -1.231 | 0.12039 | 1.0133 | 1.4545 | -0.28992 | 0.082811 |  | 0
3 | 4.7708 | 7.7365 | 4.6988 | 0.12234 | 0.57452 | 0.67282 | -0.38244 | 0.14759 | 4.9609 | 8.0332 | 4.7723 | 0.063813 | 0.6272 | 0.69956 | -0.27443 | 0.11056 |  | 0
4 | 0.063595 | 0.31542 | 0.0084333 | 0.016213 | 0.13646 | 0.27132 | 0.10591 | 0.035878 | -0.28511 | 0.35377 | -0.023033 | 0.020476 | -0.060351 | 0.3628 | -0.10752 | 0.028995 |  | 0
5 | -0.14133 | 1.3549 | 0.040204 | 0.096573 | 0.32963 | 0.37552 | -0.023732 | 0.048033 | 0.10913 | 1.3933 | 0.064686 | 0.01181 | 0.16112 | 0.43785 | -0.033109 | 0.0078239 |  | 0
6 | 0.051578 | 0.11035 | -0.021568 | 0.0031121 | -0.040939 | 0.040871 | 0.017476 | 0.039363 | -0.0073988 | 0.11697 | 0.019342 | 0.006215 | 0.053584 | 0.053742 | -0.0050657 | 0.0069438 |  | 0
7 | 1.1697 | 0.31297 | 0.005659 | 0.016109 | 0.034643 | 0.30946 | -0.012571 | 0.024254 | 1.0995 | 0.36941 | 0.0026407 | 0.0035484 | 0.00084166 | 0.31653 | -0.022341 | 0.0088073 |  | 0
8 | -0.018829 | 0.02601 | -0.0033393 | 0.0070664 | -0.015281 | 0.018517 | -0.0038529 | 0.015163 | -0.013643 | 0.025155 | 0.005149 | 0.0018946 | -0.033608 | 0.019771 | 0.0026423 | 0.0023604 |  | 0
9 | 0.46611 | 0.039657 | 0.02303 | 0.0049222 | -0.0066498 | 0.066179 | -0.0050604 | 0.0044271 | 0.48391 | 0.036031 | 0.026698 | 0.0050258 | -0.021206 | 0.067495 | -0.0098182 | 0.007214 |  | 0
10 | 0 | 0.044337 | -0.0002396 | 0.0012315 | 0 | 0.031664 | -8.2657e-05 | 0.0015788 | -0.00033548 | 0.043484 | 4.0724e-05 | 0.00097541 | 2.6446e-05 | 0.018345 | 3.8249e-05 | 0.001549 |  | 0
11 | 0.65069 | 0.023757 | 0.0068804 | 0.0023289 | -0.027394 | 0.038345 | -0.01488 | 0.0078062 | 0.65969 | 0.021996 | 0.0080419 | 0.0011886 | -0.038849 | 0.045742 | -0.020317 | 0.0060618 |  | 0
12 | -0.0012691 | 0.0053505 | 3.9582e-05 | 0.00071388 | -0.004954 | 0.0038796 | 0.0025153 | 0.00082125 | 0.0011026 | 0.0051719 | 0.0013433 | 0.00085189 | 0.00027484 | 0.0037022 | 0.0014037 | 0.00084444 |  | 0
13 | 0.068432 | 0.0041431 | 0.0036801 | 0.0011061 | -0.00096714 | 0.0055178 | -0.0020145 | 0.0011327 | 0.065634 | 0.0052714 | 0.0041984 | 0.00074112 | -0.0021632 | 0.0057478 | -0.0023367 | 0.0012363 |  | 0
14 | -0.002116 | 0.0014733 | -0.001391 | 0.00047717 | -0.0060133 | 0.0037944 | 0.010504 | 0.00096045 | -0.00076352 | 0.0013059 | -0.00057445 | 0.00090222 | -0.0057617 | 0.0040199 | 0.01365 | 0.0010196 |  | 0
15 | 0.018379 | 0.0073068 | -0.015042 | 0.0016048 | -0.0031006 | 0.0031236 | -0.0025612 | 0.0026544 | 0.017542 | 0.0075136 | -0.017888 | 0.0032213 | -0.00038474 | 0.0028223 | 0.00075541 | 0.0041491 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -1.3802 | -2.3041
3 | [12 empty] | 3 | 9 | -2.0071 | -2.477
4 | [12 empty] | 4 | 16 | -3.7248 | -3.7146
5 | [12 empty] | 5 | 25 | -2.6867 | -4.6447
6 | [12 empty] | 6 | 36 | -4.5037 | -5.0253
7 | [12 empty] | 7 | 49 | -3.9238 | -5.1867
8 | [12 empty] | 8 | 64 | -4.5706 | -6.1589
9 | [12 empty] | 9 | 81 | -5.367 | -5.1124
10 | [12 empty] | sigma (mm) |  | 0.006506 | 0.0029016
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05 | [5 empty] | Aperture 1 | Aperture 2
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | Magnetic length | -0.080914 | -0.077853
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 4.396 | 107.1391 | 107.9758 | 107.9318 | 107.6804 | 108.0021 | 108.027 | 108.0672 | 108.0582 | 108.1103 | 108.1103 | 108.1278 | 108.1305 | 108.0847 | 108.1008 | 108.0261 | 108.0497 | 107.9926 | 107.398 | 30.5748
2 | 28.4661 | 1.2976 | 1.0431 | 1.1812 | 1.15 | 0.97453 | 0.93078 | 1.1761 | 1.1218 | 1.1161 | 1.0891 | 1.0439 | 1.1385 | 1.0146 | 1.0847 | 0.87629 | 0.83916 | 0.97213 | 1.104 | 23.6546 | 2.2266
3 | 7.6199 | 4.684 | 4.7421 | 4.7209 | 4.2864 | 4.7356 | 4.6934 | 4.6455 | 4.7315 | 4.8155 | 4.7256 | 4.8433 | 4.7719 | 4.7409 | 4.7101 | 4.7625 | 4.7758 | 4.6359 | 4.5571 | 9.4586 | 4.7708
4 | 2.0711 | 0.029941 | -0.002212 | 0.018473 | 0.013287 | 0.0057289 | -0.0011943 | 0.02533 | 0.032078 | 0.026391 | 0.025893 | -0.00053507 | -0.0052384 | -0.0059923 | 0.012234 | -0.010553 | -0.019971 | -0.011764 | 0.019902 | 0.72708 | 0.063595
5 | -0.95626 | 0.071963 | 0.047159 | 0.048713 | -0.34551 | 0.061548 | 0.06731 | 0.055826 | 0.061806 | 0.070824 | 0.06332 | 0.069807 | 0.073208 | 0.071681 | 0.059558 | 0.066108 | 0.065069 | 0.052903 | 0.062367 | 0.045529 | -0.14133
6 | 0.37573 | -0.023053 | -0.018518 | -0.013716 | -0.01928 | -0.018247 | -0.023755 | -0.019794 | -0.019169 | -0.020645 | -0.020424 | -0.025803 | -0.026224 | -0.022861 | -0.024426 | -0.023815 | -0.023069 | -0.022689 | -0.022733 | -0.0030335 | 0.051578
7 | 0.15726 | -0.0087074 | -0.0085353 | 0.0032995 | 0.067535 | 0.0054619 | 0.0063291 | -0.0014321 | 0.0018638 | 0.0038654 | 0.0066093 | 0.0015501 | 0.0051715 | 0.0054249 | 0.0079319 | -0.00025381 | 0.001438 | 0.0026411 | 0.001669 | 0.039637 | 1.1697
8 | 0.071966 | -0.0056862 | -0.0046499 | -0.0045576 | 0.024144 | -0.0023346 | -0.0065147 | -0.0022117 | -0.0044308 | -0.0071225 | -0.0045092 | -0.006886 | -0.0071775 | -0.0036389 | -0.0056865 | -0.0070861 | -0.0017598 | -0.0055672 | -0.0044326 | -0.0022288 | -0.018829
9 | -0.0084501 | 0.020646 | 0.025863 | 0.024535 | 0.0054387 | 0.02149 | 0.025235 | 0.021392 | 0.022728 | 0.026035 | 0.026196 | 0.026979 | 0.024764 | 0.026158 | 0.025954 | 0.025459 | 0.024418 | 0.021334 | 0.019909 | 0.039915 | 0.46611
10 | -7.8475e-05 | 1.8624e-05 | 1.4262e-06 | -1.4163e-05 | -0.00013829 | -7.761e-05 | -0.0031282 | 0.00012833 | -0.00075904 | -0.0030213 | 5.6464e-06 | 5.7102e-05 | -2.7968e-05 | -2.3364e-05 | 3.024e-05 | 4.2629e-06 | 0.0026972 | -6.332e-05 | -2.3327e-06 | -0.00047623 | 0
11 | 0.0070283 | 0.0091623 | 0.0080872 | 0.0088267 | -0.001555 | 0.0064936 | 0.0065806 | 0.0065866 | 0.0075163 | 0.0062734 | 0.0082294 | 0.0075288 | 0.0084525 | 0.0073347 | 0.0071167 | 0.0065162 | 0.0083365 | 0.0053772 | 0.0069834 | 0.024252 | 0.65069
12 | -0.013224 | 0.011204 | -0.00030719 | 0.00018972 | 0.0095234 | 0.0045353 | 0.00058996 | -0.0017042 | 0.0014681 | 0.012924 | 0.0015882 | -0.0031806 | -0.010947 | -0.015998 | 0.0039063 | -0.00057312 | 0.0022808 | -0.0069788 | -0.0013969 | 0.00013863 | -0.012691 | ten times!
13 | 0.0097819 | 0.03304 | 0.035472 | 0.033999 | 0.078568 | 0.032116 | 0.032409 | 0.027215 | 0.037693 | 0.036168 | 0.041021 | 0.035391 | 0.03042 | 0.03178 | 0.04303 | 0.036069 | 0.032298 | 0.033532 | 0.032201 | 0.040747 | 0.68432 | ten times!
14 | -0.021538 | -0.010707 | -0.015902 | -0.0097355 | -0.01184 | -0.0043928 | -0.010865 | -0.021555 | -0.011781 | -0.013525 | -0.018642 | -0.013167 | -0.020919 | -0.015371 | -0.014377 | -0.011819 | -0.0072019 | -0.017485 | -0.021093 | -0.0045591 | -0.02116 | ten times!
15 | -0.14836 | -0.1508 | -0.15056 | -0.14798 | -0.14425 | -0.16121 | -0.15414 | -0.13866 | -0.14932 | -0.16013 | -0.14225 | -0.14085 | -0.16538 | -0.16814 | -0.18871 | -0.10902 | -0.14416 | -0.14881 | -0.1432 | -0.035379 | 0.18379 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1.5764 | 1.7168 | 1.8349 | 1.7579 | 1.2597 | 0.63046 | 0.36555 | -0.2007 | -0.35347 | 0.010882 | 0.30929 | -0.54139 | -0.5512 | -0.99617 | -1.049 | -1.0426 | -0.48378 | -0.36683 | -0.73562 | -0.79802
2 | -4.0119 | 0.42562 | 0.21657 | 0.13585 | -2.1003 | 0.12654 | 0.16181 | 0.12799 | 0.19191 | 0.26888 | 0.30165 | 0.23108 | 0.15958 | 0.30336 | 0.20872 | 0.066998 | 0.062475 | 0.21868 | 0.15395 | 1.8444 | -0.19957
3 | -2.1492 | -0.13 | -0.26293 | -0.23456 | -0.21623 | -0.4369 | -0.419 | -0.30219 | -0.27167 | -0.2994 | -0.3124 | -0.47056 | -0.59461 | -0.46545 | -0.39019 | -0.57375 | -0.6808 | -0.51399 | -0.30934 | -0.37032 | 0.57452
4 | -1.2868 | 0.17789 | 0.11772 | 0.17658 | 0.054127 | 0.11765 | 0.086297 | 0.11335 | 0.11806 | 0.13194 | 0.10786 | 0.095311 | 0.061133 | 0.10318 | 0.07802 | 0.069256 | 0.062995 | 0.093775 | 0.14121 | 0.32973 | 0.13646
5 | -0.11281 | -0.018139 | -0.033155 | -0.038993 | 0.16612 | -0.03653 | -0.033184 | -0.03454 | -0.030827 | -0.03262 | -0.022501 | -0.046774 | -0.048659 | -0.040482 | -0.033653 | -0.040508 | -0.04448 | -0.032609 | -0.025647 | 0.035401 | 0.32963
6 | -0.2521 | 0.013463 | 0.0064532 | 0.018314 | 0.17416 | 0.012966 | 0.005173 | 0.013414 | 0.0106 | 0.0051991 | 0.0081912 | 0.0095771 | 0.0042571 | 0.0056232 | 0.0077627 | 0.0043567 | -0.00017895 | 0.0037259 | 0.011519 | 0.028696 | -0.040939
7 | 0.051222 | -0.017993 | -0.0032491 | -0.010072 | 0.081807 | -0.012674 | -0.014961 | -0.026224 | -0.02143 | -0.020448 | -0.012155 | -0.021274 | -0.023795 | -0.02359 | -0.020772 | -0.024833 | -0.017779 | -0.017675 | -0.019164 | -0.040583 | 0.034643
8 | -0.025623 | -0.0011578 | -0.001517 | -0.0038548 | -0.063799 | -0.0042049 | 0.0019813 | 0.002888 | 0.00073778 | -0.0030626 | -0.0024436 | 0.0028013 | 0.00026493 | -0.0012277 | -4.4466e-05 | 0.0048947 | 0.00099786 | -0.0026121 | 5.2033e-06 | 0.00045292 | -0.015281
9 | -0.03789 | -0.0039143 | -0.0010348 | -0.00042169 | 0.0052511 | -0.0055717 | -0.002025 | -0.0074618 | -0.011644 | -0.0090514 | -0.00089234 | -0.0047126 | -0.0079897 | -0.01157 | -0.0036629 | -0.010445 | -0.0055895 | -0.0025674 | -0.0077824 | -0.034366 | -0.0066498
10 | 0.00014716 | 0.00025827 | -1.0821e-06 | -2.4269e-05 | -0.0030054 | -0.0016835 | 0.0024995 | 0.0031242 | 1.5807e-05 | 0.00016112 | -1.7565e-05 | 0.00099804 | -1.2198e-05 | -3.9966e-05 | -1.9777e-05 | 2.947e-05 | -7.4529e-05 | -0.0036775 | -1.8507e-05 | 0.00034703 | 0
11 | 0.013763 | -0.0060376 | -0.0043695 | -0.0035315 | 0.003847 | -0.011892 | -0.01305 | -0.019364 | -0.018527 | -0.015693 | -0.012343 | -0.019334 | -0.020528 | -0.024217 | -0.022396 | -0.022485 | -0.019489 | -0.018019 | -0.020419 | -0.019244 | -0.027394 | ten times!
12 | 0.02714 | 0.027538 | 0.0281 | 0.021534 | 0.049794 | 0.028002 | 0.01993 | 0.022357 | 0.017633 | 0.023929 | 0.024033 | 0.031116 | 0.023499 | 0.027958 | 0.033389 | 0.019388 | 0.012586 | 0.01509 | 0.026881 | 0.044777 | -0.04954 | ten times!
13 | -0.046896 | -0.0080139 | 0.00043113 | -0.010862 | -0.0082789 | -0.018415 | -0.0071959 | -0.028595 | -0.025612 | -0.029579 | -0.0016538 | -0.031131 | -0.03108 | -0.036801 | -0.021959 | -0.025118 | -0.023213 | -0.029929 | -0.025608 | -0.018642 | -0.0096714 | ten times!
14 | 0.014123 | 0.12932 | 0.10978 | 0.09668 | 0.083721 | 0.094193 | 0.10161 | 0.09919 | 0.10068 | 0.10174 | 0.11169 | 0.10287 | 0.10279 | 0.1097 | 0.11357 | 0.11109 | 0.1102 | 0.10243 | 0.10948 | 0.098589 | -0.060133 | ten times!
15 | 0.057535 | -0.0052535 | -0.0026189 | -0.039813 | -0.034058 | -0.015677 | -0.014952 | -0.059759 | -0.01048 | -0.0076599 | -0.017753 | -0.046713 | -0.081467 | -0.060991 | -0.017733 | 0.023961 | -0.0077712 | -0.054101 | -0.0081793 | -0.0062706 | -0.031006 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -13.3808 | 106.6961 | 107.885 | 107.9279 | 108.0072 | 107.9765 | 107.9785 | 108.0117 | 108.0165 | 107.9711 | 107.9706 | 108.0654 | 108.0439 | 107.9468 | 107.9742 | 107.9765 | 107.9222 | 107.886 | 107.3023 | 34.4061
2 | -24.8054 | -1.5679 | -1.3372 | -1.2368 | -1.108 | -1.1853 | -1.161 | -1.1659 | -1.3832 | -1.265 | -1.1872 | -1.1815 | -1.0231 | -1.2298 | -1.2771 | -1.0992 | -1.2781 | -1.2076 | -1.2641 | -23.2179 | -2.7182
3 | 10.492 | 4.7388 | 4.7666 | 4.6897 | 4.9061 | 4.8292 | 4.6638 | 4.7849 | 4.7351 | 4.7796 | 4.6944 | 4.8091 | 4.812 | 4.7854 | 4.7408 | 4.8342 | 4.8402 | 4.8039 | 4.6879 | 9.6435 | 4.9609
4 | -0.4317 | -0.034706 | -0.0089109 | -0.039669 | 0.0092292 | -0.012196 | -0.077521 | -0.019521 | -0.04512 | -0.028798 | -0.028625 | -0.032946 | 0.0031887 | -0.011693 | -0.022864 | -0.0071581 | -0.010778 | -0.0073436 | -0.039169 | -0.96137 | -0.28511
5 | -0.68244 | 0.068168 | 0.061643 | 0.029662 | 0.079582 | 0.050387 | 0.066212 | 0.050238 | 0.06874 | 0.066485 | 0.069481 | 0.060478 | 0.067603 | 0.077939 | 0.066546 | 0.077344 | 0.068353 | 0.0725 | 0.06298 | 0.049362 | 0.10913
6 | 0.22092 | 0.021442 | 0.031768 | 0.0062912 | 0.029042 | 0.0096137 | 0.016283 | 0.017899 | 0.017412 | 0.02001 | 0.019901 | 0.01423 | 0.022142 | 0.021945 | 0.015757 | 0.026337 | 0.022468 | 0.019511 | 0.016099 | -0.0034518 | -0.0073988
7 | 0.42866 | -0.002762 | 0.0032695 | -0.0020259 | 0.0044344 | 0.0014811 | 0.0060934 | -0.003448 | 0.0099462 | -0.001097 | 0.005879 | 0.0053697 | 0.0022671 | 0.0056557 | 0.0019832 | 0.0048735 | 0.0010296 | 0.0040755 | 0.00050733 | 0.024901 | 1.0995
8 | 0.027173 | 0.0020643 | 0.003657 | 0.0065506 | 0.0042104 | 0.0072091 | 0.0012575 | 0.0074355 | 0.0052099 | 0.0093644 | 0.0050026 | 0.0048127 | 0.0068306 | 0.004404 | 0.0044978 | 0.0052318 | 0.0050107 | 0.0050308 | 0.0049012 | 0.0073132 | -0.013643
9 | 0.043334 | 0.018447 | 0.032084 | 0.02249 | 0.04045 | 0.024636 | 0.030238 | 0.02122 | 0.030883 | 0.028161 | 0.030266 | 0.0243 | 0.025418 | 0.02468 | 0.029484 | 0.025408 | 0.024091 | 0.025833 | 0.022468 | 0.0099338 | 0.48391
10 | -2.8231e-05 | 0.00019904 | 9.7008e-05 | 9.3566e-06 | -6.6002e-05 | -2.3786e-05 | -0.00091896 | -1.2885e-06 | -3.2832e-06 | 0.0035144 | -1.588e-05 | 3.4128e-06 | -8.9992e-05 | -0.00025242 | -1.9934e-05 | -1.0095e-06 | -0.0016976 | 3.4134e-06 | -3.4919e-06 | 5.3914e-05 | -0.00033548
11 | 0.0243 | 0.0082321 | 0.0090529 | 0.0080188 | 0.011903 | 0.0073935 | 0.0087036 | 0.0074898 | 0.0087733 | 0.0067027 | 0.0074332 | 0.0077877 | 0.008164 | 0.0075875 | 0.00868 | 0.0074619 | 0.0073257 | 0.0075584 | 0.0064874 | 0.012691 | 0.65969
12 | -0.029139 | 0.0075148 | 0.0035813 | 0.010795 | 0.010222 | 0.0044578 | 0.013231 | 0.01307 | 0.036468 | 0.01339 | 0.022757 | -0.00071589 | 0.013544 | 0.014999 | 0.0070352 | 0.015561 | 0.022747 | 0.02054 | 0.012593 | -0.0011079 | 0.011026 | ten times!
13 | 0.057622 | 0.035668 | 0.046789 | 0.042456 | 0.055334 | 0.045021 | 0.037883 | 0.038069 | 0.056942 | 0.036793 | 0.042115 | 0.039675 | 0.03976 | 0.047427 | 0.049763 | 0.044799 | 0.030976 | 0.030518 | 0.035719 | -0.00713 | 0.65634 | ten times!
14 | -0.039018 | -0.0028739 | 0.0005004 | -0.00010801 | -0.033881 | -0.014529 | -0.0058201 | -0.014192 | -0.0082667 | -0.0029486 | -0.00073634 | 0.0073795 | -0.0019222 | 0.00064788 | -0.007991 | 0.0025235 | -0.0064234 | -0.010412 | -0.0043468 | -0.034906 | -0.0076352 | ten times!
15 | -0.20856 | -0.17178 | -0.1798 | -0.19436 | -0.089321 | -0.21595 | -0.20786 | -0.19271 | -0.12083 | -0.20121 | -0.18653 | -0.15757 | -0.17856 | -0.17112 | -0.21677 | -0.17423 | -0.21046 | -0.17696 | -0.17386 | -0.27282 | 0.17542 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 3.2361 | -0.19998 | 1.2234 | 1.408 | 1.153 | 0.067735 | 0.38653 | -0.69476 | -1.0066 | 0.33805 | 0.56927 | 0.29399 | -1.1853 | 0.025727 | -0.75557 | -0.3316 | -1.1979 | -1.2322 | -0.009953 | -0.58392
2 | -2.4969 | -0.1124 | -0.32579 | -0.3324 | -0.2554 | -0.14209 | -0.42334 | -0.28474 | -0.31641 | -0.31372 | -0.31928 | -0.27722 | -0.31505 | -0.21979 | -0.16839 | -0.37777 | -0.34677 | -0.31482 | -0.37315 | 0.4498 | 1.0133
3 | -4.3616 | -0.050389 | -0.30923 | -0.13609 | -0.44708 | -0.26103 | -0.29027 | -0.20075 | -0.13676 | -0.21947 | -0.2127 | -0.32138 | -0.41526 | -0.34994 | -0.32738 | -0.34673 | -0.40654 | -0.35871 | -0.14997 | -0.33216 | 0.6272
4 | -0.33965 | -0.13173 | -0.10874 | -0.14166 | -0.028095 | -0.14702 | -0.10501 | -0.13611 | -0.12303 | -0.13361 | -0.11704 | -0.091466 | -0.09321 | -0.073951 | -0.091597 | -0.099152 | -0.092926 | -0.091911 | -0.12911 | 0.12873 | -0.060351
5 | -0.655 | -0.015529 | -0.031811 | -0.026277 | -0.039804 | -0.033207 | -0.033575 | -0.04104 | -0.04478 | -0.042683 | -0.025346 | -0.036872 | -0.041881 | -0.0287 | -0.033298 | -0.025328 | -0.035335 | -0.037131 | -0.02337 | 0.028882 | 0.16112
6 | 0.082953 | 0.0090357 | -0.021259 | -0.0039336 | -0.0029412 | 0.0089639 | -0.0091328 | -0.0024416 | -0.012152 | -0.0081945 | -0.0098326 | -0.006693 | -0.0092877 | -0.0043474 | -0.0031113 | -0.0024429 | -0.0016709 | -0.0042403 | -0.0075011 | 0.0058384 | 0.053584
7 | 0.026272 | -0.036338 | -0.0045796 | -0.014931 | -0.0071944 | -0.012495 | -0.021141 | -0.031299 | -0.028438 | -0.022308 | -0.018301 | -0.018388 | -0.031008 | -0.018074 | -0.027431 | -0.023521 | -0.029914 | -0.031859 | -0.024919 | -0.019044 | 0.00084166
8 | 0.036665 | -0.0018681 | 0.0076277 | 0.0012823 | 0.00082937 | 0.0026677 | 0.0049387 | 0.00012512 | 0.0048519 | 0.0056816 | 0.0046361 | 0.0025598 | 0.0029808 | 0.00071859 | 0.0014828 | 0.0045049 | 0.002506 | 0.00040809 | 0.0016279 | -0.0020004 | -0.033608
9 | -0.064441 | -0.019121 | -0.002165 | -0.0072698 | 0.012116 | -0.0099813 | -0.0069619 | -0.014119 | -0.0211 | -0.011958 | -0.0068829 | -0.008091 | -0.013359 | -0.0075378 | -0.0082402 | -0.0098313 | -0.012994 | -0.01628 | -0.012951 | -0.01762 | -0.021206
10 | 0.00011976 | 0.0022691 | 0.0019263 | -6.7392e-06 | 0.0010624 | -1.2225e-05 | 0.00091206 | -9.4422e-06 | 5.5651e-05 | -0.00017653 | -3.5611e-05 | 3.6501e-07 | -0.0036162 | -0.0036185 | -1.1566e-05 | -1.1062e-06 | 0.0019467 | 3.3175e-06 | 5.7157e-07 | 0.0001591 | 2.6446e-05
11 | 0.031142 | -0.022172 | -0.011236 | -0.010908 | -0.0088271 | -0.019869 | -0.018711 | -0.024038 | -0.026781 | -0.017078 | -0.016775 | -0.018012 | -0.028007 | -0.018684 | -0.025922 | -0.021852 | -0.028165 | -0.028917 | -0.019759 | -0.019893 | -0.038849
12 | -0.0035896 | 0.016271 | 0.015123 | 0.017493 | -0.0058743 | 0.024366 | 0.015559 | 0.015418 | -0.0008565 | 0.014269 | 0.018546 | 0.011108 | 0.0090018 | 0.0087295 | 0.030149 | 0.02087 | 0.021282 | 0.01036 | 0.010849 | 0.021374 | 0.0027484 | ten times!
13 | -0.025044 | -0.023996 | -0.0084656 | -0.0096818 | -0.01448 | -0.018962 | -0.012933 | -0.034215 | -0.022147 | -0.025405 | -0.0010172 | -0.02426 | -0.043869 | -0.019642 | -0.019516 | -0.025166 | -0.039717 | -0.047953 | -0.029185 | -0.0021632 | -0.021632 | ten times!
14 | -0.012316 | 0.15338 | 0.15472 | 0.13425 | 0.15622 | 0.13561 | 0.13653 | 0.13288 | 0.14595 | 0.1248 | 0.14016 | 0.12591 | 0.12413 | 0.13404 | 0.13333 | 0.13363 | 0.12452 | 0.13853 | 0.12841 | 0.088273 | -0.057617 | ten times!
15 | 0.060849 | 0.020261 | 0.048716 | -0.014439 | -0.11272 | -0.0013674 | 0.008884 | -0.024946 | 0.073214 | 0.017114 | 0.070426 | -0.0080877 | -0.018324 | 0.021692 | 0.026785 | 0.02149 | 0.020635 | -0.02555 | 0.012194 | 0.020775 | -0.0038474 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.54 | 0.66 | 3.5 | 7 | 0.58 | 0.68 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | 0.06 | 0.3 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.056 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | 0 | 0.09 | 3.5 | 7 | -0.027 | 0.097 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0 | 0.0052 | 3.5 | 7 | 0 | 0.0049 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0052 | 0.0041 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | -0.7 | 2.2 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -3.15 | 1.42 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.38 | 0.6 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | 1.57 | 0.36 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | -0.1 | 0.3 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | 1.5 | 0.12 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0 | 0.065 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.194 | 0.043 | 3.5 | 7 | -0.012 | 0.031 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0.012 | 0.095 | 3.5 | 7
a11 (units) | 0.008 | 0.039 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | 0.19 | 0.043 | 3.5 | 7 | 0.013 | 0.038 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0 | 0.02 | 3.5 | 7 | 0 | 0.0097 | 3.5 | 7
a13 (units) | 0.0015 | 0.0054 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.0079 | 0.0035 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0012 | 0.0059 | 3.5 | 7 | -0.0023 | 0.0041 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | 0.0037 | 0.0075 | 3.5 | 7 | 0 | 0.009 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -1.77 | 2.85 | -4.08 | 5.16 | -1.8 | 2.96 | -4.18 | 5.34
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -0.99 | 1.11 | -2.04 | 2.16 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.196 | 0.196 | -0.392 | 0.392 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.315 | 0.315 | -0.63 | 0.63 | -0.3665 | 0.3125 | -0.706 | 0.652
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0182 | 0.0182 | -0.0364 | 0.0364 | -0.01715 | 0.01715 | -0.0343 | 0.0343
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.01955 | 0.00915 | -0.0339 | 0.0235 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -8.4 | 7 | -16.1 | 14.7 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -8.12 | 1.82 | -13.09 | 6.79 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -2.48 | 1.72 | -4.58 | 3.82 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | 0.31 | 2.83 | -0.95 | 4.09 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.15 | 0.95 | -2.2 | 2 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | 1.08 | 1.92 | 0.66 | 2.34 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.2275 | 0.2275 | -0.455 | 0.455 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.3445 | -0.0435 | -0.495 | 0.107 | -0.1205 | 0.0965 | -0.229 | 0.205
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.35 | 0.35 | -0.7 | 0.7 | -0.3205 | 0.3445 | -0.653 | 0.677
a11 (units) | -0.1285 | 0.1445 | -0.265 | 0.281 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | 0.0395 | 0.3405 | -0.111 | 0.491 | -0.12 | 0.146 | -0.253 | 0.279
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.07 | 0.07 | -0.14 | 0.14 | -0.03395 | 0.03395 | -0.0679 | 0.0679
a13 (units) | -0.0174 | 0.0204 | -0.0363 | 0.0393 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.02015 | 0.00435 | -0.0324 | 0.0166 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.01945 | 0.02185 | -0.0401 | 0.0425 | -0.01665 | 0.01205 | -0.031 | 0.0264
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.02255 | 0.02995 | -0.0488 | 0.0562 | -0.0315 | 0.0315 | -0.063 | 0.063
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.82075 | 3.6489 | 0.0055744 | 0.10155 | 1.1348 | -0.79576 | -0.57892 | 0.68119 | 0.1145 | 17.2165 | 0.7043 | -1.4109 | 3.8241 | -0.22127 | 0.34705 | 1.0594 | -0.90438 | 0.84696 | 0.7271 | -0.1015 | 19.0316 | 0.70426 | 2.2266 | 4.7708 | 0.063595 | -0.14133 | 1.1697 | 0 | -0.19957 | 0.57452 | 0.13646 | 14356.926 | -2.7182 | 4.9609 | -0.28511 | 0.10913 | 1.0995 | 0 | 1.0133 | 0.6272 | -0.060351 | 14354.609
Shims
Aperture 1 | [7 empty] | Aperture 2
Aperture 1 b2: Upper; Aperture 1 b7: Lower; Aperture 1 a4: Upper |  |  |  | Lower
Aperture 1 b2: Inner layer; Aperture 1 b4: Outer layer; Aperture 1 b7: Inner layer; Aperture 1 a2: Outer layer; Aperture 1 a4: Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | English | 0 | 0