## Alarm sheet
File name | HCMBALA001-3000199_cm.xls
Component ID | HCMBALA001 |  |  |  | Serial Number | 3000199
Date of test Ap 1 | 2004-07-08 |  |  |  | Date of test Ap 2 | 2004-07-08
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | yellow alarm | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 |  |  |  |  | b10
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | status ok | status ok | status ok | b15 | status ok | status ok | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 |  |  |  |  | a10
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity
Twist Integral | status ok |  |  |  | Twist Integral | status ok

## Assembly Data
Magnet name
| HCMBALA001 | 3000199 |  | Additional lamin. (mm) |  | 0
COMMENTS:
| HOLDING POINT | OK | Data at CERN |  | 2004-07-08 | Analysis by | P. Hagen | Answer to MMS-LD |  | 2004-07-08 | Answer to manufacturer |  | 2004-07-10
1 | Fill in the yellow fields in Original Data | [10 empty] | English
2 | CTRL-g to load the measurement files .txt
3 | Save as … with the "official" name in Original Data C1
4 | Open the file of the collared coil measurement of the same magnet
5 | Go back to cold_mass file. Tools - Macro - Macros - Make difference - Run (to make the difference between collared coil and cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (to perform the analysis - results in Alarm sheet)
7 | Save the file
| Cold mass template - version 11.10.2002 | [7 empty] | Version : | cern
1 | Remplir les champs jaunes dans Original Data | [10 empty] | Francais
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec le nom "officiel" dans Original Data C1
4 | Ouvrir le fichier avec les mesures de la collared coil
5 | Revenir au ficher cold_mass. Tools - Macro - Macros - Make difference - Run (pour faire la difference entre collared coil et cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (pour faire l'analyse - resultats dans Alarm sheet)
7 | Sauver le fichier
1 | Riempire I campi gialli in Original Data | [10 empty] | Italiano
2 | CTRL-g per caricare I file di misura .txt
3 | Salvare il file con il nome "ufficiale" in Original Data C1
4 | Aprire il file delle misure della collared coil corrispondente
5 | Tornare sul file cold_mass. Tools - Macro - Macros - Make difference - Run (per fare la differenza tra collared coil e cold mass)
6 | Tools - Macro - Macros - Collared coil analysis - Run (per fare l'analisi - risultati in Alarm sheet)
7 | Salvare il file
1 | Die gelben Felder in Original Data ausfüllen | [10 empty] | Deutsch
2 | Ctrl-g um die Messdatei ( Datei .txt) zu laden.
3 | Die Datei mit der ”offiziellen” bezeichnung im Original Data C1 speichern mit “Save as…”.
4 | Öffnen “Collared Coils” Messdatei (gleiche Magnete versteht sich)
5 | Zurück zum cold mass Datei. Tools - Macro - Macros - Make difference - Run (Zeigt allen unterschieden in Messdatei zw. cc und cm)
6 | Tools - Macro - Macros - Collared coil analysis - Run (Um Maßdaten zu analysieren - Resultaten ins Alarm Sheet)
7 | Datei speichern
| Last modif on | by | Modifications:
| 2002-10-20 | E. Todesco | Cell for iron additional iron laminations - check on magnetic length includes additional iron laminations
| 2002-11-13 | E. Wildner | German version of instructions added
| 2003-10-01 | E. Wildner | Update of bounds: all heads tested on differences
| 2003-12-02 | E. Todesco | Removal of correction for magnetic length and main field implemented in 12/20/2002
| 2003-12-17 | P. Hagen | Added Twist Integral - I0
| 2004-01-20 | P. Hagen | Fields which must be filled-in or check in yellow and with initial value ?
| 2004-01-20 | P. Hagen | Predefined some fields for convenience (cross section=3, analyser, controller)
| 2004-01-20 | P. Hagen | VB code changes. Macro "All" runs all macros. Cleaned up module names, macro names and shortcuts
| 2004-01-20 | P. Hagen | GetSetData changed to extract the dipole type and correctly reformat magnet name and serial number
| 2004-01-20 | P. Hagen | GetSetData changed to reformat mole name according to CERN AT/MAS convention
| 2004-01-20 | P. Hagen | Compute sigma units for cells with alarm (yellow or red)
| 2004-01-20 | P. Hagen | Using same date format in this table and holding point as in Oracle
| 2004-01-28 | P. Hagen | Changed notation "mbp" with "Heads CS/NCS" in Work Sheet
| 2004-01-28 | P. Hagen | Fixed issue in GetSetMacro macro that "Original data" must be current sheet before starting macro
| 2004-01-28 | P. Hagen | Changed the way CopyPaste macro determines CC or CM
| 2004-01-28 | P. Hagen | Sufficient to enter only first meas file name in "Original data" P4
| 2004-01-28 | P. Hagen | Copied DIMM specific mole data into "Original Data" for completeness
| 2004-01-28 | P. Hagen | Replace "Exit Sub" with "End" in macros to ensure abort execution when problems
| 2004-01-29 | P. Hagen | Processing of measurement info from web based holding point (leave "Original data" P4 and P5 empty)
| 2004-01-29 | P. Hagen | Changed ModifTitre macro to correctly handle 3 digits magnet numbers
| 2004-02-03 | E. Todesco | Correction of integrals in worksheet, worksheet diff, MTF, and computation of dx dy from multipoles
| 2004-02-12 | P. Hagen | Detection of suspect meas date, corrected processing of non-nominal shims
| 2004-03-08 | P. Hagen | Workaround for meas date wrong if Regional Settings short date format not US (Excel OpenText inconsistent with "best practice")
| 2004-03-30 | P. Hagen | Corrected "Time of test" in Original data to contain start time - end time
| 2004-04-02 | P. Hagen | Added more code to reformat magnet serial no which is followed by _ space or -
| 2004-07-02 | P. Hagen | Added Inclinometer data into separate sheet if present in ASCII files

## Original data
File name |  | HCMBALA001-03000199_cm.xls
Component ID |  | HCMBALA001 |  |  | Serial Number |  |  | 3000199 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Noell | [9 empty] | File name for aperture 1 |  |  | 08-JUL-2004_3199cm1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 34 |  |  |  | File name for aperture 2 |  |  | 08-JUL-2004_3199cm2.txt |  |  |  | a even | 1
Test Operator |  | Krause |  |  | Test Controller |  |  | P. Hagen
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2004-07-08 |  |  | Time of test |  |  | 07:04 - 08:55 |  |  |  | Date of test |  | 2004-07-08 |  |  | Time of test |  |  | 07:04 - 08:55
Mole name |  | Dipole 14 | [9 empty] | Mole name |  | Dipole 16
Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 24 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature (Celsius) |  |  | 24
Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 10 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-) |  |  |  | N. points |  | 20 |  |  | N. measure |  |  | 6 (3+,3-)
Rotation speed |  | 0.49 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.5 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00017292 |  |  | Dy average |  |  | 0.00046474 |  |  |  | Dx average |  | 0.000169 |  |  | Dy average |  |  | 0.00064231
Magnetic length |  | 14.3569 |  |  | Tranfs. Function (mT/KA) |  |  | 704.4172 |  |  |  | Magnetic length |  | 14.3569 |  |  | Tranfs. Function (mT/KA) |  |  | 704.4733
Phase |  | -1.1025 |  |  | Integral field (T.m) |  |  | 101.1546 |  |  |  | Phase |  | -0.14226 |  |  | Integral field (T.m) |  |  | 101.1755
File | Aperture 1 - Cold mass | [19 empty] | Integrals | File | Aperture 2 - Cold mass | [19 empty] | Integrals
C1 (mT) | 3.8763 | 7.038 | 7.0438 | 7.0449 | 7.0439 | 7.0445 | 7.0454 | 7.0452 | 7.0445 | 7.0444 | 7.0439 | 7.0449 | 7.0445 | 7.0449 | 7.0454 | 7.0444 | 7.045 | 7.045 | 7.0425 | 4.2188 | 101.1546 | C1 (mT) | 3.7925 | 7.0378 | 7.0447 | 7.0459 | 7.0445 | 7.0446 | 7.045 | 7.0452 | 7.0452 | 7.0451 | 7.0448 | 7.0461 | 7.0451 | 7.0458 | 7.0464 | 7.0454 | 7.0456 | 7.0452 | 7.0429 | 4.3225 | 101.1755
Angle (mrad) | 6.1044 | 1.2856 | 1.7213 | 2.345 | 1.7976 | 1.2036 | -0.43968 | -0.6789 | -0.93932 | -0.12151 | -0.69716 | -0.80698 | -0.28006 | -1.0082 | -1.5458 | -0.83679 | -0.78412 | -1.4533 | -1.4573 | -1.1047 | -1.1025 | Angle (mrad) | 2.8289 | -1.7489 | -1.437 | -0.42992 | -0.16319 | 0.062918 | 0.039498 | 0.33145 | 0.25471 | 0.76057 | 0.22736 | 0.22758 | 0.51759 | 0.3651 | -0.028057 | 0.23752 | 0.2594 | -0.33822 | -0.15575 | -0.82721 | -0.14226
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 28.3131 | 0.97768 | 0.75841 | 1.0857 | 1.1346 | 0.57953 | 1.4424 | 1.4076 | 1.4634 | 1.2984 | 0.91025 | 0.93323 | 0.80241 | 0.35665 | 0.60995 | 1.0957 | -0.12845 | 0.8035 | 1.0685 | 24.0593 | 2.4329 | b2 | -20.5033 | -0.61399 | -0.86136 | -0.85124 | -0.99494 | -0.56163 | -0.67392 | -1.0576 | -0.5963 | -1.2801 | -1.1266 | -0.66507 | -1.0141 | -0.77576 | -0.93536 | -0.3666 | -1.1165 | -0.82746 | -1.3926 | -21.7988 | -2.095
b3 | 41.8958 | 3.0047 | 3.066 | 2.5838 | 3.0266 | 1.6347 | 2.8966 | 3.055 | 3.2056 | 2.372 | 2.0657 | 2.8188 | 2.8097 | 1.6352 | 2.3923 | 3.3369 | 3.4044 | 3.2057 | 3.4425 | 9.6293 | 4.1139 | b3 | 41.5112 | 2.4285 | 2.7829 | 2.0518 | 2.9816 | 1.8784 | 3.2339 | 3.5442 | 3.375 | 2.2308 | 1.7237 | 3.554 | 2.1858 | 0.59245 | 1.0303 | 3.2784 | 3.4588 | 2.8518 | 3.2824 | 9.5136 | 3.8977
b4 | 0.51172 | 0.095798 | 0.05265 | 0.10873 | 0.20675 | 0.0332 | -0.049683 | 0.048058 | 0.0033524 | 0.023882 | -0.006796 | 0.0017793 | 0.080858 | -0.06686 | -0.056673 | -0.12031 | -0.14761 | -0.088197 | 0.061776 | 0.34016 | 0.034773 | b4 | -1.2976 | 0.18002 | 0.22338 | 0.20272 | 0.10527 | 0.029037 | 0.073216 | -0.061408 | 0.24698 | 0.15704 | 0.1023 | 0.31315 | 0.27438 | -0.053553 | -0.16868 | 0.12076 | 0.18307 | 0.12777 | -0.094799 | -1.5875 | 0.015039
b5 | -4.8955 | -0.45214 | -0.40061 | -0.025152 | 0.029564 | 0.28504 | 0.15664 | 0.23123 | -0.016876 | 0.13033 | 0.41014 | 0.1726 | 0.14267 | 0.042591 | -0.0010092 | -0.11837 | -0.19142 | -0.066616 | -0.14355 | -3.159 | -0.22979 | b5 | -4.9289 | -0.40265 | -0.56329 | -0.014178 | -0.15376 | 0.094639 | 0.018752 | -0.015904 | -0.24654 | -0.0054604 | 0.22481 | 0.13263 | 0.11562 | 0.2192 | 0.0078287 | -0.21567 | -0.30254 | -0.043873 | -0.2879 | -3.0566 | -0.31154
b6 | 0.64734 | 0.0074707 | 0.022631 | 0.0048091 | -0.045033 | 0.02676 | -0.031432 | -0.10418 | -0.083182 | -0.037686 | -0.038877 | -0.036732 | -0.07836 | -0.049171 | -0.11966 | -0.098651 | -0.055663 | -0.087154 | -0.083176 | 0.00512 | -0.027576 | b6 | 0.72219 | 0.0087073 | 0.019697 | 0.015891 | 0.019393 | 0.037821 | -0.0022735 | 0.059449 | -0.039838 | -0.028505 | -0.0050795 | 0.06099 | 0.024828 | 0.052887 | 0.081214 | 0.12168 | 0.0083821 | -0.13696 | -0.05555 | 0.0044314 | 0.033119
b7 | 2.5988 | 0.8163 | 0.82749 | 0.93267 | 0.88728 | 0.91447 | 0.88502 | 0.89279 | 0.89702 | 0.90123 | 0.91416 | 0.95156 | 0.94092 | 0.91145 | 0.90546 | 0.92883 | 0.93057 | 0.89904 | 0.93424 | 0.8149 | 0.94984 | b7 | 2.6759 | 0.81107 | 0.87847 | 0.95426 | 0.91779 | 0.95647 | 0.94374 | 0.98158 | 0.97666 | 0.96359 | 0.96365 | 0.94796 | 0.90783 | 0.93182 | 0.90223 | 0.95101 | 0.88528 | 0.96647 | 0.97098 | 0.88989 | 0.9815
b8 | 0.14835 | 0.026588 | 0.018757 | 0.02589 | 0.018723 | 0.015265 | -0.0083043 | 0.021544 | 0.024717 | -0.0040326 | -0.004939 | -0.0005332 | -0.016759 | -0.0068555 | -0.032875 | -0.007979 | 0.01367 | 0.008852 | -0.012321 | 0.022799 | 0.0091219 | b8 | 0.19307 | 0.013107 | -0.014125 | 0.02737 | 0.0022249 | 0.0047953 | -0.021024 | -0.021819 | -0.00035208 | -0.017688 | -0.0031762 | 0.03474 | -0.028275 | 0.0022052 | -0.010338 | -0.0075453 | -0.024351 | -0.012826 | -0.030056 | 0.032494 | 0.00087393
b9 | 0.43586 | 0.40908 | 0.39375 | 0.39107 | 0.41914 | 0.3688 | 0.38956 | 0.38612 | 0.40955 | 0.39979 | 0.38033 | 0.38062 | 0.39396 | 0.39362 | 0.39737 | 0.40368 | 0.39383 | 0.40126 | 0.40557 | 0.38161 | 0.39613 | b9 | 0.40136 | 0.3851 | 0.38519 | 0.39375 | 0.41225 | 0.37001 | 0.39334 | 0.40062 | 0.41035 | 0.38016 | 0.38265 | 0.3855 | 0.39807 | 0.37798 | 0.37222 | 0.41602 | 0.39958 | 0.40778 | 0.40395 | 0.40353 | 0.3936
b10 | 0.17544 | 0.037582 | 0.021126 | 0.015385 | -0.0083419 | 0.032009 | -0.00279 | -0.022597 | 0.01192 | 0.0039484 | -0.00089817 | -0.010116 | -0.07179 | -0.0080063 | -0.065191 | -0.0038047 | 0.021909 | -0.037783 | -0.060461 | 0.085808 | 0 | b10 | 0.17969 | -0.011523 | -0.014691 | 0.016432 | -0.0031633 | 0.0086911 | -0.021559 | -0.016781 | 0.019553 | -0.0057732 | 0.0045493 | 0.018501 | -0.065626 | 0.037336 | -0.0031618 | 0.021505 | 0.010447 | -0.065672 | -0.085408 | 0.097096 | 0
b11 | 0.56352 | 0.63895 | 0.63935 | 0.64592 | 0.64501 | 0.65059 | 0.6428 | 0.64246 | 0.63906 | 0.64501 | 0.64828 | 0.63925 | 0.63888 | 0.64818 | 0.65064 | 0.64057 | 0.63592 | 0.64284 | 0.63789 | 0.60289 | 0.63934 | b11 | 0.57966 | 0.64008 | 0.63776 | 0.65108 | 0.64533 | 0.65049 | 0.64292 | 0.64238 | 0.64406 | 0.64389 | 0.65504 | 0.64795 | 0.65275 | 0.6651 | 0.66066 | 0.64369 | 0.64265 | 0.64175 | 0.63529 | 0.59307 | 0.64322
b12 | 0.029808 | 0.0014944 | 0.00063307 | 0.0036724 | -0.0014352 | 0.004086 | 0.0011355 | 0.0015704 | 0.006022 | 0.0031758 | 0.00062534 | 0.00069136 | -0.008407 | 0.00074045 | -0.0058387 | 0.0017353 | 0.0029901 | -0.0012213 | -0.0061754 | 0.008217 | 0.0014006 | b12 | 0.039931 | 0.0014943 | -0.0025253 | 0.0026003 | -0.0037998 | -0.00089767 | -0.0049572 | -0.0079847 | -0.0028371 | -0.0028559 | -0.0028905 | 0.003967 | -0.011031 | 0.0044629 | 0.001594 | 0.0018072 | -0.002094 | -0.0072604 | -0.0091319 | 0.0026431 | -0.0010036
b13 | 0.078994 | 0.061107 | 0.058993 | 0.057153 | 0.056501 | 0.054425 | 0.056081 | 0.054732 | 0.055053 | 0.056008 | 0.055395 | 0.056883 | 0.056436 | 0.057105 | 0.05998 | 0.05783 | 0.058621 | 0.057572 | 0.058907 | 0.050757 | 0.057582 | b13 | 0.088482 | 0.060897 | 0.061865 | 0.056948 | 0.059669 | 0.058272 | 0.057433 | 0.057452 | 0.05946 | 0.058364 | 0.056139 | 0.057889 | 0.056716 | 0.055912 | 0.056563 | 0.057041 | 0.058052 | 0.057267 | 0.059841 | 0.046713 | 0.058588
b14 | -0.0087969 | -0.00033398 | -0.0010576 | -0.001493 | -0.0032877 | -0.00072166 | -0.0014192 | -0.00234 | -0.00011678 | -0.001204 | -0.002073 | -0.0019038 | -0.005494 | -0.0019563 | -0.0045211 | -0.0017567 | -0.00088037 | -0.0048301 | -0.0050882 | -0.0035459 | -0.0024776 | b14 | -0.0044692 | -0.007425 | -0.0064376 | -0.0042255 | -0.0051398 | -0.0033419 | -0.0068423 | -0.0061296 | -0.0029765 | -0.0043048 | -0.0055109 | -0.0039439 | -0.0094325 | -0.0028889 | -0.0046319 | -0.0022564 | -0.0040074 | -0.0098591 | -0.0087269 | -0.0057451 | -0.0054307
b15 | -0.01753 | 0.02281 | 0.022134 | 0.022083 | 0.021843 | 0.021853 | 0.021189 | 0.020805 | 0.021831 | 0.021902 | 0.023247 | 0.021002 | 0.017876 | 0.020319 | 0.017747 | 0.021283 | 0.021652 | 0.023516 | 0.018997 | 0.0087786 | 0.019724 | b15 | -0.012639 | 0.029993 | 0.026348 | 0.028991 | 0.025672 | 0.028684 | 0.027525 | 0.024225 | 0.028952 | 0.026136 | 0.031173 | 0.029769 | 0.025684 | 0.030075 | 0.026731 | 0.027845 | 0.028365 | 0.028247 | 0.024143 | 0.0092155 | 0.025972
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 61.0442 | 12.8561 | 17.213 | 23.4501 | 17.9757 | 12.0365 | -4.3968 | -6.789 | -9.3932 | -1.2151 | -6.9716 | -8.0698 | -2.8006 | -10.0822 | -15.4576 | -8.3679 | -7.8412 | -14.5332 | -14.5731 | -11.0469 | 0 | a1 | 28.2894 | -17.4886 | -14.3702 | -4.2992 | -1.6319 | 0.62918 | 0.39498 | 3.3145 | 2.5471 | 7.6057 | 2.2736 | 2.2758 | 5.1759 | 3.651 | -0.28057 | 2.3752 | 2.594 | -3.3823 | -1.5575 | -8.2721 | 0
a2 | 0.57483 | -0.48138 | -0.19875 | -0.29298 | -0.43068 | -0.82355 | -0.7329 | -0.48108 | -0.61292 | -1.0508 | -0.72148 | -1.1199 | -1.5994 | -0.28176 | -0.75309 | -1.3166 | -1.5277 | -0.63118 | -1.5785 | 8.5154 | -0.48141 | a2 | 1.7196 | 0.18563 | -0.36269 | 0.45569 | 0.18875 | 0.34319 | 0.43025 | 1.0961 | 0.21131 | 1.0296 | 0.64843 | -0.022887 | -0.82978 | 0.21864 | 0.16837 | 0.80049 | 0.79461 | -0.1555 | 0.022692 | 8.5704 | 0.59561
a3 | -1.4731 | 0.63495 | 0.29851 | 0.2202 | 0.43304 | 0.67853 | 0.28247 | -0.11541 | 0.69123 | 0.33171 | 1.1345 | 0.4115 | 0.16111 | 0.10032 | 0.075589 | 0.41533 | 0.14964 | 0.49277 | 0.65525 | 0.049493 | 0.32742 | a3 | -1.352 | 0.42147 | 0.19002 | 0.32741 | 0.58622 | 0.44782 | -0.21134 | 0.12683 | 0.74587 | 0.60781 | 0.78722 | 0.4055 | 0.50221 | 0.66057 | 0.36056 | 0.83312 | 0.47205 | 0.47803 | 1.0905 | 0.33285 | 0.4338
a4 | 0.47614 | -0.25941 | -0.31726 | -0.27605 | -0.24855 | -0.43088 | -0.30564 | -0.46433 | -0.26288 | -0.59715 | -0.58284 | -0.66365 | -0.45917 | -0.63068 | -0.69532 | -0.27147 | -0.14249 | -0.32307 | -0.42973 | 0.37956 | -0.35883 | a4 | 0.3178 | -0.19806 | -0.10952 | -0.4371 | -0.20437 | -0.18999 | -0.49206 | 0.048403 | -0.14483 | -0.63926 | -0.56305 | -0.37985 | -0.61746 | -0.24729 | -0.35643 | -0.28784 | -0.20129 | -0.27691 | -0.2875 | 0.36383 | -0.271
a5 | 1.8638 | -0.10952 | -0.039875 | -0.0025195 | 0.34375 | 0.054883 | 0.0024697 | 0.076091 | 0.28509 | 0.18211 | 0.3138 | 0.071528 | 0.09921 | 0.067555 | 0.12769 | 0.15229 | 0.06537 | 0.18754 | 0.34399 | -0.037698 | 0.1684 | a5 | 1.9994 | 0.049671 | -0.018571 | 0.084971 | 0.289 | 0.19269 | 0.017999 | 0.090628 | 0.17074 | 0.14923 | 0.19845 | -0.042374 | -0.012429 | 0.082162 | 0.20352 | 0.17603 | 0.00036807 | 0.018109 | 0.21413 | -0.12541 | 0.14953
a6 | 0.45407 | 0.023732 | 0.010432 | -0.045085 | -0.20895 | -0.17032 | -0.21858 | -0.16065 | -0.13713 | -0.093503 | -0.15303 | -0.1732 | -0.11665 | -0.14458 | -0.18424 | -0.13208 | -0.14395 | -0.073063 | -0.030694 | -0.020994 | -0.099974 | a6 | 0.49069 | 0.046469 | 0.038129 | 0.0321 | 0.11195 | 0.12164 | 0.034001 | -0.0019927 | -0.094006 | 0.030764 | 0.065437 | -0.0016995 | -0.033614 | -0.0093207 | 0.026019 | -0.033629 | 0.054796 | -0.049715 | -0.074653 | -0.02196 | 0.026803
a7 | 1.5054 | -0.024287 | 0.011403 | 0.0065507 | 0.062424 | 0.041342 | 0.060514 | 0.042794 | 0.062783 | 0.048065 | 0.044361 | -0.027374 | 0.034018 | -0.0089638 | -0.010082 | 0.029417 | 0.048352 | 0.059052 | 0.050167 | -0.037701 | 0.06979 | a7 | 1.4114 | -0.076039 | -0.034534 | -0.10462 | 0.0055149 | 0.041489 | -0.085269 | 0.051722 | 0.048207 | 0.039667 | 0.04217 | -0.025202 | 0.032148 | 0.044696 | -0.029428 | 0.028212 | -0.0082253 | 0.028182 | 0.0099105 | -0.0084944 | 0.039854
a8 | -0.0489 | -0.0073361 | -0.0098433 | -0.0018745 | -0.07078 | -0.042206 | -0.03539 | -0.0042518 | -0.009339 | -0.010159 | -0.065773 | -0.046196 | -0.012208 | -0.015946 | -0.05492 | -0.040126 | -0.030056 | -0.069576 | -0.013278 | -0.049716 | -0.031122 | a8 | 0.016504 | -0.0064911 | 0.0058046 | 0.0064524 | 0.017081 | 0.0017491 | -0.028738 | -0.032085 | -0.077279 | -0.032272 | -0.042156 | -0.049077 | -0.089081 | -0.097346 | -0.088627 | -0.05128 | -0.022574 | -0.050749 | -0.030159 | 0.0076337 | -0.034112
a9 | -0.17361 | 0.0093979 | 0.011997 | 0.0075219 | 0.029588 | 0.010034 | 0.0030195 | 0.019362 | 0.0077342 | 0.0050853 | 0.031594 | -0.0067147 | 0.0036052 | -0.0032987 | -0.013642 | 0.010598 | 0.01425 | 0.023302 | 0.021467 | -0.021784 | 0.0039849 | a9 | -0.24086 | -0.0039382 | -0.01578 | -0.0013551 | 0.00098179 | -0.0050126 | -0.036158 | 0.0016499 | -0.0068371 | -0.015077 | 0.013918 | -0.01808 | -0.021928 | -0.029285 | -0.012194 | -0.0059537 | -0.029628 | 0.0014443 | -0.013276 | -0.01051 | -0.017368
a10 | 0.10294 | 0.0047874 | 0.0085762 | 0.00154 | -0.020705 | -0.0077735 | 0.015218 | -0.00518 | 0.015672 | 0.010947 | -0.037192 | -0.0077468 | 0.010698 | 0.010658 | 0.0091888 | -0.015921 | -0.01257 | -0.035414 | 0.022259 | -0.039539 | 0 | a10 | 0.15289 | -3.8706e-05 | 0.017825 | 0.0050081 | 0.020651 | 0.0014928 | -0.01514 | 0.013925 | -0.037559 | 0.013306 | -0.02378 | -0.013845 | -0.008258 | -0.025134 | 0.0064517 | -0.0029191 | -0.0086761 | -0.025139 | -0.0013124 | 0.001375 | 0
a11 | 0.19143 | -0.0082596 | 0.0025075 | 0.0084068 | 0.017967 | 0.0018782 | -0.0083492 | -0.0053592 | -0.0047101 | -0.00037812 | -0.0011206 | -0.016105 | -0.0042808 | -0.015031 | -0.015243 | -0.0098709 | -0.0067324 | -0.011018 | -0.0086978 | -0.012753 | 0.00069483 | a11 | 0.1534 | -0.037993 | -0.02565 | -0.014717 | -0.0070817 | -0.0051365 | -0.014888 | -0.0026699 | -0.0059182 | -0.010065 | -0.0065501 | -0.019266 | -0.0095608 | -0.018066 | -0.021977 | -0.0096258 | -0.012769 | -0.01637 | -0.013108 | -0.0098004 | -0.0091282
a12 | 0.0089716 | 0.0019283 | 0.0024854 | -0.0016884 | -0.0095339 | -0.0072384 | -0.0044685 | -0.0055216 | -0.0027687 | -0.0033893 | -0.0083748 | -0.0056009 | -0.0050136 | -0.0035697 | -0.0074154 | -0.0071277 | -0.0055418 | -0.0064177 | -0.0014014 | -0.00774 | -0.0041967 | a12 | 0.013528 | -0.00055518 | 0.00050617 | -0.0011983 | 0.0014727 | 0.0021707 | -0.0015796 | 0.00045556 | -0.0075116 | 0.0015768 | -0.0020236 | -0.00444 | -0.0050489 | -0.0063046 | -0.0023947 | -0.0030434 | -0.0023267 | -0.005808 | -0.0035141 | -0.0059597 | -0.0018768
a13 | -0.007899 | 0.0015414 | 0.00051823 | -0.00020234 | 0.0073324 | 0.002597 | 0.00094179 | 0.001485 | 0.0025384 | 0.0029348 | 0.0062276 | 0.0011715 | 0.0016197 | 0.00098066 | 0.00016503 | 0.0017077 | 0.0024836 | 0.0039755 | 0.0054374 | 0.0014653 | 0.0020881 | a13 | -0.017202 | 0.0011224 | -0.0011323 | -0.0019065 | 0.0012289 | 0.002328 | -0.002015 | 0.0014838 | 0.0016925 | 0.0021931 | 0.0046654 | 0.00012862 | 0.0018346 | 0.0015906 | 0.0016664 | 0.0024281 | -0.001393 | 0.0025287 | 0.0042696 | 0.00042013 | 0.00071586
a14 | 0.014272 | -0.0085387 | -0.0072319 | -0.00712 | -0.0093939 | -0.0078872 | -0.0067268 | -0.0056227 | -0.0066621 | -0.0081932 | -0.0098045 | -0.0077954 | -0.0063583 | -0.0068899 | -0.0080473 | -0.0090355 | -0.0078891 | -0.01127 | -0.0069769 | -0.0031284 | -0.007074 | a14 | 0.013434 | -0.012603 | -0.010062 | -0.011813 | -0.0096041 | -0.012193 | -0.012406 | -0.0092973 | -0.014054 | -0.012887 | -0.016867 | -0.015105 | -0.016277 | -0.016646 | -0.01478 | -0.013903 | -0.013095 | -0.013838 | -0.011755 | -0.0048578 | -0.012163
a15 | 0.012127 | -0.0021101 | -0.0021632 | -0.0017996 | -0.0033038 | 0.00062906 | -0.0021418 | -0.0034756 | -0.00069021 | -0.0026407 | -0.0015139 | -0.0027826 | -0.0057643 | -0.0024559 | -0.0062062 | -0.0035225 | -0.0025685 | -0.0059498 | -0.0032151 | 0.0034144 | -0.0022433 | a15 | 0.011835 | -0.0083127 | -0.0081062 | -0.0047719 | -0.0067566 | -0.003593 | -0.0087075 | -0.0071056 | -0.0035493 | -0.0046127 | -0.005899 | -0.0044293 | -0.01021 | -0.0025543 | -0.0063317 | -0.0029935 | -0.0059698 | -0.01203 | -0.0094707 | 0.004039 | -0.0055635
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00057247 | -9.9807e-05 | -5.6257e-05 | -4.0541e-05 | 2.3475e-05 | -8.3578e-05 | 0 | 5.9666e-05 | -3.1403e-05 | -1.039e-05 | 0 | 2.6363e-05 | 0.00019134 | 2.1638e-05 | 0.00017087 | 0 | -5.8911e-05 | 9.8237e-05 | 0.00016197 | -0.00024417 | 0.00017292 | Dx (m) | -0.00060703 | 3.0488e-05 | 4.1004e-05 | -4.2591e-05 | 0 | -2.2681e-05 | 5.6048e-05 | 4.4564e-05 | -5.2498e-05 | 1.5786e-05 | -1.2427e-05 | -4.9572e-05 | 0.00017067 | -9.7081e-05 | 0 | -5.6897e-05 | -2.8078e-05 | 0.00017216 | 0.00022853 | -0.00027818 | 0.000169
Dy (m) | -0.00011897 | -1.4036e-05 | -2.2586e-05 | 0 | 5.3909e-05 | 2.0582e-05 | -4.0147e-05 | 1.4211e-05 | -4.1931e-05 | -2.886e-05 | 9.7499e-05 | 2.1266e-05 | -2.712e-05 | -2.7451e-05 | -1.9907e-05 | 4.2391e-05 | 3.299e-05 | 9.5337e-05 | -5.7126e-05 | 0.00010664 | 0.00046474 | Dy (m) | -0.0002945 | 0 | -4.5867e-05 | -1.4039e-05 | -5.4305e-05 | 0 | 4.134e-05 | -3.6678e-05 | 9.8614e-05 | -3.4884e-05 | 6.1587e-05 | 3.4859e-05 | 2.4108e-05 | 6.1638e-05 | -1.6313e-05 | 0 | 2.2392e-05 | 7.1107e-05 | 0 | 0 | 0.00064231
Magnetic length (m) |  | 14.3569 | [19 empty] | Magnetic length (m) |  | 14.3569
Transfer funct. (mT/KA) |  | 704.4172 | [19 empty] | Transfer funct. (mT/KA) |  | 704.4733
Coil positioning (mm) |  | 0.016513 | [19 empty] | Coil positioning (mm) |  | 0.018673
Current (A) |  | 10 | [19 empty] | Current (A) |  | 10
Soft version | DIMM V2.1.0 | [20 empty] | Soft version | DIMM V2.1.0
Level Average | -0.37414 | -0.26487 | -0.15569 | -0.22576 | -0.14891 | -0.16137 | -0.35844 | -0.33034 | -0.10791 | 0.92995 | 0.10104 | -0.04583 | 0.040728 | 0.27746 | 0.27475 | 0.24846 | 0.22964 | 0.10344 | 0.39039 | 0.096389 |  | Level Average | -0.056068 | -0.12477 | -0.10323 | -0.092938 | -0.066798 | -0.11817 | -0.013086 | -0.016149 | 0.019328 | 0.033876 | -0.021105 | -0.031454 | -0.063127 | -0.03267 | -0.25318 | -0.059183 | -0.004132 | -0.11716 | 0.064842 | -0.11926
Level Sigma2 | 0.1312 | 0.28391 | 0.41286 | 0.63475 | 1.3641 | 0.92576 | 0.12784 | 0.14406 | 0.29633 | 0.009783 | 0.084468 | 0.61308 | 0.093432 | 0.001977 | 0.060843 | 0.016112 | 0.17186 | 0.15939 | 0.052771 | 0.29754 |  | Level Sigma2 | 0.28162 | 0.8571 | 0.69417 | 0.44834 | 0.27805 | 0.20404 | 0.002469 | 0.009463 | 0.001156 | 0.000173 | 0.13657 | 0.10546 | 0.25018 | 0.03466 | 0.027252 | 0.59746 | 0.186 | 0.55094 | 0.058221 | 1.4119
Position (mm) | 0 | 748 | 1496 | 2248 | 2997 | 3746 | 4497 | 5247 | 5997 | 6746 | 7497 | 8247 | 8996 | 9746 | 10496 | 11246 | 11996 | 12745 | 13495 | 14245 |  | Position (mm) | 0 | 745 | 1494 | 2244 | 2994 | 3743 | 4492 | 5243 | 5994 | 6744 | 7494 | 8244 | 8994 | 9744 | 10494 | 11244 | 11993 | 12743 | 13493 | 14243
Fist move (mm) | First move Ap1 | 875.1815 | First move Ap2 | 869.769 | [17 empty] | Fist move (mm) | First move Ap1 |  | First move Ap2

## Summary Data
File | HCMBALA001-03000199 |  |  | Aperture 1 - Cold mass | [16 empty] | Integrals | File | HCMBALA001-3000199 |  |  | Aperture 2 - Cold mass | [16 empty] | Integrals
C1/i (mT/KA) | 387.6332 | 703.7956 | 704.3803 | 704.4899 | 704.3919 | 704.4531 | 704.5402 | 704.5162 | 704.4522 | 704.4366 | 704.3936 | 704.4927 | 704.4482 | 704.4873 | 704.5442 | 704.4441 | 704.496 | 704.5028 | 704.2453 | 421.877 | 101.1546 | C1/i (mT/KA) | 379.2533 | 703.7763 | 704.4735 | 704.5888 | 704.4471 | 704.4583 | 704.4967 | 704.5241 | 704.5156 | 704.514 | 704.4843 | 704.6076 | 704.51 | 704.5782 | 704.639 | 704.5406 | 704.5571 | 704.5187 | 704.2894 | 432.2517 | 101.1755
Angle (mrad) | 6.1044 | 1.2856 | 1.7213 | 2.345 | 1.7976 | 1.2036 | -0.43968 | -0.6789 | -0.93932 | -0.12151 | -0.69716 | -0.80698 | -0.28006 | -1.0082 | -1.5458 | -0.83679 | -0.78412 | -1.4533 | -1.4573 | -1.1047 | -1.1025 | Angle (mrad) | 2.8289 | -1.7489 | -1.437 | -0.42992 | -0.16319 | 0.062918 | 0.039498 | 0.33145 | 0.25471 | 0.76057 | 0.22736 | 0.22758 | 0.51759 | 0.3651 | -0.028057 | 0.23752 | 0.2594 | -0.33822 | -0.15575 | -0.82721 | -0.14226
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 |  | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | 28.3131 | 0.97768 | 0.75841 | 1.0857 | 1.1346 | 0.57953 | 1.4424 | 1.4076 | 1.4634 | 1.2984 | 0.91025 | 0.93323 | 0.80241 | 0.35665 | 0.60995 | 1.0957 | -0.12845 | 0.8035 | 1.0685 | 24.0593 |  | b2 | -20.5033 | -0.61399 | -0.86136 | -0.85124 | -0.99494 | -0.56163 | -0.67392 | -1.0576 | -0.5963 | -1.2801 | -1.1266 | -0.66507 | -1.0141 | -0.77576 | -0.93536 | -0.3666 | -1.1165 | -0.82746 | -1.3926 | -21.7988
b3 | 41.8958 | 3.0047 | 3.066 | 2.5838 | 3.0266 | 1.6347 | 2.8966 | 3.055 | 3.2056 | 2.372 | 2.0657 | 2.8188 | 2.8097 | 1.6352 | 2.3923 | 3.3369 | 3.4044 | 3.2057 | 3.4425 | 9.6293 |  | b3 | 41.5112 | 2.4285 | 2.7829 | 2.0518 | 2.9816 | 1.8784 | 3.2339 | 3.5442 | 3.375 | 2.2308 | 1.7237 | 3.554 | 2.1858 | 0.59245 | 1.0303 | 3.2784 | 3.4588 | 2.8518 | 3.2824 | 9.5136
b4 | 0.51172 | 0.095798 | 0.05265 | 0.10873 | 0.20675 | 0.0332 | -0.049683 | 0.048058 | 0.0033524 | 0.023882 | -0.006796 | 0.0017793 | 0.080858 | -0.06686 | -0.056673 | -0.12031 | -0.14761 | -0.088197 | 0.061776 | 0.34016 |  | b4 | -1.2976 | 0.18002 | 0.22338 | 0.20272 | 0.10527 | 0.029037 | 0.073216 | -0.061408 | 0.24698 | 0.15704 | 0.1023 | 0.31315 | 0.27438 | -0.053553 | -0.16868 | 0.12076 | 0.18307 | 0.12777 | -0.094799 | -1.5875
b5 | -4.8955 | -0.45214 | -0.40061 | -0.025152 | 0.029564 | 0.28504 | 0.15664 | 0.23123 | -0.016876 | 0.13033 | 0.41014 | 0.1726 | 0.14267 | 0.042591 | -0.0010092 | -0.11837 | -0.19142 | -0.066616 | -0.14355 | -3.159 |  | b5 | -4.9289 | -0.40265 | -0.56329 | -0.014178 | -0.15376 | 0.094639 | 0.018752 | -0.015904 | -0.24654 | -0.0054604 | 0.22481 | 0.13263 | 0.11562 | 0.2192 | 0.0078287 | -0.21567 | -0.30254 | -0.043873 | -0.2879 | -3.0566
b6 | 0.64734 | 0.0074707 | 0.022631 | 0.0048091 | -0.045033 | 0.02676 | -0.031432 | -0.10418 | -0.083182 | -0.037686 | -0.038877 | -0.036732 | -0.07836 | -0.049171 | -0.11966 | -0.098651 | -0.055663 | -0.087154 | -0.083176 | 0.00512 |  | b6 | 0.72219 | 0.0087073 | 0.019697 | 0.015891 | 0.019393 | 0.037821 | -0.0022735 | 0.059449 | -0.039838 | -0.028505 | -0.0050795 | 0.06099 | 0.024828 | 0.052887 | 0.081214 | 0.12168 | 0.0083821 | -0.13696 | -0.05555 | 0.0044314
b7 | 2.5988 | 0.8163 | 0.82749 | 0.93267 | 0.88728 | 0.91447 | 0.88502 | 0.89279 | 0.89702 | 0.90123 | 0.91416 | 0.95156 | 0.94092 | 0.91145 | 0.90546 | 0.92883 | 0.93057 | 0.89904 | 0.93424 | 0.8149 |  | b7 | 2.6759 | 0.81107 | 0.87847 | 0.95426 | 0.91779 | 0.95647 | 0.94374 | 0.98158 | 0.97666 | 0.96359 | 0.96365 | 0.94796 | 0.90783 | 0.93182 | 0.90223 | 0.95101 | 0.88528 | 0.96647 | 0.97098 | 0.88989
b8 | 0.14835 | 0.026588 | 0.018757 | 0.02589 | 0.018723 | 0.015265 | -0.0083043 | 0.021544 | 0.024717 | -0.0040326 | -0.004939 | -0.0005332 | -0.016759 | -0.0068555 | -0.032875 | -0.007979 | 0.01367 | 0.008852 | -0.012321 | 0.022799 |  | b8 | 0.19307 | 0.013107 | -0.014125 | 0.02737 | 0.0022249 | 0.0047953 | -0.021024 | -0.021819 | -0.00035208 | -0.017688 | -0.0031762 | 0.03474 | -0.028275 | 0.0022052 | -0.010338 | -0.0075453 | -0.024351 | -0.012826 | -0.030056 | 0.032494
b9 | 0.43586 | 0.40908 | 0.39375 | 0.39107 | 0.41914 | 0.3688 | 0.38956 | 0.38612 | 0.40955 | 0.39979 | 0.38033 | 0.38062 | 0.39396 | 0.39362 | 0.39737 | 0.40368 | 0.39383 | 0.40126 | 0.40557 | 0.38161 |  | b9 | 0.40136 | 0.3851 | 0.38519 | 0.39375 | 0.41225 | 0.37001 | 0.39334 | 0.40062 | 0.41035 | 0.38016 | 0.38265 | 0.3855 | 0.39807 | 0.37798 | 0.37222 | 0.41602 | 0.39958 | 0.40778 | 0.40395 | 0.40353
b10 | 0.17544 | 0.037582 | 0.021126 | 0.015385 | -0.0083419 | 0.032009 | -0.00279 | -0.022597 | 0.01192 | 0.0039484 | -0.00089817 | -0.010116 | -0.07179 | -0.0080063 | -0.065191 | -0.0038047 | 0.021909 | -0.037783 | -0.060461 | 0.085808 |  | b10 | 0.17969 | -0.011523 | -0.014691 | 0.016432 | -0.0031633 | 0.0086911 | -0.021559 | -0.016781 | 0.019553 | -0.0057732 | 0.0045493 | 0.018501 | -0.065626 | 0.037336 | -0.0031618 | 0.021505 | 0.010447 | -0.065672 | -0.085408 | 0.097096
b11 | 0.56352 | 0.63895 | 0.63935 | 0.64592 | 0.64501 | 0.65059 | 0.6428 | 0.64246 | 0.63906 | 0.64501 | 0.64828 | 0.63925 | 0.63888 | 0.64818 | 0.65064 | 0.64057 | 0.63592 | 0.64284 | 0.63789 | 0.60289 |  | b11 | 0.57966 | 0.64008 | 0.63776 | 0.65108 | 0.64533 | 0.65049 | 0.64292 | 0.64238 | 0.64406 | 0.64389 | 0.65504 | 0.64795 | 0.65275 | 0.6651 | 0.66066 | 0.64369 | 0.64265 | 0.64175 | 0.63529 | 0.59307
b12 | 0.029808 | 0.0014944 | 0.00063307 | 0.0036724 | -0.0014352 | 0.004086 | 0.0011355 | 0.0015704 | 0.006022 | 0.0031758 | 0.00062534 | 0.00069136 | -0.008407 | 0.00074045 | -0.0058387 | 0.0017353 | 0.0029901 | -0.0012213 | -0.0061754 | 0.008217 |  | b12 | 0.039931 | 0.0014943 | -0.0025253 | 0.0026003 | -0.0037998 | -0.00089767 | -0.0049572 | -0.0079847 | -0.0028371 | -0.0028559 | -0.0028905 | 0.003967 | -0.011031 | 0.0044629 | 0.001594 | 0.0018072 | -0.002094 | -0.0072604 | -0.0091319 | 0.0026431
b13 | 0.078994 | 0.061107 | 0.058993 | 0.057153 | 0.056501 | 0.054425 | 0.056081 | 0.054732 | 0.055053 | 0.056008 | 0.055395 | 0.056883 | 0.056436 | 0.057105 | 0.05998 | 0.05783 | 0.058621 | 0.057572 | 0.058907 | 0.050757 |  | b13 | 0.088482 | 0.060897 | 0.061865 | 0.056948 | 0.059669 | 0.058272 | 0.057433 | 0.057452 | 0.05946 | 0.058364 | 0.056139 | 0.057889 | 0.056716 | 0.055912 | 0.056563 | 0.057041 | 0.058052 | 0.057267 | 0.059841 | 0.046713
b14 | -0.0087969 | -0.00033398 | -0.0010576 | -0.001493 | -0.0032877 | -0.00072166 | -0.0014192 | -0.00234 | -0.00011678 | -0.001204 | -0.002073 | -0.0019038 | -0.005494 | -0.0019563 | -0.0045211 | -0.0017567 | -0.00088037 | -0.0048301 | -0.0050882 | -0.0035459 |  | b14 | -0.0044692 | -0.007425 | -0.0064376 | -0.0042255 | -0.0051398 | -0.0033419 | -0.0068423 | -0.0061296 | -0.0029765 | -0.0043048 | -0.0055109 | -0.0039439 | -0.0094325 | -0.0028889 | -0.0046319 | -0.0022564 | -0.0040074 | -0.0098591 | -0.0087269 | -0.0057451
b15 | -0.01753 | 0.02281 | 0.022134 | 0.022083 | 0.021843 | 0.021853 | 0.021189 | 0.020805 | 0.021831 | 0.021902 | 0.023247 | 0.021002 | 0.017876 | 0.020319 | 0.017747 | 0.021283 | 0.021652 | 0.023516 | 0.018997 | 0.0087786 |  | b15 | -0.012639 | 0.029993 | 0.026348 | 0.028991 | 0.025672 | 0.028684 | 0.027525 | 0.024225 | 0.028952 | 0.026136 | 0.031173 | 0.029769 | 0.025684 | 0.030075 | 0.026731 | 0.027845 | 0.028365 | 0.028247 | 0.024143 | 0.0092155
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 61.0442 | 12.8561 | 17.213 | 23.4501 | 17.9757 | 12.0365 | -4.3968 | -6.789 | -9.3932 | -1.2151 | -6.9716 | -8.0698 | -2.8006 | -10.0822 | -15.4576 | -8.3679 | -7.8412 | -14.5332 | -14.5731 | -11.0469 |  | a1 | 28.2894 | -17.4886 | -14.3702 | -4.2992 | -1.6319 | 0.62918 | 0.39498 | 3.3145 | 2.5471 | 7.6057 | 2.2736 | 2.2758 | 5.1759 | 3.651 | -0.28057 | 2.3752 | 2.594 | -3.3823 | -1.5575 | -8.2721
a2 | 0.57483 | -0.48138 | -0.19875 | -0.29298 | -0.43068 | -0.82355 | -0.7329 | -0.48108 | -0.61292 | -1.0508 | -0.72148 | -1.1199 | -1.5994 | -0.28176 | -0.75309 | -1.3166 | -1.5277 | -0.63118 | -1.5785 | 8.5154 |  | a2 | 1.7196 | 0.18563 | -0.36269 | 0.45569 | 0.18875 | 0.34319 | 0.43025 | 1.0961 | 0.21131 | 1.0296 | 0.64843 | -0.022887 | -0.82978 | 0.21864 | 0.16837 | 0.80049 | 0.79461 | -0.1555 | 0.022692 | 8.5704
a3 | -1.4731 | 0.63495 | 0.29851 | 0.2202 | 0.43304 | 0.67853 | 0.28247 | -0.11541 | 0.69123 | 0.33171 | 1.1345 | 0.4115 | 0.16111 | 0.10032 | 0.075589 | 0.41533 | 0.14964 | 0.49277 | 0.65525 | 0.049493 |  | a3 | -1.352 | 0.42147 | 0.19002 | 0.32741 | 0.58622 | 0.44782 | -0.21134 | 0.12683 | 0.74587 | 0.60781 | 0.78722 | 0.4055 | 0.50221 | 0.66057 | 0.36056 | 0.83312 | 0.47205 | 0.47803 | 1.0905 | 0.33285
a4 | 0.47614 | -0.25941 | -0.31726 | -0.27605 | -0.24855 | -0.43088 | -0.30564 | -0.46433 | -0.26288 | -0.59715 | -0.58284 | -0.66365 | -0.45917 | -0.63068 | -0.69532 | -0.27147 | -0.14249 | -0.32307 | -0.42973 | 0.37956 |  | a4 | 0.3178 | -0.19806 | -0.10952 | -0.4371 | -0.20437 | -0.18999 | -0.49206 | 0.048403 | -0.14483 | -0.63926 | -0.56305 | -0.37985 | -0.61746 | -0.24729 | -0.35643 | -0.28784 | -0.20129 | -0.27691 | -0.2875 | 0.36383
a5 | 1.8638 | -0.10952 | -0.039875 | -0.0025195 | 0.34375 | 0.054883 | 0.0024697 | 0.076091 | 0.28509 | 0.18211 | 0.3138 | 0.071528 | 0.09921 | 0.067555 | 0.12769 | 0.15229 | 0.06537 | 0.18754 | 0.34399 | -0.037698 |  | a5 | 1.9994 | 0.049671 | -0.018571 | 0.084971 | 0.289 | 0.19269 | 0.017999 | 0.090628 | 0.17074 | 0.14923 | 0.19845 | -0.042374 | -0.012429 | 0.082162 | 0.20352 | 0.17603 | 0.00036807 | 0.018109 | 0.21413 | -0.12541
a6 | 0.45407 | 0.023732 | 0.010432 | -0.045085 | -0.20895 | -0.17032 | -0.21858 | -0.16065 | -0.13713 | -0.093503 | -0.15303 | -0.1732 | -0.11665 | -0.14458 | -0.18424 | -0.13208 | -0.14395 | -0.073063 | -0.030694 | -0.020994 |  | a6 | 0.49069 | 0.046469 | 0.038129 | 0.0321 | 0.11195 | 0.12164 | 0.034001 | -0.0019927 | -0.094006 | 0.030764 | 0.065437 | -0.0016995 | -0.033614 | -0.0093207 | 0.026019 | -0.033629 | 0.054796 | -0.049715 | -0.074653 | -0.02196
a7 | 1.5054 | -0.024287 | 0.011403 | 0.0065507 | 0.062424 | 0.041342 | 0.060514 | 0.042794 | 0.062783 | 0.048065 | 0.044361 | -0.027374 | 0.034018 | -0.0089638 | -0.010082 | 0.029417 | 0.048352 | 0.059052 | 0.050167 | -0.037701 |  | a7 | 1.4114 | -0.076039 | -0.034534 | -0.10462 | 0.0055149 | 0.041489 | -0.085269 | 0.051722 | 0.048207 | 0.039667 | 0.04217 | -0.025202 | 0.032148 | 0.044696 | -0.029428 | 0.028212 | -0.0082253 | 0.028182 | 0.0099105 | -0.0084944
a8 | -0.0489 | -0.0073361 | -0.0098433 | -0.0018745 | -0.07078 | -0.042206 | -0.03539 | -0.0042518 | -0.009339 | -0.010159 | -0.065773 | -0.046196 | -0.012208 | -0.015946 | -0.05492 | -0.040126 | -0.030056 | -0.069576 | -0.013278 | -0.049716 |  | a8 | 0.016504 | -0.0064911 | 0.0058046 | 0.0064524 | 0.017081 | 0.0017491 | -0.028738 | -0.032085 | -0.077279 | -0.032272 | -0.042156 | -0.049077 | -0.089081 | -0.097346 | -0.088627 | -0.05128 | -0.022574 | -0.050749 | -0.030159 | 0.0076337
a9 | -0.17361 | 0.0093979 | 0.011997 | 0.0075219 | 0.029588 | 0.010034 | 0.0030195 | 0.019362 | 0.0077342 | 0.0050853 | 0.031594 | -0.0067147 | 0.0036052 | -0.0032987 | -0.013642 | 0.010598 | 0.01425 | 0.023302 | 0.021467 | -0.021784 |  | a9 | -0.24086 | -0.0039382 | -0.01578 | -0.0013551 | 0.00098179 | -0.0050126 | -0.036158 | 0.0016499 | -0.0068371 | -0.015077 | 0.013918 | -0.01808 | -0.021928 | -0.029285 | -0.012194 | -0.0059537 | -0.029628 | 0.0014443 | -0.013276 | -0.01051
a10 | 0.10294 | 0.0047874 | 0.0085762 | 0.00154 | -0.020705 | -0.0077735 | 0.015218 | -0.00518 | 0.015672 | 0.010947 | -0.037192 | -0.0077468 | 0.010698 | 0.010658 | 0.0091888 | -0.015921 | -0.01257 | -0.035414 | 0.022259 | -0.039539 |  | a10 | 0.15289 | -3.8706e-05 | 0.017825 | 0.0050081 | 0.020651 | 0.0014928 | -0.01514 | 0.013925 | -0.037559 | 0.013306 | -0.02378 | -0.013845 | -0.008258 | -0.025134 | 0.0064517 | -0.0029191 | -0.0086761 | -0.025139 | -0.0013124 | 0.001375
a11 | 0.19143 | -0.0082596 | 0.0025075 | 0.0084068 | 0.017967 | 0.0018782 | -0.0083492 | -0.0053592 | -0.0047101 | -0.00037812 | -0.0011206 | -0.016105 | -0.0042808 | -0.015031 | -0.015243 | -0.0098709 | -0.0067324 | -0.011018 | -0.0086978 | -0.012753 |  | a11 | 0.1534 | -0.037993 | -0.02565 | -0.014717 | -0.0070817 | -0.0051365 | -0.014888 | -0.0026699 | -0.0059182 | -0.010065 | -0.0065501 | -0.019266 | -0.0095608 | -0.018066 | -0.021977 | -0.0096258 | -0.012769 | -0.01637 | -0.013108 | -0.0098004
a12 | 0.0089716 | 0.0019283 | 0.0024854 | -0.0016884 | -0.0095339 | -0.0072384 | -0.0044685 | -0.0055216 | -0.0027687 | -0.0033893 | -0.0083748 | -0.0056009 | -0.0050136 | -0.0035697 | -0.0074154 | -0.0071277 | -0.0055418 | -0.0064177 | -0.0014014 | -0.00774 |  | a12 | 0.013528 | -0.00055518 | 0.00050617 | -0.0011983 | 0.0014727 | 0.0021707 | -0.0015796 | 0.00045556 | -0.0075116 | 0.0015768 | -0.0020236 | -0.00444 | -0.0050489 | -0.0063046 | -0.0023947 | -0.0030434 | -0.0023267 | -0.005808 | -0.0035141 | -0.0059597
a13 | -0.007899 | 0.0015414 | 0.00051823 | -0.00020234 | 0.0073324 | 0.002597 | 0.00094179 | 0.001485 | 0.0025384 | 0.0029348 | 0.0062276 | 0.0011715 | 0.0016197 | 0.00098066 | 0.00016503 | 0.0017077 | 0.0024836 | 0.0039755 | 0.0054374 | 0.0014653 |  | a13 | -0.017202 | 0.0011224 | -0.0011323 | -0.0019065 | 0.0012289 | 0.002328 | -0.002015 | 0.0014838 | 0.0016925 | 0.0021931 | 0.0046654 | 0.00012862 | 0.0018346 | 0.0015906 | 0.0016664 | 0.0024281 | -0.001393 | 0.0025287 | 0.0042696 | 0.00042013
a14 | 0.014272 | -0.0085387 | -0.0072319 | -0.00712 | -0.0093939 | -0.0078872 | -0.0067268 | -0.0056227 | -0.0066621 | -0.0081932 | -0.0098045 | -0.0077954 | -0.0063583 | -0.0068899 | -0.0080473 | -0.0090355 | -0.0078891 | -0.01127 | -0.0069769 | -0.0031284 |  | a14 | 0.013434 | -0.012603 | -0.010062 | -0.011813 | -0.0096041 | -0.012193 | -0.012406 | -0.0092973 | -0.014054 | -0.012887 | -0.016867 | -0.015105 | -0.016277 | -0.016646 | -0.01478 | -0.013903 | -0.013095 | -0.013838 | -0.011755 | -0.0048578
a15 | 0.012127 | -0.0021101 | -0.0021632 | -0.0017996 | -0.0033038 | 0.00062906 | -0.0021418 | -0.0034756 | -0.00069021 | -0.0026407 | -0.0015139 | -0.0027826 | -0.0057643 | -0.0024559 | -0.0062062 | -0.0035225 | -0.0025685 | -0.0059498 | -0.0032151 | 0.0034144 |  | a15 | 0.011835 | -0.0083127 | -0.0081062 | -0.0047719 | -0.0067566 | -0.003593 | -0.0087075 | -0.0071056 | -0.0035493 | -0.0046127 | -0.005899 | -0.0044293 | -0.01021 | -0.0025543 | -0.0063317 | -0.0029935 | -0.0059698 | -0.01203 | -0.0094707 | 0.004039
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.56909 | 0.099809 | 0.05626 | 0.040538 | -0.023488 | 0.08358 | -0.0079 | -0.059675 | 0.031401 | 0.01039 | -0.0021867 | -0.026367 | -0.19121 | -0.021635 | -0.1708 | -0.0094437 | 0.058918 | -0.098284 | -0.16191 | 0.24421 |  | Dx (mm) | 0.60338 | -0.030489 | -0.041004 | 0.042588 | -0.008929 | 0.022681 | -0.056047 | -0.044561 | 0.052518 | -0.015788 | 0.012422 | 0.049576 | -0.17056 | 0.097103 | -0.0086786 | 0.056897 | 0.02808 | -0.17216 | -0.22838 | 0.27818
Dy (mm) | 0.11722 | 0.014028 | 0.022583 | 0.0035255 | -0.053917 | -0.020554 | 0.040143 | -0.014205 | 0.041922 | 0.028857 | -0.097533 | -0.021266 | 0.027186 | 0.027451 | 0.020007 | -0.042398 | -0.03298 | -0.095337 | 0.057114 | -0.10632 |  | Dy (mm) | 0.28872 | -0.0019125 | 0.045864 | 0.014039 | 0.054304 | 0.0040805 | -0.04133 | 0.036666 | -0.098654 | 0.034883 | -0.061592 | -0.034851 | -0.024005 | -0.061604 | 0.016313 | -0.0068587 | -0.022393 | -0.070986 | -0.0082242 | 0.0085384
Mag. len. (m) | 14.3569 | [20 empty] | Mag. len. (m) | 14.3569
TF (mT/KA) | 704.4172 | [20 empty] | TF (mT/KA) | 704.4733
Coil pos. (mm) | 0.016513 | [20 empty] | Coil pos. (mm) | 0.018673
Current (A) | 10 | [20 empty] | Current (A) | 10
Twist (m2 rad) | 0.071769 | [20 empty] | Twist (m2 rad) | -0.0068407

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 (with local feed-down) | [7 empty] | Measured harmonics- Aperture 2 (with local feed-down) | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 704.4172 | 0.17039 | -0.0070155 |  | -1.4996 | 12.3966 |  |  | 704.4733 | 0.18974 | 0.0028382 |  | -0.56635 | 6.3257 |  | 0 | 4
2 | 2.3498 |  | 0.9293 | 0.40716 | -0.48962 |  | -0.81108 | 0.45377 | -2.1818 |  | -0.86497 | 0.25399 | 0.56295 |  | 0.29539 | 0.47867 |  | 0 | 5
3 | 4.1122 |  | 2.7756 | 0.55697 | 0.32288 |  | 0.38903 | 0.29668 | 3.9056 |  | 2.5825 | 0.87791 | 0.43306 |  | 0.48943 | 0.29244 |  | 0 | 6
4 | 0.061235 |  | 0.010508 | 0.089443 | -0.3624 |  | -0.40666 | 0.16536 | 0.043731 |  | 0.10765 | 0.13568 | -0.26839 |  | -0.30948 | 0.18382 |  | 0 | 7
5 | -0.2331 |  | 0.009586 | 0.22214 | 0.16453 |  | 0.12225 | 0.13178 | -0.31436 |  | -0.080299 | 0.21823 | 0.14486 |  | 0.10299 | 0.09822 |  | 0 | 8
6 | -0.036137 |  | -0.04206 | 0.030865 | -0.10875 |  | -0.11747 | 0.069667 | 0.025369 |  | 0.020963 | 0.04726 | 0.014731 |  | 0.018956 | 0.051918 |  | 0 | 9
7 | 0.94878 |  | 0.90392 | 0.035564 | 0.069772 |  | 0.029124 | 0.030626 | 0.98019 |  | 0.93405 | 0.043626 | 0.038725 |  | 0.00011905 | 0.049521 |  | 0 | 10
8 | 0.0078766 |  | 0.00837 | 0.012643 | -0.028861 |  | -0.028959 | 0.018504 | -0.00071413 |  | -0.0014513 | 0.015658 | -0.031618 |  | -0.034798 | 0.030752 |  | 0 | 11
9 | 0.39349 |  | 0.39431 | 0.012106 | 0.003217 |  | 0.010406 | 0.011655 | 0.39108 |  | 0.39185 | 0.013453 | -0.019544 |  | -0.010929 | 0.012751 |  | 0 | 12
10 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 13
11 | 0.63893 |  | 0.64277 | 0.0045807 | 0.00049645 |  | -0.0047891 | 0.0087306 | 0.64269 |  | 0.64667 | 0.0078724 | -0.0095098 |  | -0.014028 | 0.0086006 |  | 0 | 14
12 | 9.9043e-05 |  | 0.0011536 | 0.00189 | -0.00039422 |  | -0.0042258 | 0.0025781 | -0.00015705 |  | -0.0014106 | 0.0025861 | -0.00016714 |  | -0.0016471 | 0.0016519 |  | 0 | 15
13 | 0.0057693 |  | 0.057086 | 0.0017692 | 0.00018117 |  | 0.0022761 | 0.0020433 | 0.0058717 |  | 0.058045 | 0.0015709 | 3.5748e-05 |  | 0.0009999 | 0.0018702 |  | 0 | 16
14 | -0.00019718 |  | -0.0019291 | 0.00057892 | -0.00072061 |  | -0.0078792 | 0.00095209 | -0.00047661 |  | -0.0049427 | 0.00080668 | -0.0012235 |  | -0.013134 | 0.0017829 |  | 0 | 17
15 | 0.0019724 |  | 0.021227 | 0.0016184 | -0.00022433 |  | -0.0028708 | 0.001749 | 0.0025972 |  | 0.027698 | 0.0020342 | -0.00055634 |  | -0.0064113 | 0.0026649 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.84435 | -1.0536
3 | [12 empty] | 3 | 9 | -0.90018 | -0.67985
4 | [12 empty] | 4 | 16 | -2.1069 | -1.8456
5 | [12 empty] | 5 | 25 | -1.7655 | -1.9214
6 | [12 empty] | 6 | 36 | -3.0711 | -3.0051
7 | [12 empty] | 7 | 49 | -3.4112 | -3.0687
8 | [12 empty] | 8 | 64 | -4.1802 | -3.8193
9 | [12 empty] | 9 | 81 | -4.4331 | -4.3353
10 | [12 empty] | sigma (mm) |  | 0.016513 | 0.018673
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05
17 | 0.00044619 | 0.00011155 | 2.6771e-05
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 387.6332 | 703.7956 | 704.3803 | 704.4899 | 704.3919 | 704.4531 | 704.5402 | 704.5162 | 704.4522 | 704.4366 | 704.3936 | 704.4927 | 704.4482 | 704.4873 | 704.5442 | 704.4441 | 704.496 | 704.5028 | 704.2453 | 421.877
2 | 25.4877 | 0.94345 | 0.73891 | 1.0735 | 1.1402 | 0.56181 | 1.4464 | 1.4293 | 1.455 | 1.2967 | 0.89777 | 0.94095 | 0.86613 | 0.36114 | 0.6582 | 1.0973 | -0.15263 | 0.83504 | 1.1384 | 23.782 | 2.3498
3 | 41.8543 | 3.0023 | 3.0642 | 2.5829 | 3.0298 | 1.6358 | 2.8944 | 3.0567 | 3.2036 | 2.3689 | 2.0757 | 2.8213 | 2.8102 | 1.6319 | 2.3881 | 3.3387 | 3.4067 | 3.2096 | 3.4399 | 9.6075 | 4.1122
4 | 1.2186 | 0.10605 | 0.057741 | 0.10897 | 0.20255 | 0.027329 | -0.049368 | 0.051051 | 0.0062892 | 0.0248 | -0.013786 | 0.0024922 | 0.087912 | -0.066206 | -0.056112 | -0.12209 | -0.14546 | -0.093945 | 0.06093 | 0.52261 | 0.061235
5 | -4.9882 | -0.45226 | -0.40091 | -0.025257 | 0.032567 | 0.28541 | 0.15398 | 0.23007 | -0.017798 | 0.12965 | 0.41451 | 0.1734 | 0.13733 | 0.04111 | -0.0081046 | -0.11699 | -0.18906 | -0.067087 | -0.14802 | -3.1587 | -0.2331
6 | 0.18765 | -0.021405 | 0.0062902 | -0.0085269 | -0.038865 | -0.00051634 | -0.028107 | -0.085588 | -0.092195 | -0.040501 | -0.039698 | -0.027672 | -0.014536 | -0.042298 | -0.065149 | -0.095995 | -0.075577 | -0.057955 | -0.028778 | -0.063702 | -0.036137
7 | 2.5616 | 0.81516 | 0.82697 | 0.93224 | 0.88903 | 0.91431 | 0.8844 | 0.89334 | 0.89654 | 0.90112 | 0.91679 | 0.95196 | 0.93947 | 0.9112 | 0.9027 | 0.9295 | 0.93064 | 0.90213 | 0.9331 | 0.81478 | 0.94878
8 | 0.022044 | 0.0074355 | 0.0084596 | 0.018442 | 0.022605 | 0.00066237 | -0.006799 | 0.032258 | 0.018818 | -0.0059183 | -0.0059977 | 0.0042566 | 0.018736 | -0.0028905 | -0.001064 | -0.0063964 | 0.0025297 | 0.026365 | 0.019158 | -0.019966 | 0.0078766
9 | 0.38939 | 0.40713 | 0.39322 | 0.39074 | 0.41963 | 0.36747 | 0.38987 | 0.38544 | 0.4097 | 0.39994 | 0.38224 | 0.38056 | 0.38685 | 0.39369 | 0.39157 | 0.40402 | 0.39336 | 0.40108 | 0.40106 | 0.37274 | 0.39349
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.55323 | 0.63887 | 0.63937 | 0.64582 | 0.64532 | 0.65047 | 0.64269 | 0.64257 | 0.63886 | 0.64493 | 0.64881 | 0.63934 | 0.63775 | 0.64812 | 0.6499 | 0.64078 | 0.63593 | 0.64316 | 0.63719 | 0.60213 | 0.63893
12 | -0.025786 | -0.027955 | -0.017015 | 0.020364 | -0.0077749 | 0.0083733 | 0.014749 | 0.03861 | 0.048769 | 0.028249 | 0.0028209 | 0.017325 | -0.0075874 | 0.016316 | 0.013951 | 0.020697 | 0.0049432 | 0.025053 | 0.0077623 | -0.006426 | 0.00099043 |  | ten times!
13 | 0.84102 | 0.61041 | 0.58913 | 0.5718 | 0.56829 | 0.54595 | 0.55866 | 0.54686 | 0.54842 | 0.55837 | 0.56123 | 0.56972 | 0.555 | 0.56928 | 0.59266 | 0.5811 | 0.5886 | 0.58031 | 0.57973 | 0.51674 | 0.057693 |  | ten times!
14 | 0.0058958 | -0.022332 | -0.021233 | -0.022354 | -0.027184 | -0.022365 | -0.013522 | -0.012769 | -0.0070515 | -0.014541 | -0.019095 | -0.013991 | -0.028082 | -0.016497 | -0.021272 | -0.014682 | -0.018612 | -0.024596 | -0.027064 | -0.056103 | -0.0019718 |  | ten times!
15 | -0.1753 | 0.2281 | 0.22134 | 0.22083 | 0.21843 | 0.21853 | 0.21189 | 0.20805 | 0.21831 | 0.21902 | 0.23247 | 0.21002 | 0.17876 | 0.20319 | 0.17747 | 0.21283 | 0.21652 | 0.23516 | 0.18997 | 0.087786 | 0.019724 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 60.8298 | 12.8581 | 17.2126 | 23.4506 | 17.9787 | 12.0412 | -4.4005 | -6.7895 | -9.3957 | -1.2167 | -6.9665 | -8.0704 | -2.8199 | -10.0831 | -15.4659 | -8.3659 | -7.8362 | -14.5323 | -14.5917 | -11.0188 | -0.0070155
2 | 0.095673 | -0.49379 | -0.20887 | -0.29511 | -0.41029 | -0.82627 | -0.74631 | -0.47679 | -0.63129 | -1.0592 | -0.69748 | -1.1116 | -1.6048 | -0.28678 | -0.75721 | -1.2995 | -1.5155 | -0.58953 | -1.5891 | 8.6344 | -0.48962
3 | -1.5315 | 0.63928 | 0.30145 | 0.22211 | 0.43398 | 0.68501 | 0.28239 | -0.12018 | 0.69266 | 0.33268 | 1.1342 | 0.40842 | 0.14523 | 0.098232 | 0.054832 | 0.41398 | 0.15026 | 0.48568 | 0.64235 | 0.039518 | 0.32288
4 | 0.36159 | -0.25534 | -0.3146 | -0.276 | -0.24628 | -0.43058 | -0.30712 | -0.46249 | -0.26482 | -0.59848 | -0.57326 | -0.66234 | -0.45562 | -0.63061 | -0.69019 | -0.27231 | -0.14488 | -0.32022 | -0.41469 | 0.3027 | -0.3624
5 | 1.7655 | -0.11025 | -0.040198 | -0.001987 | 0.34159 | 0.059232 | 0.0023329 | 0.072836 | 0.28738 | 0.18272 | 0.31258 | 0.069956 | 0.093277 | 0.067032 | 0.11914 | 0.1507 | 0.067325 | 0.18299 | 0.34393 | -0.03603 | 0.16453
6 | 0.044183 | 0.020546 | 0.0036097 | -0.046339 | -0.19155 | -0.1649 | -0.23095 | -0.15527 | -0.1511 | -0.10286 | -0.12153 | -0.16631 | -0.12338 | -0.15348 | -0.19124 | -0.11808 | -0.13412 | -0.040764 | -0.04666 | 0.012835 | -0.10875
7 | 1.5097 | -0.024139 | 0.011456 | 0.0065444 | 0.062155 | 0.042924 | 0.060536 | 0.042815 | 0.062477 | 0.048157 | 0.044103 | -0.02788 | 0.033245 | -0.0090283 | -0.013674 | 0.029122 | 0.049267 | 0.056583 | 0.049572 | -0.031704 | 0.069772
8 | -0.02645 | -0.010478 | -0.014345 | -0.0026668 | -0.059819 | -0.039033 | -0.042737 | -0.0011271 | -0.017533 | -0.015613 | -0.048284 | -0.04247 | -0.016924 | -0.021064 | -0.059758 | -0.032025 | -0.024339 | -0.050497 | -0.022543 | -0.028119 | -0.028861
9 | -0.21551 | 0.0088659 | 0.011489 | 0.0074601 | 0.029093 | 0.010726 | 0.0031424 | 0.019028 | 0.007209 | 0.0049648 | 0.031504 | -0.0069367 | 0.0057214 | -0.0030602 | -0.012121 | 0.010433 | 0.015025 | 0.019552 | 0.025203 | -0.011842 | 0.003217
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.18586 | -0.0083977 | 0.0024078 | 0.0084427 | 0.017772 | 0.002324 | -0.0084016 | -0.0055579 | -0.0048172 | -0.00041464 | -0.001093 | -0.016191 | -0.0047532 | -0.015094 | -0.015987 | -0.0098669 | -0.0064573 | -0.011502 | -0.0086164 | -0.010965 | 0.00049645
12 | 0.056082 | 0.012147 | 0.015244 | -0.018248 | -0.07262 | -0.06602 | -0.060524 | -0.049102 | -0.044541 | -0.045517 | -0.045514 | -0.047252 | -0.05878 | -0.046612 | -0.082426 | -0.053856 | -0.042804 | -0.022675 | -0.031549 | -0.041832 | -0.0039422 |  | ten times!
13 | -0.13322 | 0.021967 | 0.0084763 | 0.00022401 | 0.070281 | 0.030898 | 0.0094472 | 0.01203 | 0.027021 | 0.030265 | 0.060566 | 0.0098339 | 0.0080425 | 0.0090773 | -0.0081687 | 0.015855 | 0.028168 | 0.027764 | 0.047958 | 0.017612 | 0.0018117 |  | ten times!
14 | 0.10281 | -0.086288 | -0.075433 | -0.07124 | -0.084879 | -0.075606 | -0.074412 | -0.055501 | -0.073979 | -0.086911 | -0.0794 | -0.07488 | -0.076661 | -0.07393 | -0.092127 | -0.083198 | -0.071764 | -0.099053 | -0.082991 | -0.030465 | -0.0072061 |  | ten times!
15 | 0.12127 | -0.021101 | -0.021632 | -0.017996 | -0.033038 | 0.0062906 | -0.021418 | -0.034756 | -0.0069021 | -0.026407 | -0.015139 | -0.027826 | -0.057642 | -0.024559 | -0.062062 | -0.035225 | -0.025685 | -0.059498 | -0.032151 | 0.034144 | -0.0022433 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 379.2533 | 703.7763 | 704.4735 | 704.5888 | 704.4471 | 704.4583 | 704.4967 | 704.5241 | 704.5156 | 704.514 | 704.4843 | 704.6076 | 704.51 | 704.5782 | 704.639 | 704.5406 | 704.5571 | 704.5187 | 704.2894 | 432.2517
2 | -23.4959 | -0.60537 | -0.84691 | -0.86098 | -0.98806 | -0.56643 | -0.65157 | -1.0385 | -0.62581 | -1.2735 | -1.1348 | -0.68746 | -0.97164 | -0.78731 | -0.93362 | -0.38921 | -1.1292 | -0.77369 | -1.3054 | -22.1098 | -2.1818
3 | 41.6655 | 2.4295 | 2.7836 | 2.0492 | 2.9798 | 1.8781 | 3.2382 | 3.5441 | 3.3753 | 2.2273 | 1.7296 | 3.5536 | 2.1966 | 0.59606 | 1.0291 | 3.2775 | 3.4587 | 2.8592 | 3.2789 | 9.5921 | 3.9056
4 | -0.46205 | 0.17711 | 0.21774 | 0.20314 | 0.10864 | 0.028717 | 0.073289 | -0.060793 | 0.24607 | 0.15825 | 0.098771 | 0.31195 | 0.27909 | -0.059752 | -0.16789 | 0.12336 | 0.18506 | 0.12569 | -0.11068 | -1.3877 | 0.043731
5 | -5.0154 | -0.4026 | -0.56254 | -0.014245 | -0.15192 | 0.094533 | 0.018302 | -0.015146 | -0.24319 | -0.0052771 | 0.22364 | 0.13175 | 0.1171 | 0.21786 | 0.0081608 | -0.21763 | -0.30297 | -0.04977 | -0.29145 | -3.057 | -0.31436
6 | 0.29615 | 0.017487 | 0.031852 | 0.0010289 | 0.022391 | 0.030224 | 0.017639 | 0.075556 | -0.05962 | -0.022647 | -0.010221 | 0.044713 | 0.079206 | 0.01998 | 0.083808 | 0.10251 | -0.00032637 | -0.07894 | 0.022686 | -0.082965 | 0.025369
7 | 2.6299 | 0.81124 | 0.87834 | 0.95382 | 0.91818 | 0.95643 | 0.94375 | 0.98069 | 0.97981 | 0.96301 | 0.96473 | 0.94796 | 0.90673 | 0.93421 | 0.9016 | 0.95133 | 0.88576 | 0.96704 | 0.96826 | 0.88619 | 0.98019
8 | 0.046387 | 0.018636 | -0.0070333 | 0.01947 | 0.0039822 | 0.00083627 | -0.0099464 | -0.013389 | -0.010176 | -0.015111 | -0.0058165 | 0.026043 | 0.0039246 | -0.014218 | -0.0089116 | -0.018665 | -0.029319 | 0.020162 | 0.013409 | -0.020374 | -0.00071413
9 | 0.36733 | 0.38492 | 0.3853 | 0.39342 | 0.41282 | 0.36991 | 0.39303 | 0.4005 | 0.41177 | 0.38036 | 0.3834 | 0.38527 | 0.39225 | 0.37688 | 0.37227 | 0.41538 | 0.39953 | 0.40274 | 0.39363 | 0.38924 | 0.39108
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.5666 | 0.64011 | 0.63771 | 0.651 | 0.64536 | 0.65051 | 0.64279 | 0.64216 | 0.64463 | 0.6439 | 0.65514 | 0.64793 | 0.65161 | 0.66507 | 0.66064 | 0.64364 | 0.64272 | 0.6412 | 0.63396 | 0.59256 | 0.64269
12 | -0.012608 | 0.028034 | -0.0077131 | 0.0086946 | -0.033766 | -0.018239 | -0.026262 | -0.061392 | -0.051592 | -0.021514 | -0.035856 | 0.019381 | -0.042332 | 0.0056135 | 0.019597 | -0.0049547 | -0.032226 | -0.0042782 | 0.0049013 | -0.06527 | -0.0015705 |  | ten times!
13 | 0.9351 | 0.60742 | 0.61311 | 0.56959 | 0.59235 | 0.58292 | 0.57531 | 0.56982 | 0.6064 | 0.57968 | 0.56985 | 0.58441 | 0.55784 | 0.56911 | 0.56348 | 0.57212 | 0.58362 | 0.56721 | 0.58391 | 0.47903 | 0.058717 |  | ten times!
14 | 0.046249 | -0.066588 | -0.058541 | -0.052975 | -0.052531 | -0.038897 | -0.052754 | -0.054551 | -0.039403 | -0.040975 | -0.055306 | -0.050322 | -0.05623 | -0.051643 | -0.045259 | -0.035443 | -0.045532 | -0.05151 | -0.04122 | -0.078279 | -0.0047661 |  | ten times!
15 | -0.12639 | 0.29993 | 0.26348 | 0.28991 | 0.25672 | 0.28684 | 0.27525 | 0.24225 | 0.28952 | 0.26136 | 0.31173 | 0.29769 | 0.25684 | 0.30075 | 0.26731 | 0.27845 | 0.28365 | 0.28247 | 0.24143 | 0.092155 | 0.025972 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 28.5766 | -17.4883 | -14.3687 | -4.2997 | -1.6287 | 0.62886 | 0.39476 | 3.3196 | 2.543 | 7.6093 | 2.269 | 2.2745 | 5.1661 | 3.647 | -0.27959 | 2.3724 | 2.5912 | -3.3873 | -1.5579 | -8.4014 | 0.0028382
2 | 0.40554 | 0.18769 | -0.37679 | 0.45067 | 0.17031 | 0.34109 | 0.44458 | 1.0815 | 0.24587 | 1.0215 | 0.65977 | -0.01068 | -0.81353 | 0.21539 | 0.16676 | 0.79755 | 0.80216 | -0.122 | 0.055168 | 8.5499 | 0.56295
3 | -1.3197 | 0.42046 | 0.18742 | 0.3302 | 0.58488 | 0.44856 | -0.21567 | 0.12761 | 0.75151 | 0.60506 | 0.78957 | 0.41075 | 0.48479 | 0.66422 | 0.3605 | 0.83616 | 0.47377 | 0.47122 | 1.0788 | 0.31738 | 0.43306
4 | 0.36877 | -0.19788 | -0.10362 | -0.43791 | -0.2018 | -0.19111 | -0.49164 | 0.049491 | -0.15266 | -0.63866 | -0.56037 | -0.37827 | -0.6173 | -0.24599 | -0.35604 | -0.29055 | -0.20289 | -0.27691 | -0.27655 | 0.37818 | -0.26839
5 | 1.851 | 0.050093 | -0.018377 | 0.084503 | 0.28898 | 0.19183 | 0.018532 | 0.08996 | 0.17104 | 0.14966 | 0.19812 | -0.041724 | -0.01394 | 0.083387 | 0.20319 | 0.17684 | -2.9265e-05 | 0.012733 | 0.20898 | -0.12362 | 0.14486
6 | -0.082554 | 0.046198 | 0.023409 | 0.028945 | 0.094378 | 0.11993 | 0.046081 | -0.013882 | -0.060893 | 0.019122 | 0.0862 | 0.010402 | -0.023987 | 0.0094078 | 0.020734 | -0.031893 | 0.061874 | -0.023789 | -0.071036 | -0.023808 | 0.014731
7 | 1.3844 | -0.07611 | -0.034169 | -0.10489 | 0.0055279 | 0.041465 | -0.08629 | 0.051463 | 0.049864 | 0.039712 | 0.042305 | -0.023702 | 0.025613 | 0.048644 | -0.029675 | 0.029392 | -0.0081888 | 0.024209 | 0.0069726 | -0.0094831 | 0.038725
8 | 0.030362 | -0.006201 | -0.0028133 | 0.0038782 | 0.0065508 | 0.001092 | -0.022042 | -0.038963 | -0.058059 | -0.038625 | -0.031147 | -0.042332 | -0.086344 | -0.08505 | -0.091535 | -0.049778 | -0.017972 | -0.03701 | -0.030022 | 0.0073881 | -0.031618
9 | -0.31716 | -0.0039505 | -0.015036 | -0.0015901 | 0.0011704 | -0.0050493 | -0.037079 | 0.0023042 | -0.0047716 | -0.014859 | 0.014223 | -0.017375 | -0.023508 | -0.026775 | -0.012137 | -0.0057877 | -0.029375 | -0.003315 | -0.013806 | -0.011151 | -0.019544
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.14066 | -0.038002 | -0.025561 | -0.014708 | -0.0069397 | -0.005166 | -0.015078 | -0.0024673 | -0.005844 | -0.0099847 | -0.006649 | -0.019034 | -0.010289 | -0.017492 | -0.022008 | -0.0095057 | -0.012757 | -0.01735 | -0.013676 | -0.0087423 | -0.0095098
12 | 0.028217 | -0.0044881 | -0.015295 | -0.017054 | -0.0080682 | 0.019655 | 0.00016235 | -0.0098473 | -0.034337 | 0.0016411 | 0.0037617 | -0.030204 | -0.038669 | -0.039823 | -0.030358 | -0.028647 | -0.013815 | -0.026312 | -0.024784 | -0.063238 | -0.0016714 |  | ten times!
13 | -0.22414 | 0.0081772 | -0.01222 | -0.014764 | 0.013767 | 0.025499 | -0.02763 | 0.013389 | 0.020324 | 0.021524 | 0.045661 | 0.0059616 | -0.0046156 | 0.026906 | 0.016261 | 0.030212 | -0.011805 | 0.0017178 | 0.021618 | 0.01491 | 0.00035748 |  | ten times!
14 | 0.10559 | -0.12764 | -0.1133 | -0.11981 | -0.10802 | -0.12223 | -0.11871 | -0.1029 | -0.11548 | -0.13698 | -0.15226 | -0.1407 | -0.17203 | -0.14916 | -0.15184 | -0.13606 | -0.12434 | -0.13892 | -0.13373 | -0.058479 | -0.012235 |  | ten times!
15 | 0.11835 | -0.083127 | -0.081062 | -0.047719 | -0.067566 | -0.03593 | -0.087075 | -0.071056 | -0.035493 | -0.046127 | -0.05899 | -0.044293 | -0.1021 | -0.025543 | -0.063317 | -0.029935 | -0.059698 | -0.1203 | -0.094707 | 0.04039 | -0.0055634 |  | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | -0.00088248 | -5.2431e-05 | 0.00010316 | -3.5964e-05 | 5.0917e-05 | 0.00017457 | 0.00014049 | 4.9639e-05 | 2.7493e-05 | -3.355e-05 | 0.00010713 | 4.3961e-05 | 9.9468e-05 | 0.00018024 | 3.814e-05 | 0.00011182 | 0.00012147 | -0.00024408 |  | 704.4172
|  | -0.00098938 | 2.9179e-07 | 0.00016396 | -3.7183e-05 | -2.1285e-05 | 3.3224e-05 | 7.2118e-05 | 6.0053e-05 | 5.7782e-05 | 1.5622e-05 | 0.00019065 | 5.2104e-05 | 0.00014891 | 0.00023522 | 9.554e-05 | 0.00011896 | 6.4453e-05 | -0.00026104 |  | 704.4733
Check Computation for Aperture1
Magnetic Length | -0.064552
dB/B Heads CS+NCS | -4254.0432
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.2893 | -4497.1079 | -8.8248 | -0.52431 | 1.0316 | -0.35964 | 0.50917 | 1.7457 | 1.4049 | 0.49639 | 0.27493 | -0.3355 | 1.0713 | 0.43961 | 0.99468 | 1.8024 | 0.3814 | 1.1182 | 1.2147 | -2.4408 | -4010.9784
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 6.2541 | 1.4354 | 1.8711 | 2.4948 | 1.9473 | 1.3534 | -0.2899 | -0.52912 | -0.78954 | 0.028263 | -0.54738 | -0.65721 | -0.13029 | -0.85844 | -1.396 | -0.68701 | -0.63434 | -1.3035 | -1.3075 | -0.95491
db2 | 1.1846 | 26.1173 | 0.18477 | -0.0067991 | 0.012248 | 0.071266 | 0.027099 | 0.032695 | -0.010119 | 0.062113 | -0.0092506 | 0.10706 | 0.051167 | -0.096243 | -0.16623 | -0.086899 | -0.05636 | -0.10703 | -0.066458 | 0.05697 | 23.9336
db3 | 4.5075 | 6.4735 | -0.0023819 | 0.043302 | -0.00057431 | -0.031357 | -0.013399 | -0.010924 | 0.012454 | 0.037627 | 0.051265 | -0.060646 | -0.020463 | -0.032401 | 0.002475 | -0.032926 | -0.0087831 | 0.0053332 | 0.065655 | -0.004257 | 9.2286
db4 | 0.026764 | 1.5887 | 0.021545 | -0.00030465 | 0.0022954 | -9.8973e-05 | 0.0079988 | 0.0090148 | -0.00031736 | 0.014457 | -0.0087487 | 0.0092625 | 0.0019617 | -0.002061 | -0.0013576 | -0.0050562 | -0.0063074 | -0.023311 | -0.021191 | 0.0022184 | 0.65435
db5 | 0.030174 | 0.22011 | -0.0013251 | 0.0043907 | -0.0026584 | -0.0045426 | 0.0018319 | 0.00026394 | 0.0026843 | -0.0047131 | 0.0045644 | 0.0013445 | 0.0053081 | 0.0060438 | -0.00068407 | 0.0010827 | -0.00036387 | -0.0030652 | -0.0035032 | -0.0066587 | 0.19782
db6 | -0.015687 | 0.4368 | -0.003185 | 0.0013272 | 0.0021297 | 0.0028555 | 0.00097707 | 0.0063599 | -9.6099e-05 | -0.0033306 | -0.001412 | -0.00049727 | 0.00061792 | 0.00065136 | 0.00083011 | 0.0018771 | -0.0026603 | -0.00060226 | -0.0041962 | -0.0016462 | -0.0081576
db7 | -0.0061096 | 0.046813 | 0.0021657 | 0.0029076 | -0.0027727 | 0.00092479 | -0.0029333 | -0.00011669 | -4.1671e-05 | -0.0017842 | -0.00035181 | -0.0026089 | 0.0029868 | 0.00013369 | 0.0015287 | 0.0035463 | 0.00043505 | -0.0026874 | 0.0018259 | -0.0031579 | 0.034169
db8 | -0.0053432 | 0.12129 | 0.0002569 | -0.00073971 | -0.00046501 | 0.001695 | -0.0013082 | 0.0020859 | -0.0041953 | 0.0014379 | -0.00080627 | 0.0043684 | 0.00056912 | -0.00016734 | 1.8993e-05 | 0.00047071 | -6.615e-05 | -0.0012439 | -0.0018414 | -6.959e-05 | 8.6326e-06
db9 | 0.013561 | 0.036819 | 0.00036768 | 0.0019641 | 4.4124e-05 | 0.00038153 | -0.0011789 | -0.0012091 | -0.0010588 | -0.0017454 | 0.0056701 | -0.00080906 | 0.00082171 | -0.0025035 | 0.0045183 | -0.00015719 | -0.00093957 | 0.00064957 | -0.00083343 | -0.0039821 | 0.02511
db10 | -9.67e-19 | 4.6325e-17 | 7.9059e-18 | 4.4364e-18 | -2.3848e-19 | -7.6772e-19 | 9.67e-19 | -5.3471e-18 | -5.7881e-18 | -2.2378e-18 | 9.67e-19 | 3.9072e-18 | 4.1718e-18 | -1.0794e-17 | 6.8474e-18 | -7.0304e-18 | -2.8406e-18 | 4.9657e-18 | 9.67e-19 | -9.1473e-20 | 1.5571e-17
db11 | 0.00637 | -0.0021384 | 0.0030782 | 0.00055066 | 0.00072767 | -0.0015987 | -0.00016236 | -0.00037294 | -0.00052412 | -0.00085807 | 0.00020201 | 6.2091e-05 | -0.00016336 | -0.0010468 | 0.00075488 | -0.00086297 | 0.000169 | -0.00011885 | 0.00071613 | -0.0005524 | 0.029368
db12 | -7.3946e-06 | -0.001672 | -0.0011792 | -0.001049 | 0.00082399 | 8.691e-05 | 7.7516e-06 | 0.00036978 | -9.2695e-05 | 0.00023956 | 0.00013611 | 0.00036646 | 0.0004263 | 0.0003713 | 0.00011291 | -9.8992e-05 | 0.00065385 | -0.00011094 | -0.00024217 | -0.00082189 | -0.0031979
db13 | -0.00013517 | 0.0011245 | 0.00025233 | 0.0006795 | -0.00064995 | -0.00016941 | -1.7255e-05 | -0.00067771 | 0.00056429 | -0.00028043 | 0.00069275 | 0.00033275 | 0.0009144 | -0.00037999 | -0.00010125 | 0.00018299 | -0.00014358 | 0.00024347 | 7.5912e-05 | -0.0015188 | 0.00062743
db14 | -0.00049129 | 0.0019368 | -0.00076897 | -5.2597e-05 | -2.5328e-05 | 1.2489e-05 | -8.2078e-05 | -1.7225e-06 | -0.00052018 | 0.00031988 | -7.2262e-05 | 0.00011353 | 0.00036153 | -0.00010317 | 0.00097325 | -0.00073528 | 7.4498e-05 | 0.00068456 | 0.00067971 | -0.00085785 | -0.003049
db15 | 0.0014419 | -0.0038102 | 0.0014958 | 0.00075526 | 0.0011885 | 0.00032954 | 0.00093984 | -0.00011762 | 0.00011125 | -0.00011996 | -0.00069859 | 0.0010984 | -0.00095925 | -0.0022086 | 0.00033209 | -0.0035 | 0.00099007 | 0.0010609 | 0.0016608 | -0.0023584 | 0.0058486
da2 | -0.10623 | -18.8253 | 0.03783 | -0.10202 | 0.0017093 | -0.072294 | 0.016735 | 0.0057048 | 0.13919 | 0.068062 | 0.067614 | 0.027133 | -0.039666 | -0.13117 | 0.037497 | 0.14927 | -0.1284 | -0.095094 | 0.023886 | -0.0059984 | 0.15269
da3 | 0.0044295 | 0.58903 | 0.14098 | -0.12821 | 0.066917 | -0.0087007 | 0.046213 | 0.076681 | 0.011472 | 0.10229 | -0.02831 | 0.074815 | 0.01343 | -0.070016 | -0.067826 | -0.014663 | -0.029661 | -0.20571 | -0.030639 | 0.050944 | -0.019879
da4 | 0.0062513 | -3.9939 | 0.027188 | -0.026702 | 0.00757 | -0.0076771 | 0.0013526 | -0.0081955 | 0.0063289 | 0.023352 | 0.0082849 | 0.017016 | -0.0037802 | -0.025292 | -0.020422 | 0.013675 | -0.029154 | -0.024505 | 0.016192 | 0.024768 | -0.18221
da5 | -0.00092259 | -0.28469 | -0.00086722 | -0.018942 | -0.0007291 | -0.003352 | 0.0074493 | 0.0049761 | 0.0072724 | 0.0065931 | 0.00025104 | 0.013246 | 0.0082591 | -0.0048808 | -0.004844 | 0.0033171 | 0.00013014 | -0.010709 | -0.0074512 | 0.00028057 | -0.052227
da6 | 0.0032132 | -0.64543 | -6.4096e-05 | -0.0015919 | -0.00055076 | -0.00080453 | -0.00084068 | 0.001811 | 0.00015925 | 0.00036732 | 0.00017433 | 0.0033937 | -0.0029215 | -0.0025691 | -0.0015059 | 0.00069656 | 0.0018179 | -0.00076711 | -0.0005529 | 0.0037484 | -0.014145
da7 | 0.027801 | 0.15901 | 0.0035265 | 0.010556 | 0.016875 | 0.0067394 | 0.0075685 | -0.0067476 | -0.0065767 | -0.0069186 | 0.00090902 | 0.0018709 | -0.005193 | -0.0014187 | -0.0096609 | -0.0080461 | 0.00011757 | 0.002817 | -0.0035153 | -0.0029024 | 0.038716
da8 | 0.0016056 | -0.055745 | 0.00021111 | 0.0033991 | -0.00032983 | 0.0009204 | -0.0028791 | -0.00024879 | -0.0010987 | 0.00016724 | 0.00011106 | 0.00069887 | -0.00054078 | 0.00066221 | 8.1188e-06 | -0.0024935 | 0.00060495 | -0.00022187 | 0.0013138 | -0.00028432 | -0.0065845
da9 | 0.015193 | 0.046421 | 0.0042673 | 0.010052 | 0.0070515 | 0.0085097 | 0.00091761 | -0.0053866 | -0.0057172 | -0.005105 | 0.00090526 | -0.0024434 | -0.004029 | -0.00038579 | -0.0050315 | -0.0033132 | 0.0019831 | 0.0040515 | -0.0029351 | -0.0033908 | 0.035617
da10 | 3.5119e-19 | 0 | -2.4093e-20 | 2.589e-18 | -8.694e-19 | -4.1882e-18 | -8.6573e-20 | 4.3237e-18 | -3.5119e-19 | -8.8043e-19 | -3.5119e-19 | -6.2316e-18 | -3.2914e-18 | -3.5119e-19 | -2.0859e-18 | -1.0862e-18 | -1.8213e-18 | 1.1189e-18 | 1.1189e-18 | 1.2468e-17 | -3.8246e-17
da11 | 0.0031332 | 0.048889 | 0.010007 | 0.015113 | 0.018057 | 0.011451 | 0.0068058 | -0.0059119 | -0.0073962 | -0.0074882 | 0.00081018 | -0.001418 | -0.0067169 | -0.0029641 | -0.0079488 | -0.0080472 | -0.0019725 | 0.00049541 | -0.0065983 | -0.006278 | 0.03731
da12 | 0.00037626 | 0.005959 | -0.0003903 | -0.00053843 | -0.00070297 | 0.00026883 | -0.00023548 | 0.00099354 | -0.00052504 | 0.00018528 | 0.00044688 | -0.00072733 | 0.00020417 | -0.00069338 | 0.00017723 | -0.0003995 | 2.11e-05 | 0.00085423 | 0.00066728 | 0.00039388 | -0.00078383
da13 | 0.0035659 | 0.0082277 | 0.0014594 | 0.0011638 | 0.0010707 | 0.0020202 | 0.0012117 | -0.00036059 | -0.001241 | -0.0014883 | -0.00039691 | 0.00014378 | 6.3902e-05 | -0.00075078 | -0.00061485 | -0.00162 | -0.00028672 | 0.0013659 | -0.00065433 | -0.001086 | 0.0049843
da14 | -0.0072958 | 0.0081238 | -0.00072823 | -5.7705e-05 | 6.7461e-05 | -0.00018258 | 1.6115e-05 | 0.00049399 | 2.2024e-05 | 0.0007042 | -0.0008516 | -0.00013541 | -9.6463e-05 | -9.4575e-05 | 0.00036731 | -0.00038086 | 0.00010998 | 0.00054238 | -0.00029995 | 0.00050393 | -0.0030615
da15 | -0.0026557 | 0.0068481 | 0.00283 | 0.0021632 | 0.0017428 | -0.00056414 | 0.0034135 | 0.00045732 | -0.0007079 | 0.00088978 | 0.00054781 | 0.00032502 | 0.00070997 | -0.0030203 | 9.697e-05 | -0.0037659 | -0.00030162 | 0.00083943 | -0.0027963 | -0.0028597 | 0.0038754
Check Computation for Aperture2
Magnetic Length | -0.064605
dB/B Heads CS+NCS | -4240.3423
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 108.3858 | -4616.4986 | -9.8938 | 0.0029179 | 1.6396 | -0.37183 | -0.21285 | 0.33224 | 0.72118 | 0.60053 | 0.57782 | 0.15622 | 1.9065 | 0.52104 | 1.4891 | 2.3522 | 0.9554 | 1.1896 | 0.64453 | -2.6104 | -3864.1861
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle |  | 2.8854 | -1.6923 | -1.3805 | -0.3734 | -0.10668 | 0.11944 | 0.096016 | 0.38797 | 0.31123 | 0.81709 | 0.28388 | 0.2841 | 0.57411 | 0.42162 | 0.028461 | 0.29404 | 0.31592 | -0.28171 | -0.099232 | -0.77069
db2 | 1.2466 | 20.6484 | 0.14235 | -0.050337 | 0.090237 | -0.027197 | -0.0047387 | 0.079993 | -0.018701 | -0.02076 | 0.056523 | 0.079701 | 0.030097 | -0.087993 | -0.059644 | -0.050226 | -0.078994 | -0.0016479 | -0.12409 | 0.045433 | 22.1098
db3 | 4.5403 | -4.825 | 0.0018891 | 0.018917 | -0.0052651 | 0.0089859 | -0.022386 | -0.031707 | -0.020914 | -0.011126 | 0.046014 | -0.074874 | 0.038861 | 0.043756 | 0.0014044 | -0.084059 | 0.030223 | 0.043697 | 0.080111 | -0.063526 | 9.2138
db4 | 0.018638 | -0.74717 | 0.0062655 | -0.0023255 | 0.01083 | 0.00012919 | 0.0059538 | -0.00086519 | -0.0028299 | -0.01109 | -0.0037334 | 0.0087512 | -0.012708 | 0.0067729 | -0.0019971 | 0.0029994 | 0.0090221 | -0.0043595 | -0.010143 | -0.00067242 | 1.3877
db5 | 0.032951 | 0.077536 | 0.00077955 | -0.0018212 | -0.0025529 | 0.0032882 | 0.0013006 | 0.0045309 | 0.0026488 | 0.0032568 | -0.0044535 | 0.0059716 | -0.0026442 | 0.0040889 | -0.005646 | 0.0029479 | 0.0032862 | 0.0025297 | -0.00026102 | -0.01725 | 0.21733
db6 | -0.021751 | -0.48659 | 0.0040152 | -0.0033671 | -0.0084679 | -0.0012103 | -0.0039873 | 0.0014038 | 0.0080479 | -0.0063285 | 0.0025729 | 0.0016041 | -0.0036087 | 0.0023002 | 0.00041066 | 7.3129e-05 | -0.00011383 | 0.0022174 | -4.6539e-05 | 0.0044848 | -0.0035106
db7 | -0.0053354 | -0.10429 | 0.0044175 | -0.0009101 | -0.0018582 | -0.00094023 | 0.0033839 | 0.00089748 | 0.0010434 | 0.0020093 | -0.002891 | -0.0023122 | 0.0022699 | -0.0026777 | 0.0023014 | -0.0017916 | 0.003054 | -0.0014676 | 0.001869 | -0.0063971 | 0.04966
db8 | -0.0043907 | -0.08799 | -0.0025245 | 0.0010069 | -0.0006543 | -0.001056 | -0.0010225 | 0.00027931 | 0.0029401 | -0.0021129 | -0.00069192 | 0.00039456 | -0.0014675 | 0.0027903 | 0.0041031 | -0.0027613 | -0.00088592 | 0.0025537 | 0.0025484 | -0.0034395 | -0.0055558
db9 | 0.012392 | -0.026564 | -0.0003537 | -4.4255e-05 | -0.00018284 | 0.0012995 | -0.00065629 | 0.0011398 | -0.00096419 | 0.00047034 | -0.00012692 | -0.0016909 | 0.00062733 | -0.0026037 | -0.00044432 | 0.00099018 | 0.0021601 | 0.0024315 | 0.0019766 | -0.0040281 | 0.027933
db10 | 2.6119e-18 | 5.5511e-17 | -8.7716e-19 | 3.2685e-18 | -2.6119e-18 | 1.1957e-18 | -2.9794e-18 | 9.149e-18 | 8.5756e-19 | -9.0215e-18 | 2.4012e-18 | -1.7445e-18 | 6.738e-19 | -5.5521e-18 | -5.5521e-18 | -1.8107e-18 | 8.5756e-19 | -1.1418e-18 | -4.9494e-19 | 1.3383e-17 | 0
db11 | 0.0098923 | -0.024221 | 0.00024769 | -0.0015173 | -0.0006309 | 0.0010145 | -7.3034e-05 | 0.00038421 | 0.0014734 | 0.00073816 | -0.00068577 | -0.00017232 | 0.00011369 | -0.00038823 | -0.00059056 | 0.00034565 | 0.00047491 | 0.00027545 | 0.00044757 | -0.0014572 | 0.03445
db12 | 0.0016693 | -0.003524 | -0.00059047 | -8.2343e-05 | -0.00037835 | 0.00069866 | -0.0010381 | 0.0013478 | 0.0011986 | 0.0003852 | -0.0011334 | 0.00083204 | -0.0021611 | 0.0014366 | 0.0011734 | 6.5551e-06 | -0.00044771 | -0.0015236 | 0.00013296 | 0.00014317 | 0.006527
db13 | -0.00065367 | 0.0090818 | -2.1871e-05 | 0.00043137 | 0.00076296 | 0.00070336 | 0.00039721 | -0.00059383 | 0.00042992 | 0.0017972 | -0.00042854 | -0.00038957 | 0.00056676 | -0.0016584 | -8.593e-05 | -0.001785 | -0.0010203 | 0.00050003 | 0.00060221 | -0.00020767 | -4.1056e-05
db14 | 0.00091662 | -9.0718e-06 | 0.0014341 | 0.0010142 | 0.00077107 | -0.00036163 | -0.00058303 | 0.0001463 | 0.00048252 | -0.00068504 | -0.0006275 | -0.00036751 | -0.00038062 | -0.00044379 | -0.00018989 | 0.00088066 | -0.00090684 | 5.4403e-05 | -1.7321e-05 | -0.00022005 | 0.0078279
db15 | 0.0092786 | 0.00061513 | 0.00061323 | -0.0034233 | 0.00035156 | -0.0016099 | 0.0017122 | 0.002522 | -0.0028164 | 0.0017573 | -0.0029928 | 0.002991 | 0.00077437 | -0.0018754 | 0.0027523 | -0.002149 | -0.00024635 | 0.0012273 | 0.0031674 | -0.0027555 | 0.01019
da2 | 0.088942 | -15.102 | 0.11899 | -0.19422 | 0.047064 | -0.056558 | 0.082696 | 0.059755 | 0.086974 | 0.094265 | 0.090521 | 0.14807 | 0.014371 | -0.16514 | -0.038028 | 0.088717 | -0.17074 | -0.12554 | -0.032228 | -0.048958 | -0.28337
da3 | 0.049372 | -1.0868 | -0.0011525 | -0.12069 | 0.018385 | 0.007224 | -0.0093211 | 0.025383 | -0.016264 | 0.082422 | -0.041357 | 0.044577 | -0.021232 | -0.03358 | -0.064531 | -0.0029618 | 0.043754 | -0.023983 | 0.015695 | 0.097632 | -0.069242
da4 | -0.017274 | -2.7829 | 0.0020529 | -0.0013034 | 0.0037412 | -0.010963 | 0.0094243 | 0.0082249 | 0.0069963 | 0.019951 | 0.011247 | -0.017169 | 0.022253 | -0.01778 | 0.0082515 | 0.0068859 | -0.016279 | -0.014907 | -0.01284 | -0.007787 | -0.3235
da5 | 0.0031197 | 0.032429 | 0.011723 | 0.00088358 | -0.0010495 | 0.0036061 | -0.0068276 | 0.0053429 | -0.00036775 | 0.0064399 | -0.0027047 | 0.0013218 | 0.00030666 | -0.010781 | -0.016401 | 0.0016858 | 0.0057977 | -0.00094843 | -0.0041995 | 0.0061731 | 0.031195
da6 | -0.0057267 | -0.47259 | -0.00041083 | 0.0057318 | 0.0003042 | -0.0016295 | 0.0030965 | -0.00021857 | 0.00066239 | 0.0024282 | -0.0014465 | -0.0023166 | -0.0089527 | 0.0027491 | 0.0010312 | -0.0002579 | -0.0013728 | 0.0017284 | 0.0029331 | -0.0040594 | -0.018975
da7 | -0.0094271 | -0.068746 | -0.0089776 | -0.0029354 | -0.00048578 | 6.6865e-05 | 0.00018817 | -0.001095 | -0.002747 | 0.0028962 | 0.0055537 | 0.0030194 | 0.0011277 | -0.0021555 | 0.0014567 | 3.783e-06 | 0.0031435 | 0.0063759 | 0.00057807 | -0.0060138 | 0.0050435
da8 | 0.00011305 | -0.01722 | -0.00037752 | 0.0040267 | 0.0027248 | -0.0017267 | 0.0040519 | 0.0016699 | 0.00023056 | -0.0023438 | -0.0015377 | -0.00035376 | -0.00263 | 0.00036825 | -0.0046517 | -0.0009723 | 2.2268e-05 | -0.0014124 | 0.00088359 | 0.0020279 | -0.00035577
da9 | -0.0022386 | -0.066431 | -0.0055849 | -0.0014927 | -0.00097083 | 0.0038318 | -0.0036136 | 0.00099967 | 0.00049632 | -0.00044536 | 0.0032992 | 0.0033392 | -0.00050645 | 0.00099434 | 0.00082746 | 0.0044797 | -0.0022774 | 0.003699 | -0.0021482 | -0.0049272 | -0.011151
da10 | -3.4459e-19 | -3.095e-18 | 3.3104e-19 | 3.4459e-19 | 1.0796e-18 | 3.814e-18 | 3.4459e-19 | -4.3303e-18 | -4.3303e-18 | 6.5484e-18 | 3.2848e-18 | -2.5956e-18 | 3.2848e-18 | -3.1249e-18 | -3.1249e-18 | 2.0793e-18 | 7.1212e-19 | -2.5956e-18 | 3.814e-18 | -5.5358e-18 | 0
da11 | -0.0096422 | -0.026801 | -0.0084366 | -0.0077367 | -0.0012284 | -0.0018994 | -0.0011058 | -0.00015385 | -0.0014253 | -0.0012933 | 0.0051287 | 0.0037533 | -0.0005583 | 0.0018733 | -0.00062968 | -0.00089157 | 0.0047064 | 0.007393 | 0.0023281 | 0.00017597 | -0.0087423
da12 | 0.0011141 | 0.0070322 | 0.00064396 | 0.00073757 | -0.00052098 | -0.00099406 | 0.00053943 | 0.00027836 | 0.0006913 | 0.0012282 | 0.00042371 | -0.00078274 | -0.00066616 | -0.00036856 | -0.0011895 | -1.7309e-05 | -0.00059609 | -0.00043715 | 0.00064291 | 0.00038716 | -0.0063238
da13 | 5.3247e-05 | -0.003748 | -0.00038163 | -0.00031453 | 0.00041362 | -0.00049938 | -5.5806e-05 | -0.00041931 | 0.0004285 | -0.0006866 | 0.00048559 | 0.0002043 | -0.00010892 | -0.00018761 | 0.00013804 | 0.00057054 | 0.00028974 | 0.0008467 | -0.00084116 | 0.00011792 | 0.0022482
da14 | -0.012059 | 0.0096365 | -0.00056274 | -0.00039358 | -0.00057842 | 0.00027131 | 0.00052538 | 0.00091008 | 0.0012338 | 0.0013085 | 0.00030462 | -0.00033882 | -0.00013739 | -0.0013077 | 0.00050942 | 0.00029412 | -0.00028207 | -0.00027252 | -0.00079683 | -0.00068721 | -0.0058479
da15 | -0.0085428 | 0.0089282 | -0.00075203 | -0.00037853 | 0.0033936 | -0.0011008 | 0.0019973 | -0.0013656 | -0.0016331 | 0.0015373 | 0.0019211 | 0.00073049 | 0.0016639 | -0.0037946 | 0.0044871 | 0.00031684 | 0.002967 | 0.0018818 | -0.0066352 | -0.0052365 | 0.004039
Field Colinearity | -0.96021
Twist Integral Aperture 1
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 6.1044 | 1.2856 | 1.7213 | 2.345 | 1.7976 | 1.2036 | -0.43968 | -0.6789 | -0.93932 | -0.12151 | -0.69716 | -0.80698 | -0.28006 | -1.0082 | -1.5458 | -0.83679 | -0.78412 | -1.4533 | -1.4573 | -1.1047
Twist (m2 mrad) | 45.7826 | 18.32 | 23.2375 | 29.8988 | 21.5708 | 13.541 | -4.6166 | -6.6193 | -8.4539 | -1.0025 | -5.2287 | -5.4471 | -1.6804 | -5.2932 | -6.9559 | -3.138 | -2.3524 | -3.27 | -2.186 | -0.41426
Twist Integral (I0) (m2 rad) | 0.071769
Twist Integral Aperture 2
Effective length (m) | 15 | 14.25 | 13.5 | 12.75 | 12 | 11.25 | 10.5 | 9.75 | 9 | 8.25 | 7.5 | 6.75 | 6 | 5.25 | 4.5 | 3.75 | 3 | 2.25 | 1.5 | 0.75
a1 angle (mrad) | 2.8289 | -1.7489 | -1.437 | -0.42992 | -0.16319 | 0.062918 | 0.039498 | 0.33145 | 0.25471 | 0.76057 | 0.22736 | 0.22758 | 0.51759 | 0.3651 | -0.028057 | 0.23752 | 0.2594 | -0.33822 | -0.15575 | -0.82721
Twist (m2 mrad) | 21.217 | -24.9212 | -19.3997 | -5.4815 | -1.9583 | 0.70783 | 0.41473 | 3.2316 | 2.2924 | 6.2747 | 1.7052 | 1.5362 | 3.1055 | 1.9168 | -0.12626 | 0.8907 | 0.77819 | -0.76101 | -0.23362 | -0.3102
Twist Integral (I0) (m2 rad) | -0.0068407
Note on D189: 3.9s

## Work sheet diff
| HCMBALA001-3000199 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector | Integer vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles | [5 empty] | 1
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  |  |  | 2
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma |  |  | 3
1 |  |  | 108.1995 | 0.26953 | 0 |  | -0.080501 | 1.2034 |  |  | 108.2878 | 0.30194 | 0 |  | 0.026989 | 0.57021 |  | 0 | 4
2 | 0 |  | 1.1846 | 0.08612 | 0 |  | -0.10623 | 0.082354 | 0 |  | -1.2466 | 0.071481 | 0 |  | 0.088942 | 0.10831 |  | 0 | 5
3 | 0 |  | 4.5075 | 0.03257 | 0 |  | 0.0044295 | 0.084889 | 0 |  | 4.5403 | 0.044741 | 0 |  | 0.049372 | 0.051142 |  | 0 | 6
4 | 0 |  | 0.026764 | 0.011005 | 0 |  | 0.0062513 | 0.019043 | 0 |  | -0.018638 | 0.0070451 | 0 |  | -0.017274 | 0.012844 |  | 0 | 7
5 | 0 |  | 0.030174 | 0.0037315 | 0 |  | -0.00092259 | 0.0077469 | 0 |  | 0.032951 | 0.0054 | 0 |  | 0.0031197 | 0.0066711 |  | 0 | 8
6 | 0 |  | -0.015687 | 0.0025681 | 0 |  | 0.0032132 | 0.0018039 | 0 |  | 0.021751 | 0.0040296 | 0 |  | -0.0057267 | 0.0032358 |  | 0 | 9
7 | 0 |  | -0.0061096 | 0.0022359 | 0 |  | 0.027801 | 0.0071459 | 0 |  | -0.0053354 | 0.0027917 | 0 |  | -0.0094271 | 0.0037703 |  | 0 | 10
8 | 0 |  | -0.0053432 | 0.0017936 | 0 |  | 0.0016056 | 0.001368 | 0 |  | 0.0043907 | 0.0022206 | 0 |  | 0.00011305 | 0.0023077 |  | 0 | 11
9 | 0 |  | 0.013561 | 0.0022907 | 0 |  | 0.015193 | 0.0050288 | 0 |  | 0.012392 | 0.0016568 | 0 |  | -0.0022386 | 0.0030228 |  | 0 | 12
10 | 0 |  | -9.67e-19 | 5.035e-18 | 0 |  | 3.5119e-19 | 3.9207e-18 | 0 |  | -2.6119e-18 | 5.1449e-18 | 0 |  | -3.4459e-19 | 3.439e-18 |  | 0 | 13
11 | 0 |  | 0.00637 | 0.0010112 | 0 |  | 0.031332 | 0.0085728 | 0 |  | 0.0098923 | 0.00078278 | 0 |  | -0.0096422 | 0.0039761 |  | 0 | 14
12 | 0 |  | -7.3946e-06 | 0.00054499 | 0 |  | 0.00037626 | 0.00054954 | 0 |  | -0.0016693 | 0.0010322 | 0 |  | 0.0011141 | 0.00070585 |  | 0 | 15
13 | 0 |  | -0.00013517 | 0.00058614 | 0 |  | 0.0035659 | 0.0011212 | 0 |  | -0.00065367 | 0.00089066 | 0 |  | 5.3247e-05 | 0.00046593 |  | 0 | 16
14 | 0 |  | -0.00049129 | 0.00050258 | 0 |  | -0.0072958 | 0.00042038 | 0 |  | -0.00091662 | 0.00066111 | 0 |  | -0.012059 | 0.00071792 |  | 0 | 17
15 | 0 |  | 0.0014419 | 0.0014457 | 0 |  | -0.0026557 | 0.0020323 | 0 |  | 0.0092786 | 0.002282 | 0 |  | -0.0085428 | 0.0029679 |  | 0 | 18
16 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 19
17 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | 20
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -2.4744 | -2.4306
3 | [12 empty] | 3 | 9 | -2.9454 | -3.04
4 | [12 empty] | 4 | 16 | -4.2352 | -4.6552
5 | [12 empty] | 5 | 25 | -5.2257 | -5.1157
6 | [12 empty] | 6 | 36 | -6.1412 | -5.6238
7 | [12 empty] | 7 | 49 | -5.5222 | -5.7309
8 | [12 empty] | 8 | 64 | -6.459 | -6.0908
9 | [12 empty] | 9 | 81 | -5.6857 | -6.1022
10 | [12 empty] | sigma (mm) |  | 0.0017456 | 0.0017239
11
12
13
14
15
16
17
Random harmonics
| d (mm) |  |  |  | sig(n)=d alpha beta^n
| 0.1 | 0.025 | 0.006
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939
5 | 0.61258 | 0.15315 | 0.036755
6 | 0.35886 | 0.089715 | 0.021532
7 | 0.20767 | 0.051917 | 0.01246
8 | 0.11871 | 0.029678 | 0.0071228
9 | 0.067037 | 0.016759 | 0.0040222
10 | 0.037395 | 0.0093488 | 0.0022437
11 | 0.020607 | 0.0051516 | 0.0012364
12 | 0.011217 | 0.0028042 | 0.00067302
13 | 0.0060316 | 0.0015079 | 0.0003619
14 | 0.0032039 | 0.00080097 | 0.00019223
15 | 0.0016811 | 0.00042029 | 0.00010087
16 | 0.0008714 | 0.00021785 | 5.2284e-05 | [5 empty] | Aperture 1 | Aperture 2
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | Magnetic length | -0.064552 | -0.064605
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 38.611 | 107.3241 | 108.2221 | 108.259 | 108.2551 | 108.25 | 108.3091 | 108.315 | 108.3326 | 108.2932 | 108.2463 | 108.3003 | 108.3543 | 108.3429 | 108.3397 | 108.2921 | 108.261 | 108.2562 | 107.6382 | 54.2842
2 | 26.1173 | 1.3694 | 1.1778 | 1.1968 | 1.2558 | 1.2117 | 1.2173 | 1.1745 | 1.2467 | 1.1753 | 1.2916 | 1.2357 | 1.0883 | 1.0184 | 1.0977 | 1.1282 | 1.0776 | 1.1181 | 1.2416 | 23.9336 | 0
3 | 6.4735 | 4.5051 | 4.5508 | 4.5069 | 4.4761 | 4.4941 | 4.4966 | 4.5199 | 4.5451 | 4.5587 | 4.4468 | 4.487 | 4.4751 | 4.51 | 4.4745 | 4.4987 | 4.5128 | 4.5731 | 4.5032 | 9.2286 | 0
4 | 1.5887 | 0.048308 | 0.026459 | 0.029059 | 0.026665 | 0.034762 | 0.035778 | 0.026446 | 0.041221 | 0.018015 | 0.036026 | 0.028725 | 0.024703 | 0.025406 | 0.021707 | 0.020456 | 0.0034525 | 0.005573 | 0.028982 | 0.65435 | 0
5 | 0.22011 | 0.028849 | 0.034565 | 0.027516 | 0.025632 | 0.032006 | 0.030438 | 0.032858 | 0.025461 | 0.034738 | 0.031519 | 0.035482 | 0.036218 | 0.02949 | 0.031257 | 0.02981 | 0.027109 | 0.026671 | 0.023515 | 0.19782 | 0
6 | 0.4368 | -0.018872 | -0.01436 | -0.013557 | -0.012831 | -0.01471 | -0.009327 | -0.015783 | -0.019018 | -0.017099 | -0.016184 | -0.015069 | -0.015036 | -0.014857 | -0.01381 | -0.018347 | -0.016289 | -0.019883 | -0.017333 | -0.0081576 | 0
7 | 0.046813 | -0.0039439 | -0.003202 | -0.0088823 | -0.0051848 | -0.0090429 | -0.0062263 | -0.0061513 | -0.0078938 | -0.0064614 | -0.0087185 | -0.0031228 | -0.0059759 | -0.0045809 | -0.0025633 | -0.0056746 | -0.008797 | -0.0042837 | -0.0092675 | 0.034169 | 0
8 | 0.12129 | -0.0050863 | -0.0060829 | -0.0058082 | -0.0036482 | -0.0066514 | -0.0032573 | -0.0095386 | -0.0039053 | -0.0061495 | -0.00097486 | -0.0047741 | -0.0055106 | -0.0053242 | -0.0048725 | -0.0054094 | -0.0065871 | -0.0071846 | -0.0054128 | 8.6326e-06 | 0
9 | 0.036819 | 0.013928 | 0.015525 | 0.013605 | 0.013942 | 0.012382 | 0.012352 | 0.012502 | 0.011815 | 0.019231 | 0.012752 | 0.014382 | 0.011057 | 0.018079 | 0.013403 | 0.012621 | 0.01421 | 0.012727 | 0.0095785 | 0.02511 | 0
10 | 4.6325e-17 | 6.9389e-18 | 3.4694e-18 | -1.2055e-18 | -1.7347e-18 | 0 | -6.3141e-18 | -6.7551e-18 | -3.2048e-18 | 0 | 2.9402e-18 | 3.2048e-18 | -1.1761e-17 | 5.8804e-18 | -7.9974e-18 | -3.8076e-18 | 3.9987e-18 | 0 | -1.0585e-18 | 1.5571e-17 | 0
11 | -0.0021384 | 0.0094482 | 0.0069207 | 0.0070977 | 0.0047713 | 0.0062077 | 0.0059971 | 0.0058459 | 0.005512 | 0.006572 | 0.0064321 | 0.0062067 | 0.0053232 | 0.0071249 | 0.0055071 | 0.006539 | 0.0062512 | 0.0070862 | 0.0058176 | 0.029368 | 0
12 | -0.01672 | -0.011866 | -0.010564 | 0.0081659 | 0.00079515 | 3.5696e-06 | 0.0036239 | -0.0010009 | 0.0023216 | 0.0012871 | 0.0035906 | 0.004189 | 0.0036391 | 0.0010552 | -0.0010639 | 0.0064646 | -0.0011834 | -0.0024957 | -0.0082928 | -0.031979 | 0 | ten times!
13 | 0.011245 | 0.0011716 | 0.0054433 | -0.0078512 | -0.0030458 | -0.0015243 | -0.0081288 | 0.0042912 | -0.004156 | 0.0055758 | 0.0019758 | 0.0077923 | -0.0051516 | -0.0023642 | 0.00047821 | -0.0027875 | 0.001083 | -0.00059259 | -0.01654 | 0.0062743 | 0 | ten times!
14 | 0.019368 | -0.012603 | -0.0054388 | -0.0051662 | -0.004788 | -0.0057337 | -0.0049301 | -0.010115 | -0.0017141 | -0.0056355 | -0.0037776 | -0.0012976 | -0.0059446 | 0.0048196 | -0.012266 | -0.0041679 | 0.0019327 | 0.0018843 | -0.013491 | -0.03049 | 0 | ten times!
15 | -0.038102 | 0.029377 | 0.021971 | 0.026303 | 0.017714 | 0.023817 | 0.013242 | 0.015531 | 0.013219 | 0.0074326 | 0.025402 | 0.0048261 | -0.007667 | 0.017739 | -0.020582 | 0.024319 | 0.025027 | 0.031027 | -0.0091655 | 0.058486 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 2.3008 | 1.5748 | 2.0196 | 2.424 | 1.4185 | 0.84048 | -0.80117 | -1.1489 | -1.1805 | -0.007106 | -0.47684 | -1.0359 | -0.50833 | -1.1373 | -1.3348 | -0.37616 | -0.016421 | -0.86377 | -0.83913 | 0.48312
2 | -18.8253 | -0.068402 | -0.20825 | -0.10452 | -0.17853 | -0.089497 | -0.10053 | 0.032961 | -0.03817 | -0.038618 | -0.079099 | -0.1459 | -0.2374 | -0.068734 | 0.043042 | -0.23463 | -0.20133 | -0.082346 | -0.11223 | 0.15269 | 0
3 | 0.58903 | 0.14541 | -0.12378 | 0.071346 | -0.0042712 | 0.050642 | 0.081111 | 0.015901 | 0.10671 | -0.023881 | 0.079244 | 0.01786 | -0.065587 | -0.063397 | -0.010233 | -0.025231 | -0.20128 | -0.026209 | 0.055373 | -0.019879 | 0
4 | -3.9939 | 0.033439 | -0.02045 | 0.013821 | -0.0014258 | 0.0076039 | -0.0019442 | 0.01258 | 0.029603 | 0.014536 | 0.023267 | 0.0024711 | -0.019041 | -0.01417 | 0.019926 | -0.022903 | -0.018254 | 0.022444 | 0.03102 | -0.18221 | 0
5 | -0.28469 | -0.0017898 | -0.019865 | -0.0016517 | -0.0042746 | 0.0065268 | 0.0040535 | 0.0063498 | 0.0056705 | -0.00067155 | 0.012323 | 0.0073365 | -0.0058034 | -0.0057665 | 0.0023945 | -0.00079244 | -0.011631 | -0.0083738 | -0.00064201 | -0.052227 | 0
6 | -0.64543 | 0.0031491 | 0.0016212 | 0.0026624 | 0.0024086 | 0.0023725 | 0.0050242 | 0.0033724 | 0.0035805 | 0.0033875 | 0.0066068 | 0.00029169 | 0.00064409 | 0.0017072 | 0.0039097 | 0.0050311 | 0.002446 | 0.0026603 | 0.0069616 | -0.014145 | 0
7 | 0.15901 | 0.031327 | 0.038357 | 0.044675 | 0.03454 | 0.035369 | 0.021053 | 0.021224 | 0.020882 | 0.02871 | 0.029672 | 0.022608 | 0.026382 | 0.01814 | 0.019755 | 0.027918 | 0.030618 | 0.024285 | 0.024898 | 0.038716 | 0
8 | -0.055745 | 0.0018167 | 0.0050047 | 0.0012758 | 0.002526 | -0.0012735 | 0.0013568 | 0.0005069 | 0.0017728 | 0.0017166 | 0.0023045 | 0.0010648 | 0.0022678 | 0.0016137 | -0.00088792 | 0.0022105 | 0.0013837 | 0.0029194 | 0.0013213 | -0.0065845 | 0
9 | 0.046421 | 0.01946 | 0.025244 | 0.022244 | 0.023702 | 0.01611 | 0.0098061 | 0.0094756 | 0.010088 | 0.016098 | 0.012749 | 0.011164 | 0.014807 | 0.010161 | 0.01188 | 0.017176 | 0.019244 | 0.012258 | 0.011802 | 0.035617 | 0
10 | 0 | 3.271e-19 | 2.9402e-18 | -5.1821e-19 | -3.837e-18 | 2.6462e-19 | 4.6749e-18 | 0 | -5.2924e-19 | 0 | -5.8804e-18 | -2.9402e-18 | 0 | -1.7347e-18 | -7.3505e-19 | -1.4701e-18 | 1.4701e-18 | 1.4701e-18 | 1.2819e-17 | -3.8246e-17 | 0
11 | 0.048889 | 0.04134 | 0.046446 | 0.049389 | 0.042784 | 0.038138 | 0.02542 | 0.023936 | 0.023844 | 0.032142 | 0.029914 | 0.024615 | 0.028368 | 0.023383 | 0.023285 | 0.02936 | 0.031828 | 0.024734 | 0.025054 | 0.03731 | 0 | ten times!
12 | 0.05959 | -0.00014039 | -0.0016217 | -0.0032672 | 0.0064509 | 0.0014078 | 0.013698 | -0.0014878 | 0.0056154 | 0.0082314 | -0.0035108 | 0.0058043 | -0.0031713 | 0.0055349 | -0.00023243 | 0.0039736 | 0.012305 | 0.010435 | 0.0077014 | -0.0078383 | 0 | ten times!
13 | 0.082277 | 0.050253 | 0.047297 | 0.046367 | 0.055861 | 0.047776 | 0.032053 | 0.023249 | 0.020776 | 0.03169 | 0.037097 | 0.036298 | 0.028151 | 0.029511 | 0.019459 | 0.032792 | 0.049318 | 0.029116 | 0.024799 | 0.049843 | 0 | ten times!
14 | 0.081238 | -0.08024 | -0.073535 | -0.072283 | -0.074784 | -0.072797 | -0.068018 | -0.072738 | -0.065916 | -0.081474 | -0.074312 | -0.073922 | -0.073904 | -0.069285 | -0.076766 | -0.071858 | -0.067534 | -0.075957 | -0.067919 | -0.030615 | 0 | ten times!
15 | 0.068481 | 0.0017431 | -0.0049251 | -0.0091286 | -0.032198 | 0.0075782 | -0.021984 | -0.033636 | -0.017659 | -0.021079 | -0.023307 | -0.019457 | -0.05676 | -0.025587 | -0.064215 | -0.029573 | -0.018162 | -0.054519 | -0.055154 | 0.038754 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 37.5041 | 107.244 | 108.3211 | 108.3789 | 108.4029 | 108.3493 | 108.3272 | 108.359 | 108.4352 | 108.4171 | 108.3811 | 108.4279 | 108.4439 | 108.4162 | 108.4552 | 108.4002 | 108.3297 | 108.3282 | 107.7628 | 62.1095
2 | -20.6484 | -1.3889 | -1.1962 | -1.3368 | -1.2194 | -1.2418 | -1.3266 | -1.2279 | -1.2258 | -1.3031 | -1.3263 | -1.2767 | -1.1586 | -1.1869 | -1.1963 | -1.1676 | -1.2449 | -1.1225 | -1.292 | -22.5237 | 0
3 | 4.825 | 4.5422 | 4.5593 | 4.5351 | 4.5493 | 4.518 | 4.5086 | 4.5194 | 4.5292 | 4.5864 | 4.4655 | 4.5792 | 4.5841 | 4.5418 | 4.4563 | 4.5706 | 4.584 | 4.6205 | 4.4768 | 9.2138 | 0
4 | 0.74717 | -0.024903 | -0.016312 | -0.029468 | -0.018767 | -0.024592 | -0.017773 | -0.015808 | -0.0075484 | -0.014905 | -0.027389 | -0.0059295 | -0.025411 | -0.016641 | -0.021637 | -0.02766 | -0.014278 | -0.0084949 | -0.017966 | -1.0804 | 0
5 | -0.077536 | 0.033731 | 0.03113 | 0.030398 | 0.036239 | 0.034252 | 0.037482 | 0.0356 | 0.036208 | 0.028498 | 0.038923 | 0.030307 | 0.03704 | 0.027305 | 0.035899 | 0.036237 | 0.035481 | 0.03269 | 0.015701 | 0.21733 | 0
6 | 0.48659 | 0.017736 | 0.025118 | 0.030219 | 0.022961 | 0.025738 | 0.020347 | 0.013703 | 0.028079 | 0.019178 | 0.020147 | 0.02536 | 0.019451 | 0.02134 | 0.021678 | 0.021865 | 0.019533 | 0.021797 | 0.017266 | 0.0035106 | 0
7 | 0.10429 | -0.00091797 | -0.0062455 | -0.0071937 | -0.0062757 | -0.0019515 | -0.004438 | -0.0042921 | -0.0033261 | -0.0082265 | -0.0076476 | -0.0030655 | -0.0080131 | -0.0030341 | -0.007127 | -0.0022814 | -0.0068031 | -0.0034665 | -0.011733 | 0.04966 | 0
8 | 0.08799 | 0.0069152 | 0.0033839 | 0.005045 | 0.0054467 | 0.0054133 | 0.0041114 | 0.0014507 | 0.0065037 | 0.0050827 | 0.0039962 | 0.0058582 | 0.0016005 | 0.00028765 | 0.007152 | 0.0052767 | 0.0018371 | 0.0018423 | 0.0078302 | 0.0055558 | 0
9 | 0.026564 | 0.012038 | 0.012347 | 0.012209 | 0.013691 | 0.011735 | 0.013531 | 0.011427 | 0.012862 | 0.012265 | 0.010701 | 0.013019 | 0.009788 | 0.011947 | 0.013382 | 0.014552 | 0.014823 | 0.014368 | 0.0083635 | 0.027933 | 0
10 | -5.5511e-17 | -1.7347e-18 | -5.8804e-18 | 0 | -3.8076e-18 | 3.6753e-19 | -1.1761e-17 | -3.4694e-18 | 6.4097e-18 | -5.0131e-18 | -8.6736e-19 | -3.2857e-18 | 2.9402e-18 | 2.9402e-18 | -8.0121e-19 | -3.4694e-18 | -1.4701e-18 | -2.117e-18 | -1.5995e-17 | 1.3878e-17 | 0
11 | 0.024221 | 0.01014 | 0.008375 | 0.0092614 | 0.010907 | 0.0098193 | 0.010277 | 0.011366 | 0.01063 | 0.0092065 | 0.00972 | 0.010006 | 0.0095041 | 0.0093017 | 0.010238 | 0.010367 | 0.010168 | 0.01034 | 0.0084351 | 0.03445 | 0
12 | -0.03524 | -0.010789 | -0.01587 | -0.01291 | -0.02368 | -0.0063125 | -0.030171 | -0.02868 | -0.020545 | -0.0053596 | -0.025014 | 0.0049173 | -0.031059 | -0.028428 | -0.016759 | -0.012216 | -0.0014568 | -0.018023 | -0.018125 | -0.051401 | 0 | ten times!
13 | 0.090818 | -0.0067554 | -0.002223 | 0.0010929 | 0.00049691 | -0.0025646 | -0.012475 | -0.0022375 | 0.011435 | -0.010822 | -0.010432 | -0.0008691 | -0.02312 | -0.007396 | -0.024387 | -0.016739 | -0.0015364 | -0.0005146 | -0.0086134 | -0.00041056 | 0 | ten times!
14 | -9.0718e-05 | -0.023507 | -0.019308 | -0.016877 | -0.0055498 | -0.0033359 | -0.010629 | -0.013991 | -0.0023158 | -0.0028912 | -0.0054911 | -0.0053599 | -0.0047283 | -0.0072673 | -0.017973 | -9.7729e-05 | -0.0097102 | -0.0089929 | -0.0069656 | -0.052737 | 0 | ten times!
15 | 0.0061513 | 0.098918 | 0.058553 | 0.096302 | 0.076688 | 0.10991 | 0.11801 | 0.064622 | 0.11036 | 0.062858 | 0.1227 | 0.10053 | 0.074032 | 0.12031 | 0.071296 | 0.090323 | 0.10506 | 0.12446 | 0.065232 | 0.1019 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | -1.1411 | -1.2029 | -1.064 | -0.1295 | -0.30536 | -0.15868 | -0.16519 | -0.168 | -0.17479 | 0.81329 | 0.47682 | 0.026081 | 0.38544 | 0.018725 | -0.053408 | 0.64188 | 1.0827 | 0.29345 | 0.16925 | 0.41614
2 | -15.102 | 0.20793 | -0.10528 | 0.13601 | 0.032384 | 0.17164 | 0.1487 | 0.17592 | 0.18321 | 0.17946 | 0.23701 | 0.10331 | -0.076203 | 0.050914 | 0.17766 | -0.081799 | -0.036597 | 0.056714 | 0.039984 | -0.28337 | 0
3 | -1.0868 | 0.04822 | -0.071317 | 0.067757 | 0.056596 | 0.040051 | 0.074755 | 0.033108 | 0.13179 | 0.0080148 | 0.093949 | 0.02814 | 0.015792 | -0.015159 | 0.04641 | 0.093126 | 0.025389 | 0.065067 | 0.147 | -0.069242 | 0
4 | -2.7829 | -0.015221 | -0.018578 | -0.013533 | -0.028237 | -0.00785 | -0.0090494 | -0.010278 | 0.0026771 | -0.006027 | -0.034443 | 0.0049783 | -0.035054 | -0.0090228 | -0.010388 | -0.033554 | -0.032182 | -0.030114 | -0.025061 | -0.3235 | 0
5 | 0.032429 | 0.014842 | 0.0040033 | 0.0020703 | 0.0067259 | -0.0037079 | 0.0084627 | 0.002752 | 0.0095597 | 0.00041509 | 0.0044415 | 0.0034264 | -0.0076616 | -0.013282 | 0.0048056 | 0.0089175 | 0.0021713 | -0.0010798 | 0.0092928 | 0.031195 | 0
6 | -0.47259 | -0.0061375 | 5.1613e-06 | -0.0054225 | -0.0073562 | -0.0026302 | -0.0059452 | -0.0050643 | -0.0032984 | -0.0071732 | -0.0080433 | -0.014679 | -0.0029775 | -0.0046955 | -0.0059846 | -0.0070995 | -0.0039983 | -0.0027936 | -0.009786 | -0.018975 | 0
7 | -0.068746 | -0.018405 | -0.012363 | -0.0099129 | -0.0093602 | -0.0092389 | -0.010522 | -0.012174 | -0.0065308 | -0.0038734 | -0.0064077 | -0.0082994 | -0.011583 | -0.0079703 | -0.0094233 | -0.0062836 | -0.0030512 | -0.008849 | -0.015441 | 0.0050435 | 0
8 | -0.01722 | -0.00026447 | 0.0041397 | 0.0028379 | -0.0016137 | 0.0041649 | 0.001783 | 0.0003436 | -0.0022307 | -0.0014247 | -0.00024071 | -0.0025169 | 0.0004813 | -0.0045387 | -0.00085925 | 0.00013532 | -0.0012994 | 0.00099664 | 0.0021409 | -0.00035577 | 0
9 | -0.066431 | -0.0078235 | -0.0037313 | -0.0032094 | 0.0015932 | -0.0058522 | -0.0012389 | -0.0017423 | -0.0026839 | 0.0010606 | 0.0011006 | -0.002745 | -0.0012442 | -0.0014111 | 0.0022411 | -0.004516 | 0.0014604 | -0.0043868 | -0.0071657 | 0.0087614 | 0
10 | -3.095e-18 | -1.3553e-20 | 0 | 7.3505e-19 | 3.4694e-18 | 0 | -4.6749e-18 | -4.6749e-18 | 6.2038e-18 | 2.9402e-18 | -2.9402e-18 | 2.9402e-18 | -3.4694e-18 | -3.4694e-18 | 1.7347e-18 | 3.6753e-19 | -2.9402e-18 | 3.4694e-18 | -5.8804e-18 | -2.1684e-19 | 0
11 | -0.026801 | -0.018079 | -0.017379 | -0.010871 | -0.011542 | -0.010748 | -0.009796 | -0.011067 | -0.010935 | -0.0045135 | -0.0058889 | -0.0102 | -0.0077689 | -0.010272 | -0.010534 | -0.0049358 | -0.0022491 | -0.0073141 | -0.0094662 | -0.00014808 | 0
12 | 0.070322 | 0.017581 | 0.018517 | 0.0059312 | 0.0012004 | 0.016535 | 0.013925 | 0.018054 | 0.023423 | 0.015378 | 0.0033136 | 0.0044794 | 0.0074554 | -0.00075419 | 0.010968 | 0.0051801 | 0.0067695 | 0.01757 | 0.015013 | -0.035606 | 0 | ten times!
13 | -0.03748 | -0.0032838 | -0.0026128 | 0.0046686 | -0.0044613 | -2.5588e-05 | -0.0036607 | 0.0048175 | -0.0063336 | 0.0053884 | 0.0025755 | -0.00055669 | -0.0013436 | 0.0019128 | 0.0062378 | 0.0034299 | 0.0089995 | -0.0078791 | 0.0017117 | 0.022482 | 0 | ten times!
14 | 0.096365 | -0.12622 | -0.12453 | -0.12637 | -0.11788 | -0.11534 | -0.11149 | -0.10825 | -0.10751 | -0.11754 | -0.12398 | -0.12196 | -0.13367 | -0.1155 | -0.11765 | -0.12341 | -0.12332 | -0.12856 | -0.12746 | -0.059333 | 0 | ten times!
15 | 0.089282 | -0.092949 | -0.089214 | -0.051493 | -0.096436 | -0.065455 | -0.099084 | -0.10176 | -0.070055 | -0.066218 | -0.078123 | -0.068789 | -0.12337 | -0.040557 | -0.08226 | -0.055758 | -0.066611 | -0.15178 | -0.13779 | 0.022774 | 0 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | -0.058 | 0.008 | 3.5 | 7 | Field Colinearity |  | 0 | 0.28 | 3.5 | 7
dB/B Heads CS+NCS (units) | -4070 | 100 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 108.3 | 0.38 | 4 | 8 | 0.34 | 1.52 | 4 | 8 | -4140 | 380 | 3.5 | 7 | -4022 | 211 | 3.5 | 7 | -2.1 | 2.02 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | -0.16 | 0.96 | 4 | 8 | 3.29 | 0.88 | 3.5 | 7 | 1.08 | 1.77 | 3.5 | 7
b2 (units) | 1.3 | 0.15 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 32 | 9 | 3.5 | 7 | 30 | 8 | 3.5 | 7
b3 (units) | 4.6 | 0.16 | 3.5 | 7 | 0 | 0.50847 | 4 | 8 | 1.44 | 5.5 | 3.5 | 7 | 7.55 | 1.64 | 3.5 | 7
b4 (units) | 0.06 | 0.034 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | 0.65 | 1.15 | 3.5 | 7 | 0.73 | 0.6 | 3.5 | 7
b5 (units) | 0.04 | 0.024 | 3.5 | 7 | 0 | 0.13559 | 4 | 8 | -0.7 | 0.82 | 3.5 | 7 | 0 | 0.23 | 3.5 | 7
b6 (units) | -0.019774 | 0.0054189 | 3.5 | 7 | 0 | 0.059322 | 4 | 8 | -0.0032 | 0.62 | 3.5 | 7 | 0 | 0.07 | 3.5 | 7
b7 (units) | -0.010837 | 0.01 | 3.5 | 7 | 0 | 0.045763 | 4 | 8 | -0.13 | 0.37 | 3.5 | 7 | 0 | 0.03 | 3.5 | 7
b8 (units) | -0.0031174 | 0.0034365 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | -0.004 | 0.097 | 3.5 | 7 | 0 | 0.038 | 3.5 | 7
b9 (units) | 0.0097838 | 0.014 | 3.5 | 7 | 0 | 0.016949 | 4 | 8 | 0.032 | 0.073 | 3.5 | 7 | 0.029 | 0.025 | 3.5 | 7
b10 (units) | 0.00014752 | 0.001858 | 3.5 | 7 | 0 | 0.033898 | 4 | 8 | -0.0029 | 0.146 | 3.5 | 7 | -0.0045 | 0.113 | 3.5 | 7
b11 (units) | 0 | 0.021 | 3.5 | 7 | 0 | 0.005 | 4 | 8 | 0.016 | 0.032 | 3.5 | 7 | 0.016 | 0.009 | 3.5 | 7
b12 (units) | 2.0565e-05 | 0.001 | 3.5 | 7 | 0 | 0.0051695 | 4 | 8 | -0.0024 | 0.02 | 3.5 | 7 | 0 | 0.015 | 3.5 | 7
b13 (units) | 0 | 0.0023601 | 3.5 | 7 | 0 | 0.0021186 | 4 | 8 | 0.0017 | 0.0066 | 3.5 | 7 | 0 | 0.0031 | 3.5 | 7
b14 (units) | -0.00049241 | 0.0062 | 3.5 | 7 | 0 | 0.0028814 | 4 | 8 | 0.0016 | 0.0084 | 3.5 | 7 | -0.0022 | 0.0071 | 3.5 | 7
b15 (units) | 0.001 | 0.0039 | 3.5 | 7 | 0 | 0.0022881 | 4 | 8 | -0.0075 | 0.0115 | 3.5 | 7 | -0.004 | 0.0067 | 3.5 | 7
a2 (units) | -0.03 | 0.11 | 3.5 | 7 | 0 | 0.9322 | 4 | 8 | 0.81 | 6.49 | 3.5 | 7 | -0.4 | 0.62 | 3.5 | 7
a3 (units) | 0 | 0.083 | 3.5 | 7 | 0 | 0.29661 | 4 | 8 | -1.63 | 2.13 | 3.5 | 7 | 0 | 0.2 | 3.5 | 7
a4 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.22881 | 4 | 8 | -0.6 | 1.67 | 3.5 | 7 | -0.32 | 0.26 | 3.5 | 7
a5 (units) | 0 | 0.045 | 3.5 | 7 | 0 | 0.10169 | 4 | 8 | -0.62 | 0.76 | 3.5 | 7 | 0 | 0.16 | 3.5 | 7
a6 (units) | 0 | 0.0049064 | 3.5 | 7 | 0 | 0.057627 | 4 | 8 | 0.35 | 0.66 | 3.5 | 7 | 0.02 | 0.07 | 3.5 | 7
a7 (units) | 0 | 0.026 | 3.5 | 7 | 0 | 0.047458 | 4 | 8 | -0.21 | 0.23 | 3.5 | 7 | 0 | 0.019 | 3.5 | 7
a8 (units) | 0.001 | 0.0021936 | 3.5 | 7 | 0 | 0.021186 | 4 | 8 | 0.067 | 0.115 | 3.5 | 7 | 0.026 | 0.037 | 3.5 | 7
a9 (units) | 0 | 0.018 | 3.5 | 7 | 0 | 0.017797 | 4 | 8 | -0.045 | 0.062 | 3.5 | 7 | -0.0084 | 0.034 | 3.5 | 7
a10 (units) | -1.8748e-05 | 0.00094989 | 3.5 | 7 | 0 | 0.028814 | 4 | 8 | 0.152 | 0.157 | 3.5 | 7 | 0.098 | 0.116 | 3.5 | 7
a11 (units) | -0.005 | 0.0511 | 3.5 | 7 | 0 | 0.0076271 | 4 | 8 | -0.0157 | 0.048 | 3.5 | 7 | 0.004 | 0.053 | 3.5 | 7
a12 (units) | 0 | 0.0008 | 3.5 | 7 | 0 | 0.0047458 | 4 | 8 | 0.032 | 0.029 | 3.5 | 7 | 0.0076 | 0.0107 | 3.5 | 7
a13 (units) | 0.0015 | 0.006 | 3.5 | 7 | 0 | 0.0018644 | 4 | 8 | 0 | 0.007 | 3.5 | 7 | 0.0015 | 0.0054 | 3.5 | 7
a14 (units) | -0.002 | 0.0039 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.003 | 0.008 | 3.5 | 7 | 0.0021 | 0.0063 | 3.5 | 7
a15 (units) | -0.0031 | 0.0066 | 3.5 | 7 | 0 | 0.0025424 | 4 | 8 | -0.0032 | 0.011 | 3.5 | 7 | 0.0004 | 0.0106 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  |  | Inf. R limit | Sup. R limit
Magnetic Length (mm) | -0.086 | -0.03 | -0.114 | -0.002 | Field Colinearity |  | -0.98 | 0.98 | -1.96 | 1.96 | Twist Integral |  |  | -0.18 | 0.18
dB/B Heads CS+NCS (units) | -4420 | -3720 | -4770 | -3370
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 106.78 | 109.82 | 105.26 | 111.34 | -5.74 | 6.42 | -11.82 | 12.5 | -5470 | -2810 | -6800 | -1480 | -4760.5 | -3283.5 | -5499 | -2545 | -9.17 | 4.97 | -16.24 | 12.04
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -4 | 3.68 | -7.84 | 7.52 | 0.21 | 6.37 | -2.87 | 9.45 | -5.115 | 7.275 | -11.31 | 13.47
b2 (units) | 0.775 | 1.825 | 0.25 | 2.35 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | 0.5 | 63.5 | -31 | 95 | 2 | 58 | -26 | 86
b3 (units) | 4.04 | 5.16 | 3.48 | 5.72 | -2.0339 | 2.0339 | -4.0678 | 4.0678 | -17.81 | 20.69 | -37.06 | 39.94 | 1.81 | 13.29 | -3.93 | 19.03
b4 (units) | -0.059 | 0.179 | -0.178 | 0.298 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.375 | 4.675 | -7.4 | 8.7 | -1.37 | 2.83 | -3.47 | 4.93
b5 (units) | -0.044 | 0.124 | -0.128 | 0.208 | -0.54237 | 0.54237 | -1.0847 | 1.0847 | -3.57 | 2.17 | -6.44 | 5.04 | -0.805 | 0.805 | -1.61 | 1.61
b6 (units) | -0.03874 | -0.00080783 | -0.057707 | 0.018158 | -0.23729 | 0.23729 | -0.47458 | 0.47458 | -2.1732 | 2.1668 | -4.3432 | 4.3368 | -0.245 | 0.245 | -0.49 | 0.49
b7 (units) | -0.045837 | 0.024163 | -0.080837 | 0.059163 | -0.18305 | 0.18305 | -0.3661 | 0.3661 | -1.425 | 1.165 | -2.72 | 2.46 | -0.105 | 0.105 | -0.21 | 0.21
b8 (units) | -0.015145 | 0.0089104 | -0.027173 | 0.020938 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3435 | 0.3355 | -0.683 | 0.675 | -0.133 | 0.133 | -0.266 | 0.266
b9 (units) | -0.039216 | 0.058784 | -0.088216 | 0.10778 | -0.067797 | 0.067797 | -0.13559 | 0.13559 | -0.2235 | 0.2875 | -0.479 | 0.543 | -0.0585 | 0.1165 | -0.146 | 0.204
b10 (units) | -0.0063556 | 0.0066507 | -0.012859 | 0.013154 | -0.13559 | 0.13559 | -0.27119 | 0.27119 | -0.5139 | 0.5081 | -1.0249 | 1.0191 | -0.4 | 0.391 | -0.7955 | 0.7865
b11 (units) | -0.0735 | 0.0735 | -0.147 | 0.147 | -0.02 | 0.02 | -0.04 | 0.04 | -0.096 | 0.128 | -0.208 | 0.24 | -0.0155 | 0.0475 | -0.047 | 0.079
b12 (units) | -0.0034794 | 0.0035206 | -0.0069794 | 0.0070206 | -0.020678 | 0.020678 | -0.041356 | 0.041356 | -0.0724 | 0.0676 | -0.1424 | 0.1376 | -0.0525 | 0.0525 | -0.105 | 0.105
b13 (units) | -0.0082604 | 0.0082604 | -0.016521 | 0.016521 | -0.0084746 | 0.0084746 | -0.016949 | 0.016949 | -0.0214 | 0.0248 | -0.0445 | 0.0479 | -0.01085 | 0.01085 | -0.0217 | 0.0217
b14 (units) | -0.022192 | 0.021208 | -0.043892 | 0.042908 | -0.011525 | 0.011525 | -0.023051 | 0.023051 | -0.0278 | 0.031 | -0.0572 | 0.0604 | -0.02705 | 0.02265 | -0.0519 | 0.0475
b15 (units) | -0.01265 | 0.01465 | -0.0263 | 0.0283 | -0.0091525 | 0.0091525 | -0.018305 | 0.018305 | -0.04775 | 0.03275 | -0.088 | 0.073 | -0.02745 | 0.01945 | -0.0509 | 0.0429
a2 (units) | -0.415 | 0.355 | -0.8 | 0.74 | -3.7288 | 3.7288 | -7.4576 | 7.4576 | -21.905 | 23.525 | -44.62 | 46.24 | -2.57 | 1.77 | -4.74 | 3.94
a3 (units) | -0.2905 | 0.2905 | -0.581 | 0.581 | -1.1864 | 1.1864 | -2.3729 | 2.3729 | -9.085 | 5.825 | -16.54 | 13.28 | -0.7 | 0.7 | -1.4 | 1.4
a4 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.91525 | 0.91525 | -1.8305 | 1.8305 | -6.445 | 5.245 | -12.29 | 11.09 | -1.23 | 0.59 | -2.14 | 1.5
a5 (units) | -0.1575 | 0.1575 | -0.315 | 0.315 | -0.40678 | 0.40678 | -0.81356 | 0.81356 | -3.28 | 2.04 | -5.94 | 4.7 | -0.56 | 0.56 | -1.12 | 1.12
a6 (units) | -0.017172 | 0.017172 | -0.034345 | 0.034345 | -0.23051 | 0.23051 | -0.46102 | 0.46102 | -1.96 | 2.66 | -4.27 | 4.97 | -0.225 | 0.265 | -0.47 | 0.51
a7 (units) | -0.091 | 0.091 | -0.182 | 0.182 | -0.18983 | 0.18983 | -0.37966 | 0.37966 | -1.015 | 0.595 | -1.82 | 1.4 | -0.0665 | 0.0665 | -0.133 | 0.133
a8 (units) | -0.0066776 | 0.0086776 | -0.014355 | 0.016355 | -0.084746 | 0.084746 | -0.16949 | 0.16949 | -0.3355 | 0.4695 | -0.738 | 0.872 | -0.1035 | 0.1555 | -0.233 | 0.285
a9 (units) | -0.063 | 0.063 | -0.126 | 0.126 | -0.071186 | 0.071186 | -0.14237 | 0.14237 | -0.262 | 0.172 | -0.479 | 0.389 | -0.1274 | 0.1106 | -0.2464 | 0.2296
a10 (units) | -0.0033434 | 0.0033059 | -0.006668 | 0.0066305 | -0.11525 | 0.11525 | -0.23051 | 0.23051 | -0.3975 | 0.7015 | -0.947 | 1.251 | -0.308 | 0.504 | -0.714 | 0.91
a11 (units) | -0.18385 | 0.17385 | -0.3627 | 0.3527 | -0.030508 | 0.030508 | -0.061017 | 0.061017 | -0.1837 | 0.1523 | -0.3517 | 0.3203 | -0.1815 | 0.1895 | -0.367 | 0.375
a12 (units) | -0.0028 | 0.0028 | -0.0056 | 0.0056 | -0.018983 | 0.018983 | -0.037966 | 0.037966 | -0.0695 | 0.1335 | -0.171 | 0.235 | -0.02985 | 0.04505 | -0.0673 | 0.0825
a13 (units) | -0.0195 | 0.0225 | -0.0405 | 0.0435 | -0.0074576 | 0.0074576 | -0.014915 | 0.014915 | -0.0245 | 0.0245 | -0.049 | 0.049 | -0.0174 | 0.0204 | -0.0363 | 0.0393
a14 (units) | -0.01565 | 0.01165 | -0.0293 | 0.0253 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.031 | 0.025 | -0.059 | 0.053 | -0.01995 | 0.02415 | -0.042 | 0.0462
a15 (units) | -0.0262 | 0.02 | -0.0493 | 0.0431 | -0.010169 | 0.010169 | -0.020339 | 0.020339 | -0.0417 | 0.0353 | -0.0802 | 0.0738 | -0.0367 | 0.0375 | -0.0738 | 0.0746
Coil positioning | 0 | 0.01 | 0 | 0.02
Blue: test on cold mass
Black: test on cm-collared coil

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.9293 | 2.7756 | 0.010508 | 0.009586 | 0.90392 | -1.4996 | -0.81108 | 0.38903 | -0.40666 | 16.5131 | 0.70442 | -0.86497 | 2.5825 | 0.10765 | -0.080299 | 0.93405 | -0.56635 | 0.29539 | 0.48943 | -0.30948 | 18.6734 | 0.70447 | 2.3498 | 4.1122 | 0.061235 | -0.2331 | 0.94878 | -0.0070155 | -0.48962 | 0.32288 | -0.3624 | 14356.879 | -2.1818 | 3.9056 | 0.043731 | -0.31436 | 0.98019 | 0.0028382 | 0.56295 | 0.43306 | -0.26839 | 14356.933
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | English | 0 | 0

## Inclinometer
Inclinometer (rad)
| Aperture 1 | Aperture 2
jig_position_NCS | -0.00027914 | -0.00027914
jig_position_CS | -0.00036085 | -0.00036085
jig_position_NCS_error | -5.761e-05 | -5.761e-05
jig_position_CS_error | -3.8214e-05 | -3.8214e-05